## Test Conditions
DESIREE-FIRE Test Number Japanese 3 - Circuit Simulator Data Only
Test Date: | 2009-09-23
DCSim DAQ Start Time | 8:47
DCSim Time for Synchronization | 10:18
Heat Source Synchronization | 1:20 | minutes:seconds
Heat Source DAQ Start Time | 10:17
Fire Time | 10:22
Offset | 1:35 | min:sec | This is the time after the DCSim DAQ was started and the heat source was initiated.
| 5700 | s
Heat Source Offset | 0:05 | minutes:seconds | This is the time after the Heat Source DAQ was started and the heat source was initiated.
| 320 | seconds
Heat Source Stop Time | 11:11
Test Duration | 0:49 | min:sec
| 2930 | s
CT500 turns | 2 |  | This is the number of times the lead conductor within the DCSim Panels were wrapped around the Hall Effect Device of the current transducer.
CT35 turns | 5
Pre-Test Baseline: | Gathered
Cable Information: | Record info about cable type on each circuit and which of the cable thermocouples corresponds to cable and air temperature
Unit number | Cable Type |  | Cable TC channel | Air TC Channel
MOV 1 | Mitsubishi, 6c, 14AWG |  | Channel 4, labeled as Cable 2 | Channel 4 of the first system in the Temperature Data File
MOV 2 | Mitsubishi, 6c, 14AWG |  | Channel 3, labeled as Cable 1 | Channel 4 of the first system in the Temperature Data File
A Note on Data Conversion and Accuracy:
Two different types of voltage transducers were used for each circuit. A bidirectional transducer (Ohio Semitronics, model number VT7-005X5-11) produces a 0 - 5V output signal which is directly proportional to the DC input ranging from 0 - 150VDC. The accuracy of this device is +/-0.25% full scale at DC and the response time is 50 mircoseconds. The unidirectional DC voltage isolator (Ohio Semitronics, model number VTU-005X5) used throughout tests also produces a 0 - 5V output over a 0 - 150VDC range. The accuracy of this device is +/-0.25% Rdg., +/-1% full scale at DC and the response time is 500 microseconds. Two different current transducers were for each circuit. One current transducer (Ohio Semitronics, model number CTL-51/35) was used to detect lower (0 - 35A) amperage transient behavior through a paired and calibrated signal conditioner (CTA201RX5). The other transducer (Ohio Semitronics, model number CTL-601/500) was set to detect transient current increases from 0 - 500A. These were also paired and calibrated with a signal conditioner (CTA2-1X5).
The readings were monitored using a National Instruments SCXI-1100 modules attached to SCXI-1300 multiplexers and housed within a SCXI-1001, 12-slot chassis. The "Raw Data" worksheet has been previously converted by LabView program created for this test series. The nominal zero point readings for each current transducer was taken before each test due to an electromagnetic effect created by direct current power. Zero Point values are provided, and the full run is documented in a separate data sheet (see "Current Offsets" worksheet). The converted data are presented as nominal indications of circuit path voltage and current. The data gathering equipment (transducers and DAQ card) have not been subject to specific calibration beyond that provided by the manufacturers. The accuracy of the converted data is estimated at +/-2%.

## Electrical Data
Fire Time (s) | Ground_VT7_1 | Ground_VTU_1 | Ground_VTU_2 | MOV-1_CT500 | MOV-1a_CT35_1_P | MOV-1a_CT35_2_N | MOV-1b_CT35_1_G | MOV-1b_CT35_2_R | MOV-1b_CT35_3_YO | MOV-1b_CT35_4_YC | MOV-2_CT500 | MOV-2a_CT35_1_P | MOV-2a_CT35_2_N | MOV-2b_CT35_1_G | MOV-2b_CT35_2_R | MOV-2b_CT35_3_YO | MOV-2b_CT35_4_YC1 | MOV-1a_VT7_1_SP | MOV-1a_VTU_1_P | MOV-1a_VTU_2_N | MOV-1b_VT7_1_G | MOV-1b_VT7_2_R | MOV-1b_VT7_3_YO1 | MOV-1b_VT7_4_YC1 | MOV-2a_VT7_1_SP | MOV-2a_VTU_1_P | MOV-2a_VTU_2_N | MOV-2b_VT7_1_G | MOV-2b_VT7_2_R | 2b_VT7_3_YO1 | MOV-2b_VT7_4_YC1 | {MOV1} V(N) - V(Batt -) | {MOV1} V(P) - V(Batt +) | {MOV1} V(G) - V(Batt +) | {MOV1} V(G) - V(P) | {MOV1} V(R) - V(Batt -) | {MOV1} V(R) - V(N) | {MOV1} V(YC1) - V(Batt -) | {MOV1} V(YO1) - V(Batt -) | {MOV1} V(YC1) - V(N) | {MOV1} V(YO1) - V(N) | {MOV2} V(N) - V(Batt -) | {MOV2} V(P) - V(Batt +) | {MOV2} V(G) - V(Batt +) | {MOV2} V(G) - V(P) | {MOV2} V(R) - V(Batt -) | {MOV2} V(R) - V(N) | {MOV2} V(YC1) - V(Batt -) | {MOV2} V(YO1) - V(Batt -) | {MOV2} V(YC1) - V(N) | {MOV2} V(YO1) - V(N)
-0.21875 | 122.6074 | -59.1934 | 63.0615 | -0.18394 | -0.00079462 | 0.069827 | 0.067424 | -0.00088229 | -0.00026063 | 0.002179 | -0.047089 | 0.0016685 | 0.066843 | 0.066396 | 3.2078e-05 | -0.0028347 | -0.0040439 | -0.050354 | 63.0478 | -59.3353 | 63.0569 | -59.0469 | -59.4818 | -59.3903 | -0.013733 | 63.0432 | -59.2667 | 63.1668 | -59.2026 | -59.3124 | -59.0607 | -0.14191 | -0.013732 | -0.004577 | 0.009155 | 0.14648 | 0.28839 | -0.19684 | -0.28839 | -0.054932 | -0.14648 | -0.073243 | -0.01831 | 0.10529 | 0.1236 | -0.009156 | 0.064087 | 0.13275 | -0.11902 | 0.20599 | -0.045776
0.78125 | 122.5388 | -59.2026 | 63.1165 | -0.084763 | -0.00036742 | 0.067691 | 0.065074 | 0.0010403 | 0.0023028 | -0.00059819 | -0.013519 | 0.00017309 | 0.062784 | 0.058065 | -0.0029587 | -0.0019803 | 0.0036465 | -0.07782 | 63.0524 | -59.3674 | 63.0249 | -59.0057 | -59.6054 | -59.5184 | -0.091553 | 63.0386 | -59.2529 | 63.176 | -59.1843 | -59.317 | -59.0424 | -0.16479 | -0.064087 | -0.091553 | -0.027466 | 0.19684 | 0.36163 | -0.31586 | -0.40283 | -0.15106 | -0.23804 | -0.050354 | -0.07782 | 0.059509 | 0.13733 | 0.018311 | 0.068665 | 0.16022 | -0.11444 | 0.21057 | -0.064087
1.7812 | 122.5388 | -59.1751 | 63.0798 | -0.10765 | 5.9776e-05 | 0.068545 | 0.067851 | -0.00066869 | 0.0010212 | 0.0011108 | -0.025726 | 0.00060029 | 0.065347 | 0.056569 | 0.00045928 | -0.0011257 | -0.00019868 | -0.086975 | 63.0615 | -59.3765 | 63.0249 | -59.0744 | -59.491 | -59.4131 | -0.036621 | 63.0615 | -59.2529 | 63.1348 | -59.0515 | -59.0927 | -58.8867 | -0.20142 | -0.018311 | -0.054932 | -0.036621 | 0.10071 | 0.30212 | -0.23804 | -0.31586 | -0.036621 | -0.11444 | -0.07782 | -0.018311 | 0.054932 | 0.073243 | 0.1236 | 0.20142 | 0.28839 | 0.082398 | 0.36621 | 0.16022
2.7812 | 122.5983 | -59.1934 | 63.1119 | -0.10765 | -0.00058102 | 0.065982 | 0.067637 | -0.0021641 | 0.0012348 | 0.00047001 | 0.0047913 | -0.00046791 | 0.065561 | 0.065755 | 3.2078e-05 | -0.0017667 | 0.0030057 | -0.073242 | 63.0753 | -59.3582 | 63.0341 | -59.0744 | -59.4772 | -59.4177 | -0.032043 | 63.0386 | -59.2392 | 63.1714 | -59.1705 | -59.285 | -59.0469 | -0.16479 | -0.036621 | -0.077819 | -0.041198 | 0.11902 | 0.28381 | -0.2243 | -0.28381 | -0.05951 | -0.11902 | -0.045777 | -0.073242 | 0.05951 | 0.13275 | 0.022888 | 0.068665 | 0.14648 | -0.091553 | 0.19226 | -0.045776
3.7812 | 122.6028 | -59.2163 | 63.089 | -0.061874 | 0.00048698 | 0.066836 | 0.069133 | 0.00039951 | -0.001756 | -0.001239 | -0.025726 | 0.00060029 | 0.063211 | 0.066396 | -0.00082252 | -0.0028347 | 0.00044212 | -0.082397 | 63.0661 | -59.3857 | 63.0386 | -59.0836 | -59.4864 | -59.4086 | -0.036621 | 63.0569 | -59.2712 | 63.1622 | -59.1934 | -59.2941 | -59.0469 | -0.16937 | -0.022888 | -0.050354 | -0.027466 | 0.13275 | 0.30212 | -0.19226 | -0.27008 | -0.022888 | -0.10071 | -0.054931 | -0.032043 | 0.073242 | 0.10528 | 0.022889 | 0.07782 | 0.16937 | -0.077819 | 0.2243 | -0.022888
4.7812 | 122.5479 | -59.1888 | 63.1119 | -0.092392 | 0.00070078 | 0.0694 | 0.068278 | -0.00066869 | -0.00047423 | 0.0019654 | -0.053192 | -0.00046791 | 0.063211 | 0.064473 | -0.0010361 | -0.0017667 | -0.00083948 | -0.07782 | 63.0341 | -59.3491 | 63.0341 | -59.079 | -59.4864 | -59.4086 | -0.050354 | 63.0524 | -59.2529 | 63.1485 | -59.198 | -59.317 | -59.0515 | -0.16022 | -0.077819 | -0.077819 | 0 | 0.10986 | 0.27008 | -0.21973 | -0.29755 | -0.059509 | -0.13733 | -0.064087 | -0.059509 | 0.036622 | 0.096131 | -0.009155 | 0.054932 | 0.13733 | -0.12817 | 0.20142 | -0.064087
5.7812 | 122.5983 | -59.198 | 63.1165 | -0.11528 | -0.0020764 | 0.068973 | 0.065715 | -0.0019503 | 0.0018756 | 0.00047001 | -0.028778 | 0.00017309 | 0.064279 | 0.062978 | -0.0010361 | -0.00069854 | -0.0040439 | -0.082397 | 63.0661 | -59.3353 | 63.0249 | -59.0881 | -59.4864 | -59.4131 | -0.041199 | 63.0707 | -59.2392 | 63.1622 | -59.1751 | -59.3033 | -59.0378 | -0.13733 | -0.050354 | -0.091553 | -0.041199 | 0.10986 | 0.24719 | -0.21515 | -0.28839 | -0.07782 | -0.15106 | -0.041199 | -0.045776 | 0.045776 | 0.091552 | 0.022888 | 0.064087 | 0.16022 | -0.10529 | 0.20142 | -0.064087
6.7812 | 122.612 | -59.2255 | 63.0798 | -0.12291 | 0.00091438 | 0.068973 | 0.069773 | -0.0019503 | 0.00016657 | 4.2808e-05 | -0.089813 | 0.0016685 | 0.065561 | 0.062337 | -0.0012497 | -0.0015529 | -0.00062588 | -0.054932 | 63.0341 | -59.3811 | 63.0753 | -59.0286 | -59.4498 | -59.3811 | 0.013733 | 63.0569 | -59.2712 | 63.1943 | -59.198 | -59.285 | -59.0149 | -0.15564 | -0.045776 | -0.004578 | 0.041198 | 0.19684 | 0.35248 | -0.15564 | -0.2243 | 0 | -0.068664 | -0.045776 | -0.022888 | 0.11444 | 0.13733 | 0.027466 | 0.073242 | 0.21057 | -0.059509 | 0.25635 | -0.013733
7.7812 | 122.5891 | -59.2072 | 63.0707 | -0.1458 | 0.0017688 | 0.065768 | 0.068064 | 0.00082671 | 0.0012348 | -0.001239 | -0.068451 | 0.0018821 | 0.064706 | 0.065328 | 0.00024568 | -0.0030483 | 1.4916e-05 | -0.059509 | 63.0615 | -59.3536 | 63.0844 | -59.0286 | -59.4498 | -59.3765 | 0.022888 | 63.0432 | -59.2346 | 63.1622 | -59.0744 | -59.3033 | -59.0744 | -0.14648 | -0.009156 | 0.013733 | 0.022889 | 0.17853 | 0.32501 | -0.16937 | -0.24262 | -0.022888 | -0.09613 | -0.027466 | -0.027466 | 0.091552 | 0.11902 | 0.13275 | 0.16022 | 0.13275 | -0.096131 | 0.16022 | -0.068665
8.7812 | 122.6395 | -59.1751 | 63.0936 | -0.069504 | 0.002196 | 0.067264 | 0.067424 | -2.7692e-05 | 0.00038017 | -0.002948 | -0.016571 | 0.0010275 | 0.06492 | 0.059133 | -0.0021041 | -0.0011257 | -0.0036167 | -0.009155 | 63.1027 | -59.3262 | 63.0936 | -59.0012 | -59.4223 | -59.3353 | 0.09613 | 63.0661 | -59.2072 | 63.2217 | -59.1019 | -59.2163 | -58.9737 | -0.15106 | 0.009155 | 0 | -0.009155 | 0.17395 | 0.32501 | -0.16022 | -0.24719 | -0.009155 | -0.09613 | -0.032043 | -0.027466 | 0.12817 | 0.15564 | 0.073242 | 0.10528 | 0.20142 | -0.041199 | 0.23346 | -0.009156
9.7812 | 122.6257 | -59.1934 | 63.0707 | -0.092392 | 0.001128 | 0.067264 | 0.068278 | 0.0023221 | 0.0010212 | 0.00068361 | -0.056244 | 0.0029501 | 0.06492 | 0.068746 | -0.00039512 | -0.00069854 | -0.0019077 | -0.041199 | 63.0295 | -59.3582 | 63.0524 | -59.0286 | -59.4727 | -59.3857 | -0.009155 | 63.0386 | -59.2438 | 63.1989 | -59.1522 | -59.2758 | -59.0012 | -0.16479 | -0.041199 | -0.018311 | 0.022888 | 0.1648 | 0.32959 | -0.19226 | -0.27924 | -0.027466 | -0.11444 | -0.050354 | -0.032044 | 0.12817 | 0.16022 | 0.041198 | 0.091552 | 0.19226 | -0.082398 | 0.24261 | -0.032044
10.7812 | 122.6166 | -59.1934 | 63.089 | -0.13054 | 5.9776e-05 | 0.068332 | 0.069773 | 0.0014675 | 0.0012348 | 0.003247 | -0.065399 | -0.00068151 | 0.062997 | 0.060628 | -0.0010361 | -0.00027134 | -0.00083948 | -0.064087 | 63.0386 | -59.3445 | 63.0569 | -59.024 | -59.4498 | -59.3903 | 0 | 63.0524 | -59.2621 | 63.1897 | -59.1705 | -59.2896 | -59.0195 | -0.15106 | -0.050354 | -0.032043 | 0.018311 | 0.16937 | 0.32043 | -0.19684 | -0.25635 | -0.045777 | -0.10529 | -0.068665 | -0.036621 | 0.10071 | 0.13733 | 0.022888 | 0.091553 | 0.17395 | -0.096131 | 0.24262 | -0.027466
11.7812 | 122.5662 | -59.2117 | 63.0936 | -0.077133 | 0.00091438 | 0.068332 | 0.066996 | -0.0010959 | 0.0025166 | -0.0016662 | -0.0043641 | 0.00038669 | 0.06492 | 0.063405 | 0.00088648 | -0.00048494 | 0.00022852 | -0.086975 | 63.0844 | -59.3399 | 63.0203 | -59.079 | -59.4772 | -59.4086 | -0.054932 | 63.0432 | -59.2529 | 63.1485 | -59.1888 | -59.317 | -59.0561 | -0.12817 | -0.009155 | -0.073242 | -0.064087 | 0.13275 | 0.26093 | -0.19684 | -0.2655 | -0.068664 | -0.13733 | -0.041199 | -0.050354 | 0.054932 | 0.10529 | 0.022888 | 0.064087 | 0.15564 | -0.10529 | 0.19684 | -0.064087
12.7812 | 122.5708 | -59.198 | 63.0707 | -0.12291 | 5.9776e-05 | 0.068332 | 0.067424 | 0.0012539 | 0.00038017 | 0.00068361 | -0.089813 | 0.00038669 | 0.063638 | 0.063405 | -0.0010361 | -0.00069854 | 0.00022852 | -0.045776 | 63.0524 | -59.3445 | 63.0798 | -59.024 | -59.4589 | -59.3765 | -0.009155 | 63.0386 | -59.2667 | 63.1989 | -59.1614 | -59.285 | -59.024 | -0.14648 | -0.018311 | 0.009155 | 0.027466 | 0.17395 | 0.32043 | -0.17853 | -0.26093 | -0.032044 | -0.11444 | -0.068665 | -0.032044 | 0.12817 | 0.16022 | 0.036621 | 0.10529 | 0.17395 | -0.086975 | 0.24262 | -0.01831
13.7812 | 122.6395 | -59.2072 | 63.0707 | -0.084763 | 0.0015552 | 0.070041 | 0.070842 | 0.0012539 | 0.00016657 | 0.00047001 | -0.10507 | -0.00046791 | 0.063425 | 0.067891 | 3.2078e-05 | 0.0010105 | 0.0010831 | -0.07782 | 63.0615 | -59.3719 | 63.0569 | -59.0469 | -59.4727 | -59.3994 | -0.027466 | 63.0478 | -59.2529 | 63.2034 | -59.166 | -59.2941 | -58.9966 | -0.1648 | -0.009156 | -0.013733 | -0.004577 | 0.16022 | 0.32501 | -0.19226 | -0.2655 | -0.027466 | -0.10071 | -0.045777 | -0.022888 | 0.13275 | 0.15564 | 0.041198 | 0.086975 | 0.21057 | -0.086975 | 0.25635 | -0.041198
14.7812 | 122.5937 | -59.2117 | 63.0661 | -0.12291 | -0.00036742 | 0.068332 | 0.069987 | 0.00061311 | -0.00047423 | 0.0019654 | -0.068451 | 0.0014547 | 0.064066 | 0.066182 | -0.0010361 | -0.0013393 | -0.0042575 | -0.10529 | 63.1027 | -59.3674 | 63.0112 | -59.0927 | -59.4955 | -59.4269 | -0.064087 | 63.0478 | -59.2712 | 63.1805 | -59.1568 | -59.3033 | -59.0057 | -0.15564 | 0.036621 | -0.054932 | -0.091553 | 0.11902 | 0.27466 | -0.21515 | -0.28381 | -0.059509 | -0.12817 | -0.059509 | -0.01831 | 0.11444 | 0.13275 | 0.054932 | 0.11444 | 0.20599 | -0.091553 | 0.2655 | -0.032044
15.7812 | 122.5342 | -59.1934 | 63.0936 | -0.10002 | -0.0018628 | 0.06705 | 0.068492 | -0.0010959 | -0.00026063 | -0.00038459 | -0.010467 | 0.00081389 | 0.06257 | 0.063405 | -0.0012497 | -0.0011257 | -0.0027621 | -0.082397 | 63.0707 | -59.3079 | 63.0112 | -59.0515 | -59.4589 | -59.3994 | -0.036621 | 63.0569 | -59.2255 | 63.1577 | -59.1843 | -59.2941 | -59.0515 | -0.11444 | -0.022888 | -0.082398 | -0.05951 | 0.14191 | 0.25635 | -0.20599 | -0.2655 | -0.091553 | -0.15106 | -0.032044 | -0.036621 | 0.064087 | 0.10071 | 0.009155 | 0.041199 | 0.14191 | -0.10071 | 0.17395 | -0.068664
16.7812 | 122.5937 | -59.2072 | 63.0661 | -0.092392 | 0.00048698 | 0.068759 | 0.069987 | -0.0025913 | 0.001662 | 0.0011108 | -0.095917 | -0.00068151 | 0.064706 | 0.065969 | -0.00060892 | 0.0022923 | -0.00019868 | -0.10529 | 63.0798 | -59.3811 | 63.0295 | -59.0607 | -59.491 | -59.4269 | -0.032043 | 63.0386 | -59.2621 | 63.176 | -59.1705 | -59.3033 | -59.0195 | -0.17395 | 0.013733 | -0.036621 | -0.050354 | 0.14648 | 0.32044 | -0.21973 | -0.28381 | -0.045776 | -0.10986 | -0.054932 | -0.027466 | 0.10986 | 0.13733 | 0.036621 | 0.091553 | 0.18768 | -0.096131 | 0.24262 | -0.041199
17.7812 | 122.5937 | -59.2163 | 63.0753 | -0.12291 | 0.0013416 | 0.068332 | 0.070414 | -0.00045509 | 0.00038017 | -0.0018798 | -0.092865 | 0.0020957 | 0.063852 | 0.069387 | -0.00082252 | -0.00048494 | -0.00062588 | -0.068665 | 63.089 | -59.3674 | 62.9791 | -59.1293 | -59.523 | -59.4681 | -0.07782 | 63.0432 | -59.2529 | 63.1668 | -59.1888 | -59.3079 | -59.0515 | -0.15106 | 0.013733 | -0.09613 | -0.10986 | 0.086976 | 0.23804 | -0.25177 | -0.3067 | -0.10071 | -0.15564 | -0.036621 | -0.032043 | 0.091553 | 0.1236 | 0.027466 | 0.064087 | 0.1648 | -0.091552 | 0.20142 | -0.054931
18.7812 | 122.5708 | -59.2163 | 63.121 | -0.13817 | -0.00079462 | 0.066623 | 0.065715 | 0.0016813 | 0.00059397 | -0.00038459 | 0.02005 | 0.0014547 | 0.063425 | 0.059347 | -0.00060892 | -0.0011257 | -0.00019868 | -0.082397 | 63.0569 | -59.3445 | 63.0157 | -59.079 | -59.4772 | -59.3994 | -0.054932 | 63.0386 | -59.2575 | 63.1439 | -59.1751 | -59.3033 | -59.0469 | -0.12817 | -0.064087 | -0.10529 | -0.041199 | 0.13733 | 0.2655 | -0.1831 | -0.26093 | -0.054932 | -0.13275 | -0.041198 | -0.082398 | 0.022888 | 0.10529 | 0.041199 | 0.082397 | 0.16937 | -0.086975 | 0.21057 | -0.045777
19.7812 | 122.5571 | -59.2346 | 63.0981 | -0.11528 | -0.00079462 | 0.068759 | 0.070201 | 0.0023221 | -0.0021832 | 0.003888 | -0.080658 | 0.0018821 | 0.06492 | 0.065542 | -0.0025315 | 0.00036966 | 0.0017239 | -0.10986 | 63.0844 | -59.404 | 62.9517 | -59.1476 | -59.5779 | -59.5001 | -0.13733 | 63.0569 | -59.2712 | 63.1165 | -59.2667 | -59.3948 | -59.0973 | -0.16937 | -0.013733 | -0.14649 | -0.13275 | 0.086975 | 0.25635 | -0.2655 | -0.34332 | -0.09613 | -0.17395 | -0.036621 | -0.041199 | 0.01831 | 0.059509 | -0.032044 | 0.004577 | 0.13733 | -0.16022 | 0.17395 | -0.1236
20.7812 | 122.4976 | -59.2392 | 63.1119 | -0.061874 | 0.0017688 | 0.066836 | 0.069133 | 0.0031765 | 0.00080757 | 0.0017518 | -0.10812 | 0.00081389 | 0.066843 | 0.061696 | -0.00039512 | -0.00048494 | -0.0021213 | -0.091553 | 63.089 | -59.3674 | 62.9379 | -59.198 | -59.5917 | -59.5139 | -0.15106 | 63.0936 | -59.285 | 63.0936 | -59.2804 | -59.4086 | -59.1202 | -0.12817 | -0.022888 | -0.17395 | -0.15106 | 0.041199 | 0.16937 | -0.27466 | -0.35248 | -0.14648 | -0.2243 | -0.045776 | -0.01831 | -0.01831 | 0 | -0.041199 | 0.004577 | 0.11902 | -0.16937 | 0.16479 | -0.1236
21.7812 | 122.4884 | -59.2392 | 63.1119 | -0.1458 | 0.001128 | 0.0694 | 0.070842 | -0.0015231 | 0.0031574 | 0.00025641 | -0.11423 | 0.0010275 | 0.06492 | 0.067678 | -0.00060892 | -0.0024075 | 1.4916e-05 | -0.10986 | 63.1027 | -59.404 | 62.9333 | -59.1888 | -59.5825 | -59.5184 | -0.14191 | 63.089 | -59.2712 | 63.0798 | -59.2896 | -59.4315 | -59.1385 | -0.1648 | -0.009155 | -0.17853 | -0.16937 | 0.050354 | 0.21515 | -0.27924 | -0.34332 | -0.11444 | -0.17853 | -0.032043 | -0.022888 | -0.032043 | -0.009155 | -0.050354 | -0.018311 | 0.10071 | -0.19226 | 0.13275 | -0.16022
22.7812 | 122.525 | -59.2392 | 63.0981 | -0.092392 | 0.00070078 | 0.068545 | 0.068919 | -0.0021641 | -0.00047423 | 0.0026062 | -0.086761 | 0.0010275 | 0.065347 | 0.065114 | 0.00024568 | -0.00069854 | -0.0010531 | -0.10529 | 63.1073 | -59.3811 | 62.9379 | -59.1705 | -59.5871 | -59.5001 | -0.14648 | 63.0844 | -59.285 | 63.0981 | -59.285 | -59.4131 | -59.1293 | -0.14191 | 0.009155 | -0.16022 | -0.16937 | 0.068665 | 0.21057 | -0.26093 | -0.3479 | -0.11902 | -0.20599 | -0.045776 | -0.013733 | 0 | 0.013733 | -0.045776 | 0 | 0.10986 | -0.17395 | 0.15564 | -0.12817
23.7812 | 122.5159 | -59.2346 | 63.1073 | -0.092392 | -0.00015382 | 0.068332 | 0.069346 | -0.0043003 | 0.00016657 | 0.0034606 | -0.092865 | -0.0015359 | 0.065775 | 0.070241 | 0.00067288 | 0.00058326 | 0.0012967 | -0.10986 | 63.1027 | -59.3719 | 62.915 | -59.198 | -59.5688 | -59.5047 | -0.16937 | 63.089 | -59.2896 | 63.0844 | -59.2941 | -59.4223 | -59.111 | -0.13733 | -0.004578 | -0.19226 | -0.18768 | 0.036621 | 0.17395 | -0.27008 | -0.33417 | -0.13275 | -0.19684 | -0.054932 | -0.018311 | -0.022888 | -0.004577 | -0.059509 | -0.004577 | 0.1236 | -0.18768 | 0.17853 | -0.13275
24.7812 | 122.4426 | -59.2255 | 63.121 | -0.069504 | -0.00079462 | 0.066836 | 0.068492 | -0.00024129 | 0.001662 | 0.00025641 | -0.08371 | 0.0014547 | 0.06492 | 0.062978 | 0.0015275 | -0.0011257 | -0.00041228 | -0.086975 | 63.0661 | -59.3582 | 62.9471 | -59.1843 | -59.5871 | -59.523 | -0.16022 | 63.0707 | -59.2896 | 63.0661 | -59.285 | -59.4315 | -59.1202 | -0.13275 | -0.054932 | -0.17395 | -0.11902 | 0.041199 | 0.17395 | -0.29755 | -0.36163 | -0.16479 | -0.22888 | -0.064087 | -0.050354 | -0.054932 | -0.004578 | -0.059509 | 0.004578 | 0.10529 | -0.20599 | 0.16937 | -0.14191
25.7812 | 122.4884 | -59.2575 | 63.1119 | -0.046616 | 0.0015552 | 0.067477 | 0.069133 | -2.7692e-05 | 0.00016657 | -0.001239 | -0.062347 | 0.0031637 | 0.065775 | 0.066396 | 0.00088648 | -0.00069854 | 0.0027921 | -0.10529 | 63.1073 | -59.3948 | 62.9333 | -59.166 | -59.5734 | -59.5047 | -0.16479 | 63.0981 | -59.2804 | 63.0753 | -59.2896 | -59.4131 | -59.1202 | -0.13733 | -0.004577 | -0.17853 | -0.17395 | 0.091552 | 0.22888 | -0.24719 | -0.31586 | -0.10986 | -0.17853 | -0.022889 | -0.013732 | -0.036621 | -0.022889 | -0.032044 | -0.009155 | 0.13733 | -0.15564 | 0.16022 | -0.13275
26.7812 | 122.5159 | -59.2438 | 63.1027 | -0.084763 | 0.0013416 | 0.068332 | 0.067424 | 0.0038175 | 0.0010212 | 0.00047001 | -0.05014 | 0.0018821 | 0.065561 | 0.067678 | 0.00024568 | -0.0030483 | 0.00086952 | -0.036621 | 63.0981 | -59.3399 | 63.0203 | -59.0836 | -59.4864 | -59.4406 | 0.009155 | 63.0844 | -59.2621 | 63.121 | -59.23 | -59.3536 | -59.0973 | -0.096131 | -0.004577 | -0.082397 | -0.07782 | 0.16022 | 0.25635 | -0.19684 | -0.24262 | -0.10071 | -0.14648 | -0.018311 | -0.01831 | 0.018311 | 0.036621 | 0.013732 | 0.032043 | 0.14648 | -0.10986 | 0.16479 | -0.091553
27.7812 | 122.4701 | -59.2346 | 63.1256 | -0.1458 | 0.0013416 | 0.067477 | 0.068492 | 0.0023221 | 0.00080757 | 0.0011108 | -0.062347 | 0.00081389 | 0.063638 | 0.067251 | 3.2078e-05 | -0.0011257 | 0.00065592 | -0.09613 | 63.089 | -59.3628 | 62.9562 | -59.1797 | -59.5642 | -59.5093 | -0.15106 | 63.0981 | -59.2575 | 63.0661 | -59.317 | -59.4406 | -59.1476 | -0.12817 | -0.036621 | -0.16937 | -0.13275 | 0.054932 | 0.18311 | -0.27466 | -0.32959 | -0.14648 | -0.20142 | -0.022888 | -0.027465 | -0.059509 | -0.032044 | -0.082398 | -0.05951 | 0.086975 | -0.20599 | 0.10986 | -0.18311
28.7812 | 122.4655 | -59.2529 | 63.1348 | -0.077133 | 0.00070078 | 0.068545 | 0.069133 | 0.00018591 | 0.0014484 | -0.00059819 | -0.065399 | 0.00060029 | 0.066415 | 0.069814 | 0.00045928 | 0.0018649 | -0.0016941 | -0.07782 | 63.0981 | -59.3811 | 62.97 | -59.1522 | -59.5642 | -59.491 | -0.13275 | 63.121 | -59.2804 | 63.0615 | -59.2896 | -59.4223 | -59.1202 | -0.12817 | -0.036621 | -0.16479 | -0.12817 | 0.10071 | 0.22888 | -0.23804 | -0.31128 | -0.10986 | -0.1831 | -0.027466 | -0.013733 | -0.073243 | -0.05951 | -0.036621 | -0.009155 | 0.13275 | -0.16937 | 0.16022 | -0.14191
29.7812 | 122.5571 | -59.2163 | 63.1577 | -0.061874 | 0.0019824 | 0.068332 | 0.068492 | 0.00039951 | 0.00038017 | 0.00025641 | -0.040985 | 0.00017309 | 0.067697 | 0.065969 | 0.00024568 | 0.0014377 | 0.0027921 | -0.013733 | 63.0936 | -59.3399 | 63.002 | -59.0744 | -59.491 | -59.436 | -0.004578 | 63.089 | -59.2529 | 63.1348 | -59.2117 | -59.3628 | -59.0698 | -0.1236 | -0.064087 | -0.15564 | -0.091553 | 0.14191 | 0.2655 | -0.21973 | -0.27466 | -0.09613 | -0.15106 | -0.036621 | -0.068665 | -0.022888 | 0.045777 | 0.004578 | 0.041199 | 0.14649 | -0.14648 | 0.18311 | -0.10986
30.7812 | 122.4838 | -59.2026 | 63.1577 | -0.10002 | 0.0013416 | 0.070041 | 0.069346 | -0.00024129 | 0.0014484 | -0.00059819 | -0.025726 | 0.0012411 | 0.068124 | 0.07195 | -0.00018152 | -5.7742e-05 | 0.0038601 | -0.045776 | 63.1119 | -59.3582 | 62.97 | -59.1064 | -59.5276 | -59.4589 | -0.068665 | 63.0981 | -59.23 | 63.0981 | -59.2575 | -59.3857 | -59.111 | -0.15564 | -0.045777 | -0.18768 | -0.14191 | 0.096131 | 0.25177 | -0.25635 | -0.32501 | -0.10071 | -0.16937 | -0.027466 | -0.059509 | -0.059509 | 0 | -0.054931 | -0.027465 | 0.091553 | -0.1831 | 0.11902 | -0.15564
31.7812 | 122.5204 | -59.2209 | 63.1622 | -0.092392 | 0.00048698 | 0.068545 | 0.069987 | 0.0031765 | -0.00026063 | 0.002179 | -0.034882 | 0.0020957 | 0.066415 | 0.060201 | -0.00060892 | -0.0015529 | 0.0025785 | -0.036621 | 63.1027 | -59.3491 | 62.9974 | -59.0973 | -59.5139 | -59.4498 | -0.032043 | 63.1027 | -59.2346 | 63.1073 | -59.2484 | -59.3948 | -59.0881 | -0.12817 | -0.059509 | -0.16479 | -0.10529 | 0.1236 | 0.25177 | -0.22888 | -0.29297 | -0.10071 | -0.16479 | -0.013733 | -0.059509 | -0.054931 | 0.004578 | -0.027466 | -0.013733 | 0.13275 | -0.17395 | 0.14648 | -0.16022
32.7812 | 122.5113 | -59.23 | 63.121 | -0.061874 | 0.0013416 | 0.068118 | 0.068064 | -0.0023777 | 0.00080757 | -0.001239 | -0.034882 | 0.0014547 | 0.063211 | 0.062551 | 0.00088648 | -0.00027134 | -0.0014805 | -0.013733 | 63.1165 | -59.3536 | 62.9974 | -59.0698 | -59.5047 | -59.436 | -0.004578 | 63.089 | -59.2438 | 63.1073 | -59.2255 | -59.3674 | -59.1156 | -0.1236 | -0.004578 | -0.1236 | -0.11902 | 0.16022 | 0.28381 | -0.20599 | -0.27466 | -0.082397 | -0.15106 | -0.013732 | -0.032044 | -0.013733 | 0.018311 | 0.004578 | 0.01831 | 0.11444 | -0.13733 | 0.12817 | -0.1236
33.7812 | 122.5937 | -59.1843 | 63.0844 | -0.13817 | 0.0015552 | 0.068332 | 0.066569 | -0.0017367 | 0.0035846 | -0.0025208 | 0.02005 | -4.0508e-05 | 0.067697 | 0.061696 | 0.00024568 | -0.0013393 | 0.0036465 | -0.054932 | 63.0478 | -59.3353 | 63.0432 | -59.0286 | -59.4589 | -59.3948 | -0.018311 | 63.0203 | -59.23 | 63.1577 | -59.1614 | -59.2758 | -58.9966 | -0.15106 | -0.036621 | -0.041199 | -0.004578 | 0.15564 | 0.3067 | -0.21057 | -0.27466 | -0.059509 | -0.1236 | -0.045777 | -0.064087 | 0.073242 | 0.13733 | 0.022888 | 0.068665 | 0.18768 | -0.091553 | 0.23346 | -0.045776
34.7812 | 122.5296 | -59.2072 | 63.1348 | -0.084763 | 0.00048698 | 0.069827 | 0.069346 | 0.0012539 | -0.00026063 | -0.001239 | -0.028778 | -0.00068151 | 0.063211 | 0.066182 | -0.00018152 | -0.00048494 | -0.0010531 | 0.009155 | 63.1165 | -59.3262 | 63.0295 | -59.0652 | -59.4589 | -59.4131 | 0.036621 | 63.0936 | -59.2438 | 63.1393 | -59.2163 | -59.3491 | -59.0607 | -0.11902 | -0.018311 | -0.10529 | -0.086975 | 0.14191 | 0.26093 | -0.20599 | -0.25177 | -0.086975 | -0.13275 | -0.036621 | -0.041199 | 0.004577 | 0.045776 | -0.009156 | 0.027465 | 0.14648 | -0.14191 | 0.18311 | -0.10529
35.7812 | 122.5479 | -59.2163 | 63.1393 | -0.13054 | -0.00058102 | 0.068973 | 0.068919 | 0.0021085 | -0.0023968 | 0.00025641 | -0.03183 | 0.0016685 | 0.065134 | 0.065755 | -0.0016769 | 0.00079686 | 0.00044212 | -0.032043 | 63.089 | -59.3491 | 62.9974 | -59.1202 | -59.4955 | -59.4543 | -0.027466 | 63.0936 | -59.2667 | 63.1027 | -59.2575 | -59.3811 | -59.0973 | -0.13275 | -0.050354 | -0.14191 | -0.091552 | 0.096131 | 0.22888 | -0.23804 | -0.27923 | -0.10529 | -0.14648 | -0.050354 | -0.045776 | -0.036621 | 0.009155 | -0.041198 | 0.009156 | 0.11902 | -0.16479 | 0.16937 | -0.11444
36.7812 | 122.5204 | -59.2117 | 63.1348 | -0.061874 | -0.0014356 | 0.068545 | 0.067424 | -0.0049411 | 0.0014484 | 0.00047001 | -0.03183 | 0.0012411 | 0.066415 | 0.067037 | 0.00045928 | -0.0011257 | -0.0016941 | 0.009155 | 63.1073 | -59.3628 | 63.0112 | -59.0698 | -59.4818 | -59.4223 | 0.027466 | 63.0981 | -59.2163 | 63.1393 | -59.2026 | -59.3262 | -59.0607 | -0.15106 | -0.027466 | -0.1236 | -0.096131 | 0.14191 | 0.29297 | -0.21057 | -0.27008 | -0.059509 | -0.11902 | -0.004578 | -0.036621 | 0.004577 | 0.041198 | 0.009155 | 0.013733 | 0.15106 | -0.11444 | 0.15564 | -0.10986
37.7812 | 122.493 | -59.2438 | 63.1302 | -0.11528 | 0.00048698 | 0.066623 | 0.068064 | 0.0046719 | 0.00059397 | 0.0028198 | -0.040985 | 0.0025229 | 0.066415 | 0.068105 | 0.00067288 | -0.00048494 | 0.0019375 | -0.10529 | 63.1027 | -59.3628 | 62.9333 | -59.198 | -59.5825 | -59.5322 | -0.14191 | 63.0753 | -59.2712 | 63.0569 | -59.2987 | -59.4177 | -59.1156 | -0.11902 | -0.027466 | -0.19684 | -0.16937 | 0.045776 | 0.1648 | -0.28839 | -0.33875 | -0.16937 | -0.21973 | -0.027466 | -0.054932 | -0.073242 | -0.01831 | -0.054932 | -0.027466 | 0.12817 | -0.17395 | 0.15564 | -0.14648
38.7812 | 122.6074 | -59.1888 | 63.0936 | -0.061874 | 0.00070078 | 0.068973 | 0.069133 | 0.0010403 | 0.00080757 | 0.0019654 | -0.010467 | -0.00068151 | 0.065347 | 0.058278 | -0.0010361 | -0.00091214 | -0.0019077 | -0.10071 | 63.0844 | -59.3491 | 62.9791 | -59.1385 | -59.5551 | -59.4543 | -0.10529 | 63.0386 | -59.2484 | 63.1256 | -59.2255 | -59.3399 | -59.0561 | -0.16022 | -0.009155 | -0.11444 | -0.10529 | 0.050354 | 0.21057 | -0.2655 | -0.36621 | -0.10529 | -0.20599 | -0.059509 | -0.054932 | 0.032043 | 0.086975 | -0.036621 | 0.022888 | 0.13275 | -0.15106 | 0.19226 | -0.091553
39.7812 | 122.5662 | -59.1934 | 63.1348 | -0.10002 | 0.0015552 | 0.065768 | 0.068278 | 0.0038175 | 0.0018756 | 0.00025641 | -0.013519 | -4.0508e-05 | 0.065134 | 0.059987 | 0.00088648 | 0.00058326 | 0.00065592 | 0.013733 | 63.1073 | -59.3491 | 63.0066 | -59.0698 | -59.4727 | -59.4223 | 0.022888 | 63.1027 | -59.2392 | 63.1531 | -59.2163 | -59.3353 | -59.0607 | -0.15564 | -0.027466 | -0.12817 | -0.10071 | 0.1236 | 0.27924 | -0.22888 | -0.27924 | -0.073242 | -0.1236 | -0.045777 | -0.032044 | 0.01831 | 0.050354 | -0.022889 | 0.022888 | 0.13275 | -0.14191 | 0.17853 | -0.09613
40.7812 | 122.5616 | -59.198 | 63.0936 | -0.092392 | 0.0015552 | 0.06705 | 0.06721 | -0.0055819 | 0.00038017 | 0.0019654 | 0.056671 | -0.0013223 | 0.064279 | 0.059987 | 0.0013139 | -0.0021939 | 1.4916e-05 | -0.054932 | 63.0386 | -59.3445 | 63.0478 | -59.0057 | -59.4543 | -59.3857 | 0.050354 | 63.0524 | -59.198 | 63.1348 | -59.1843 | -59.3079 | -59.0378 | -0.14648 | -0.054932 | -0.045776 | 0.009156 | 0.19226 | 0.33874 | -0.18768 | -0.25635 | -0.041199 | -0.10986 | 0 | -0.041199 | 0.041199 | 0.082398 | 0.013733 | 0.013733 | 0.16022 | -0.10986 | 0.16022 | -0.10986
41.7812 | 122.5983 | -59.2072 | 63.1393 | -0.069504 | -0.00079462 | 0.069186 | 0.069987 | -0.0010959 | -0.001115 | 0.001538 | -0.034882 | 0.0014547 | 0.065347 | 0.067251 | -0.00082252 | -0.0024075 | -0.0010531 | 0 | 63.1302 | -59.3491 | 63.0203 | -59.079 | -59.4727 | -59.4223 | 0.045776 | 63.089 | -59.2255 | 63.1485 | -59.2072 | -59.3308 | -59.0561 | -0.14191 | -0.009155 | -0.11902 | -0.10986 | 0.12817 | 0.27008 | -0.21515 | -0.2655 | -0.073242 | -0.1236 | -0.018311 | -0.050354 | 0.009156 | 0.05951 | 0 | 0.018311 | 0.15106 | -0.1236 | 0.16937 | -0.10529
42.7812 | 122.5433 | -59.2117 | 63.1393 | -0.13817 | 0.001128 | 0.068759 | 0.070201 | 0.0025357 | -0.00047423 | -0.0010254 | -0.053192 | 0.0023093 | 0.066843 | 0.059987 | -0.0010361 | -0.0021939 | -0.0036167 | 0 | 63.1027 | -59.3582 | 63.0112 | -59.0836 | -59.4818 | -59.4177 | 0.009155 | 63.089 | -59.2392 | 63.1348 | -59.2163 | -59.3536 | -59.0927 | -0.14648 | -0.036621 | -0.12817 | -0.091553 | 0.12817 | 0.27466 | -0.20599 | -0.27008 | -0.05951 | -0.1236 | -0.027466 | -0.050354 | -0.004577 | 0.045777 | -0.004578 | 0.022888 | 0.11902 | -0.14191 | 0.14648 | -0.11444
43.7812 | 122.4518 | -59.2392 | 63.121 | -0.11528 | 0.0013416 | 0.06705 | 0.068064 | 0.0031765 | 0.0023028 | 0.00047001 | -0.044037 | 0.00017309 | 0.064066 | 0.059987 | 0.00045928 | 0.00036966 | 0.0025785 | -0.10529 | 63.0753 | -59.3628 | 62.9288 | -59.1888 | -59.5734 | -59.5139 | -0.14648 | 63.0569 | -59.2804 | 63.0615 | -59.2941 | -59.4315 | -59.1248 | -0.1236 | -0.045777 | -0.19226 | -0.14648 | 0.050354 | 0.17395 | -0.27466 | -0.33417 | -0.15106 | -0.21057 | -0.041199 | -0.064087 | -0.05951 | 0.004577 | -0.054931 | -0.013732 | 0.11444 | -0.19226 | 0.15564 | -0.15106
44.7812 | 122.5159 | -59.1888 | 63.1439 | -0.054245 | 0.002837 | 0.066623 | 0.069346 | 0.0018949 | 0.0012348 | 0.0013244 | -0.037933 | 0.00017309 | 0.065134 | 0.064473 | -0.0010361 | -0.00091214 | -0.0036167 | -0.004578 | 63.1302 | -59.3308 | 63.0112 | -59.0607 | -59.4635 | -59.4269 | 0.041199 | 63.0844 | -59.23 | 63.1393 | -59.2072 | -59.3399 | -59.0652 | -0.14191 | -0.013733 | -0.13275 | -0.11902 | 0.12817 | 0.27008 | -0.23804 | -0.27466 | -0.09613 | -0.13275 | -0.041199 | -0.059509 | -0.004578 | 0.054931 | -0.01831 | 0.022889 | 0.1236 | -0.15106 | 0.16479 | -0.10986
45.7969 | 122.5754 | -59.198 | 63.1165 | -0.092392 | -0.00036742 | 0.069827 | 0.066355 | -0.0017367 | 0.0014484 | -0.00081179 | 0.065826 | -0.0011087 | 0.065561 | 0.06896 | 0.00024568 | -0.0026211 | -0.0019077 | -0.082397 | 63.0432 | -59.3491 | 62.9974 | -59.0881 | -59.4727 | -59.3994 | -0.036621 | 63.0432 | -59.2484 | 63.1165 | -59.1934 | -59.3216 | -59.0561 | -0.15106 | -0.073242 | -0.11902 | -0.045776 | 0.10986 | 0.26093 | -0.20142 | -0.27466 | -0.050354 | -0.1236 | -0.050354 | -0.073242 | 0 | 0.073242 | 0.004578 | 0.054932 | 0.14191 | -0.1236 | 0.19226 | -0.073242
46.7812 | 122.5204 | -59.1705 | 63.1485 | -0.00083923 | 5.9776e-05 | 0.067264 | 0.069133 | -0.0013095 | -0.00026063 | 0.00068361 | -0.0043641 | 0.0020957 | 0.065347 | 0.065114 | -0.00060892 | -0.00027134 | 0.00044212 | 0.004578 | 63.1302 | -59.3262 | 63.0295 | -59.0607 | -59.4727 | -59.4315 | 0.032043 | 63.0844 | -59.2346 | 63.1348 | -59.2026 | -59.3399 | -59.0561 | -0.15564 | -0.018311 | -0.11902 | -0.10071 | 0.10986 | 0.2655 | -0.26093 | -0.30212 | -0.10529 | -0.14648 | -0.064087 | -0.064087 | -0.013733 | 0.050354 | -0.032044 | 0.032043 | 0.11444 | -0.16937 | 0.17853 | -0.10529
47.7812 | 122.5525 | -59.2072 | 63.1622 | -0.023727 | 0.00048698 | 0.065982 | 0.068919 | -0.00024129 | -0.0019696 | 0.00089721 | -0.010467 | 0.0010275 | 0.066629 | 0.057638 | -0.00060892 | -0.0026211 | 0.0034329 | -0.013733 | 63.1577 | -59.3399 | 62.9929 | -59.1019 | -59.4727 | -59.436 | -0.013733 | 63.1073 | -59.2438 | 63.089 | -59.2072 | -59.3124 | -59.0561 | -0.13275 | -0.004577 | -0.16937 | -0.16479 | 0.10528 | 0.23804 | -0.22888 | -0.2655 | -0.09613 | -0.13275 | -0.036621 | -0.054931 | -0.073242 | -0.018311 | 0 | 0.036621 | 0.15106 | -0.10529 | 0.18768 | -0.068665
48.7812 | 122.5113 | -59.2438 | 63.1119 | -0.10765 | 0.00070078 | 0.067264 | 0.067424 | -0.0023777 | 0.00038017 | -0.00081179 | -0.047089 | 0.00081389 | 0.066629 | 0.065755 | -0.00060892 | -5.7742e-05 | 0.00044212 | -0.10071 | 63.0798 | -59.3491 | 62.9333 | -59.1797 | -59.5825 | -59.5001 | -0.14648 | 63.0707 | -59.2804 | 63.0661 | -59.285 | -59.4269 | -59.111 | -0.10529 | -0.032043 | -0.17853 | -0.14648 | 0.064087 | 0.16937 | -0.25635 | -0.33875 | -0.15106 | -0.23346 | -0.036622 | -0.041198 | -0.045776 | -0.004578 | -0.041199 | -0.004577 | 0.13275 | -0.18311 | 0.16937 | -0.14648
49.7812 | 122.5296 | -59.1797 | 63.1348 | -0.046616 | 0.0015552 | 0.067691 | 0.068705 | -0.00088229 | 0.00059397 | 0.0017518 | -0.025726 | 0.0012411 | 0.065347 | 0.067037 | -0.0012497 | -0.00069854 | -0.00019868 | -0.009155 | 63.1073 | -59.3445 | 63.0386 | -59.0561 | -59.4727 | -59.4223 | 0.036621 | 63.0707 | -59.23 | 63.1348 | -59.2209 | -59.3262 | -59.0469 | -0.16479 | -0.027466 | -0.096131 | -0.068665 | 0.1236 | 0.28839 | -0.24262 | -0.29297 | -0.07782 | -0.12817 | -0.050355 | -0.064087 | 0 | 0.064087 | -0.041199 | 0.009156 | 0.13275 | -0.14648 | 0.18311 | -0.09613
50.7812 | 122.525 | -59.2346 | 63.0753 | -0.10765 | 0.0015552 | 0.068118 | 0.070201 | -0.0010959 | 0.0020892 | 0.0019654 | -0.047089 | 0.0016685 | 0.065988 | 0.067037 | -0.00060892 | 0.0010105 | 0.0015103 | -0.11444 | 63.0798 | -59.3674 | 62.9425 | -59.1385 | -59.5642 | -59.4772 | -0.11902 | 63.0753 | -59.2804 | 63.0981 | -59.198 | -59.4269 | -59.1248 | -0.13275 | 0.004578 | -0.13275 | -0.13733 | 0.09613 | 0.22888 | -0.24262 | -0.32959 | -0.10986 | -0.19684 | -0.045777 | 0 | 0.022889 | 0.022889 | 0.036621 | 0.082398 | 0.10986 | -0.19226 | 0.15564 | -0.14648
51.7812 | 122.4838 | -59.2255 | 63.1348 | -0.15343 | -0.0014356 | 0.070682 | 0.069773 | -0.00024129 | -4.7032e-05 | 0.00025641 | -0.044037 | 0.0010275 | 0.064493 | 0.068746 | -0.0014633 | -0.00027134 | 0.0012967 | -0.009155 | 63.1119 | -59.3216 | 62.9974 | -59.0836 | -59.4772 | -59.4498 | 0.009155 | 63.1073 | -59.2575 | 63.1165 | -59.2209 | -59.3582 | -59.0744 | -0.09613 | -0.022889 | -0.13733 | -0.11444 | 0.14191 | 0.23804 | -0.2243 | -0.25177 | -0.12817 | -0.15564 | -0.032043 | -0.027466 | -0.018311 | 0.009155 | 0.004578 | 0.036621 | 0.15106 | -0.13275 | 0.1831 | -0.10071
52.7812 | 122.525 | -59.2072 | 63.1256 | -0.069504 | -0.0016492 | 0.067904 | 0.067637 | 0.0016813 | 0.0014484 | 4.2808e-05 | -0.013519 | 0.0012411 | 0.065561 | 0.065542 | -0.0014633 | -0.00069854 | 0.0059963 | 0 | 63.1302 | -59.3399 | 63.0203 | -59.0744 | -59.4635 | -59.4223 | 0.050354 | 63.0798 | -59.1888 | 63.1348 | -59.2026 | -59.3216 | -59.0515 | -0.13275 | 0.004578 | -0.10529 | -0.10986 | 0.13275 | 0.2655 | -0.21515 | -0.25635 | -0.082397 | -0.1236 | 0.01831 | -0.045776 | 0.009156 | 0.054932 | 0.004577 | -0.013733 | 0.15564 | -0.11444 | 0.13733 | -0.13275
53.7812 | 122.5571 | -59.2255 | 63.089 | -0.10765 | -0.00058102 | 0.067264 | 0.066142 | -0.00024129 | 0.00080757 | 0.0013244 | -0.037933 | -4.0508e-05 | 0.064279 | 0.065328 | 0.0013139 | -0.00027134 | -0.0034031 | -0.10986 | 63.0844 | -59.3536 | 62.9608 | -59.1248 | -59.4543 | -59.4498 | -0.091553 | 63.0386 | -59.2484 | 63.1119 | -59.2255 | -59.3491 | -59.0744 | -0.12817 | -0.004577 | -0.12817 | -0.1236 | 0.10071 | 0.22888 | -0.2243 | -0.22888 | -0.09613 | -0.10071 | -0.022888 | -0.050354 | 0.022888 | 0.073242 | 0 | 0.022888 | 0.15106 | -0.1236 | 0.17395 | -0.10071
54.7812 | 122.4655 | -59.198 | 63.1668 | -0.023727 | 5.9776e-05 | 0.068545 | 0.068919 | 0.00039951 | 0.0012348 | -0.00017099 | 0.0047913 | 0.0010275 | 0.065347 | 0.064473 | -0.00039512 | -5.7742e-05 | 1.4916e-05 | -0.022888 | 63.1439 | -59.317 | 62.97 | -59.111 | -59.4818 | -59.436 | -0.009155 | 63.1073 | -59.2346 | 63.0936 | -59.2163 | -59.3445 | -59.0698 | -0.11902 | -0.022888 | -0.19684 | -0.17395 | 0.086975 | 0.20599 | -0.23804 | -0.28381 | -0.11902 | -0.16479 | -0.036621 | -0.059509 | -0.073242 | -0.013733 | -0.018311 | 0.01831 | 0.12817 | -0.14648 | 0.1648 | -0.10986
55.7812 | 122.4976 | -59.2255 | 63.0936 | -0.054245 | -0.00058102 | 0.067264 | 0.068492 | 0.00082671 | 0.00016657 | 0.00089721 | -0.047089 | 0.00060029 | 0.066843 | 0.06426 | 0.00088648 | -0.0011257 | 0.0023647 | -0.10529 | 63.0753 | -59.3857 | 62.9471 | -59.1568 | -59.5596 | -59.4955 | -0.10986 | 63.0753 | -59.2712 | 63.0936 | -59.2438 | -59.3857 | -59.1431 | -0.16022 | -0.018311 | -0.14648 | -0.12817 | 0.068665 | 0.22888 | -0.27008 | -0.33417 | -0.10986 | -0.17395 | -0.045776 | -0.018311 | 0 | 0.018311 | -0.01831 | 0.027466 | 0.082398 | -0.16022 | 0.12817 | -0.11444
56.7812 | 122.5067 | -59.2117 | 63.1531 | -0.061874 | 0.0019824 | 0.06705 | 0.069773 | 0.0038175 | -0.00047423 | 0.0023926 | -0.013519 | 0.0012411 | 0.070688 | 0.060628 | 0.00045928 | -0.0021939 | -0.0025485 | -0.027466 | 63.1439 | -59.3445 | 62.9745 | -59.1248 | -59.4681 | -59.4315 | -0.045776 | 63.089 | -59.2712 | 63.0753 | -59.079 | -59.3262 | -59.1019 | -0.13275 | -0.009155 | -0.17853 | -0.16937 | 0.086975 | 0.21973 | -0.21973 | -0.25635 | -0.086976 | -0.1236 | -0.059509 | -0.064087 | -0.07782 | -0.013733 | 0.13275 | 0.19226 | 0.10986 | -0.11444 | 0.16937 | -0.054932
57.7812 | 122.5296 | -59.1934 | 63.1256 | -0.038986 | 5.9776e-05 | 0.0694 | 0.068492 | 0.0012539 | -4.7032e-05 | 0.00047001 | -0.034882 | 0.00038669 | 0.065561 | 0.061269 | -0.00018152 | 0.00036966 | 1.4916e-05 | 0.004578 | 63.1348 | -59.3399 | 63.0157 | -59.0652 | -59.4681 | -59.4269 | 0.045776 | 63.0707 | -59.2438 | 63.1348 | -59.1934 | -59.317 | -59.0652 | -0.14648 | 0.009156 | -0.10986 | -0.11902 | 0.12817 | 0.27466 | -0.23346 | -0.27466 | -0.086975 | -0.12817 | -0.050354 | -0.054931 | 0.009156 | 0.064087 | 0 | 0.050354 | 0.12817 | -0.1236 | 0.17853 | -0.073243
58.7812 | 122.5525 | -59.2026 | 63.1485 | 0.060196 | 5.9776e-05 | 0.068973 | 0.068705 | 0.0021085 | -0.00026063 | -0.0014526 | -0.019623 | 0.0029501 | 0.063425 | 0.066182 | -0.00018152 | -0.0019803 | -0.0029757 | -0.018311 | 63.1577 | -59.3399 | 63.0157 | -59.0652 | -59.4406 | -59.404 | 0.064087 | 63.1027 | -59.2209 | 63.1256 | -59.198 | -59.3491 | -59.0561 | -0.13733 | 0.009155 | -0.13275 | -0.14191 | 0.13733 | 0.27466 | -0.20142 | -0.23804 | -0.064087 | -0.10071 | -0.01831 | -0.045777 | -0.022889 | 0.022888 | 0.004578 | 0.022888 | 0.14648 | -0.14648 | 0.16479 | -0.12817
59.7812 | 122.5388 | -59.198 | 63.1393 | -0.092392 | 0.00091438 | 0.066836 | 0.06956 | -0.00024129 | 0.0014484 | 4.2808e-05 | -0.040985 | 0.00017309 | 0.065347 | 0.067464 | -0.0025315 | -0.00069854 | -0.0025485 | -0.013733 | 63.1027 | -59.3628 | 63.0157 | -59.0652 | -59.4681 | -59.4269 | 0.032043 | 63.0753 | -59.23 | 63.121 | -59.2163 | -59.3399 | -59.0561 | -0.1648 | -0.036621 | -0.1236 | -0.086975 | 0.13275 | 0.29755 | -0.22888 | -0.27008 | -0.064087 | -0.10529 | -0.032044 | -0.064087 | -0.01831 | 0.045777 | -0.018311 | 0.013733 | 0.14191 | -0.14191 | 0.17395 | -0.10986
60.7812 | 122.5342 | -59.2026 | 63.1393 | -0.046616 | 0.0026234 | 0.069613 | 0.068919 | 0.0029629 | 0.00059397 | -0.0014526 | -0.019623 | 0.00017309 | 0.066843 | 0.063619 | 0.00024568 | -0.00048494 | 0.0015103 | -0.2655 | 63.0798 | -59.3674 | 63.0524 | -59.1156 | -59.4589 | -59.4223 | 0.045776 | 63.1256 | -59.2209 | 63.1256 | -59.1888 | -59.317 | -59.0378 | -0.16479 | -0.059509 | -0.086975 | -0.027466 | 0.086975 | 0.25177 | -0.21973 | -0.25635 | -0.054931 | -0.091552 | -0.01831 | -0.013733 | -0.013733 | 0 | 0.013733 | 0.032043 | 0.16479 | -0.11444 | 0.1831 | -0.096131
61.7812 | 122.5296 | -59.2072 | 63.1531 | -0.0084687 | -0.00015382 | 0.06705 | 0.068278 | -0.0021641 | -0.0013288 | 0.00089721 | 0.010895 | 0.0016685 | 0.065988 | 0.070241 | 0.0015275 | -0.0019803 | 0.0032193 | -0.013733 | 63.1485 | -59.3308 | 62.9654 | -59.1248 | -59.4864 | -59.436 | 0 | 63.1165 | -59.2484 | 63.0707 | -59.2209 | -59.3445 | -59.079 | -0.1236 | -0.004577 | -0.18768 | -0.18311 | 0.082397 | 0.20599 | -0.22888 | -0.27924 | -0.10528 | -0.15564 | -0.041199 | -0.036621 | -0.082397 | -0.045776 | -0.013733 | 0.027466 | 0.12817 | -0.13733 | 0.16937 | -0.09613
62.7812 | 122.5067 | -59.1797 | 63.1851 | -0.038986 | -0.00015382 | 0.065768 | 0.069773 | 0.00039951 | -0.00026063 | -0.0023072 | 0.044464 | 0.0029501 | 0.066629 | 0.056997 | 3.2078e-05 | -0.0015529 | -0.0040439 | -0.022888 | 63.1439 | -59.3399 | 62.9745 | -59.1064 | -59.4772 | -59.436 | -0.013733 | 63.1302 | -59.2346 | 63.089 | -59.2117 | -59.3491 | -59.0927 | -0.16022 | -0.041199 | -0.21057 | -0.16937 | 0.073242 | 0.23346 | -0.25635 | -0.29755 | -0.09613 | -0.13733 | -0.054932 | -0.054932 | -0.096131 | -0.041199 | -0.032044 | 0.022888 | 0.086975 | -0.16937 | 0.14191 | -0.11444
63.7812 | 122.4701 | -59.2209 | 63.1073 | -0.084763 | 0.00048698 | 0.067691 | 0.069987 | 0.0012539 | -0.00026063 | -0.0016662 | -0.077606 | 0.0023093 | 0.065988 | 0.066396 | -0.00039512 | -0.0015529 | -0.00041228 | -0.082397 | 63.089 | -59.3719 | 62.9425 | -59.1888 | -59.5505 | -59.5139 | -0.14191 | 63.0844 | -59.2484 | 63.0569 | -59.2804 | -59.4223 | -59.1202 | -0.15106 | -0.018311 | -0.1648 | -0.14648 | 0.032043 | 0.18311 | -0.29297 | -0.32959 | -0.14191 | -0.17853 | -0.027466 | -0.022888 | -0.050354 | -0.027466 | -0.05951 | -0.032044 | 0.10071 | -0.20142 | 0.12817 | -0.17395
64.7812 | 122.5433 | -59.1888 | 63.1439 | -0.031357 | 5.9776e-05 | 0.065768 | 0.068064 | -0.00024129 | 0.0018756 | -0.0014526 | 0.0047913 | 0.00060029 | 0.06492 | 0.062124 | -0.00018152 | 0.00015586 | -0.0027621 | -0.004578 | 63.1439 | -59.3399 | 62.9745 | -59.1248 | -59.5047 | -59.4223 | -0.004578 | 63.121 | -59.2346 | 63.1165 | -59.1888 | -59.3353 | -59.0652 | -0.15106 | 0 | -0.16937 | -0.16937 | 0.064087 | 0.21515 | -0.23346 | -0.31586 | -0.082397 | -0.16479 | -0.045776 | -0.022888 | -0.027466 | -0.004578 | 0 | 0.045776 | 0.1236 | -0.14648 | 0.16937 | -0.10071
65.7812 | 122.5113 | -59.2026 | 63.1393 | -0.031357 | 0.00091438 | 0.066623 | 0.068492 | -0.00088229 | 0.0014484 | 0.001538 | -0.044037 | 0.0023093 | 0.064493 | 0.060628 | 0.00045928 | -0.0024075 | 0.0038601 | 0.004578 | 63.1027 | -59.3399 | 63.0249 | -59.0515 | -59.4635 | -59.4269 | 0.036621 | 63.0981 | -59.2438 | 63.1165 | -59.2209 | -59.3491 | -59.0607 | -0.13733 | -0.036621 | -0.11444 | -0.07782 | 0.15106 | 0.28839 | -0.2243 | -0.26093 | -0.086975 | -0.1236 | -0.041198 | -0.041198 | -0.022888 | 0.01831 | -0.01831 | 0.022888 | 0.14191 | -0.14648 | 0.18311 | -0.10529
66.7812 | 122.5525 | -59.1934 | 63.1302 | -0.046616 | 0.00027338 | 0.065768 | 0.068064 | -0.0025913 | -0.00026063 | -0.0020936 | -0.0074159 | 0.00038669 | 0.066415 | 0.065328 | -0.0010361 | -0.00091214 | -0.00062588 | -0.018311 | 63.1439 | -59.3399 | 63.0112 | -59.0836 | -59.4772 | -59.4086 | 0.036621 | 63.1027 | -59.2255 | 63.1348 | -59.1934 | -59.3033 | -59.0424 | -0.14648 | 0.013733 | -0.11902 | -0.13275 | 0.10986 | 0.25635 | -0.21515 | -0.28381 | -0.068664 | -0.13733 | -0.032044 | -0.027466 | 0.004578 | 0.032044 | 0 | 0.032044 | 0.15106 | -0.10986 | 0.18311 | -0.07782
67.7812 | 122.5067 | -59.198 | 63.1622 | -0.054245 | 0.0015552 | 0.066836 | 0.06721 | 0.00061311 | 0.0012348 | 0.00089721 | -0.0074159 | 0.0012411 | 0.062784 | 0.059347 | 0.00045928 | -0.0026211 | -0.0012667 | -0.009155 | 63.1439 | -59.3216 | 63.002 | -59.1064 | -59.4543 | -59.4223 | 0.004578 | 63.0936 | -59.2255 | 63.1165 | -59.1888 | -59.317 | -59.0515 | -0.1236 | -0.01831 | -0.16022 | -0.14191 | 0.091553 | 0.21515 | -0.2243 | -0.25635 | -0.10071 | -0.13275 | -0.027466 | -0.068664 | -0.045776 | 0.022888 | 0.009155 | 0.036621 | 0.14648 | -0.11902 | 0.17395 | -0.091553
68.7812 | 122.5067 | -59.2712 | 63.1119 | -0.13054 | 0.00091438 | 0.065982 | 0.069133 | 0.0016813 | 0.00080757 | -0.00017099 | -0.068451 | 0.0016685 | 0.064706 | 0.067251 | -0.00039512 | -0.0039029 | 0.0036465 | -0.10071 | 63.1027 | -59.3857 | 62.9379 | -59.1888 | -59.5825 | -59.5001 | -0.14191 | 63.0798 | -59.2392 | 63.0661 | -59.198 | -59.761 | -59.3765 | -0.11444 | -0.009155 | -0.17395 | -0.16479 | 0.082397 | 0.19684 | -0.22888 | -0.31128 | -0.11444 | -0.19684 | 0.032043 | -0.032043 | -0.045776 | -0.013733 | 0.073242 | 0.041199 | -0.10529 | -0.48981 | -0.13733 | -0.52185
69.7812 | 122.5204 | -59.2026 | 63.1393 | -0.046616 | 0.00027338 | 0.064273 | 0.066783 | 0.00082671 | 0.00038017 | -0.00017099 | 0.035309 | -0.00025411 | 0.065134 | 0.05721 | 0.00024568 | -0.0043301 | 0.0021511 | -0.041199 | 63.1531 | -59.3308 | 62.9745 | -59.0973 | -59.4681 | -59.4223 | -0.009155 | 63.1165 | -59.2346 | 63.1851 | -59.3536 | -59.5093 | -59.1476 | -0.12817 | 0.013733 | -0.16479 | -0.17853 | 0.10529 | 0.23346 | -0.21973 | -0.2655 | -0.091552 | -0.13733 | -0.032043 | -0.022888 | 0.045777 | 0.068665 | -0.15106 | -0.11902 | 0.054932 | -0.3067 | 0.086975 | -0.27466
70.7812 | 122.5616 | -59.1934 | 63.1485 | -0.077133 | 0.00091438 | 0.070254 | 0.069346 | 0.00061311 | 0.0014484 | 0.0023926 | -0.010467 | -4.0508e-05 | 0.066415 | 0.066396 | -0.00039512 | 0.00036966 | -0.0038303 | -0.004578 | 63.1027 | -59.3445 | 63.0112 | -59.0744 | -59.4498 | -59.4131 | 0.013733 | 63.1073 | -59.2346 | 63.1165 | -59.1843 | -59.317 | -59.0515 | -0.15106 | -0.045777 | -0.13733 | -0.091553 | 0.11902 | 0.27008 | -0.21973 | -0.25635 | -0.068665 | -0.10529 | -0.041199 | -0.041199 | -0.032044 | 0.009155 | 0.009155 | 0.050354 | 0.14191 | -0.1236 | 0.18311 | -0.082398
71.7812 | 122.5342 | -59.1705 | 63.1531 | -0.077133 | 0.001128 | 0.06705 | 0.069133 | 0.0014675 | -0.00068783 | 0.00047001 | 0.013947 | 0.0012411 | 0.06727 | 0.064901 | -0.00082252 | -0.00091214 | -0.0010531 | -0.009155 | 63.1348 | -59.3536 | 62.9791 | -59.1156 | -59.4635 | -59.4177 | 0.013733 | 63.1302 | -59.2346 | 63.1119 | -59.1843 | -59.3308 | -59.0607 | -0.18311 | -0.01831 | -0.17395 | -0.15564 | 0.054931 | 0.23804 | -0.24719 | -0.29297 | -0.064087 | -0.10986 | -0.064087 | -0.022888 | -0.041199 | -0.018311 | -0.013733 | 0.050354 | 0.10986 | -0.16022 | 0.17395 | -0.096131
72.7812 | 122.5342 | -59.1843 | 63.1439 | -0.15343 | -0.00229 | 0.067264 | 0.06721 | 0.0023221 | 0.0010212 | 0.0017518 | -0.0013123 | 0.00060029 | 0.064706 | 0.05956 | -0.0010361 | -0.0021939 | -0.0010531 | -0.004578 | 63.121 | -59.3353 | 62.9929 | -59.111 | -59.5001 | -59.4406 | -0.004578 | 63.0936 | -59.2438 | 63.1119 | -59.1888 | -59.3124 | -59.0607 | -0.15106 | -0.022888 | -0.15106 | -0.12817 | 0.073242 | 0.2243 | -0.25635 | -0.31586 | -0.10529 | -0.16479 | -0.059509 | -0.050354 | -0.032044 | 0.01831 | -0.004578 | 0.054931 | 0.1236 | -0.12817 | 0.18311 | -0.068665
73.7812 | 122.5479 | -59.2026 | 63.1302 | -0.077133 | 0.001128 | 0.07175 | 0.06721 | -0.0021641 | -0.00090143 | -0.0025208 | -0.053192 | -4.0508e-05 | 0.064279 | 0.05956 | -0.0010361 | -5.7742e-05 | -0.00019868 | -0.009155 | 63.089 | -59.3628 | 63.0157 | -59.0698 | -59.4772 | -59.4223 | 0.036621 | 63.0981 | -59.2255 | 63.1302 | -59.2117 | -59.3216 | -59.0607 | -0.16022 | -0.041199 | -0.11444 | -0.073242 | 0.13275 | 0.29297 | -0.21973 | -0.27466 | -0.059509 | -0.11444 | -0.022888 | -0.032043 | 0 | 0.032043 | -0.009155 | 0.013733 | 0.14191 | -0.11902 | 0.1648 | -0.09613
74.7812 | 122.5067 | -59.2117 | 63.1531 | -0.038986 | 5.9776e-05 | 0.067264 | 0.070414 | -0.0040867 | -0.0015424 | 0.00068361 | 0.016998 | 0.0010275 | 0.066843 | 0.06191 | -0.00018152 | -0.0024075 | -0.00019868 | -0.022888 | 63.1531 | -59.3262 | 62.97 | -59.111 | -59.4955 | -59.4223 | -0.013733 | 63.1027 | -59.23 | 63.1531 | -59.0698 | -59.4864 | -59.2026 | -0.11444 | 0 | -0.1831 | -0.1831 | 0.10071 | 0.21515 | -0.21057 | -0.28381 | -0.09613 | -0.16937 | -0.018311 | -0.050354 | 0 | 0.050354 | 0.14191 | 0.16022 | 0.009155 | -0.27466 | 0.027466 | -0.25635
75.7812 | 122.5159 | -59.1888 | 63.176 | -0.061874 | 0.00048698 | 0.065127 | 0.066142 | 0.00082671 | -0.0021832 | -0.00081179 | 0.065826 | 0.00038669 | 0.065134 | 0.062978 | 3.2078e-05 | -0.0017667 | -0.0038303 | 0.022888 | 63.1348 | -59.3262 | 63.0295 | -59.0607 | -59.4223 | -59.3948 | 0.050354 | 63.089 | -59.2346 | 63.0753 | -59.2346 | -59.3262 | -59.0881 | -0.13733 | -0.041198 | -0.14648 | -0.10529 | 0.12817 | 0.2655 | -0.20599 | -0.23346 | -0.068664 | -0.09613 | -0.045776 | -0.086975 | -0.10071 | -0.013733 | -0.045776 | 0 | 0.10071 | -0.13733 | 0.14648 | -0.091553
76.7812 | 122.5708 | -59.1705 | 63.121 | -0.077133 | 0.002196 | 0.06705 | 0.067637 | -0.0015231 | 0.0023028 | 4.2808e-05 | 0.047516 | 0.00060029 | 0.065134 | 0.05721 | -0.0012497 | -0.0019803 | -0.0012667 | 0.036621 | 63.0936 | -59.3124 | 63.0478 | -59.0149 | -59.4086 | -59.3536 | 0.068665 | 63.0981 | -59.2117 | 63.121 | -59.1568 | -59.2896 | -59.0561 | -0.14191 | -0.027466 | -0.073242 | -0.045776 | 0.15564 | 0.29755 | -0.18311 | -0.23804 | -0.041199 | -0.09613 | -0.041199 | -0.022888 | 0 | 0.022888 | 0.013733 | 0.054932 | 0.11444 | -0.11902 | 0.15564 | -0.07782
77.7812 | 122.4838 | -59.1888 | 63.1622 | -0.046616 | 0.00048698 | 0.064273 | 0.066783 | -0.0023777 | -4.7032e-05 | -0.0055114 | 0.053619 | -0.00089511 | 0.066202 | 0.064901 | -0.00082252 | -0.0019803 | 0.0027921 | 0.027466 | 63.1027 | -59.3262 | 62.9974 | -59.0744 | -59.4452 | -59.3903 | 0.027466 | 63.121 | -59.2438 | 63.0798 | -59.2072 | -59.3353 | -59.0927 | -0.13733 | -0.059509 | -0.16479 | -0.10529 | 0.11444 | 0.25177 | -0.20142 | -0.25635 | -0.064087 | -0.11902 | -0.054931 | -0.041198 | -0.082397 | -0.041199 | -0.01831 | 0.036621 | 0.096131 | -0.14648 | 0.15106 | -0.091553
78.7812 | 122.5021 | -59.1751 | 63.1851 | -0.092392 | -0.00036742 | 0.066623 | 0.067637 | 0.0023221 | 0.0010212 | -0.0010254 | 0.013947 | -4.0508e-05 | 0.067697 | 0.065328 | -0.0016769 | -0.0045437 | 0.0032193 | -0.032043 | 63.1302 | -59.3399 | 62.9837 | -59.111 | -59.4772 | -59.4315 | 0 | 63.1027 | -59.2438 | 63.0981 | -59.23 | -59.3491 | -59.0836 | -0.1648 | -0.054932 | -0.20142 | -0.14648 | 0.064087 | 0.22888 | -0.25635 | -0.30212 | -0.091553 | -0.13733 | -0.068664 | -0.082398 | -0.086975 | -0.004577 | -0.054932 | 0.013732 | 0.091553 | -0.17395 | 0.16022 | -0.10529
79.7812 | 122.4655 | -59.1843 | 63.1805 | -0.038986 | 0.00048698 | 0.068118 | 0.06721 | -0.0030185 | -0.00047423 | -0.0010254 | -0.0074159 | 0.0016685 | 0.066415 | 0.059347 | -0.0016769 | -0.0024075 | 0.0017239 | 0.009155 | 63.1256 | -59.3079 | 62.9745 | -59.1019 | -59.4772 | -59.4315 | 0.018311 | 63.0936 | -59.2346 | 63.0798 | -59.2163 | -59.3399 | -59.1064 | -0.1236 | -0.054932 | -0.20599 | -0.15106 | 0.082397 | 0.20599 | -0.24719 | -0.29297 | -0.1236 | -0.16937 | -0.050354 | -0.086975 | -0.10071 | -0.013733 | -0.032044 | 0.01831 | 0.07782 | -0.15564 | 0.12817 | -0.10529
80.7812 | 122.5067 | -59.166 | 63.1714 | -0.11528 | -0.00036742 | 0.067477 | 0.06721 | 0.00061311 | -0.0019696 | -0.00059819 | 0.032257 | 0.0018821 | 0.066202 | 0.059774 | -0.0018905 | -0.0043301 | 0.0030057 | 0.027466 | 63.1119 | -59.3308 | 63.0203 | -59.0469 | -59.4681 | -59.4269 | 0.054932 | 63.1027 | -59.2072 | 63.0798 | -59.2117 | -59.3308 | -59.1064 | -0.1648 | -0.05951 | -0.15106 | -0.091552 | 0.11902 | 0.28381 | -0.26093 | -0.30212 | -0.09613 | -0.13733 | -0.041198 | -0.068665 | -0.091553 | -0.022888 | -0.045776 | -0.004578 | 0.05951 | -0.1648 | 0.10071 | -0.1236
81.7812 | 122.5067 | -59.1705 | 63.1668 | -0.046616 | 5.9776e-05 | 0.065555 | 0.067851 | -0.00066869 | 0.0020892 | 0.0011108 | 0.0017395 | 0.0014547 | 0.067056 | 0.062978 | -0.0012497 | -0.00069854 | 0.0019375 | -0.027466 | 63.1256 | -59.3308 | 62.97 | -59.1019 | -59.4772 | -59.4315 | 0.004578 | 63.1027 | -59.2346 | 63.0798 | -59.2163 | -59.3445 | -59.0836 | -0.16022 | -0.041199 | -0.19684 | -0.15564 | 0.068664 | 0.22888 | -0.26093 | -0.3067 | -0.10071 | -0.14648 | -0.064087 | -0.064087 | -0.086975 | -0.022888 | -0.045777 | 0.01831 | 0.086975 | -0.17395 | 0.15106 | -0.10986
82.7812 | 122.6257 | -59.1614 | 63.1393 | -0.046616 | 0.00027338 | 0.066623 | 0.066996 | -0.00024129 | 0.0012348 | -0.001239 | 0.053619 | -0.0015359 | 0.064706 | 0.061269 | -0.00018152 | -0.0021939 | 0.0025785 | 0.036621 | 63.1027 | -59.3033 | 63.1027 | -58.96 | -59.3582 | -59.3216 | 0.15106 | 63.0524 | -59.2026 | 63.176 | -59.1064 | -59.2209 | -58.9828 | -0.14191 | -0.036621 | -0.036621 | 0 | 0.20142 | 0.34332 | -0.16022 | -0.19684 | -0.01831 | -0.054931 | -0.041199 | -0.086975 | 0.036621 | 0.1236 | 0.054932 | 0.096131 | 0.17853 | -0.059509 | 0.21973 | -0.01831
83.7812 | 122.58 | -59.2117 | 63.1485 | 0.060196 | 0.00027338 | 0.066623 | 0.069133 | -0.00045509 | -0.00047423 | -0.0023072 | -0.025726 | 0.0018821 | 0.062997 | 0.067464 | 0.00067288 | -0.00048494 | -0.0029757 | -0.068665 | 63.121 | -59.3491 | 62.9929 | -59.1248 | -59.4864 | -59.4589 | 0.004578 | 63.1165 | -59.2255 | 63.121 | -59.1934 | -59.3262 | -59.0561 | -0.13733 | -0.027466 | -0.15564 | -0.12817 | 0.086975 | 0.2243 | -0.24719 | -0.27466 | -0.10986 | -0.13733 | -0.013733 | -0.032044 | -0.027466 | 0.004578 | 0.018311 | 0.032044 | 0.15564 | -0.11444 | 0.16937 | -0.10071
84.7812 | 122.5937 | -59.166 | 63.1256 | -0.069504 | 0.0019824 | 0.068332 | 0.066142 | -0.00088229 | 0.00059397 | -0.0020936 | 0.02005 | -0.0013223 | 0.064066 | 0.072378 | -0.00018152 | -0.0034757 | -0.0042575 | 0.036621 | 63.0753 | -59.2987 | 63.089 | -58.992 | -59.3765 | -59.3308 | 0.11444 | 63.0753 | -59.1843 | 63.1851 | -59.1705 | -59.23 | -58.9737 | -0.13275 | -0.050354 | -0.036621 | 0.013733 | 0.17395 | 0.3067 | -0.1648 | -0.21057 | -0.032044 | -0.07782 | -0.01831 | -0.050354 | 0.05951 | 0.10986 | -0.004577 | 0.013733 | 0.19226 | -0.064087 | 0.21057 | -0.045777
85.7812 | 122.5296 | -59.1934 | 63.1668 | -0.16869 | 0.002196 | 0.066836 | 0.065501 | 0.0016813 | 0.0020892 | 0.00025641 | 0.010895 | -0.00046791 | 0.064279 | 0.065328 | 0.00024568 | -0.0032619 | 0.0045011 | 0.013733 | 63.0844 | -59.3308 | 63.0707 | -59.0149 | -59.3948 | -59.3491 | 0.07782 | 63.089 | -59.2072 | 63.1393 | -59.1568 | -59.2804 | -59.0424 | -0.13733 | -0.082397 | -0.09613 | -0.013733 | 0.17853 | 0.31586 | -0.15564 | -0.20142 | -0.01831 | -0.064086 | -0.013733 | -0.07782 | -0.027466 | 0.050354 | 0.036621 | 0.050354 | 0.15106 | -0.086976 | 0.16479 | -0.073243
86.7812 | 122.5754 | -59.1751 | 63.1439 | -0.084763 | 0.00027338 | 0.065982 | 0.066996 | -2.7692e-05 | 0.00016657 | -0.0020936 | 0.016998 | -4.0508e-05 | 0.066415 | 0.062124 | -0.00082252 | -0.0019803 | -0.0012667 | 0.045776 | 63.0798 | -59.285 | 63.1119 | -59.0286 | -59.436 | -59.3445 | 0.10986 | 63.0753 | -59.2163 | 63.1439 | -59.1339 | -59.2438 | -59.0195 | -0.10986 | -0.064087 | -0.032044 | 0.032043 | 0.14648 | 0.25635 | -0.16937 | -0.26093 | -0.059509 | -0.15106 | -0.041199 | -0.068665 | 0 | 0.068665 | 0.041199 | 0.082398 | 0.15564 | -0.068664 | 0.19684 | -0.027465
87.7812 | 122.5891 | -59.1705 | 63.1348 | -0.031357 | 0.001128 | 0.066836 | 0.066783 | -0.0036593 | 0.0023028 | 0.0011108 | 0.17264 | -0.0011087 | 0.067056 | 0.059774 | -0.0014633 | -0.0011257 | 0.0021511 | 0.032043 | 63.0981 | -59.3079 | 63.1027 | -58.9874 | -59.3811 | -59.3216 | 0.036621 | 63.089 | -59.2163 | 63.1668 | -59.1248 | -59.23 | -59.0103 | -0.13733 | -0.036621 | -0.032044 | 0.004577 | 0.18311 | 0.32043 | -0.15106 | -0.21057 | -0.013733 | -0.073243 | -0.045777 | -0.045777 | 0.032043 | 0.07782 | 0.045776 | 0.091553 | 0.16022 | -0.05951 | 0.20599 | -0.013733
88.7812 | 122.493 | -59.2026 | 63.1531 | -0.031357 | -0.0014356 | 0.067691 | 0.068705 | 0.0012539 | 0.00038017 | -0.00081179 | -0.053192 | 0.00060029 | 0.065561 | 0.061269 | -0.0010361 | -5.7742e-05 | -0.00062588 | -0.009155 | 63.1302 | -59.3491 | 62.9929 | -59.111 | -59.5001 | -59.436 | 0.16479 | 63.1577 | -59.2117 | 63.1027 | -59.2209 | -59.3445 | -59.0881 | -0.14648 | -0.022888 | -0.16022 | -0.13733 | 0.091553 | 0.23804 | -0.23346 | -0.29755 | -0.086975 | -0.15106 | -0.009155 | 0.004578 | -0.050354 | -0.054932 | -0.01831 | -0.009155 | 0.11444 | -0.14191 | 0.1236 | -0.13275
89.7812 | 122.612 | -59.1339 | 63.1256 | -0.061874 | 0.0017688 | 0.067691 | 0.066996 | -0.00024129 | 0.00038017 | -0.0025208 | 0.023102 | -0.00046791 | 0.062143 | 0.059774 | 0.00045928 | -0.0026211 | 0.0034329 | 0.032043 | 63.0615 | -59.3079 | 63.121 | -58.96 | -59.3628 | -59.3216 | 0.13275 | 63.0569 | -59.2026 | 63.1897 | -59.079 | -59.2117 | -58.9828 | -0.17395 | -0.064087 | -0.004577 | 0.05951 | 0.17395 | 0.3479 | -0.18768 | -0.22888 | -0.013733 | -0.054932 | -0.068665 | -0.068664 | 0.064087 | 0.13275 | 0.054932 | 0.1236 | 0.15106 | -0.07782 | 0.21973 | -0.009155
90.7812 | 122.6532 | -59.1705 | 63.0936 | -0.023727 | 5.9776e-05 | 0.065127 | 0.065287 | -0.00088229 | 0.0020892 | -0.0010254 | 0.02005 | -0.00089511 | 0.063425 | 0.060628 | 3.2078e-05 | -0.0021939 | -0.00041228 | -0.018311 | 63.0798 | -59.3079 | 63.0478 | -59.0103 | -59.4315 | -59.3491 | 0.086975 | 63.0798 | -59.23 | 63.176 | -59.1248 | -59.2117 | -58.9783 | -0.13733 | -0.013733 | -0.045776 | -0.032043 | 0.16022 | 0.29755 | -0.17853 | -0.26093 | -0.041199 | -0.1236 | -0.05951 | -0.013733 | 0.082397 | 0.09613 | 0.045776 | 0.10529 | 0.19226 | -0.041199 | 0.25177 | 0.018311
91.7812 | 122.5891 | -59.166 | 63.121 | -0.084763 | 5.9776e-05 | 0.06705 | 0.069346 | 0.00018591 | 0.00080757 | -0.00059819 | 0.026154 | -0.0013223 | 0.064279 | 0.065542 | -0.0040267 | 0.00079686 | 0.0012967 | 0.050354 | 63.0753 | -59.2804 | 63.1119 | -58.9783 | -59.3628 | -59.3216 | 0.13733 | 63.0798 | -59.198 | 63.0936 | -59.1522 | -59.2209 | -59.0103 | -0.11444 | -0.045777 | -0.009156 | 0.036621 | 0.18768 | 0.30213 | -0.15564 | -0.19684 | -0.041198 | -0.082397 | -0.032043 | -0.041199 | -0.027466 | 0.013733 | 0.013733 | 0.045776 | 0.15564 | -0.054931 | 0.18768 | -0.022888
92.7812 | 122.5616 | -59.166 | 63.1348 | -0.023727 | -0.00036742 | 0.067264 | 0.068492 | -2.7692e-05 | 0.0018756 | -0.00059819 | -0.0043641 | -0.0015359 | 0.065347 | 0.05956 | -0.00018152 | -0.0030483 | -0.0010531 | 0.045776 | 63.0569 | -59.285 | 63.1073 | -58.9417 | -59.3536 | -59.3216 | 0.16022 | 63.0661 | -59.2163 | 63.2034 | -59.0836 | -59.198 | -58.9828 | -0.11902 | -0.07782 | -0.027466 | 0.050354 | 0.2243 | 0.34332 | -0.15564 | -0.18768 | -0.036621 | -0.068665 | -0.050354 | -0.068665 | 0.068664 | 0.13733 | 0.082398 | 0.13275 | 0.18311 | -0.032043 | 0.23346 | 0.018311
93.7812 | 122.5571 | -59.1888 | 63.2263 | -0.038986 | -0.00058102 | 0.066409 | 0.070842 | 0.0010403 | -0.0023968 | -0.0016662 | -0.05014 | 0.0010275 | 0.066629 | 0.067678 | -0.0014633 | -0.0019803 | -0.00083948 | -0.009155 | 63.1531 | -59.3445 | 62.9837 | -59.1019 | -59.491 | -59.4223 | 0.018311 | 63.1348 | -59.2438 | 63.0981 | -59.2346 | -59.3582 | -59.1614 | -0.15564 | -0.073242 | -0.24261 | -0.16937 | 0.086975 | 0.24261 | -0.23346 | -0.30212 | -0.07782 | -0.14648 | -0.054931 | -0.091552 | -0.12817 | -0.036621 | -0.045776 | 0.009155 | 0.027466 | -0.16937 | 0.082397 | -0.11444
94.7812 | 122.6028 | -59.1568 | 63.1073 | -0.061874 | 0.00027338 | 0.065555 | 0.068278 | -0.0015231 | -0.001115 | 0.00089721 | 0.023102 | 0.00060029 | 0.065347 | 0.065542 | -0.00018152 | -0.0026211 | -0.0019077 | 0.032043 | 63.0753 | -59.3033 | 63.1119 | -58.9417 | -59.3811 | -59.3033 | 0.16937 | 63.0615 | -59.2026 | 63.2172 | -59.0881 | -59.198 | -58.9737 | -0.14648 | -0.032044 | 0.004577 | 0.036621 | 0.21515 | 0.36163 | -0.14648 | -0.22431 | 0 | -0.07782 | -0.045777 | -0.045777 | 0.10986 | 0.15564 | 0.068664 | 0.11444 | 0.1831 | -0.041199 | 0.22888 | 0.004578
95.7812 | 122.6303 | -59.1431 | 63.1302 | -0.084763 | -0.00079462 | 0.064914 | 0.065928 | -0.0015231 | 0.00080757 | 0.0013244 | 0.023102 | -4.0508e-05 | 0.064279 | 0.062764 | -0.0016769 | -0.0013393 | -0.0034031 | 0.013733 | 63.0569 | -59.3262 | 63.1027 | -58.9508 | -59.3674 | -59.3124 | 0.13733 | 63.0478 | -59.2117 | 63.2309 | -59.0836 | -59.1843 | -58.9508 | -0.18311 | -0.073242 | -0.027466 | 0.045776 | 0.19226 | 0.37537 | -0.16937 | -0.22431 | 0.013733 | -0.041199 | -0.068665 | -0.082397 | 0.10071 | 0.18311 | 0.059509 | 0.12817 | 0.19226 | -0.041199 | 0.26093 | 0.027466
96.7812 | 122.5708 | -59.1797 | 63.121 | -0.13817 | 5.9776e-05 | 0.06705 | 0.066996 | -0.0013095 | 0.0010212 | -0.00038459 | 0.0047913 | -0.00046791 | 0.064066 | 0.059774 | -0.00082252 | -0.0019803 | 0.0042875 | -0.027466 | 63.0753 | -59.3445 | 63.0295 | -59.0469 | -59.436 | -59.3628 | 0.013733 | 63.0753 | -59.23 | 63.1439 | -59.1431 | -59.2758 | -59.0012 | -0.16479 | -0.045777 | -0.091553 | -0.045776 | 0.13275 | 0.29755 | -0.18311 | -0.25635 | -0.018311 | -0.091553 | -0.050355 | -0.045777 | 0.022888 | 0.068665 | 0.036621 | 0.086976 | 0.17853 | -0.096131 | 0.22888 | -0.045776
97.7812 | 122.5937 | -59.1843 | 63.1073 | -0.13054 | 0.00091438 | 0.067264 | 0.067851 | 0.0021085 | 0.0025166 | -0.00081179 | 0.010895 | -0.00025411 | 0.063425 | 0.067678 | -0.0014633 | -0.0021939 | 0.00065592 | -0.027466 | 63.0661 | -59.3399 | 63.0844 | -59.0012 | -59.3903 | -59.3216 | 0.09613 | 63.0707 | -59.2484 | 63.208 | -59.1019 | -59.2117 | -58.9783 | -0.15564 | -0.041199 | -0.022888 | 0.018311 | 0.18311 | 0.33875 | -0.13733 | -0.20599 | 0.018311 | -0.050354 | -0.064087 | -0.036621 | 0.10071 | 0.13733 | 0.082397 | 0.14648 | 0.20599 | -0.027466 | 0.27008 | 0.036621
98.7812 | 122.4747 | -59.2072 | 63.1989 | -0.061874 | -0.0016492 | 0.065555 | 0.066142 | -0.0023777 | 0.0012348 | 0.003888 | -0.068451 | -0.00089511 | 0.061929 | 0.06661 | -0.00082252 | -0.00091214 | -0.0025485 | -0.045776 | 63.1485 | -59.3216 | 62.9929 | -59.0927 | -59.4681 | -59.4406 | -0.004578 | 63.1302 | -59.2346 | 63.0936 | -59.2163 | -59.3353 | -59.0836 | -0.11444 | -0.050354 | -0.20599 | -0.15564 | 0.11444 | 0.22888 | -0.23346 | -0.26093 | -0.11902 | -0.14649 | -0.027466 | -0.068665 | -0.10529 | -0.036621 | -0.009156 | 0.01831 | 0.1236 | -0.12817 | 0.15106 | -0.10071
99.7812 | 122.6532 | -59.1705 | 63.1119 | -0.069504 | 5.9776e-05 | 0.067691 | 0.067424 | 0.0016813 | 0.0020892 | 0.00089721 | 0.023102 | 0.00017309 | 0.063211 | 0.061483 | 0.0013139 | -0.0017667 | 0.0019375 | 0.018311 | 63.0798 | -59.285 | 63.1027 | -58.9508 | -59.3536 | -59.3079 | 0.14191 | 63.0615 | -59.2026 | 63.2401 | -59.0698 | -59.1843 | -58.9783 | -0.11444 | -0.032043 | -0.009155 | 0.022888 | 0.21973 | 0.33417 | -0.13733 | -0.18311 | -0.022888 | -0.068665 | -0.032044 | -0.050354 | 0.12817 | 0.17853 | 0.10071 | 0.13275 | 0.19226 | -0.013733 | 0.22431 | 0.018311
100.7812 | 122.6257 | -59.1385 | 63.0936 | -0.054245 | 0.0013416 | 0.066623 | 0.067637 | 0.00039951 | 0.0012348 | 0.0034606 | -0.013519 | 0.0014547 | 0.064066 | 0.058492 | -0.0012497 | -0.0028347 | -0.0012667 | 0.009155 | 63.0798 | -59.3124 | 63.0753 | -58.9874 | -59.3903 | -59.3262 | 0.14648 | 63.0661 | -59.2163 | 63.2126 | -59.079 | -59.1934 | -58.9554 | -0.17395 | -0.013733 | -0.018311 | -0.004578 | 0.15106 | 0.32501 | -0.18768 | -0.25177 | -0.013733 | -0.07782 | -0.07782 | -0.027466 | 0.11902 | 0.14648 | 0.05951 | 0.13733 | 0.18311 | -0.054931 | 0.26093 | 0.022889
101.7812 | 122.6212 | -59.1797 | 63.1027 | -0.10765 | 0.00091438 | 0.065127 | 0.066783 | -0.0017367 | -4.7032e-05 | -0.002948 | -0.016571 | -0.00046791 | 0.062997 | 0.064687 | -0.00018152 | -0.00091214 | 0.00022852 | -0.018311 | 63.0569 | -59.3445 | 63.0661 | -59.0012 | -59.3994 | -59.3353 | 0.10071 | 63.0615 | -59.2392 | 63.208 | -59.1019 | -59.2117 | -58.9828 | -0.16479 | -0.045776 | -0.036621 | 0.009155 | 0.17853 | 0.34332 | -0.15564 | -0.21973 | 0.009155 | -0.054932 | -0.05951 | -0.041199 | 0.10529 | 0.14648 | 0.077819 | 0.13733 | 0.19684 | -0.032044 | 0.25635 | 0.027466
102.7812 | 122.6028 | -59.1522 | 63.0981 | -0.092392 | -0.001222 | 0.067477 | 0.066569 | 0.00018591 | -0.00090143 | -0.0018798 | -0.0043641 | -0.00025411 | 0.065134 | 0.057851 | 3.2078e-05 | -0.0015529 | -0.0025485 | -0.022888 | 63.0844 | -59.3399 | 63.0707 | -59.024 | -59.4086 | -59.3628 | 0.064087 | 63.0615 | -59.23 | 63.2034 | -59.1019 | -59.2392 | -58.9966 | -0.18768 | -0.013733 | -0.027466 | -0.013733 | 0.12817 | 0.31586 | -0.21057 | -0.25635 | -0.022888 | -0.068664 | -0.07782 | -0.036622 | 0.10528 | 0.14191 | 0.050354 | 0.12817 | 0.15564 | -0.086975 | 0.23346 | -0.009155
103.7812 | 122.5159 | -59.198 | 63.1348 | -0.077133 | -0.00015382 | 0.06705 | 0.066996 | -0.0013095 | 0.00038017 | -0.001239 | -0.040985 | 0.00060029 | 0.062357 | 0.063192 | -0.0010361 | -0.0013393 | 0.0017239 | 0.036621 | 63.0707 | -59.3262 | 63.0707 | -59.0057 | -59.3948 | -59.3399 | 0.10071 | 63.1027 | -59.2209 | 63.1485 | -59.1476 | -59.2621 | -59.0195 | -0.12817 | -0.064087 | -0.064087 | 0 | 0.19226 | 0.32043 | -0.14191 | -0.19684 | -0.013733 | -0.068664 | -0.022888 | -0.032044 | 0.013733 | 0.045777 | 0.050354 | 0.073242 | 0.17853 | -0.064087 | 0.20142 | -0.041199
104.7812 | 122.5525 | -59.2072 | 63.1119 | -0.046616 | 0.00070078 | 0.067904 | 0.066996 | 0.00018591 | 0.0025166 | 0.00047001 | -0.040985 | -0.00089511 | 0.066843 | 0.065969 | -0.0016769 | -0.0032619 | -0.0019077 | -0.068665 | 63.0524 | -59.3353 | 63.0249 | -59.0607 | -59.4681 | -59.3948 | -0.022888 | 63.0478 | -59.2209 | 63.2126 | -59.2072 | -59.2896 | -59.0469 | -0.12817 | -0.059509 | -0.086975 | -0.027466 | 0.14648 | 0.27466 | -0.18768 | -0.26093 | -0.059509 | -0.13275 | -0.013733 | -0.064086 | 0.10071 | 0.16479 | 0 | 0.013733 | 0.16022 | -0.082398 | 0.17395 | -0.068665
105.7812 | 122.6349 | -59.1751 | 63.0981 | -0.10765 | -0.0016492 | 0.067691 | 0.06956 | -0.0021641 | 0.003371 | 0.00025641 | -0.0074159 | 0.00017309 | 0.06492 | 0.058492 | -0.0031723 | -0.00069854 | 0.0042875 | -0.009155 | 63.0524 | -59.3124 | 63.0615 | -58.9966 | -59.4086 | -59.3445 | 0.091553 | 63.0798 | -59.2163 | 63.1943 | -59.079 | -59.2072 | -58.992 | -0.13733 | -0.045777 | -0.036622 | 0.009155 | 0.17853 | 0.31586 | -0.16937 | -0.23346 | -0.032043 | -0.09613 | -0.041199 | -0.018311 | 0.09613 | 0.11444 | 0.096131 | 0.13733 | 0.18311 | -0.032043 | 0.22431 | 0.009156
106.7812 | 122.5754 | -59.166 | 63.0936 | -0.1458 | -0.00079462 | 0.066623 | 0.068492 | 0.0014675 | -4.7032e-05 | 0.00068361 | -0.022675 | 0.00060029 | 0.06492 | 0.061483 | 0.00088648 | -0.0019803 | 0.00086952 | -0.027466 | 63.0707 | -59.3308 | 63.0753 | -59.0195 | -59.4223 | -59.3399 | 0.064087 | 63.0798 | -59.2026 | 63.1851 | -59.1156 | -59.2438 | -58.992 | -0.1648 | -0.022888 | -0.018311 | 0.004577 | 0.14648 | 0.31128 | -0.17395 | -0.25635 | -0.009155 | -0.091552 | -0.036621 | -0.013733 | 0.091553 | 0.10529 | 0.050354 | 0.086975 | 0.17395 | -0.077819 | 0.21057 | -0.041198
107.7812 | 122.6074 | -59.1934 | 63.0936 | -0.10002 | 0.001128 | 0.068118 | 0.067424 | -0.0017367 | 0.00038017 | -0.0016662 | -0.0013123 | -4.0508e-05 | 0.060434 | 0.062764 | -0.00060892 | -0.0030483 | -0.0023349 | -0.041199 | 63.0661 | -59.3124 | 63.0432 | -59.0286 | -59.4223 | -59.3674 | 0.022888 | 63.0432 | -59.2255 | 63.1714 | -59.1248 | -59.2392 | -59.0012 | -0.11902 | -0.027466 | -0.050354 | -0.022888 | 0.1648 | 0.28381 | -0.17395 | -0.22888 | -0.054932 | -0.10986 | -0.032044 | -0.050354 | 0.07782 | 0.12817 | 0.068664 | 0.10071 | 0.19226 | -0.045777 | 0.2243 | -0.013733
108.7812 | 122.4976 | -59.198 | 63.1485 | -0.031357 | 0.001128 | 0.068118 | 0.069987 | -0.0019503 | -0.0013288 | 0.0019654 | -0.053192 | 0.0018821 | 0.065775 | 0.063833 | -0.0018905 | -0.00069854 | -0.0029757 | 0.009155 | 63.1027 | -59.3124 | 62.9974 | -59.0836 | -59.4543 | -59.3948 | 0.018311 | 63.1256 | -59.23 | 63.1165 | -59.2209 | -59.3491 | -59.0881 | -0.11444 | -0.045777 | -0.15106 | -0.10529 | 0.11444 | 0.22888 | -0.19684 | -0.25635 | -0.082397 | -0.14191 | -0.032044 | -0.022889 | -0.032044 | -0.009155 | -0.022888 | 0.009156 | 0.10986 | -0.15106 | 0.14191 | -0.11902
109.7812 | 122.5845 | -59.1843 | 63.1027 | -0.054245 | 0.001128 | 0.068545 | 0.06486 | -0.0023777 | 0.0031574 | 0.0019654 | -0.010467 | 0.0016685 | 0.06492 | 0.061056 | 0.00024568 | -0.0019803 | -0.0042575 | -0.064087 | 63.0569 | -59.3445 | 63.0341 | -59.0469 | -59.4589 | -59.3903 | -0.009155 | 63.0615 | -59.23 | 63.1531 | -59.1522 | -59.2712 | -59.0195 | -0.16022 | -0.045776 | -0.068664 | -0.022888 | 0.13733 | 0.29755 | -0.20599 | -0.27466 | -0.045777 | -0.11444 | -0.045777 | -0.041199 | 0.050354 | 0.091553 | 0.032043 | 0.07782 | 0.1648 | -0.086975 | 0.21057 | -0.041198
110.7812 | 122.5525 | -59.1843 | 63.1119 | -0.13817 | -0.0020764 | 0.066623 | 0.06721 | -0.00045509 | 0.00038017 | -0.0023072 | -0.010467 | -0.00025411 | 0.064493 | 0.060842 | -0.0023177 | -0.0030483 | -0.0027621 | 0.10071 | 63.0936 | -59.3262 | 63.0203 | -59.0744 | -59.4772 | -59.3857 | -0.032043 | 63.0524 | -59.2621 | 63.1393 | -59.166 | -59.2712 | -59.0469 | -0.14191 | -0.01831 | -0.091552 | -0.073242 | 0.10986 | 0.25177 | -0.20142 | -0.29297 | -0.059509 | -0.15106 | -0.07782 | -0.059509 | 0.027466 | 0.086975 | 0.01831 | 0.09613 | 0.13733 | -0.086975 | 0.21515 | -0.009155
111.7812 | 122.5891 | -59.1843 | 63.0936 | -0.077133 | -0.00015382 | 0.067477 | 0.066142 | 0.0010403 | 0.00038017 | 0.0017518 | -0.025726 | 0.00038669 | 0.063638 | 0.064473 | -0.00039512 | -0.00091214 | 0.00044212 | -0.068665 | 63.0569 | -59.3308 | 63.0157 | -59.0561 | -59.4635 | -59.3948 | -0.027466 | 63.0341 | -59.2484 | 63.1256 | -59.198 | -59.2941 | -59.0378 | -0.14648 | -0.036621 | -0.07782 | -0.041199 | 0.12817 | 0.27466 | -0.21057 | -0.27924 | -0.064086 | -0.13275 | -0.064087 | -0.059509 | 0.032043 | 0.091552 | -0.013733 | 0.050354 | 0.14648 | -0.10986 | 0.21057 | -0.045776
112.7812 | 122.5296 | -59.1934 | 63.0844 | -0.13054 | -0.00036742 | 0.067904 | 0.068278 | -0.00088229 | 0.0020892 | -0.0016662 | -0.037933 | 0.00038669 | 0.067484 | 0.064901 | 0.00067288 | -5.7742e-05 | 0.0045011 | -0.064087 | 63.0661 | -59.3353 | 63.0249 | -59.0698 | -59.4681 | -59.3857 | -0.018311 | 63.0524 | -59.2392 | 63.1165 | -59.1751 | -59.3216 | -59.079 | -0.14191 | -0.018311 | -0.05951 | -0.041199 | 0.1236 | 0.2655 | -0.19226 | -0.27466 | -0.050354 | -0.13275 | -0.045777 | -0.032044 | 0.032043 | 0.064087 | 0.01831 | 0.064087 | 0.11444 | -0.12817 | 0.16022 | -0.082397
113.7812 | 122.6074 | -59.1751 | 63.0936 | -0.10002 | 0.00027338 | 0.067264 | 0.066142 | -0.00045509 | 0.001662 | 0.002179 | 0.0017395 | -0.00089511 | 0.063211 | 0.062551 | -0.00039512 | -0.0013393 | 0.00022852 | -0.045776 | 63.0707 | -59.3399 | 63.0569 | -58.9966 | -59.4131 | -59.3399 | 0.054932 | 63.0615 | -59.2209 | 63.2034 | -59.1019 | -59.2255 | -58.9874 | -0.1648 | -0.022888 | -0.036621 | -0.013733 | 0.17853 | 0.34332 | -0.1648 | -0.23804 | 0 | -0.073242 | -0.045776 | -0.032044 | 0.10986 | 0.14191 | 0.073242 | 0.11902 | 0.18768 | -0.050354 | 0.23346 | -0.004578
114.7812 | 122.5571 | -59.1705 | 63.0936 | -0.092392 | 0.0017688 | 0.067477 | 0.067637 | -0.0043003 | 0.0010212 | 0.00047001 | -0.025726 | 0.00038669 | 0.064493 | 0.06426 | -0.0023177 | -0.00069854 | -0.00041228 | -0.059509 | 63.0569 | -59.3262 | 63.0157 | -59.079 | -59.4818 | -59.404 | -0.036621 | 63.0707 | -59.2255 | 63.1485 | -59.1385 | -59.2804 | -59.0286 | -0.15564 | -0.036621 | -0.07782 | -0.041199 | 0.091553 | 0.24719 | -0.23346 | -0.31128 | -0.07782 | -0.15564 | -0.054932 | -0.022888 | 0.054932 | 0.07782 | 0.032043 | 0.086975 | 0.14191 | -0.10986 | 0.19684 | -0.054932
115.7812 | 122.5708 | -59.198 | 63.089 | -0.10002 | 0.00070078 | 0.068545 | 0.066569 | -0.0010959 | 0.0010212 | 0.0011108 | -0.0013123 | 0.00017309 | 0.065561 | 0.065328 | -0.0016769 | -0.0041165 | -0.0027621 | -0.07782 | 63.0753 | -59.3353 | 63.0157 | -59.0698 | -59.4681 | -59.3903 | -0.004578 | 63.0661 | -59.2209 | 63.1439 | -59.1934 | -59.2941 | -59.0149 | -0.13733 | -0.013733 | -0.073242 | -0.059509 | 0.12817 | 0.2655 | -0.19226 | -0.27008 | -0.054932 | -0.13275 | -0.022888 | -0.022888 | 0.054932 | 0.07782 | 0.004578 | 0.027466 | 0.1831 | -0.09613 | 0.20599 | -0.073242
116.7812 | 122.5296 | -59.1843 | 63.0844 | -0.10765 | 0.0026234 | 0.067904 | 0.067851 | 0.0018949 | 0.001662 | -0.0027344 | -0.044037 | -0.0011087 | 0.063638 | 0.063405 | 0.00067288 | -0.0021939 | 0.00086952 | -0.082397 | 63.0569 | -59.3033 | 63.0249 | -59.0607 | -59.4589 | -59.3994 | -0.041199 | 63.0569 | -59.2575 | 63.089 | -59.0332 | -59.2072 | -58.9233 | -0.11902 | -0.027466 | -0.05951 | -0.032044 | 0.1236 | 0.24262 | -0.21515 | -0.27466 | -0.09613 | -0.15564 | -0.073242 | -0.027466 | 0.004577 | 0.032043 | 0.15106 | 0.2243 | 0.26093 | -0.022888 | 0.33417 | 0.050354
117.7812 | 122.5983 | -59.1705 | 63.089 | -0.038986 | 0.00070078 | 0.068332 | 0.068064 | -0.00088229 | 0.00080757 | 0.0043152 | -0.062347 | 0.0016685 | 0.064493 | 0.067891 | -0.0023177 | -0.0015529 | 0.0032193 | -0.064087 | 63.0341 | -59.3491 | 63.0524 | -59.0149 | -59.4406 | -59.3719 | 0.004578 | 63.0295 | -59.2255 | 63.1577 | -59.1751 | -59.2941 | -59.0332 | -0.17853 | -0.054931 | -0.036621 | 0.01831 | 0.15564 | 0.33417 | -0.20142 | -0.27008 | -0.022888 | -0.091553 | -0.054932 | -0.059509 | 0.068665 | 0.12817 | -0.004578 | 0.050354 | 0.13733 | -0.1236 | 0.19226 | -0.068664
118.7812 | 122.5433 | -59.1614 | 63.0844 | -0.10002 | 0.0015552 | 0.066195 | 0.067851 | 0.0021085 | 0.0018756 | -0.00017099 | -0.056244 | -0.00089511 | 0.064066 | 0.062337 | -0.00018152 | -0.0011257 | 0.0025785 | -0.082397 | 63.0478 | -59.3399 | 63.0295 | -59.0607 | -59.4864 | -59.404 | -0.032043 | 63.0569 | -59.2346 | 63.1439 | -59.1705 | -59.3079 | -59.0332 | -0.17853 | -0.036621 | -0.054932 | -0.018311 | 0.10071 | 0.27924 | -0.24262 | -0.32501 | -0.064087 | -0.14648 | -0.073242 | -0.027466 | 0.059509 | 0.086975 | -0.009155 | 0.064087 | 0.12817 | -0.14648 | 0.20142 | -0.073242
119.7812 | 122.5754 | -59.1934 | 63.0936 | -0.092392 | 0.002196 | 0.065127 | 0.066783 | 0.00082671 | -0.0015424 | -0.0035888 | -0.034882 | -0.00025411 | 0.065775 | 0.061483 | -0.0025315 | -5.7742e-05 | 0.0017239 | -0.059509 | 63.0569 | -59.3262 | 63.002 | -59.0744 | -59.4635 | -59.3811 | -0.022888 | 63.0524 | -59.2621 | 63.2401 | -59.1751 | -59.2987 | -59.0561 | -0.13275 | -0.036621 | -0.091553 | -0.054932 | 0.11902 | 0.25177 | -0.18768 | -0.27008 | -0.054932 | -0.13733 | -0.068665 | -0.041199 | 0.14648 | 0.18768 | 0.01831 | 0.086975 | 0.13733 | -0.10529 | 0.20599 | -0.036621
120.7812 | 122.5433 | -59.198 | 63.0798 | -0.077133 | -0.0016492 | 0.065341 | 0.067424 | -0.0051547 | 0.00059397 | 0.0011108 | -0.044037 | -0.00046791 | 0.064066 | 0.060415 | 0.00045928 | -0.00069854 | -0.0027621 | -0.082397 | 63.0569 | -59.3582 | 63.0157 | -59.079 | -59.4818 | -59.3948 | -0.054932 | 63.0524 | -59.2346 | 63.1531 | -59.166 | -59.3033 | -59.0469 | -0.16022 | -0.022888 | -0.064087 | -0.041199 | 0.11902 | 0.27924 | -0.19684 | -0.28381 | -0.036621 | -0.1236 | -0.036621 | -0.027466 | 0.073242 | 0.10071 | 0.032043 | 0.068664 | 0.15106 | -0.10529 | 0.18768 | -0.068665
121.7812 | 122.5479 | -59.2026 | 63.0798 | -0.11528 | 0.0019824 | 0.067477 | 0.068705 | -0.0021641 | 0.00016657 | 0.002179 | -0.044037 | -0.00025411 | 0.063425 | 0.06661 | -0.00060892 | -0.0028347 | 0.0017239 | -0.064087 | 63.0569 | -59.3216 | 63.0066 | -59.0561 | -59.4681 | -59.4177 | -0.041199 | 63.0341 | -59.2484 | 63.1256 | -59.1797 | -59.2987 | -59.0469 | -0.11902 | -0.022888 | -0.073242 | -0.050354 | 0.14648 | 0.2655 | -0.21515 | -0.2655 | -0.096131 | -0.14649 | -0.045776 | -0.045776 | 0.045776 | 0.091552 | 0.022889 | 0.068665 | 0.15564 | -0.09613 | 0.20142 | -0.050354
122.7812 | 122.5479 | -59.1797 | 63.0661 | -0.13054 | 0.0013416 | 0.06705 | 0.067851 | -0.0013095 | 0.0018756 | -0.0044434 | -0.03183 | -0.0013223 | 0.063638 | 0.065542 | -0.0018905 | 0.00058326 | 0.0017239 | -0.091553 | 63.0661 | -59.3491 | 63.0203 | -59.079 | -59.4498 | -59.404 | -0.022888 | 63.0615 | -59.2438 | 63.2904 | -59.2987 | -59.2896 | -59.024 | -0.16937 | 0 | -0.045776 | -0.045776 | 0.10071 | 0.27008 | -0.22431 | -0.27008 | -0.054932 | -0.10071 | -0.064087 | -0.004578 | 0.2243 | 0.22888 | -0.11902 | -0.054932 | 0.15564 | -0.10986 | 0.21973 | -0.045777
123.7812 | 122.5388 | -59.1797 | 63.089 | 0.22041 | 5.9776e-05 | 0.069186 | 0.06956 | -0.0028049 | 0.0018756 | 0.0017518 | -0.03183 | -0.00025411 | 0.061929 | 0.058706 | -0.0010361 | -0.0021939 | 0.0012967 | -0.091553 | 63.0661 | -59.3216 | 63.0112 | -59.0836 | -59.4543 | -59.3674 | -0.032043 | 63.0478 | -59.2255 | 63.1485 | -59.1568 | -59.2758 | -59.0286 | -0.14191 | -0.022888 | -0.07782 | -0.054932 | 0.09613 | 0.23804 | -0.18768 | -0.27466 | -0.045777 | -0.13275 | -0.045777 | -0.041198 | 0.05951 | 0.10071 | 0.022888 | 0.068665 | 0.15106 | -0.096131 | 0.19684 | -0.050354
124.7812 | 122.6166 | -59.1888 | 63.0798 | -0.16106 | 0.0013416 | 0.068973 | 0.070414 | -0.0010959 | -0.00026063 | 0.00089721 | -0.11118 | -4.0508e-05 | 0.06492 | 0.066396 | 0.00024568 | -0.0039029 | 1.4916e-05 | -0.045776 | 63.0112 | -59.3765 | 63.0478 | -59.0332 | -59.4681 | -59.3903 | 0.009155 | 63.0386 | -59.2346 | 63.2996 | -59.1614 | -59.285 | -59.0195 | -0.18768 | -0.068665 | -0.032043 | 0.036622 | 0.15564 | 0.34332 | -0.20142 | -0.27924 | -0.013733 | -0.091553 | -0.045776 | -0.041199 | 0.21973 | 0.26093 | 0.027466 | 0.073242 | 0.16937 | -0.09613 | 0.21515 | -0.050354
125.7812 | 122.5342 | -59.1934 | 63.0844 | -0.061874 | 0.001128 | 0.070254 | 0.06721 | -0.0013095 | 0.0018756 | 0.0013244 | -0.056244 | -0.00046791 | 0.062357 | 0.065755 | -0.0012497 | -0.00048494 | -0.0012667 | -0.059509 | 63.0478 | -59.317 | 63.0386 | -59.0378 | -59.4772 | -59.3811 | -0.018311 | 63.0707 | -59.2484 | 63.1439 | -59.2026 | -59.3033 | -59.0286 | -0.1236 | -0.036621 | -0.045777 | -0.009156 | 0.15564 | 0.27924 | -0.18768 | -0.28381 | -0.064087 | -0.16022 | -0.054932 | -0.013733 | 0.059509 | 0.073242 | -0.009156 | 0.045776 | 0.1648 | -0.10986 | 0.21973 | -0.054932
126.7812 | 122.5845 | -59.1522 | 63.1073 | -0.21446 | 0.001128 | 0.065555 | 0.066569 | -0.0015231 | 0.00080757 | -0.002948 | -0.013519 | 0.00060029 | 0.062784 | 0.064046 | 3.2078e-05 | -5.7742e-05 | -0.0025485 | 0.07782 | 63.0661 | -59.3033 | 63.0981 | -58.9645 | -59.3719 | -59.3353 | 0.12817 | 63.0661 | -59.2117 | 63.1943 | -59.0973 | -59.2255 | -58.9737 | -0.15106 | -0.041199 | -0.009155 | 0.032044 | 0.18768 | 0.33874 | -0.1831 | -0.21973 | -0.032043 | -0.068664 | -0.059509 | -0.041199 | 0.086975 | 0.12817 | 0.054932 | 0.11444 | 0.17853 | -0.073242 | 0.23804 | -0.013733
127.7812 | 122.5708 | -59.1843 | 63.0936 | -0.12291 | 0.00070078 | 0.068545 | 0.068064 | 0.0016813 | 0.0031574 | 0.00089721 | -0.065399 | -0.0017495 | 0.065988 | 0.077932 | -0.0018905 | -0.00091214 | 0.0021511 | -0.036621 | 63.0341 | -59.3399 | 63.0432 | -59.0195 | -59.4406 | -59.3857 | 0.004578 | 63.0341 | -59.2209 | 63.1805 | -59.1476 | -59.2758 | -59.0103 | -0.15564 | -0.059509 | -0.050354 | 0.009155 | 0.1648 | 0.32044 | -0.20142 | -0.25635 | -0.045776 | -0.10071 | -0.036621 | -0.059509 | 0.086975 | 0.14648 | 0.036621 | 0.073242 | 0.17395 | -0.091553 | 0.21057 | -0.054932
128.7812 | 122.6212 | -59.2072 | 63.0615 | -0.11528 | -0.00015382 | 0.0694 | 0.06721 | 0.0025357 | 0.003371 | 0.00047001 | -0.077606 | 0.00060029 | 0.062997 | 0.065969 | -0.0016769 | -0.00091214 | -0.0014805 | -0.1236 | 62.9974 | -59.3811 | 63.0615 | -58.9737 | -59.4681 | -59.3811 | 0.004578 | 63.0569 | -59.23 | 63.1714 | -59.1339 | -59.2667 | -58.9966 | -0.17395 | -0.064086 | 0 | 0.064086 | 0.23346 | 0.40741 | -0.17395 | -0.26093 | 0 | -0.086975 | -0.022889 | -0.004577 | 0.10986 | 0.11444 | 0.073242 | 0.096131 | 0.21057 | -0.05951 | 0.23346 | -0.036621
129.7812 | 122.5937 | -59.2026 | 63.0386 | -0.12291 | 0.0019824 | 0.071536 | 0.069987 | 0.0023221 | 0.0014484 | -0.0016662 | -0.095917 | -0.00068151 | 0.065561 | 0.057851 | -0.0010361 | -0.0013393 | 0.0019375 | -0.054932 | 63.0249 | -59.3353 | 63.0569 | -59.0378 | -59.4543 | -59.3765 | -0.027466 | 63.0157 | -59.2804 | 63.1714 | -59.0881 | -59.2804 | -59.0057 | -0.13275 | -0.013733 | 0.018311 | 0.032044 | 0.16479 | 0.29755 | -0.17395 | -0.25177 | -0.041199 | -0.11902 | -0.07782 | -0.022888 | 0.13275 | 0.15564 | 0.11444 | 0.19226 | 0.19684 | -0.07782 | 0.27466 | 0
130.7812 | 122.612 | -59.2026 | 63.0661 | -0.13054 | -0.00036742 | 0.068545 | 0.067637 | 0.0023221 | 0.0025166 | 0.0017518 | -0.077606 | -4.0508e-05 | 0.065561 | 0.066182 | 0.00045928 | 0.00058326 | -0.00019868 | -0.064087 | 63.0432 | -59.3582 | 63.0478 | -59.0378 | -59.4635 | -59.404 | -0.027466 | 63.0157 | -59.2621 | 63.1851 | -59.1614 | -59.285 | -58.9966 | -0.15564 | -0.022888 | -0.01831 | 0.004578 | 0.16479 | 0.32043 | -0.20142 | -0.26093 | -0.045777 | -0.10529 | -0.059509 | -0.050354 | 0.11902 | 0.16937 | 0.041199 | 0.10071 | 0.20599 | -0.082397 | 0.2655 | -0.022888
131.7812 | 122.4792 | -59.1888 | 63.1393 | -0.10765 | 0.0026234 | 0.063846 | 0.068492 | -0.0030185 | -0.00068783 | -0.00038459 | -0.074554 | 0.0012411 | 0.065561 | 0.074514 | -0.00082252 | -0.0011257 | -0.00062588 | 0.16022 | 63.2126 | -59.2941 | 62.9608 | -59.1156 | -59.4772 | -59.4406 | -0.032043 | 63.1348 | -59.2346 | 63.1165 | -59.2072 | -59.3399 | -59.0652 | -0.10529 | 0.073242 | -0.17853 | -0.25177 | 0.073242 | 0.17853 | -0.25177 | -0.28839 | -0.14648 | -0.18311 | -0.045776 | -0.004577 | -0.022888 | -0.018311 | -0.01831 | 0.027466 | 0.1236 | -0.15106 | 0.16937 | -0.10529
132.7812 | 122.5616 | -59.2484 | 62.9974 | -0.13054 | 0.0013416 | 0.070041 | 0.069346 | -0.00045509 | -0.00026063 | -0.00038459 | -0.056244 | 0.0016685 | 0.061929 | 0.064901 | -0.0025315 | -0.00069854 | -0.0016941 | -0.054932 | 63.0112 | -59.3582 | 63.0341 | -59.0332 | -59.4589 | -59.3719 | 0 | 63.0341 | -59.2346 | 63.1805 | -59.1293 | -59.2712 | -59.0057 | -0.10986 | 0.013732 | 0.036621 | 0.022889 | 0.21515 | 0.32501 | -0.1236 | -0.21057 | -0.013733 | -0.10071 | 0.013733 | 0.036621 | 0.18311 | 0.14648 | 0.11902 | 0.10529 | 0.24261 | -0.022888 | 0.22888 | -0.036621
133.7812 | 122.5845 | -59.2026 | 63.0569 | -0.1458 | -0.0025036 | 0.069186 | 0.069346 | 0.0025357 | 0.001662 | 4.2808e-05 | -0.074554 | 0.0018821 | 0.065347 | 0.066823 | -0.0025315 | -0.0015529 | -0.0010531 | -0.091553 | 63.0615 | -59.3536 | 63.0112 | -59.0561 | -59.4589 | -59.404 | -0.041199 | 63.002 | -59.2438 | 63.176 | -59.1751 | -59.2758 | -58.992 | -0.15106 | 0.004577 | -0.045777 | -0.050354 | 0.14648 | 0.29755 | -0.20142 | -0.25635 | -0.050354 | -0.10529 | -0.041198 | -0.054932 | 0.11902 | 0.17395 | 0.027466 | 0.068664 | 0.21057 | -0.073242 | 0.25177 | -0.032044
134.7812 | 122.5891 | -59.1934 | 63.0661 | -0.13817 | -0.00036742 | 0.068759 | 0.067851 | -0.0019503 | 0.0018756 | -0.0018798 | -0.095917 | 0.00081389 | 0.064066 | 0.06191 | 0.00045928 | 0.00058326 | 0.0012967 | -0.11444 | 63.0661 | -59.3674 | 62.9883 | -59.0927 | -59.5184 | -59.4498 | -0.073242 | 63.0386 | -59.285 | 63.1668 | -59.1751 | -59.3124 | -59.0195 | -0.17395 | 0 | -0.07782 | -0.07782 | 0.10071 | 0.27466 | -0.25635 | -0.32501 | -0.082397 | -0.15106 | -0.091553 | -0.027466 | 0.10071 | 0.12817 | 0.01831 | 0.10986 | 0.17395 | -0.11902 | 0.2655 | -0.027466
135.7812 | 122.5204 | -59.198 | 63.0753 | -0.038986 | 0.00048698 | 0.069186 | 0.069987 | -0.0021641 | 0.0027302 | -0.00038459 | -0.074554 | 0.00038669 | 0.06492 | 0.066396 | -0.00018152 | -0.00069854 | 0.00044212 | -0.11902 | 63.0798 | -59.3719 | 62.9471 | -59.1339 | -59.5367 | -59.4543 | -0.10529 | 63.0661 | -59.2667 | 63.1256 | -59.2163 | -59.3353 | -59.0424 | -0.17395 | 0.004578 | -0.12817 | -0.13275 | 0.064087 | 0.23804 | -0.25635 | -0.33875 | -0.082398 | -0.16479 | -0.068665 | -0.009155 | 0.050354 | 0.059509 | -0.018311 | 0.050354 | 0.15564 | -0.13733 | 0.22431 | -0.068664
136.7812 | 122.5342 | -59.198 | 63.089 | -0.038986 | 0.00091438 | 0.069186 | 0.066783 | -0.0053683 | 0.003371 | -0.00038459 | 0.0078429 | -0.0011087 | 0.063852 | 0.05187 | -0.0012497 | -0.00069854 | 0.0019375 | -0.086975 | 63.0753 | -59.3353 | 62.9974 | -59.0652 | -59.4635 | -59.3903 | -0.059509 | 63.0753 | -59.2484 | 63.1439 | -59.1797 | -59.3079 | -59.0332 | -0.13733 | -0.013733 | -0.091552 | -0.077819 | 0.13275 | 0.27008 | -0.19226 | -0.2655 | -0.054932 | -0.12817 | -0.050354 | -0.013733 | 0.054932 | 0.068665 | 0.018311 | 0.068665 | 0.16479 | -0.10986 | 0.21515 | -0.059509
137.7812 | 122.3694 | -59.2346 | 63.0386 | -0.11528 | 0.00048698 | 0.067477 | 0.069773 | 0.0018949 | -0.00026063 | -0.0010254 | -0.10507 | 0.00060029 | 0.06492 | 0.067678 | -0.0012497 | -0.00091214 | 0.0015103 | -0.10071 | 63.0844 | -59.3903 | 62.9333 | -59.1888 | -59.5871 | -59.5184 | -0.15564 | 63.1027 | -59.2987 | 63.0798 | -59.285 | -59.3994 | -59.0561 | -0.15564 | 0.045777 | -0.10529 | -0.15106 | 0.045776 | 0.20142 | -0.28381 | -0.35248 | -0.12817 | -0.19684 | -0.064087 | 0.064087 | 0.041199 | -0.022888 | -0.050354 | 0.013733 | 0.17853 | -0.16479 | 0.24262 | -0.10071
138.7812 | 122.5388 | -59.2072 | 63.0524 | -0.16106 | 0.00027338 | 0.068545 | 0.068492 | -0.0010959 | 0.00038017 | 0.00089721 | 0.010895 | 0.00081389 | 0.064706 | 0.061269 | 0.0015275 | -0.0013393 | -0.0027621 | -0.068665 | 63.0203 | -59.3216 | 63.0157 | -59.0378 | -59.4864 | -59.404 | -0.036621 | 63.0066 | -59.2484 | 63.1302 | -59.1751 | -59.2758 | -59.024 | -0.11444 | -0.032043 | -0.036621 | -0.004578 | 0.16937 | 0.28381 | -0.19684 | -0.27924 | -0.082398 | -0.1648 | -0.041199 | -0.045776 | 0.07782 | 0.1236 | 0.032043 | 0.073242 | 0.1831 | -0.068665 | 0.2243 | -0.027466
139.7812 | 122.4564 | -59.2209 | 63.0798 | -0.10002 | -0.00036742 | 0.069827 | 0.06956 | 0.0012539 | 0.00016657 | 0.0028198 | -0.22714 | -0.0013223 | 0.062997 | 0.065755 | -0.00082252 | -0.0017667 | 0.0027921 | -0.11444 | 63.0798 | -59.3765 | 62.9379 | -59.166 | -59.5871 | -59.5917 | 0.050354 | 63.0569 | -59.2667 | 63.0753 | -59.2529 | -59.4086 | -59.1019 | -0.15564 | 0 | -0.14191 | -0.14191 | 0.054931 | 0.21057 | -0.37079 | -0.36621 | -0.21515 | -0.21057 | -0.045777 | -0.022888 | -0.004578 | 0.01831 | -0.032044 | 0.013733 | 0.11902 | -0.18768 | 0.16479 | -0.14191
140.7812 | 122.5204 | -59.2072 | 63.1119 | -0.092392 | -0.00058102 | 0.068332 | 0.070842 | 0.0021085 | 0.0020892 | 0.00025641 | -0.065399 | 0.0020957 | 0.065988 | 0.06191 | -0.00018152 | 0.00015586 | 0.00022852 | -0.09613 | 63.1073 | -59.3399 | 62.9242 | -59.1339 | -59.5551 | -59.4864 | -0.13275 | 63.0569 | -59.2621 | 63.0753 | -59.2529 | -59.3765 | -59.0836 | -0.13275 | -0.004577 | -0.18768 | -0.18311 | 0.073242 | 0.20599 | -0.27924 | -0.3479 | -0.14648 | -0.21515 | -0.054932 | -0.054931 | -0.036621 | 0.01831 | -0.045777 | 0.009155 | 0.1236 | -0.16937 | 0.17853 | -0.11444
141.7969 | 122.5616 | -59.1568 | 63.1348 | -0.084763 | 0.0013416 | 0.068545 | 0.066569 | 0.00018591 | -0.0013288 | -0.00017099 | 0.013947 | 0.0010275 | 0.065347 | 0.062124 | -0.00082252 | -0.0028347 | -0.0014805 | 0.013733 | 63.0798 | -59.2758 | 63.0707 | -58.992 | -59.3811 | -59.3353 | 0.091553 | 63.0661 | -59.2072 | 63.1577 | -59.1385 | -59.2346 | -59.0012 | -0.11902 | -0.054932 | -0.064087 | -0.009155 | 0.16479 | 0.28381 | -0.17853 | -0.22431 | -0.059509 | -0.10529 | -0.050354 | -0.068665 | 0.022888 | 0.091553 | 0.01831 | 0.068664 | 0.15564 | -0.07782 | 0.20599 | -0.027466
142.7812 | 122.5021 | -59.2255 | 63.1027 | -0.10002 | -0.0010082 | 0.068332 | 0.069987 | 0.0029629 | -0.00090143 | -0.0016662 | -0.086761 | -0.00025411 | 0.065134 | 0.069814 | -0.00039512 | 0.0022923 | 0.0023647 | -0.09613 | 63.089 | -59.3582 | 62.7319 | -58.992 | -59.2804 | -59.4589 | -0.10071 | 63.0661 | -59.2438 | 63.0615 | -59.2667 | -59.4177 | -59.1202 | -0.13275 | -0.013733 | -0.37079 | -0.35705 | 0.23346 | 0.36621 | -0.23346 | -0.054932 | -0.10071 | 0.077819 | -0.01831 | -0.036621 | -0.041199 | -0.004578 | -0.041199 | -0.022889 | 0.10529 | -0.19226 | 0.1236 | -0.17395
143.7812 | 122.5204 | -59.2163 | 63.0798 | -0.077133 | 0.00070078 | 0.067691 | 0.069133 | -0.0017367 | -4.7032e-05 | 0.0036744 | -0.077606 | 0.00017309 | 0.064706 | 0.065755 | -0.0012497 | -0.00048494 | -0.0023349 | -0.1236 | 63.0798 | -59.3811 | 62.9517 | -59.1568 | -59.5505 | -59.491 | -0.11444 | 63.0661 | -59.2712 | 63.1073 | -59.2529 | -59.3719 | -59.0652 | -0.16479 | 0 | -0.12817 | -0.12817 | 0.05951 | 0.22431 | -0.27466 | -0.33417 | -0.10986 | -0.16937 | -0.054931 | -0.013733 | 0.027466 | 0.041199 | -0.036621 | 0.01831 | 0.15106 | -0.15564 | 0.20599 | -0.10071
144.7812 | 122.5021 | -59.2163 | 63.1073 | -0.092392 | 0.001128 | 0.0694 | 0.068705 | -0.0034457 | -0.00026063 | 0.00025641 | -0.10202 | 0.0014547 | 0.065134 | 0.069814 | -0.00018152 | 0.0014377 | 0.0015103 | -0.073242 | 63.1165 | -59.3491 | 62.97 | -59.1156 | -59.5322 | -59.4543 | -0.073242 | 63.0798 | -59.2484 | 63.0707 | -59.2575 | -59.3994 | -59.0927 | -0.13275 | 0.009155 | -0.13733 | -0.14648 | 0.10071 | 0.23346 | -0.23804 | -0.31586 | -0.10529 | -0.18311 | -0.032043 | -0.027466 | -0.036621 | -0.009155 | -0.041198 | -0.009155 | 0.1236 | -0.1831 | 0.15564 | -0.15106
145.7812 | 122.4655 | -59.2484 | 63.0981 | -0.084763 | -0.00015382 | 0.066836 | 0.069346 | 0.0031765 | 0.0029438 | 4.2808e-05 | -0.077606 | -0.00089511 | 0.06727 | 0.068105 | 0.00024568 | 0.00079686 | 1.4916e-05 | -0.10071 | 63.0981 | -59.3536 | 62.9196 | -59.1843 | -59.5642 | -59.5322 | -0.16479 | 63.0661 | -59.2896 | 63.0569 | -59.2758 | -59.4131 | -59.111 | -0.10529 | 0 | -0.17853 | -0.17853 | 0.064087 | 0.16937 | -0.28381 | -0.31586 | -0.17853 | -0.21057 | -0.041199 | -0.032044 | -0.041199 | -0.009155 | -0.027466 | 0.013733 | 0.13733 | -0.1648 | 0.17853 | -0.1236
146.7812 | 122.5067 | -59.2209 | 63.0981 | -0.038986 | -0.001222 | 0.066409 | 0.068278 | -0.0015231 | 0.001662 | 0.002179 | -0.062347 | 0.0029501 | 0.065561 | 0.059347 | 0.0013139 | -0.00091214 | 0.0025785 | -0.009155 | 63.0798 | -59.3811 | 62.8738 | -59.1293 | -59.5139 | -59.4772 | -0.14191 | 63.0661 | -59.2667 | 63.0615 | -59.2667 | -59.3948 | -59.0973 | -0.16022 | -0.018311 | -0.22431 | -0.20599 | 0.091553 | 0.25177 | -0.25635 | -0.29297 | -0.09613 | -0.13275 | -0.045777 | -0.032044 | -0.036622 | -0.004578 | -0.045777 | 0 | 0.1236 | -0.17395 | 0.16937 | -0.12817
147.7812 | 122.4655 | -59.23 | 63.1073 | -0.061874 | 0.00027338 | 0.06705 | 0.068919 | -0.0010959 | 0.001662 | 0.00047001 | 0.02005 | 0.0035911 | 0.067697 | 0.062764 | -0.00039512 | -0.0030483 | 0.00022852 | -0.11444 | 63.0798 | -59.3811 | 62.9608 | -59.4315 | -59.8343 | -59.3491 | -0.18768 | 63.089 | -59.23 | 63.0798 | -59.2896 | -59.3719 | -59.079 | -0.15106 | -0.027466 | -0.14649 | -0.11902 | -0.20142 | -0.050354 | -0.11902 | -0.60425 | 0.032044 | -0.45319 | 0 | -0.018311 | -0.027466 | -0.009155 | -0.059509 | -0.059509 | 0.15106 | -0.14191 | 0.15106 | -0.14191
148.7812 | 122.4838 | -59.2117 | 63.121 | -0.084763 | 0.00048698 | 0.067904 | 0.068492 | 0.0016813 | -4.7032e-05 | -0.00038459 | -0.071503 | -4.0508e-05 | 0.066415 | 0.064473 | -0.0018905 | -5.7742e-05 | 0.0025785 | -0.12817 | 63.089 | -59.3811 | 62.8876 | -59.1797 | -59.5596 | -59.491 | -0.009155 | 63.0569 | -59.2712 | 63.0432 | -59.2392 | -59.4131 | -59.1339 | -0.16937 | -0.032044 | -0.23346 | -0.20142 | 0.032044 | 0.20142 | -0.27924 | -0.3479 | -0.10986 | -0.17853 | -0.059509 | -0.064087 | -0.07782 | -0.013733 | -0.027466 | 0.032043 | 0.07782 | -0.20142 | 0.13733 | -0.14191
149.7812 | 122.4976 | -59.23 | 63.1027 | -0.15343 | -0.00015382 | 0.066623 | 0.068278 | 0.00061311 | 0.0018756 | 0.00047001 | -0.028778 | 0.0010275 | 0.063425 | 0.058706 | 0.00088648 | -0.00048494 | -0.0023349 | -0.11444 | 63.1119 | -59.3582 | 62.9059 | -59.1522 | -59.5596 | -59.4772 | -0.11902 | 63.0707 | -59.2575 | 63.0844 | -59.2712 | -59.3811 | -59.0744 | -0.12817 | 0.009155 | -0.19684 | -0.20599 | 0.07782 | 0.20599 | -0.24719 | -0.32959 | -0.11902 | -0.20142 | -0.027465 | -0.032043 | -0.01831 | 0.013733 | -0.041198 | -0.013733 | 0.15564 | -0.15106 | 0.1831 | -0.1236
150.7812 | 122.4655 | -59.2255 | 63.0753 | -0.10765 | 0.00091438 | 0.067904 | 0.068278 | 0.0033901 | 0.00016657 | 0.0034606 | -0.053192 | 0.0010275 | 0.064493 | 0.063833 | -0.0010361 | -0.0011257 | 0.0021511 | -0.091553 | 63.0936 | -59.3628 | 62.9196 | -59.1843 | -59.5734 | -59.523 | -0.16937 | 63.0844 | -59.2575 | 63.0524 | -59.2804 | -59.4177 | -59.1248 | -0.13733 | 0.018311 | -0.15564 | -0.17395 | 0.041199 | 0.17853 | -0.29755 | -0.3479 | -0.16022 | -0.21057 | -0.032043 | 0.009156 | -0.022888 | -0.032044 | -0.054932 | -0.022889 | 0.10071 | -0.19226 | 0.13275 | -0.16022
151.7812 | 122.4792 | -59.2072 | 63.1439 | -0.061874 | -0.00015382 | 0.067477 | 0.068919 | 0.0025357 | 0.0018756 | 4.2808e-05 | -0.037933 | 0.0014547 | 0.065134 | 0.072805 | -0.00039512 | -0.00048494 | 0.00044212 | -0.036621 | 63.0798 | -59.3536 | 62.9929 | -59.079 | -59.4818 | -59.436 | -0.036621 | 63.1027 | -59.2346 | 63.1073 | -59.2392 | -59.3674 | -59.0698 | -0.14648 | -0.064087 | -0.15106 | -0.086975 | 0.12817 | 0.27466 | -0.22888 | -0.27466 | -0.082397 | -0.12817 | -0.027466 | -0.041199 | -0.036621 | 0.004578 | -0.032044 | -0.004578 | 0.13733 | -0.16022 | 0.1648 | -0.13275
152.7812 | 122.4976 | -59.198 | 63.1165 | -0.13054 | 0.00048698 | 0.069613 | 0.069773 | 0.0018949 | 0.00038017 | 0.0045288 | -0.056244 | 0.0016685 | 0.06492 | 0.069814 | 0.0017411 | -0.00027134 | -0.0016941 | -0.018311 | 63.0798 | -59.3079 | 62.9837 | -59.0652 | -59.4818 | -59.4315 | -0.013733 | 63.0798 | -59.2529 | 63.1073 | -59.2163 | -59.3536 | -59.0607 | -0.10986 | -0.036621 | -0.13275 | -0.09613 | 0.13275 | 0.24261 | -0.23346 | -0.28381 | -0.1236 | -0.17395 | -0.054932 | -0.036621 | -0.009155 | 0.027466 | -0.018311 | 0.036621 | 0.13733 | -0.15564 | 0.19226 | -0.10071
153.7812 | 122.5204 | -59.1888 | 63.1165 | -0.077133 | 5.9776e-05 | 0.068332 | 0.068919 | -0.0010959 | 0.0012348 | 4.2808e-05 | -0.013519 | 0.00017309 | 0.065347 | 0.062124 | 0.00024568 | -0.0011257 | 0.0021511 | -0.11444 | 63.0798 | -59.3399 | 62.9105 | -59.1705 | -59.5642 | -59.4818 | -0.11902 | 63.0661 | -59.2529 | 63.0798 | -59.2575 | -59.3994 | -59.0698 | -0.15106 | -0.036621 | -0.20599 | -0.16937 | 0.018311 | 0.16937 | -0.29297 | -0.37537 | -0.14191 | -0.2243 | -0.064087 | -0.050354 | -0.036621 | 0.013733 | -0.068664 | -0.004577 | 0.11902 | -0.21057 | 0.18311 | -0.14648
154.7812 | 122.5159 | -59.2026 | 63.121 | -0.069504 | 0.00091438 | 0.066836 | 0.068919 | 0.0018949 | 0.0020892 | 0.00068361 | -0.037933 | 0.0031637 | 0.065775 | 0.065969 | -0.00039512 | -0.0011257 | 0.00022852 | -0.032043 | 63.089 | -59.3353 | 62.97 | -59.0881 | -59.491 | -59.4543 | -0.054932 | 63.0707 | -59.2392 | 63.0936 | -59.2484 | -59.3811 | -59.0973 | -0.13275 | -0.032044 | -0.15106 | -0.11902 | 0.11444 | 0.24719 | -0.25177 | -0.28839 | -0.11902 | -0.15564 | -0.036621 | -0.050354 | -0.027466 | 0.022888 | -0.045776 | -0.009155 | 0.10529 | -0.17853 | 0.14191 | -0.14191
155.7812 | 122.525 | -59.23 | 63.1165 | -0.077133 | 0.00048698 | 0.067904 | 0.068278 | -2.7692e-05 | 0.00016657 | 0.0028198 | 0.041412 | 0.0023093 | 0.065561 | 0.061269 | 3.2078e-05 | -5.7742e-05 | 0.0012967 | -0.10986 | 63.0844 | -59.3765 | 62.9379 | -59.4177 | -59.848 | -59.3445 | -0.17853 | 63.0844 | -59.2209 | 63.0661 | -59.3033 | -59.3994 | -59.0744 | -0.14648 | -0.032043 | -0.17853 | -0.14648 | -0.18768 | -0.041199 | -0.11444 | -0.61798 | 0.032044 | -0.4715 | 0.009156 | -0.032043 | -0.050354 | -0.018311 | -0.073242 | -0.082398 | 0.15564 | -0.16937 | 0.14648 | -0.17853
156.7812 | 122.4792 | -59.2255 | 63.121 | -0.069504 | -0.00058102 | 0.068332 | 0.068278 | 0.0012539 | -0.0013288 | -0.0016662 | -0.047089 | -0.00046791 | 0.065775 | 0.062124 | -0.00018152 | -0.0019803 | 0.0015103 | -0.068665 | 63.089 | -59.3765 | 62.9517 | -59.1293 | -59.5459 | -59.4772 | -0.091553 | 63.0844 | -59.2621 | 63.0661 | -59.2758 | -59.4086 | -59.1064 | -0.15106 | -0.032044 | -0.16937 | -0.13733 | 0.096131 | 0.24719 | -0.25177 | -0.32043 | -0.10071 | -0.16937 | -0.036621 | -0.036621 | -0.054932 | -0.018311 | -0.050354 | -0.013733 | 0.11902 | -0.1831 | 0.15564 | -0.14648
157.7812 | 122.4838 | -59.23 | 63.1073 | -0.061874 | -0.0018628 | 0.068545 | 0.069773 | -0.0013095 | -0.0023968 | -0.0018798 | -0.025726 | -0.0021769 | 0.06492 | 0.062764 | 0.00024568 | -5.7742e-05 | 0.0027921 | -0.10529 | 63.0707 | -59.3445 | 62.9105 | -59.1797 | -59.5596 | -59.4955 | -0.16022 | 63.0524 | -59.2529 | 63.0615 | -59.2941 | -59.4131 | -59.1019 | -0.11444 | -0.036621 | -0.19684 | -0.16022 | 0.050355 | 0.16479 | -0.2655 | -0.32959 | -0.15106 | -0.21515 | -0.022888 | -0.054932 | -0.045777 | 0.009155 | -0.064086 | -0.041198 | 0.12817 | -0.18311 | 0.15106 | -0.16022
158.7812 | 122.5296 | -59.1934 | 63.121 | -0.092392 | 0.001128 | 0.06705 | 0.068705 | -0.0019503 | 0.0012348 | 0.0017518 | -0.056244 | 0.0010275 | 0.065561 | 0.067891 | 0.00088648 | -0.0019803 | -0.0023349 | -0.009155 | 63.1119 | -59.3216 | 63.0066 | -59.0607 | -59.4452 | -59.4269 | 0.027466 | 63.0844 | -59.2621 | 63.1577 | -59.198 | -59.3262 | -59.0515 | -0.12817 | -0.009156 | -0.11444 | -0.10529 | 0.13275 | 0.26093 | -0.23346 | -0.25177 | -0.10529 | -0.1236 | -0.068665 | -0.036621 | 0.036621 | 0.073242 | -0.004578 | 0.064087 | 0.14191 | -0.13275 | 0.21057 | -0.064087
159.7812 | 122.5021 | -59.2209 | 63.1256 | -0.046616 | -0.00036742 | 0.066836 | 0.068064 | 0.0031765 | -0.0021832 | -0.00038459 | -0.019623 | -0.00046791 | 0.065988 | 0.062551 | -0.00060892 | -5.7742e-05 | 0.0021511 | -0.013733 | 63.0981 | -59.3079 | 63.0112 | -59.0607 | -59.4589 | -59.4131 | 0.022888 | 63.0661 | -59.2255 | 63.0295 | -59.166 | -59.3262 | -59.024 | -0.086975 | -0.027465 | -0.11444 | -0.086976 | 0.16022 | 0.24719 | -0.19226 | -0.23804 | -0.10529 | -0.15106 | -0.004578 | -0.059509 | -0.09613 | -0.036621 | 0.054931 | 0.059509 | 0.19684 | -0.10529 | 0.20142 | -0.10071
160.7812 | 122.4701 | -59.198 | 63.1302 | -0.054245 | -0.0010082 | 0.067904 | 0.068705 | -0.00024129 | 0.00016657 | -0.0010254 | -0.044037 | -0.00068151 | 0.065347 | 0.065114 | 0.00067288 | -0.00027134 | 0.0045011 | -0.045776 | 63.1027 | -59.3353 | 62.9471 | -59.1156 | -59.523 | -59.4727 | -0.064087 | 63.0524 | -59.2621 | 63.0707 | -59.2758 | -59.4086 | -59.1019 | -0.13733 | -0.027466 | -0.1831 | -0.15564 | 0.082397 | 0.21973 | -0.27466 | -0.32501 | -0.13733 | -0.18768 | -0.064087 | -0.07782 | -0.059509 | 0.018311 | -0.07782 | -0.013733 | 0.09613 | -0.21057 | 0.16022 | -0.14648
161.7812 | 122.5067 | -59.2072 | 63.1439 | -0.084763 | 0.0013416 | 0.068332 | 0.070414 | 0.0029629 | -0.003465 | -0.001239 | -0.034882 | -4.0508e-05 | 0.066202 | 0.064901 | -0.00082252 | -0.0021939 | 0.00086952 | -0.013733 | 63.1027 | -59.3445 | 63.0112 | -59.0744 | -59.4681 | -59.3811 | 0.027466 | 63.089 | -59.2529 | 63.1165 | -59.2026 | -59.3262 | -59.0698 | -0.13733 | -0.041199 | -0.13275 | -0.091553 | 0.13275 | 0.27008 | -0.17395 | -0.26093 | -0.036622 | -0.1236 | -0.045777 | -0.054932 | -0.027466 | 0.027466 | 0.004577 | 0.050354 | 0.13733 | -0.11902 | 0.18311 | -0.073242
162.7812 | 122.4976 | -59.1934 | 63.1256 | -0.10765 | 0.00070078 | 0.066836 | 0.06956 | 0.0023221 | 0.001662 | 0.00068361 | -0.03183 | 0.0016685 | 0.065775 | 0.058706 | 0.00045928 | -0.0011257 | -0.0025485 | -0.032043 | 63.1027 | -59.3491 | 62.9883 | -59.0881 | -59.491 | -59.4498 | -0.036621 | 63.0707 | -59.2392 | 63.0981 | -59.2484 | -59.3628 | -59.079 | -0.15564 | -0.022888 | -0.13733 | -0.11444 | 0.10529 | 0.26093 | -0.25635 | -0.29755 | -0.10071 | -0.14191 | -0.045777 | -0.054931 | -0.027465 | 0.027466 | -0.054932 | -0.009155 | 0.11444 | -0.16937 | 0.16022 | -0.1236
163.7812 | 122.493 | -59.1934 | 63.1439 | -0.092392 | 0.00091438 | 0.069186 | 0.069133 | 0.0048855 | -0.00090143 | 0.0047424 | -0.028778 | 0.0012411 | 0.066415 | 0.062551 | 0.00045928 | -0.0024075 | 0.0038601 | -0.013733 | 63.0936 | -59.3445 | 62.9791 | -59.0698 | -59.491 | -59.436 | -0.018311 | 63.0936 | -59.1888 | 63.1073 | -59.2346 | -59.3536 | -59.0698 | -0.15106 | -0.050354 | -0.16479 | -0.11444 | 0.1236 | 0.27466 | -0.24261 | -0.29755 | -0.091553 | -0.14648 | 0.004577 | -0.050354 | -0.036621 | 0.013733 | -0.041199 | -0.045776 | 0.1236 | -0.16022 | 0.11902 | -0.16479
164.7812 | 122.5388 | -59.2026 | 63.1256 | -0.016098 | -0.0014356 | 0.068545 | 0.068064 | 0.00018591 | -0.001115 | -0.00017099 | -0.0013123 | 0.00060029 | 0.06492 | 0.065755 | -0.0012497 | 0.00036966 | 1.4916e-05 | -0.018311 | 63.1073 | -59.3491 | 63.002 | -59.0652 | -59.4772 | -59.4315 | 0.009155 | 63.0478 | -59.2346 | 63.1119 | -59.2163 | -59.3399 | -59.0607 | -0.14648 | -0.01831 | -0.1236 | -0.10529 | 0.13733 | 0.28381 | -0.22888 | -0.27466 | -0.082398 | -0.12817 | -0.032043 | -0.077819 | -0.013733 | 0.064086 | -0.013733 | 0.01831 | 0.14191 | -0.13733 | 0.17395 | -0.10529
165.7812 | 122.5159 | -59.2072 | 63.1302 | -0.077133 | 0.00048698 | 0.067264 | 0.069987 | 0.00082671 | -0.0013288 | -0.0016662 | -0.022675 | 0.0016685 | 0.063211 | 0.063619 | -0.0010361 | -0.0017667 | -0.00041228 | -0.018311 | 63.1073 | -59.3445 | 63.0066 | -59.0881 | -59.4269 | -59.4086 | 0.018311 | 63.0936 | -59.2392 | 63.1165 | -59.2072 | -59.3353 | -59.0744 | -0.13733 | -0.022888 | -0.1236 | -0.10071 | 0.11902 | 0.25635 | -0.20142 | -0.21973 | -0.064087 | -0.082398 | -0.032044 | -0.036621 | -0.013733 | 0.022888 | 0 | 0.032044 | 0.13275 | -0.12817 | 0.16479 | -0.09613
166.7812 | 122.4976 | -59.23 | 63.121 | -0.10002 | -0.00058102 | 0.067264 | 0.069133 | 0.0023221 | 0.0023028 | -0.00081179 | -0.0074159 | 0.00038669 | 0.065561 | 0.057424 | 3.2078e-05 | -0.0019803 | 0.0027921 | -0.041199 | 63.0798 | -59.3491 | 62.9654 | -59.0836 | -59.5047 | -59.4543 | -0.022888 | 63.089 | -59.2484 | 63.0753 | -59.2529 | -59.3811 | -59.079 | -0.11902 | -0.041199 | -0.15564 | -0.11444 | 0.14648 | 0.2655 | -0.2243 | -0.27466 | -0.10529 | -0.15564 | -0.01831 | -0.032044 | -0.045777 | -0.013733 | -0.022888 | -0.004578 | 0.15106 | -0.15106 | 0.16937 | -0.13275
167.7812 | 122.4747 | -59.23 | 63.121 | -0.10002 | -0.0016492 | 0.068118 | 0.068064 | -0.0021641 | -0.00047423 | -0.0010254 | -0.03183 | 0.0029501 | 0.065775 | 0.070241 | -0.00060892 | -0.00069854 | -0.0019077 | -0.07782 | 63.0981 | -59.3582 | 62.9242 | -59.1522 | -59.5367 | -59.5001 | -0.14191 | 63.0753 | -59.2667 | 63.0386 | -59.285 | -59.4131 | -59.1156 | -0.12817 | -0.022888 | -0.19684 | -0.17395 | 0.07782 | 0.20599 | -0.27008 | -0.3067 | -0.14191 | -0.17853 | -0.036621 | -0.045777 | -0.082398 | -0.036621 | -0.054931 | -0.01831 | 0.11444 | -0.18311 | 0.15106 | -0.14648
168.7812 | 122.5021 | -59.198 | 63.121 | -0.092392 | 0.002196 | 0.066195 | 0.068705 | -0.00066869 | 0.001662 | -0.00081179 | -0.028778 | 0.00081389 | 0.065134 | 0.063619 | -0.0010361 | -5.7742e-05 | 0.0015103 | -0.064087 | 63.1256 | -59.3445 | 62.9425 | -59.1202 | -59.5184 | -59.4727 | -0.11902 | 63.0524 | -59.2621 | 63.0524 | -59.2941 | -59.4177 | -59.1019 | -0.14648 | 0.004577 | -0.17853 | -0.18311 | 0.07782 | 0.2243 | -0.27466 | -0.32044 | -0.12817 | -0.17395 | -0.064087 | -0.068665 | -0.068665 | 0 | -0.09613 | -0.032043 | 0.09613 | -0.21973 | 0.16022 | -0.15564
169.7812 | 122.5021 | -59.1843 | 63.121 | -0.11528 | 0.00027338 | 0.067477 | 0.06721 | -0.0010959 | 0.00016657 | -0.0025208 | -0.016571 | 0.0012411 | 0.064706 | 0.064687 | -0.0012497 | -0.0017667 | -0.0012667 | -0.027466 | 63.0936 | -59.3308 | 62.9974 | -59.0561 | -59.4589 | -59.4131 | -0.004578 | 63.1027 | -59.2575 | 63.1119 | -59.2072 | -59.3308 | -59.0744 | -0.14648 | -0.027466 | -0.1236 | -0.09613 | 0.12817 | 0.27466 | -0.22888 | -0.27466 | -0.082397 | -0.12817 | -0.073242 | -0.018311 | -0.009156 | 0.009155 | -0.022888 | 0.050354 | 0.10986 | -0.14648 | 0.1831 | -0.073243
170.7812 | 122.5204 | -59.2117 | 63.121 | -0.084763 | 0.00091438 | 0.068118 | 0.068492 | -0.0030185 | -0.0021832 | 0.0011108 | -0.016571 | 0.0023093 | 0.066843 | 0.068746 | -0.0029587 | -0.00091214 | -0.00083948 | -0.009155 | 63.0753 | -59.3353 | 62.9883 | -59.0561 | -59.5001 | -59.5001 | 0.022888 | 63.0661 | -59.2346 | 63.1165 | -59.1888 | -59.3353 | -59.079 | -0.1236 | -0.045777 | -0.13275 | -0.086975 | 0.15564 | 0.27924 | -0.28839 | -0.28839 | -0.16479 | -0.16479 | -0.022888 | -0.054932 | -0.004578 | 0.050354 | 0.022888 | 0.045776 | 0.13275 | -0.1236 | 0.15564 | -0.10071
171.7812 | 122.5159 | -59.2026 | 63.1302 | -0.069504 | 0.001128 | 0.068545 | 0.067637 | 0.0029629 | 0.0010212 | -0.0018798 | -0.0043641 | 0.0020957 | 0.064066 | 0.064687 | 0.00045928 | -0.0021939 | -0.00062588 | -0.009155 | 63.1119 | -59.3445 | 62.9974 | -59.0424 | -59.4635 | -59.4131 | 0.032043 | 63.0936 | -59.2392 | 63.1531 | -59.1843 | -59.3262 | -59.0149 | -0.14191 | -0.018311 | -0.13275 | -0.11444 | 0.16022 | 0.30212 | -0.21057 | -0.26093 | -0.068665 | -0.11902 | -0.036621 | -0.036621 | 0.022888 | 0.059509 | 0.018311 | 0.054932 | 0.18768 | -0.1236 | 0.2243 | -0.086975
172.7812 | 122.5388 | -59.1797 | 63.1348 | -0.061874 | -0.0010082 | 0.067477 | 0.068064 | -0.0030185 | -0.00047423 | 0.00047001 | 0.023102 | -0.00068151 | 0.062997 | 0.065969 | -0.0010361 | -0.0028347 | -0.0010531 | -0.027466 | 63.1348 | -59.3353 | 62.9791 | -59.0881 | -59.4543 | -59.4223 | 0.018311 | 63.1119 | -59.2072 | 63.1165 | -59.1934 | -59.317 | -59.0378 | -0.15564 | 0 | -0.15564 | -0.15564 | 0.091552 | 0.24719 | -0.24262 | -0.27466 | -0.086975 | -0.11902 | -0.027466 | -0.022889 | -0.018311 | 0.004578 | -0.013733 | 0.013733 | 0.14191 | -0.13733 | 0.16937 | -0.10986
173.7812 | 122.5296 | -59.1751 | 63.1302 | -0.016098 | 0.0015552 | 0.066836 | 0.067637 | 0.00018591 | -0.0015424 | -0.00081179 | 0.016998 | 0.0010275 | 0.064493 | 0.059774 | -0.0012497 | -0.0034757 | -0.0016941 | 0.004578 | 63.121 | -59.3216 | 62.9974 | -59.0607 | -59.4681 | -59.4086 | 0.041199 | 63.1027 | -59.2209 | 63.1027 | -59.2163 | -59.3079 | -59.024 | -0.14648 | -0.009155 | -0.13275 | -0.1236 | 0.11444 | 0.26093 | -0.23346 | -0.29297 | -0.086975 | -0.14649 | -0.045776 | -0.027466 | -0.027466 | 0 | -0.041199 | 0.004577 | 0.15106 | -0.13275 | 0.19684 | -0.086975
174.7812 | 122.5159 | -59.1888 | 63.1119 | -0.038986 | 0.00048698 | 0.067264 | 0.067851 | 0.0025357 | 0.00080757 | 0.0019654 | 0.010895 | 0.0023093 | 0.064279 | 0.056356 | 0.00045928 | -0.0011257 | -0.00041228 | 0 | 63.0524 | -59.3399 | 63.0066 | -59.0744 | -59.4635 | -59.4864 | 0.036621 | 63.0707 | -59.2529 | 63.121 | -59.1843 | -59.3536 | -59.0469 | -0.15106 | -0.059509 | -0.10529 | -0.045776 | 0.11444 | 0.2655 | -0.29755 | -0.27466 | -0.14648 | -0.1236 | -0.064087 | -0.041198 | 0.009156 | 0.050354 | 0.004578 | 0.068665 | 0.14191 | -0.16479 | 0.20599 | -0.10071
175.7812 | 122.4976 | -59.198 | 63.1302 | -0.12291 | -0.00015382 | 0.06705 | 0.067637 | -0.0010959 | 0.00016657 | -0.00017099 | -0.022675 | 0.00081389 | 0.067056 | 0.060415 | -0.0014633 | -0.0017667 | 0.0036465 | -0.009155 | 63.1302 | -59.3399 | 62.9974 | -59.0652 | -59.4727 | -59.4131 | 0.045776 | 63.0615 | -59.2163 | 63.1073 | -59.2072 | -59.2758 | -59.0149 | -0.14191 | 0 | -0.13275 | -0.13275 | 0.13275 | 0.27466 | -0.21515 | -0.27466 | -0.073242 | -0.13275 | -0.018311 | -0.068665 | -0.022888 | 0.045777 | -0.009155 | 0.009156 | 0.1831 | -0.07782 | 0.20142 | -0.059509
176.7812 | 122.5433 | -59.1888 | 63.121 | -0.077133 | -0.00036742 | 0.067477 | 0.068492 | 0.0033901 | 0.0018756 | -0.0018798 | -0.0074159 | 0.0012411 | 0.065775 | 0.068105 | 0.00045928 | -0.00069854 | 0.0023647 | 0.004578 | 63.1073 | -59.3308 | 62.9929 | -59.0515 | -59.4589 | -59.4269 | 0.027466 | 63.0753 | -59.2209 | 63.2126 | -59.3033 | -59.4223 | -59.0973 | -0.14191 | -0.013733 | -0.12817 | -0.11444 | 0.13733 | 0.27924 | -0.23804 | -0.27008 | -0.09613 | -0.12817 | -0.032043 | -0.045777 | 0.091552 | 0.13733 | -0.11444 | -0.082398 | 0.091553 | -0.23346 | 0.1236 | -0.20142
177.7812 | 122.4976 | -59.2026 | 63.1256 | -0.084763 | 0.00027338 | 0.067904 | 0.066996 | -0.0036593 | 0.0020892 | 0.00047001 | -0.0074159 | 0.0025229 | 0.065775 | 0.067251 | -0.0016769 | -0.0021939 | -0.0010531 | -0.004578 | 63.121 | -59.3445 | 62.9974 | -59.0515 | -59.4727 | -59.4177 | 0.036621 | 63.0936 | -59.2484 | 63.121 | -59.1705 | -59.317 | -59.0332 | -0.14191 | -0.004577 | -0.12817 | -0.1236 | 0.15106 | 0.29297 | -0.21515 | -0.27008 | -0.073243 | -0.12817 | -0.045776 | -0.032043 | -0.004577 | 0.027466 | 0.032044 | 0.07782 | 0.16937 | -0.11444 | 0.21515 | -0.068665
178.7812 | 122.4701 | -59.2392 | 63.1302 | -0.10002 | 0.00091438 | 0.065341 | 0.066569 | -0.0013095 | -0.00026063 | -0.00059819 | -0.0074159 | -0.0011087 | 0.065988 | 0.064901 | 0.00045928 | -0.00048494 | -0.0016941 | -0.073242 | 63.0936 | -59.3536 | 62.9196 | -59.1431 | -59.5505 | -59.4727 | -0.10986 | 63.0661 | -59.2621 | 63.0478 | -59.2804 | -59.3948 | -59.0927 | -0.11444 | -0.036621 | -0.21057 | -0.17395 | 0.096131 | 0.21057 | -0.23346 | -0.31128 | -0.11902 | -0.19684 | -0.022888 | -0.064087 | -0.082397 | -0.01831 | -0.041199 | -0.018311 | 0.14648 | -0.15564 | 0.16937 | -0.13275
179.7812 | 122.5021 | -59.1797 | 63.1348 | -0.092392 | 0.00070078 | 0.067477 | 0.068919 | 0.0029629 | -0.00026063 | -0.00038459 | 0.0078429 | 0.00038669 | 0.066629 | 0.062124 | -0.0014633 | -0.00091214 | 0.00065592 | -0.027466 | 63.1439 | -59.3216 | 62.9883 | -59.111 | -59.4727 | -59.4406 | -0.013733 | 63.1073 | -59.2255 | 63.0981 | -59.198 | -59.3216 | -59.0378 | -0.14191 | 0.009155 | -0.14648 | -0.15564 | 0.068664 | 0.21057 | -0.26093 | -0.29297 | -0.11902 | -0.15106 | -0.045777 | -0.027466 | -0.036621 | -0.009155 | -0.018311 | 0.027466 | 0.14191 | -0.14191 | 0.18768 | -0.09613
180.7812 | 122.5204 | -59.166 | 63.1165 | -0.054245 | -0.0027172 | 0.067904 | 0.06721 | -2.7692e-05 | 0.001662 | -0.0014526 | -0.040985 | 0.00038669 | 0.064066 | 0.064901 | 0.00045928 | 0.00036966 | -0.0029757 | 0.018311 | 63.1073 | -59.3399 | 63.0066 | -59.0378 | -59.4269 | -59.4086 | 0.064087 | 63.1073 | -59.2255 | 63.1165 | -59.1705 | -59.317 | -59.0424 | -0.17395 | -0.009155 | -0.10986 | -0.10071 | 0.12817 | 0.30212 | -0.24261 | -0.26093 | -0.068664 | -0.086975 | -0.059509 | -0.009155 | 0 | 0.009155 | -0.004577 | 0.054932 | 0.1236 | -0.15106 | 0.18311 | -0.091553
181.7812 | 122.5021 | -59.1888 | 63.1302 | -0.092392 | -0.00058102 | 0.069827 | 0.070201 | 0.0016813 | 0.0014484 | -0.00081179 | 0.013947 | -0.0011087 | 0.064706 | 0.062551 | -0.0021041 | -0.0013393 | 0.00065592 | -0.009155 | 63.1073 | -59.3399 | 62.9974 | -59.0607 | -59.4727 | -59.4635 | 0.036621 | 63.089 | -59.2484 | 63.1119 | -59.2026 | -59.3445 | -59.0424 | -0.15106 | -0.022888 | -0.13275 | -0.10986 | 0.12817 | 0.27924 | -0.27466 | -0.28381 | -0.1236 | -0.13275 | -0.059509 | -0.041199 | -0.018311 | 0.022888 | -0.013733 | 0.045776 | 0.14648 | -0.15564 | 0.20599 | -0.09613
182.7812 | 122.5616 | -59.1888 | 63.1531 | -0.061874 | -0.00079462 | 0.06705 | 0.070414 | 0.0053129 | 0.001662 | 0.00025641 | 0.0047913 | -0.00068151 | 0.065134 | 0.058706 | -0.0010361 | 0.00036966 | -0.0031895 | -0.004578 | 63.1302 | -59.3262 | 62.9974 | -59.0698 | -59.4498 | -59.404 | 0.027466 | 63.089 | -59.2255 | 63.1165 | -59.1888 | -59.2987 | -59.0424 | -0.13733 | -0.022888 | -0.15564 | -0.13275 | 0.11902 | 0.25635 | -0.21515 | -0.26093 | -0.07782 | -0.1236 | -0.036621 | -0.064087 | -0.036621 | 0.027466 | 0 | 0.036621 | 0.14648 | -0.10986 | 0.18311 | -0.073242
183.7812 | 122.4701 | -59.1797 | 63.1348 | -0.077133 | 0.0032642 | 0.068332 | 0.067424 | 0.0012539 | 0.0014484 | -0.0025208 | -0.03183 | 0.0027365 | 0.065134 | 0.06426 | -0.00039512 | -0.0011257 | 1.4916e-05 | -0.009155 | 63.0798 | -59.3582 | 62.9929 | -59.0881 | -59.5047 | -59.4498 | -0.013733 | 63.089 | -59.2392 | 63.0753 | -59.2392 | -59.3674 | -59.0744 | -0.17853 | -0.054932 | -0.14191 | -0.086975 | 0.091552 | 0.27008 | -0.27008 | -0.32501 | -0.091553 | -0.14648 | -0.05951 | -0.045777 | -0.05951 | -0.013733 | -0.05951 | 0 | 0.10529 | -0.18768 | 0.16479 | -0.12817
184.7812 | 122.5342 | -59.2209 | 63.1302 | -0.092392 | 0.00070078 | 0.065555 | 0.067851 | -0.00045509 | 0.0012348 | -0.00081179 | -0.019623 | 0.0020957 | 0.067484 | 0.070028 | -0.00060892 | -0.0028347 | 0.00022852 | -0.004578 | 63.0936 | -59.3308 | 63.0066 | -59.0469 | -59.4589 | -59.3948 | 0.032043 | 63.089 | -59.2438 | 63.0844 | -59.2758 | -59.3079 | -59.0561 | -0.10986 | -0.036621 | -0.1236 | -0.086975 | 0.17395 | 0.28381 | -0.17395 | -0.23804 | -0.064086 | -0.12817 | -0.022888 | -0.041199 | -0.045776 | -0.004577 | -0.054932 | -0.032044 | 0.16479 | -0.086975 | 0.18768 | -0.064087
185.7812 | 122.4197 | -59.1843 | 63.1119 | -0.069504 | 0.00070078 | 0.070254 | 0.067851 | 0.0038175 | 0.00016657 | 0.00025641 | -0.022675 | 0.0018821 | 0.065988 | 0.062124 | -0.0014633 | -0.0019803 | 0.0010831 | 0.013733 | 63.1027 | -59.3399 | 62.9883 | -59.0515 | -59.4498 | -59.4177 | 0.032043 | 63.0981 | -59.2163 | 63.1165 | -59.2026 | -59.3079 | -58.9737 | -0.15564 | -0.009155 | -0.1236 | -0.11444 | 0.13275 | 0.28839 | -0.23346 | -0.2655 | -0.07782 | -0.10986 | -0.032044 | -0.013732 | 0.004578 | 0.01831 | -0.018311 | 0.013733 | 0.21057 | -0.1236 | 0.24262 | -0.091552
186.7812 | 122.5571 | -59.1705 | 63.1393 | -0.023727 | -0.001222 | 0.0694 | 0.068278 | 0.0010403 | 0.00059397 | 4.2808e-05 | 0.02005 | -0.00046791 | 0.064706 | 0.066396 | -0.00039512 | -0.0017667 | -0.00083948 | -0.018311 | 63.1393 | -59.3353 | 62.9837 | -59.0927 | -59.436 | -59.4177 | 0.013733 | 63.1119 | -59.2209 | 63.1073 | -59.1843 | -59.3124 | -59.0378 | -0.16479 | 0 | -0.15564 | -0.15564 | 0.07782 | 0.24262 | -0.24719 | -0.2655 | -0.082398 | -0.10071 | -0.050354 | -0.027466 | -0.032043 | -0.004577 | -0.013733 | 0.036621 | 0.13275 | -0.14191 | 0.1831 | -0.091553
187.7812 | 122.5021 | -59.1705 | 63.1393 | -0.016098 | 0.00091438 | 0.0694 | 0.068278 | -0.0017367 | 0.00038017 | -0.0010254 | 0.0017395 | -0.00025411 | 0.063211 | 0.066396 | -0.00082252 | -0.00048494 | -0.0027621 | -0.004578 | 63.089 | -59.317 | 62.9974 | -59.0607 | -59.4589 | -59.4131 | 0.036621 | 63.0753 | -59.23 | 63.1348 | -59.1797 | -59.3262 | -59.0424 | -0.14648 | -0.050354 | -0.14191 | -0.091552 | 0.10986 | 0.25635 | -0.24262 | -0.28839 | -0.09613 | -0.14191 | -0.05951 | -0.064087 | -0.004577 | 0.05951 | -0.009155 | 0.050355 | 0.12817 | -0.15564 | 0.18768 | -0.09613
188.7812 | 122.5204 | -59.1751 | 63.121 | -0.36705 | -0.00058102 | 0.066623 | 0.068492 | 0.0036039 | -0.00090143 | -0.001239 | -0.028778 | 0.0012411 | 0.06492 | 0.068105 | -0.00082252 | -0.0021939 | -0.0014805 | -0.004578 | 63.0753 | -59.3399 | 62.9837 | -59.0515 | -59.4772 | -59.436 | 0.018311 | 63.089 | -59.2209 | 63.0981 | -59.2072 | -59.3262 | -59.0286 | -0.1648 | -0.045777 | -0.13733 | -0.091552 | 0.1236 | 0.28839 | -0.26093 | -0.30212 | -0.09613 | -0.13733 | -0.045776 | -0.032044 | -0.022888 | 0.009156 | -0.032043 | 0.013733 | 0.14648 | -0.15106 | 0.19226 | -0.10529
189.7812 | 122.493 | -59.2072 | 63.1393 | -0.11528 | 0.00091438 | 0.066836 | 0.066569 | 0.0040311 | 0.001662 | -0.0016662 | -0.022675 | 0.0025229 | 0.064493 | 0.059774 | -0.00082252 | -0.0028347 | 0.0017239 | -0.022888 | 63.089 | -59.317 | 63.0066 | -59.0515 | -59.4589 | -59.3994 | 0.013733 | 63.0798 | -59.2209 | 63.1073 | -59.2026 | -59.3399 | -59.0469 | -0.10986 | -0.050354 | -0.13275 | -0.082397 | 0.15564 | 0.2655 | -0.19226 | -0.25177 | -0.082397 | -0.14191 | -0.013733 | -0.059509 | -0.032043 | 0.027466 | 0.004577 | 0.01831 | 0.16022 | -0.13275 | 0.17395 | -0.11902
190.7812 | 122.5342 | -59.1888 | 63.1165 | -0.092392 | -0.0010082 | 0.068545 | 0.067851 | -0.0015231 | -0.0013288 | -0.0010254 | -0.0074159 | -4.0508e-05 | 0.06492 | 0.056356 | -0.0025315 | -0.00069854 | -0.00019868 | 0 | 63.0798 | -59.3216 | 62.9929 | -59.0469 | -59.436 | -59.3948 | 0.027466 | 63.0707 | -59.2438 | 63.1256 | -59.2072 | -59.3399 | -59.0698 | -0.13275 | -0.036621 | -0.1236 | -0.086975 | 0.14191 | 0.27466 | -0.20599 | -0.24719 | -0.073242 | -0.11444 | -0.054931 | -0.045776 | 0.009155 | 0.054931 | -0.01831 | 0.036621 | 0.11902 | -0.15106 | 0.17395 | -0.096131
191.7812 | 122.5296 | -59.1843 | 63.1165 | -0.10002 | -0.00036742 | 0.068545 | 0.06956 | -0.0025913 | 0.0031574 | -0.0016662 | 0.0047913 | 0.00081389 | 0.064706 | 0.063833 | -0.0012497 | 0.0012241 | -0.00019868 | -0.013733 | 63.0844 | -59.3399 | 62.9105 | -58.9554 | -59.491 | -59.4452 | 0.022888 | 63.0844 | -59.2346 | 63.1165 | -59.2072 | -59.3124 | -59.0469 | -0.15564 | -0.032043 | -0.20599 | -0.17395 | 0.22888 | 0.38452 | -0.26092 | -0.3067 | -0.10528 | -0.15106 | -0.050354 | -0.032043 | 0 | 0.032043 | -0.022888 | 0.027466 | 0.13733 | -0.12817 | 0.18768 | -0.07782
192.7812 | 122.5433 | -59.1888 | 63.1302 | -0.061874 | 0.0015552 | 0.066623 | 0.070201 | -0.0017367 | 0.0025166 | 0.0030334 | -0.034882 | 0.0014547 | 0.064493 | 0.062124 | -0.0012497 | -0.0024075 | 1.4916e-05 | -0.013733 | 63.089 | -59.3445 | 62.915 | -59.1476 | -59.3948 | -59.4681 | 0.036621 | 63.0707 | -59.2392 | 63.0981 | -59.2072 | -59.3262 | -59.0195 | -0.15564 | -0.041199 | -0.21515 | -0.17395 | 0.041199 | 0.19684 | -0.27924 | -0.20599 | -0.1236 | -0.050354 | -0.050354 | -0.059509 | -0.032043 | 0.027466 | -0.01831 | 0.032044 | 0.16937 | -0.13733 | 0.21973 | -0.086975
193.7812 | 122.5021 | -59.2026 | 63.1302 | -0.084763 | -0.0010082 | 0.069827 | 0.067424 | -2.7692e-05 | 0.00059397 | -0.001239 | 0.0017395 | 0.0014547 | 0.063638 | 0.065969 | 3.2078e-05 | -0.0011257 | 0.0027921 | -0.009155 | 63.1027 | -59.3353 | 62.9974 | -59.0652 | -59.4315 | -59.3903 | 0.036621 | 63.0936 | -59.2346 | 63.1165 | -59.1843 | -59.3079 | -59.0424 | -0.13275 | -0.027466 | -0.13275 | -0.10529 | 0.13733 | 0.27008 | -0.18768 | -0.22888 | -0.054932 | -0.096131 | -0.032043 | -0.036621 | -0.013733 | 0.022888 | 0.018311 | 0.050354 | 0.16022 | -0.10529 | 0.19226 | -0.073242
194.7812 | 122.5388 | -59.2026 | 63.1256 | -0.054245 | 5.9776e-05 | 0.065768 | 0.068492 | -0.0023777 | 0.00038017 | 0.00089721 | 0.038361 | 0.0016685 | 0.065561 | 0.069387 | 0.00045928 | -0.0024075 | -0.00062588 | -0.018311 | 63.0753 | -59.3491 | 63.002 | -59.0469 | -59.5459 | -59.4543 | 0.018311 | 63.0981 | -59.2392 | 63.1302 | -59.1888 | -59.3491 | -59.0515 | -0.14648 | -0.050354 | -0.1236 | -0.073242 | 0.15564 | 0.30212 | -0.25177 | -0.34332 | -0.10529 | -0.19684 | -0.036621 | -0.027465 | 0.004578 | 0.032043 | 0.013733 | 0.050354 | 0.15106 | -0.14648 | 0.18768 | -0.10986
195.7812 | 122.5159 | -59.1888 | 63.1119 | -0.061874 | 0.0019824 | 0.067904 | 0.067424 | -0.0017367 | -4.7032e-05 | -0.001239 | -0.0043641 | 0.0012411 | 0.066629 | 0.065114 | -0.0016769 | -0.00069854 | 0.0019375 | -0.018311 | 63.0981 | -59.317 | 63.002 | -59.0561 | -59.4589 | -59.4955 | 0.041199 | 63.0661 | -59.2484 | 63.1256 | -59.1843 | -59.3262 | -59.0332 | -0.12817 | -0.013732 | -0.10986 | -0.096131 | 0.13275 | 0.26093 | -0.3067 | -0.27008 | -0.17853 | -0.14191 | -0.059509 | -0.045776 | 0.013733 | 0.059509 | 0.004578 | 0.064087 | 0.15564 | -0.13733 | 0.21515 | -0.07782
196.7812 | 122.4976 | -59.2163 | 63.121 | -0.10002 | 0.0015552 | 0.067691 | 0.068705 | -0.0017367 | -0.00047423 | 0.003247 | -0.0074159 | 0.0018821 | 0.064706 | 0.061269 | 3.2078e-05 | -0.0015529 | 0.00086952 | -0.022888 | 63.0936 | -59.3582 | 62.9654 | -59.0744 | -59.4955 | -59.4315 | -0.022888 | 63.0753 | -59.2575 | 63.0844 | -59.2392 | -59.3628 | -59.0607 | -0.14191 | -0.027466 | -0.15564 | -0.12817 | 0.14191 | 0.28381 | -0.21515 | -0.27923 | -0.073243 | -0.13733 | -0.041198 | -0.045777 | -0.036621 | 0.009156 | -0.022888 | 0.01831 | 0.15564 | -0.14648 | 0.19684 | -0.10529
197.7812 | 122.4838 | -59.1751 | 63.1393 | -0.00083923 | 0.00091438 | 0.068759 | 0.067637 | -0.0010959 | -0.0015424 | -0.00017099 | -0.022675 | 0.00060029 | 0.065134 | 0.064901 | -0.0014633 | -0.0011257 | 0.0023647 | -0.027466 | 63.0936 | -59.3674 | 62.9791 | -59.0652 | -59.4864 | -59.4315 | -0.009155 | 63.0707 | -59.2529 | 63.0798 | -59.23 | -59.3628 | -59.0744 | -0.19226 | -0.045776 | -0.16022 | -0.11444 | 0.10986 | 0.30212 | -0.25635 | -0.31128 | -0.064087 | -0.11902 | -0.07782 | -0.068664 | -0.059509 | 0.009155 | -0.054932 | 0.022888 | 0.10071 | -0.18768 | 0.17853 | -0.10986
198.7812 | 122.4838 | -59.2438 | 63.1027 | -0.12291 | 0.00048698 | 0.068118 | 0.067851 | 0.0010403 | 0.00038017 | -0.00059819 | 0.11465 | 0.00081389 | 0.065775 | 0.05956 | -0.00060892 | -0.00091214 | 1.4916e-05 | -0.11444 | 63.0661 | -59.3674 | 62.9196 | -59.1568 | -59.9762 | -59.642 | -0.14191 | 63.0798 | -59.2484 | 63.0615 | -59.2941 | -59.3994 | -59.1064 | -0.1236 | -0.036621 | -0.1831 | -0.14648 | 0.086975 | 0.21057 | -0.39825 | -0.73242 | -0.27466 | -0.60883 | -0.004578 | -0.022888 | -0.041199 | -0.018311 | -0.050354 | -0.045776 | 0.13733 | -0.15564 | 0.14191 | -0.15106
199.7812 | 122.4838 | -59.1888 | 63.121 | -0.084763 | 0.00070078 | 0.070254 | 0.06956 | -0.0017367 | -4.7032e-05 | -0.001239 | -0.028778 | 0.0010275 | 0.063211 | 0.059987 | 0.00024568 | -0.0032619 | 0.0042875 | -0.032043 | 63.089 | -59.3674 | 62.9745 | -59.0652 | -59.4772 | -59.4269 | -0.013733 | 63.0707 | -59.2484 | 63.0936 | -59.2209 | -59.3445 | -59.0652 | -0.17853 | -0.032044 | -0.14648 | -0.11444 | 0.1236 | 0.30212 | -0.23804 | -0.28839 | -0.059509 | -0.10986 | -0.059509 | -0.050354 | -0.027466 | 0.022888 | -0.032043 | 0.027466 | 0.1236 | -0.15564 | 0.1831 | -0.09613
200.7812 | 122.5296 | -59.1934 | 63.1165 | 0.01442 | -0.00015382 | 0.065982 | 0.068278 | -0.00066869 | 0.00080757 | 0.00047001 | -0.025726 | 0.00081389 | 0.064066 | 0.06426 | -0.0029587 | -0.00069854 | 0.0010831 | 0.009155 | 63.0981 | -59.3033 | 62.9791 | -59.0515 | -59.4406 | -59.4177 | 0.036621 | 63.0753 | -59.2392 | 63.0432 | -59.1385 | -59.3536 | -59.0469 | -0.10986 | -0.01831 | -0.13733 | -0.11902 | 0.14191 | 0.25177 | -0.2243 | -0.24719 | -0.11444 | -0.13733 | -0.045777 | -0.041199 | -0.073242 | -0.032043 | 0.054931 | 0.10071 | 0.14648 | -0.16022 | 0.19226 | -0.11444
201.7812 | 122.4518 | -59.2209 | 63.121 | -0.084763 | 0.00070078 | 0.069827 | 0.067637 | 0.0038175 | -0.00090143 | -0.001239 | -0.03183 | -4.0508e-05 | 0.06492 | 0.063833 | -0.0010361 | -0.0021939 | 0.00086952 | -0.07782 | 63.089 | -59.3536 | 62.9379 | -59.1248 | -59.5367 | -59.4681 | -0.091553 | 63.0936 | -59.2438 | 63.0478 | -59.2758 | -59.3994 | -59.1202 | -0.13275 | -0.032044 | -0.18311 | -0.15106 | 0.09613 | 0.22888 | -0.24719 | -0.31586 | -0.11444 | -0.18311 | -0.022888 | -0.027466 | -0.073242 | -0.045776 | -0.054932 | -0.032044 | 0.10071 | -0.17853 | 0.1236 | -0.15564
202.7812 | 122.4976 | -59.1797 | 63.1119 | 0.0067903 | 0.0041186 | 0.068759 | 0.067851 | 0.0018949 | 0.00016657 | 0.00068361 | -0.040985 | 0.00060029 | 0.063638 | 0.069387 | -0.0010361 | -5.7742e-05 | 0.0023647 | -0.009155 | 63.089 | -59.3491 | 62.9837 | -59.0744 | -59.4727 | -59.4498 | -0.004578 | 63.0844 | -59.2346 | 63.0936 | -59.2072 | -59.3445 | -59.0698 | -0.16937 | -0.022888 | -0.12817 | -0.10528 | 0.10529 | 0.27466 | -0.27008 | -0.29297 | -0.10071 | -0.1236 | -0.054932 | -0.027465 | -0.01831 | 0.009155 | -0.027466 | 0.027466 | 0.10986 | -0.16479 | 0.1648 | -0.10986
203.7812 | 122.612 | -59.1934 | 63.0661 | -0.084763 | -0.001222 | 0.067904 | 0.066783 | -0.0015231 | 0.0010212 | 0.0036744 | -0.013519 | 0.00081389 | 0.065134 | 0.065114 | 3.2078e-05 | -0.0013393 | 0.00065592 | -0.10529 | 63.0524 | -59.3445 | 63.002 | -59.0607 | -59.4727 | -59.4177 | -0.032043 | 63.0112 | -59.2209 | 63.1485 | -59.166 | -59.2896 | -58.9874 | -0.15106 | -0.013733 | -0.064087 | -0.050354 | 0.13275 | 0.28381 | -0.2243 | -0.27924 | -0.073243 | -0.12817 | -0.027466 | -0.054932 | 0.082398 | 0.13733 | 0.027465 | 0.054931 | 0.20599 | -0.096131 | 0.23346 | -0.068665
204.7812 | 122.4884 | -59.2026 | 63.1165 | -0.077133 | 0.00027338 | 0.067691 | 0.067637 | 0.00018591 | -0.001115 | 0.0028198 | -0.037933 | 0.00060029 | 0.065347 | 0.063405 | 0.00024568 | -0.0017667 | 0.0025785 | -0.064087 | 63.0936 | -59.3262 | 62.9471 | -59.1248 | -59.5276 | -59.4589 | -0.091553 | 63.0707 | -59.2392 | 63.0615 | -59.2804 | -59.4086 | -59.1064 | -0.1236 | -0.022888 | -0.16937 | -0.14648 | 0.07782 | 0.20142 | -0.25635 | -0.32501 | -0.13275 | -0.20142 | -0.036621 | -0.045776 | -0.054932 | -0.009156 | -0.07782 | -0.041199 | 0.096131 | -0.20599 | 0.13275 | -0.16937
205.7812 | 122.4518 | -59.2392 | 63.1073 | -0.046616 | 5.9776e-05 | 0.069186 | 0.068705 | -0.0038729 | 0.00059397 | 0.00089721 | -0.044037 | 0.00017309 | 0.064066 | 0.06661 | -0.00082252 | 0.00036966 | 0.0051419 | -0.059509 | 63.1027 | -59.3353 | 62.9471 | -59.111 | -59.5276 | -59.4498 | -0.10986 | 63.0524 | -59.2438 | 63.0524 | -59.285 | -59.3948 | -59.0927 | -0.09613 | -0.004578 | -0.16022 | -0.15564 | 0.12817 | 0.2243 | -0.21057 | -0.28839 | -0.11444 | -0.19226 | -0.004577 | -0.054932 | -0.054932 | 0 | -0.045776 | -0.041199 | 0.14648 | -0.15564 | 0.15106 | -0.15106
206.7812 | 122.4655 | -59.2209 | 63.1073 | -0.061874 | -0.0014356 | 0.070254 | 0.069346 | 0.0027493 | 0.00038017 | 0.00025641 | -0.037933 | 0.0010275 | 0.06492 | 0.060628 | -0.0018905 | -0.0011257 | -0.0021213 | -0.064087 | 63.0844 | -59.3445 | 62.9517 | -59.1064 | -59.5139 | -59.4818 | -0.07782 | 63.0707 | -59.2255 | 63.0569 | -59.2804 | -59.404 | -59.0927 | -0.1236 | -0.022888 | -0.15564 | -0.13275 | 0.11444 | 0.23804 | -0.26093 | -0.29297 | -0.13733 | -0.16937 | -0.004578 | -0.036621 | -0.050354 | -0.013733 | -0.05951 | -0.054932 | 0.12817 | -0.18311 | 0.13275 | -0.17853
207.7812 | 122.4792 | -59.1751 | 63.1119 | -0.092392 | -0.00079462 | 0.068973 | 0.068492 | -0.00024129 | -0.0026104 | 0.00047001 | -0.047089 | 0.0025229 | 0.064279 | 0.063192 | 0.00088648 | -0.00091214 | -0.00062588 | -0.022888 | 63.0707 | -59.3445 | 62.9883 | -59.0744 | -59.4635 | -59.4315 | -0.004578 | 63.0661 | -59.2392 | 63.1027 | -59.2117 | -59.3491 | -59.0744 | -0.16937 | -0.041198 | -0.1236 | -0.082398 | 0.10071 | 0.27008 | -0.25635 | -0.28839 | -0.086976 | -0.11902 | -0.064087 | -0.045776 | -0.009155 | 0.036621 | -0.036621 | 0.027466 | 0.10071 | -0.17395 | 0.16479 | -0.10986
208.7812 | 122.5571 | -59.2072 | 63.0569 | -0.046616 | -0.00015382 | 0.069186 | 0.066996 | 0.0016813 | -0.00026063 | -0.00081179 | 0.0017395 | -0.00025411 | 0.06492 | 0.064473 | 0.00024568 | -0.0017667 | -0.0036167 | -0.10986 | 63.0707 | -59.3445 | 62.9654 | -59.0927 | -59.5001 | -59.4498 | -0.07782 | 63.0524 | -59.2621 | 63.1348 | -59.1843 | -59.3216 | -59.0195 | -0.13733 | 0.013733 | -0.091553 | -0.10529 | 0.11444 | 0.25177 | -0.24262 | -0.29297 | -0.10529 | -0.15564 | -0.054932 | -0.004578 | 0.07782 | 0.082398 | 0.022888 | 0.07782 | 0.18768 | -0.11444 | 0.24262 | -0.059509
209.7812 | 122.4518 | -59.2163 | 63.0981 | -0.077133 | 0.0024096 | 0.068332 | 0.069346 | 0.0021085 | -0.00047423 | -0.0027344 | -0.056244 | 0.0012411 | 0.065988 | 0.063192 | -0.0018905 | -0.00027134 | -0.0010531 | -0.045776 | 63.0844 | -59.3399 | 62.9517 | -59.111 | -59.5047 | -59.4406 | -0.054932 | 63.0707 | -59.2438 | 63.0615 | -59.2484 | -59.3628 | -59.0836 | -0.1236 | -0.013733 | -0.14649 | -0.13275 | 0.10529 | 0.22888 | -0.2243 | -0.28839 | -0.10071 | -0.16479 | -0.027465 | -0.027466 | -0.036622 | -0.009156 | -0.032043 | -0.004578 | 0.13275 | -0.14648 | 0.16022 | -0.11902
210.7812 | 122.4747 | -59.2255 | 63.0753 | -0.13054 | 0.00070078 | 0.06705 | 0.06956 | 0.0016813 | 0.0012348 | 0.002179 | -0.022675 | 0.0010275 | 0.065561 | 0.063619 | 0.00045928 | -0.0036893 | -0.00019868 | -0.11444 | 63.0798 | -59.3765 | 62.8876 | -59.1751 | -59.5734 | -59.4955 | -0.17395 | 63.0844 | -59.2667 | 63.0386 | -59.2667 | -59.404 | -59.1064 | -0.15106 | 0.004578 | -0.18768 | -0.19226 | 0.050354 | 0.20142 | -0.27008 | -0.3479 | -0.11902 | -0.19684 | -0.041199 | 0.009156 | -0.036621 | -0.045777 | -0.041199 | 0 | 0.11902 | -0.17853 | 0.16022 | -0.13733
211.7812 | 122.4289 | -59.2255 | 63.1073 | -0.13054 | 0.0013416 | 0.0694 | 0.067637 | 0.0036039 | -4.7032e-05 | -0.00059819 | -0.016571 | 0.0010275 | 0.065775 | 0.063192 | -0.0012497 | -0.0015529 | -0.0012667 | -0.082397 | 63.1027 | -59.3628 | 62.915 | -59.166 | -59.5688 | -59.4955 | -0.15564 | 63.089 | -59.285 | 63.0249 | -59.3079 | -59.4269 | -59.1248 | -0.13733 | -0.004578 | -0.19226 | -0.18768 | 0.059509 | 0.19684 | -0.27008 | -0.34332 | -0.13275 | -0.20599 | -0.059509 | -0.018311 | -0.082398 | -0.064087 | -0.082397 | -0.022888 | 0.10071 | -0.20142 | 0.16022 | -0.14191
212.7812 | 122.4655 | -59.2209 | 63.0936 | -0.11528 | -0.00036742 | 0.067477 | 0.068278 | 0.0018949 | -0.00068783 | 0.00089721 | -0.025726 | 0.0012411 | 0.063638 | 0.066823 | -0.0014633 | -0.0013393 | 0.0032193 | -0.07782 | 63.1027 | -59.3582 | 62.915 | -59.166 | -59.5642 | -59.5001 | -0.14191 | 63.0661 | -59.2621 | 63.0341 | -59.2987 | -59.4223 | -59.1248 | -0.13733 | 0.009155 | -0.17853 | -0.18768 | 0.054931 | 0.19226 | -0.27924 | -0.34332 | -0.14191 | -0.20599 | -0.041199 | -0.027466 | -0.059509 | -0.032043 | -0.07782 | -0.036621 | 0.09613 | -0.20142 | 0.13733 | -0.16022
213.7812 | 122.5754 | -59.1934 | 63.0844 | -0.10765 | 0.0015552 | 0.06705 | 0.066142 | 0.00082671 | 0.001662 | -0.0014526 | 0.050568 | 0.00060029 | 0.064493 | 0.061269 | -0.0012497 | -0.0026211 | 0.0045011 | -0.068665 | 63.0707 | -59.3216 | 63.0112 | -59.0195 | -59.4635 | -59.3674 | -0.009155 | 62.9929 | -59.2163 | 63.1165 | -59.1797 | -59.2941 | -59.0057 | -0.12817 | -0.013733 | -0.073243 | -0.05951 | 0.17395 | 0.30212 | -0.17395 | -0.27008 | -0.045777 | -0.14191 | -0.022889 | -0.091553 | 0.032043 | 0.1236 | 0.013733 | 0.036622 | 0.18768 | -0.10071 | 0.21057 | -0.077819
214.7812 | 122.5113 | -59.1934 | 63.089 | -0.10765 | 0.00091438 | 0.067904 | 0.069773 | 0.0018949 | 0.0018756 | -0.00017099 | -0.034882 | 0.00060029 | 0.066202 | 0.065114 | -0.0029587 | -0.0024075 | -0.0040439 | -0.032043 | 63.0844 | -59.3216 | 62.9745 | -59.079 | -59.4955 | -59.4269 | -0.022888 | 63.0524 | -59.2163 | 63.089 | -59.2163 | -59.3491 | -59.0836 | -0.12817 | -0.004577 | -0.11444 | -0.10986 | 0.11444 | 0.24262 | -0.23346 | -0.30212 | -0.10529 | -0.17395 | -0.022889 | -0.036621 | 0 | 0.036621 | -0.022889 | 0 | 0.10986 | -0.15564 | 0.13275 | -0.13275
215.7812 | 122.493 | -59.1797 | 63.1027 | -0.11528 | 0.00048698 | 0.068759 | 0.069133 | -0.0023777 | -0.001115 | 0.001538 | -0.056244 | 0.00017309 | 0.062784 | 0.063833 | -0.00018152 | -0.0015529 | -0.0012667 | -0.027466 | 63.0707 | -59.3445 | 62.9883 | -59.0607 | -59.4818 | -59.436 | -0.009155 | 63.0981 | -59.2392 | 63.089 | -59.2117 | -59.3536 | -59.0561 | -0.16479 | -0.032043 | -0.11444 | -0.082398 | 0.11902 | 0.28381 | -0.25635 | -0.30212 | -0.091553 | -0.13733 | -0.05951 | -0.004577 | -0.013733 | -0.009156 | -0.032044 | 0.027466 | 0.1236 | -0.17395 | 0.18311 | -0.11444
216.7812 | 122.4792 | -59.198 | 63.1302 | -0.038986 | -0.00079462 | 0.068118 | 0.068919 | 0.0016813 | 0.00059397 | -0.0014526 | -0.059296 | -4.0508e-05 | 0.065775 | 0.061056 | -0.0021041 | -0.00069854 | 0.0062101 | -0.11902 | 63.0798 | -59.3582 | 62.9059 | -59.1614 | -59.5642 | -59.4955 | -0.013733 | 63.0569 | -59.2804 | 63.0112 | -59.2209 | -59.3719 | -59.1064 | -0.16022 | -0.050354 | -0.2243 | -0.17395 | 0.036621 | 0.19684 | -0.29755 | -0.36621 | -0.13733 | -0.20599 | -0.082398 | -0.073242 | -0.11902 | -0.045777 | -0.022888 | 0.05951 | 0.091553 | -0.17395 | 0.17395 | -0.091552
217.7812 | 122.5708 | -59.1614 | 63.1027 | -0.092392 | -0.00036742 | 0.06705 | 0.066355 | -0.0028049 | -0.00090143 | -0.0010254 | 0.093292 | 0.00038669 | 0.064066 | 0.055929 | -0.00039512 | -0.0036893 | -0.0029757 | -0.041199 | 63.0707 | -59.2896 | 63.0295 | -59.0103 | -59.4131 | -59.3399 | 0.041199 | 63.0432 | -59.2163 | 63.1531 | -59.1064 | -59.2255 | -58.9691 | -0.12817 | -0.032043 | -0.073242 | -0.041199 | 0.15106 | 0.27924 | -0.17853 | -0.25177 | -0.050354 | -0.1236 | -0.054932 | -0.059509 | 0.050354 | 0.10986 | 0.054932 | 0.10986 | 0.19226 | -0.064087 | 0.24719 | -0.009155
218.7812 | 122.5433 | -59.2026 | 63.0615 | -0.11528 | -0.00058102 | 0.068545 | 0.066142 | -0.0015231 | -0.0023968 | -0.00017099 | 0.02005 | -0.00068151 | 0.063638 | 0.060415 | -0.0018905 | -0.0026211 | -0.0016941 | -0.050354 | 63.0295 | -59.3079 | 63.0295 | -59.0149 | -59.4223 | -59.3582 | -0.004578 | 63.0295 | -59.2484 | 63.1302 | -59.1797 | -59.2896 | -59.0195 | -0.10529 | -0.032043 | -0.032043 | 0 | 0.18768 | 0.29297 | -0.15564 | -0.21973 | -0.050354 | -0.11444 | -0.045776 | -0.032043 | 0.068665 | 0.10071 | 0.022889 | 0.068665 | 0.18311 | -0.086975 | 0.22888 | -0.041199
219.7812 | 122.5342 | -59.2163 | 63.121 | -0.061874 | -0.001222 | 0.06705 | 0.068492 | -0.0010959 | 0.0025166 | 0.00025641 | -0.086761 | -0.00089511 | 0.062357 | 0.063192 | -0.0021041 | -0.00048494 | -0.00019868 | -0.10529 | 63.089 | -59.3765 | 62.9105 | -59.1934 | -59.5825 | -59.5047 | -0.13733 | 63.0707 | -59.2575 | 63.1577 | -59.5276 | -59.4681 | -59.0286 | -0.16022 | -0.032044 | -0.21057 | -0.17853 | 0.022889 | 0.18311 | -0.28839 | -0.36621 | -0.12817 | -0.20599 | -0.041198 | -0.050354 | 0.036621 | 0.086975 | -0.31128 | -0.27008 | 0.18768 | -0.25177 | 0.22888 | -0.21057
220.7812 | 122.406 | -59.2392 | 63.0203 | -0.10002 | -0.00036742 | 0.068332 | 0.068919 | -0.0015231 | -4.7032e-05 | 0.0019654 | -0.062347 | 0.0010275 | 0.066629 | 0.062551 | -0.0010361 | -0.00091214 | -0.0034031 | -0.10071 | 63.0753 | -59.3719 | 62.8967 | -59.1797 | -59.5825 | -59.5047 | -0.13733 | 63.0844 | -59.2804 | 63.0524 | -59.2667 | -59.4131 | -59.0607 | -0.13275 | 0.054931 | -0.1236 | -0.17853 | 0.05951 | 0.19226 | -0.2655 | -0.34332 | -0.13275 | -0.21057 | -0.041199 | 0.064087 | 0.032043 | -0.032044 | -0.027466 | 0.013733 | 0.17853 | -0.17395 | 0.21973 | -0.13275
221.7812 | 122.4564 | -59.23 | 63.1027 | -0.038986 | 0.0017688 | 0.068332 | 0.069133 | 0.00061311 | 0.00080757 | -0.0010254 | -0.062347 | 0.0010275 | 0.063211 | 0.06426 | -0.00039512 | -0.00069854 | 0.00065592 | -0.10071 | 63.0936 | -59.3811 | 62.9105 | -59.166 | -59.5734 | -59.5001 | -0.15564 | 63.0707 | -59.2529 | 63.0432 | -59.2712 | -59.4177 | -59.0836 | -0.15106 | -0.009155 | -0.19226 | -0.18311 | 0.064087 | 0.21515 | -0.27008 | -0.34332 | -0.11902 | -0.19226 | -0.022888 | -0.032043 | -0.059509 | -0.027466 | -0.041198 | -0.01831 | 0.14648 | -0.18768 | 0.16937 | -0.16479
222.7812 | 122.525 | -59.2026 | 63.0478 | -0.13054 | -0.00036742 | 0.068759 | 0.068064 | -0.0010959 | -0.0019696 | -0.0042298 | 0.035309 | 0.00081389 | 0.065561 | 0.057638 | 0.00024568 | -0.0019803 | 0.0021511 | -0.073242 | 63.0844 | -59.2987 | 62.9883 | -59.0332 | -59.4635 | -59.3765 | -0.041199 | 63.0066 | -59.2209 | 63.1073 | -59.1888 | -59.2804 | -59.024 | -0.09613 | 0.036621 | -0.05951 | -0.096131 | 0.16937 | 0.2655 | -0.17395 | -0.26093 | -0.07782 | -0.16479 | -0.01831 | -0.041199 | 0.059509 | 0.10071 | 0.013733 | 0.032043 | 0.17853 | -0.07782 | 0.19684 | -0.05951
223.7812 | 122.5525 | -59.166 | 63.089 | -0.092392 | 0.00048698 | 0.068973 | 0.068064 | -0.0060093 | 0.00059397 | -0.00081179 | 0.053619 | -0.00068151 | 0.065561 | 0.058492 | 0.0011001 | 0.0012241 | 0.00022852 | -0.082397 | 63.0844 | -59.3262 | 62.9929 | -59.0561 | -59.4681 | -59.3903 | -0.045776 | 63.0341 | -59.2209 | 63.1897 | -59.166 | -59.2621 | -58.992 | -0.16022 | -0.004577 | -0.09613 | -0.091553 | 0.10986 | 0.27008 | -0.2243 | -0.30212 | -0.064087 | -0.14191 | -0.054931 | -0.054931 | 0.10071 | 0.15564 | 0 | 0.054931 | 0.17395 | -0.09613 | 0.22888 | -0.041199
224.7812 | 122.493 | -59.2163 | 63.1165 | -0.16106 | 0.0017688 | 0.068973 | 0.069987 | 0.0012539 | 0.0020892 | 0.0030334 | -0.065399 | 0.00017309 | 0.065561 | 0.068105 | -0.0021041 | -0.0011257 | -0.0027621 | -0.09613 | 63.0569 | -59.3674 | 62.8143 | -58.9142 | -59.404 | -59.4772 | -0.11902 | 63.0798 | -59.2392 | 63.0798 | -59.2621 | -59.3948 | -59.1019 | -0.15106 | -0.059509 | -0.30212 | -0.24262 | 0.30212 | 0.45319 | -0.26093 | -0.18768 | -0.10986 | -0.036621 | -0.022888 | -0.036621 | -0.036621 | 0 | -0.045776 | -0.022888 | 0.11444 | -0.17853 | 0.13733 | -0.15564
225.7812 | 122.4472 | -59.2026 | 63.121 | -0.084763 | 0.0019824 | 0.0694 | 0.071055 | 0.0042447 | 0.00016657 | 0.0013244 | -0.15695 | -0.00046791 | 0.061716 | 0.063619 | -0.0016769 | -0.0034757 | 1.4916e-05 | -0.10529 | 63.0844 | -59.3628 | 62.9288 | -59.1705 | -59.5734 | -59.6924 | 0.064087 | 63.0707 | -59.2667 | 63.0844 | -59.2621 | -59.4223 | -59.0927 | -0.16022 | -0.036621 | -0.19226 | -0.15564 | 0.032044 | 0.19226 | -0.48981 | -0.37079 | -0.32959 | -0.21057 | -0.064087 | -0.050354 | -0.036621 | 0.013733 | -0.059509 | 0.004578 | 0.10986 | -0.21973 | 0.17395 | -0.15564
226.7812 | 122.4792 | -59.198 | 63.0936 | -0.061874 | -0.0010082 | 0.067691 | 0.06956 | -0.00088229 | 0.00016657 | 0.0026062 | -0.074554 | 0.00081389 | 0.064706 | 0.067891 | -0.0012497 | -0.0030483 | 1.4916e-05 | -0.11902 | 63.0936 | -59.3628 | 62.9242 | -59.1568 | -59.5642 | -59.491 | -0.16022 | 63.0569 | -59.2712 | 63.1577 | -59.3216 | -59.3536 | -59.0469 | -0.1648 | 0 | -0.16937 | -0.16937 | 0.041199 | 0.20599 | -0.29297 | -0.36621 | -0.12817 | -0.20142 | -0.073242 | -0.036621 | 0.064087 | 0.10071 | -0.1236 | -0.050354 | 0.15106 | -0.15564 | 0.2243 | -0.082398
227.7812 | 122.5662 | -59.198 | 63.0981 | -0.077133 | 0.00091438 | 0.067691 | 0.066355 | -0.0019503 | 0.0014484 | 0.0011108 | 0.068878 | -4.0508e-05 | 0.063425 | 0.060201 | -0.00082252 | -0.0021939 | 0.0036465 | -0.045776 | 63.0707 | -59.3079 | 63.0249 | -59.0103 | -59.3948 | -59.3491 | 0.041199 | 63.0478 | -59.2117 | 63.1485 | -59.1019 | -59.2255 | -58.9691 | -0.10986 | -0.027466 | -0.073243 | -0.045777 | 0.18768 | 0.29755 | -0.15106 | -0.19684 | -0.041199 | -0.086975 | -0.013733 | -0.050354 | 0.050354 | 0.10071 | 0.09613 | 0.10986 | 0.22888 | -0.027466 | 0.24262 | -0.013733
228.7812 | 122.5525 | -59.1843 | 63.0615 | -0.17632 | -0.00036742 | 0.069613 | 0.067637 | 0.0016813 | 0.0018756 | 0.001538 | -0.013519 | -0.0013223 | 0.06257 | 0.064687 | -0.00018152 | -0.0034757 | -0.0016941 | -0.045776 | 63.0432 | -59.317 | 63.0386 | -58.9966 | -59.436 | -59.3628 | 0.013733 | 62.9974 | -59.2346 | 63.1439 | -59.1476 | -59.2575 | -59.0012 | -0.13275 | -0.01831 | -0.022888 | -0.004578 | 0.18768 | 0.32044 | -0.17853 | -0.25177 | -0.045776 | -0.11902 | -0.050354 | -0.064086 | 0.082398 | 0.14648 | 0.036621 | 0.086975 | 0.18311 | -0.073242 | 0.23346 | -0.022888
229.7812 | 122.4243 | -59.198 | 63.0981 | -0.077133 | 0.0019824 | 0.06705 | 0.070842 | 0.0042447 | -0.00047423 | 0.001538 | -0.089813 | 0.0012411 | 0.063425 | 0.063833 | -0.00039512 | -0.0024075 | -0.0019077 | -0.091553 | 63.0936 | -59.3674 | 62.9242 | -59.1751 | -59.5642 | -59.4955 | -0.15564 | 63.0981 | -59.2575 | 63.0478 | -59.285 | -59.4315 | -59.1202 | -0.16937 | -0.004578 | -0.17395 | -0.16937 | 0.022888 | 0.19226 | -0.29755 | -0.36621 | -0.12817 | -0.19684 | -0.059509 | 0 | -0.050354 | -0.050354 | -0.086975 | -0.027466 | 0.07782 | -0.23346 | 0.13733 | -0.17395
230.7812 | 122.4655 | -59.198 | 63.1119 | -0.061874 | -0.00036742 | 0.068118 | 0.070628 | -2.7692e-05 | 0.00059397 | -0.00038459 | 0.0017395 | 0.0020957 | 0.062997 | 0.059347 | -0.0021041 | -0.0015529 | 0.0023647 | -0.10071 | 63.089 | -59.3628 | 63.0066 | -59.3582 | -59.8709 | -59.3811 | -0.19684 | 63.0936 | -59.2484 | 63.0524 | -59.2804 | -59.3903 | -59.079 | -0.1648 | -0.022888 | -0.10529 | -0.082397 | -0.16022 | 0.004578 | -0.18311 | -0.67291 | -0.018311 | -0.50812 | -0.050354 | -0.01831 | -0.059509 | -0.041199 | -0.082398 | -0.032044 | 0.11902 | -0.19226 | 0.16937 | -0.14191
231.7812 | 122.4426 | -59.2072 | 63.1027 | -0.077133 | -0.0025036 | 0.068973 | 0.070201 | 0.0040311 | 0.00038017 | -0.00038459 | -0.095917 | 0.0025229 | 0.06492 | 0.065969 | -0.0018905 | -0.0017667 | -0.0014805 | -0.13275 | 63.0569 | -59.3353 | 62.9196 | -59.166 | -59.5779 | -59.4681 | -0.14648 | 63.0569 | -59.2758 | 63.0661 | -59.2621 | -59.3491 | -59.0561 | -0.12817 | -0.045776 | -0.1831 | -0.13733 | 0.041198 | 0.16937 | -0.26093 | -0.37079 | -0.13275 | -0.24262 | -0.068665 | -0.045776 | -0.036621 | 0.009155 | -0.054932 | 0.013733 | 0.15106 | -0.14191 | 0.21973 | -0.073242
232.7812 | 122.5525 | -59.1797 | 63.1119 | -0.10765 | -0.00036742 | 0.067264 | 0.066569 | -2.7692e-05 | -0.0019696 | 0.00025641 | 0.032257 | 0.00017309 | 0.06492 | 0.056356 | -0.00060892 | -0.0024075 | -0.00019868 | 0.009155 | 63.0707 | -59.3033 | 63.0753 | -58.9828 | -59.3674 | -59.3353 | 0.086975 | 63.0844 | -59.2117 | 63.1439 | -59.1156 | -59.2438 | -58.9966 | -0.1236 | -0.041198 | -0.036621 | 0.004577 | 0.19684 | 0.32043 | -0.15564 | -0.18768 | -0.032043 | -0.064087 | -0.032044 | -0.027465 | 0.032044 | 0.059509 | 0.064086 | 0.09613 | 0.18311 | -0.064087 | 0.21515 | -0.032043
233.7812 | 122.4518 | -59.2117 | 63.1348 | -0.023727 | -0.0010082 | 0.068973 | 0.069133 | 0.0055265 | 0.00059397 | -0.0014526 | -0.086761 | 0.0018821 | 0.063425 | 0.060842 | 0.0019547 | 0.0010105 | -0.0051121 | -0.10071 | 63.0798 | -59.3582 | 62.9013 | -59.1614 | -59.5505 | -59.5047 | -0.16022 | 63.0844 | -59.2758 | 63.0478 | -59.2667 | -59.4177 | -59.0927 | -0.14648 | -0.054932 | -0.23346 | -0.17853 | 0.050354 | 0.19684 | -0.29297 | -0.33875 | -0.14648 | -0.19226 | -0.064087 | -0.050354 | -0.086975 | -0.036621 | -0.054932 | 0.009155 | 0.11902 | -0.20599 | 0.18311 | -0.14191
234.7812 | 122.4609 | -59.2072 | 63.1073 | -0.10765 | 0.00091438 | 0.068545 | 0.069773 | -0.0025913 | -4.7032e-05 | -0.0014526 | -0.098968 | 0.00017309 | 0.064706 | 0.061696 | -0.0016769 | -0.00091214 | 0.0015103 | -0.086975 | 63.0844 | -59.3674 | 62.9333 | -59.1385 | -59.5505 | -59.4955 | -0.12817 | 63.0661 | -59.2575 | 63.0386 | -59.285 | -59.4223 | -59.1064 | -0.16022 | -0.022888 | -0.17395 | -0.15106 | 0.068664 | 0.22888 | -0.28839 | -0.34332 | -0.12817 | -0.18311 | -0.050354 | -0.041199 | -0.068665 | -0.027466 | -0.07782 | -0.027466 | 0.10071 | -0.21515 | 0.15106 | -0.1648
235.7812 | 122.4884 | -59.198 | 63.1348 | -0.13054 | -0.00015382 | 0.066623 | 0.069987 | 0.0029629 | -0.001115 | 4.2808e-05 | -0.056244 | 0.0014547 | 0.065561 | 0.067678 | 0.0019547 | -0.00048494 | 0.0019375 | -0.018311 | 63.0707 | -59.3262 | 62.9745 | -59.0607 | -59.4818 | -59.4315 | -0.009155 | 63.0615 | -59.2392 | 63.1073 | -59.23 | -59.3399 | -59.0332 | -0.12817 | -0.064087 | -0.16022 | -0.096131 | 0.13733 | 0.2655 | -0.23346 | -0.28381 | -0.10529 | -0.15564 | -0.041199 | -0.073243 | -0.027466 | 0.045777 | -0.032044 | 0.009155 | 0.16479 | -0.14191 | 0.20599 | -0.10071
236.7812 | 122.5296 | -59.2072 | 63.1073 | -0.077133 | 0.00027338 | 0.070468 | 0.06956 | -0.0025913 | 0.00080757 | -0.001239 | -0.059296 | 0.00038669 | 0.064706 | 0.065328 | 0.0015275 | -0.0015529 | 0.00086952 | -0.09613 | 63.0753 | -59.3628 | 62.8967 | -59.1843 | -59.5779 | -59.5184 | -0.15106 | 63.0524 | -59.2575 | 63.1393 | -59.523 | -59.4131 | -59.0103 | -0.15564 | -0.032044 | -0.21057 | -0.17853 | 0.022888 | 0.17853 | -0.31128 | -0.37079 | -0.15564 | -0.21515 | -0.050354 | -0.054932 | 0.032043 | 0.086975 | -0.31586 | -0.2655 | 0.19684 | -0.20599 | 0.24719 | -0.15564
237.7812 | 122.4564 | -59.198 | 63.1027 | -0.092392 | 0.00048698 | 0.068545 | 0.069133 | -0.0013095 | 0.001662 | -0.00059819 | -0.062347 | 0.00081389 | 0.064279 | 0.064687 | -0.00018152 | -0.0017667 | -0.0040439 | -0.041199 | 63.089 | -59.3445 | 62.9608 | -59.1064 | -59.491 | -59.4498 | -0.054932 | 63.0707 | -59.2392 | 63.0753 | -59.2255 | -59.3903 | -59.079 | -0.14648 | -0.013733 | -0.14191 | -0.12817 | 0.091553 | 0.23804 | -0.25177 | -0.29297 | -0.10529 | -0.14648 | -0.041199 | -0.032043 | -0.027466 | 0.004577 | -0.027466 | 0.013733 | 0.11902 | -0.19226 | 0.16022 | -0.15106
238.7812 | 122.4976 | -59.1614 | 63.0844 | -0.10002 | 0.0019824 | 0.066409 | 0.06721 | -0.0032321 | -0.00026063 | 4.2808e-05 | 0.029205 | -0.0015359 | 0.06492 | 0.057424 | 0.00088648 | -0.0032619 | -0.00083948 | -0.064087 | 63.0569 | -59.317 | 62.9883 | -59.0332 | -59.4452 | -59.3857 | -0.027466 | 63.0249 | -59.2484 | 63.121 | -59.1888 | -59.2896 | -59.0195 | -0.15564 | -0.027466 | -0.096131 | -0.068665 | 0.12817 | 0.28381 | -0.2243 | -0.28381 | -0.068664 | -0.12817 | -0.086975 | -0.05951 | 0.036621 | 0.096131 | -0.027466 | 0.059509 | 0.14191 | -0.12817 | 0.22888 | -0.041199
239.7812 | 122.4335 | -59.2255 | 63.1073 | -0.084763 | 0.0024096 | 0.068545 | 0.070414 | 0.00018591 | 0.00080757 | -0.00038459 | -0.074554 | 0.0023093 | 0.064279 | 0.067037 | -0.0021041 | -5.7742e-05 | -0.00019868 | -0.082397 | 63.0798 | -59.3674 | 62.9562 | -59.1339 | -59.5367 | -59.491 | -0.12817 | 63.089 | -59.2163 | 63.0386 | -59.2712 | -59.3994 | -59.1019 | -0.14191 | -0.027466 | -0.15106 | -0.1236 | 0.091553 | 0.23346 | -0.2655 | -0.31128 | -0.1236 | -0.16937 | 0.009155 | -0.018311 | -0.068665 | -0.050354 | -0.045776 | -0.054931 | 0.1236 | -0.17395 | 0.11444 | -0.1831
240.7812 | 122.4701 | -59.2117 | 63.1165 | -0.11528 | 0.00070078 | 0.071109 | 0.068919 | -0.0021641 | -0.00068783 | 0.0013244 | -0.077606 | 0.0029501 | 0.063425 | 0.064046 | -0.0018905 | -0.00027134 | 0.0021511 | -0.064087 | 63.0844 | -59.3399 | 62.9425 | -59.1293 | -59.5139 | -59.4635 | -0.10986 | 63.089 | -59.2438 | 63.0661 | -59.2529 | -59.3994 | -59.1019 | -0.12817 | -0.032043 | -0.17395 | -0.14191 | 0.082398 | 0.21057 | -0.25177 | -0.30212 | -0.1236 | -0.17395 | -0.032043 | -0.027466 | -0.050354 | -0.022888 | -0.041199 | -0.009156 | 0.10986 | -0.18768 | 0.14191 | -0.15564
241.7812 | 122.4838 | -59.1843 | 63.0981 | -0.092392 | 5.9776e-05 | 0.065768 | 0.069773 | -0.0025913 | 0.0012348 | 0.0013244 | -0.047089 | 0.0014547 | 0.064706 | 0.06661 | -0.00060892 | 0.00015586 | -0.0010531 | 0.064087 | 63.1073 | -59.3262 | 63.0112 | -59.0515 | -59.4406 | -59.3948 | 0.004578 | 63.0707 | -59.2163 | 63.1119 | -59.198 | -59.3536 | -59.0469 | -0.14191 | 0.009155 | -0.086976 | -0.096131 | 0.13275 | 0.27466 | -0.21057 | -0.25635 | -0.068664 | -0.11444 | -0.032044 | -0.027466 | 0.013732 | 0.041198 | -0.013733 | 0.018311 | 0.13733 | -0.16937 | 0.16937 | -0.13733
242.7812 | 122.4747 | -59.23 | 63.1027 | -0.10002 | 5.9776e-05 | 0.067477 | 0.068919 | 0.0010403 | -0.0019696 | -0.00081179 | -0.05014 | 0.00060029 | 0.065347 | 0.06426 | -0.00018152 | -0.0013393 | -0.00019868 | -0.07782 | 63.0798 | -59.3719 | 62.915 | -59.1614 | -59.5551 | -59.5001 | -0.16479 | 63.0798 | -59.2758 | 63.0386 | -59.2804 | -59.4177 | -59.1156 | -0.14191 | -0.022888 | -0.18768 | -0.16479 | 0.068665 | 0.21057 | -0.27008 | -0.32501 | -0.12817 | -0.18311 | -0.045776 | -0.022888 | -0.064087 | -0.041199 | -0.050354 | -0.004578 | 0.11444 | -0.18768 | 0.16022 | -0.14191
243.7812 | 122.4747 | -59.2163 | 63.1256 | -0.061874 | 0.00048698 | 0.067691 | 0.068705 | -0.0028049 | 0.0020892 | 0.0013244 | -0.047089 | 0.0016685 | 0.064279 | 0.067251 | 0.0011001 | 0.00015586 | -0.0029757 | -0.022888 | 63.0936 | -59.3216 | 62.9837 | -59.0744 | -59.4681 | -59.4315 | -0.027466 | 63.1027 | -59.2438 | 63.0844 | -59.2255 | -59.3582 | -59.0744 | -0.10528 | -0.032043 | -0.14191 | -0.10986 | 0.14191 | 0.24719 | -0.21515 | -0.25177 | -0.10986 | -0.14649 | -0.027465 | -0.022888 | -0.041198 | -0.01831 | -0.009155 | 0.01831 | 0.14191 | -0.14191 | 0.16937 | -0.11444
244.7812 | 122.4884 | -59.1934 | 63.1073 | -0.084763 | 0.0026234 | 0.068545 | 0.06721 | 0.0046719 | 0.00016657 | 0.00025641 | -0.056244 | -0.00068151 | 0.066202 | 0.068532 | 0.0015275 | -0.0024075 | 0.0025785 | -0.07782 | 63.0936 | -59.3399 | 62.9379 | -59.1156 | -59.5184 | -59.4681 | -0.11902 | 63.1027 | -59.2392 | 63.0478 | -59.2804 | -59.404 | -59.0973 | -0.14648 | -0.013733 | -0.16937 | -0.15564 | 0.077819 | 0.2243 | -0.27466 | -0.32501 | -0.12817 | -0.17853 | -0.045777 | -0.004578 | -0.059509 | -0.054931 | -0.086976 | -0.041199 | 0.09613 | -0.21057 | 0.14191 | -0.1648
245.7812 | 122.4243 | -59.2117 | 63.0844 | -0.061874 | 0.0019824 | 0.067477 | 0.068064 | -2.7692e-05 | 0.0031574 | -0.00017099 | -0.05014 | 0.0012411 | 0.063425 | 0.066823 | 0.00067288 | -5.7742e-05 | 0.0021511 | -0.064087 | 63.089 | -59.3353 | 62.9517 | -59.0973 | -59.5093 | -59.4543 | -0.07782 | 63.0707 | -59.23 | 63.0569 | -59.2392 | -59.3857 | -59.0744 | -0.1236 | 0.004577 | -0.13275 | -0.13733 | 0.11444 | 0.23804 | -0.24262 | -0.29755 | -0.11902 | -0.17395 | -0.018311 | -0.013733 | -0.027466 | -0.013733 | -0.027466 | -0.009155 | 0.13733 | -0.17395 | 0.15564 | -0.15564
246.7812 | 122.493 | -59.198 | 63.1165 | -0.084763 | -0.001222 | 0.0694 | 0.070414 | -0.0021641 | 0.0012348 | 4.2808e-05 | -0.03183 | 0.00060029 | 0.064066 | 0.065755 | -0.0010361 | -0.00027134 | 0.0015103 | -0.045776 | 63.0936 | -59.3399 | 62.9745 | -59.0836 | -59.5093 | -59.4498 | -0.050354 | 63.0844 | -59.2346 | 63.0798 | -59.2392 | -59.3674 | -59.0744 | -0.14191 | -0.022888 | -0.14191 | -0.11902 | 0.11444 | 0.25635 | -0.25177 | -0.31128 | -0.10986 | -0.16937 | -0.036621 | -0.032043 | -0.036621 | -0.004578 | -0.041199 | -0.004578 | 0.1236 | -0.16937 | 0.16022 | -0.13275
247.7812 | 122.4838 | -59.2163 | 63.1256 | -0.13817 | 0.00048698 | 0.068759 | 0.066783 | -0.0021641 | 0.0027302 | -0.0018798 | -0.05014 | 0.00060029 | 0.065561 | 0.068105 | -0.00060892 | -0.0024075 | -0.00062588 | -0.022888 | 63.0981 | -59.3399 | 62.9791 | -59.0607 | -59.4818 | -59.4315 | -0.036621 | 63.0707 | -59.2117 | 63.0844 | -59.23 | -59.3628 | -59.0652 | -0.1236 | -0.027465 | -0.14648 | -0.11902 | 0.15564 | 0.27924 | -0.21515 | -0.2655 | -0.091553 | -0.14191 | 0.004578 | -0.054931 | -0.041198 | 0.013733 | -0.013733 | -0.018311 | 0.15106 | -0.14648 | 0.14648 | -0.15106
248.7812 | 122.4197 | -59.1476 | 63.1027 | -0.031357 | -0.0010082 | 0.068332 | 0.066355 | 0.0023221 | -0.0015424 | 0.00068361 | -0.03183 | -0.0011087 | 0.064066 | 0.065328 | -0.00039512 | -0.0017667 | -0.00083948 | -0.009155 | 63.0936 | -59.3124 | 62.9974 | -59.0469 | -59.4315 | -59.4086 | 0.032043 | 63.0615 | -59.2209 | 63.121 | -59.198 | -59.3079 | -58.9096 | -0.16479 | -0.009155 | -0.10529 | -0.09613 | 0.10071 | 0.2655 | -0.26093 | -0.28381 | -0.09613 | -0.11902 | -0.073242 | -0.041199 | 0.018311 | 0.05951 | -0.050354 | 0.022888 | 0.23804 | -0.16022 | 0.31128 | -0.086975
249.7812 | 122.4747 | -59.198 | 63.1393 | -0.092392 | -0.00015382 | 0.067904 | 0.065715 | 0.00039951 | 0.00080757 | -0.00081179 | -0.016571 | 0.0012411 | 0.064706 | 0.064473 | -0.00082252 | -0.0039029 | 0.0023647 | -0.018311 | 63.0936 | -59.3216 | 62.9837 | -59.0378 | -59.4223 | -59.4086 | 0.004578 | 63.0661 | -59.2209 | 63.1073 | -59.198 | -59.3353 | -59.0561 | -0.1236 | -0.045776 | -0.15564 | -0.10986 | 0.16022 | 0.28381 | -0.21057 | -0.2243 | -0.086975 | -0.10071 | -0.022888 | -0.073242 | -0.032043 | 0.041199 | 0 | 0.022888 | 0.14191 | -0.13733 | 0.16479 | -0.11444
250.7812 | 122.525 | -59.1705 | 63.1439 | -0.054245 | -0.00015382 | 0.066836 | 0.067851 | 0.00039951 | -0.00047423 | -0.00059819 | -0.0074159 | 0.00017309 | 0.065347 | 0.066182 | -0.00082252 | -0.0024075 | 0.0021511 | 0 | 63.121 | -59.3216 | 62.9883 | -59.0424 | -59.4498 | -59.4131 | -0.091553 | 63.0524 | -59.2438 | 63.1119 | -59.1843 | -59.3262 | -59.0515 | -0.15106 | -0.022888 | -0.15564 | -0.13275 | 0.12817 | 0.27924 | -0.24262 | -0.27924 | -0.091553 | -0.12817 | -0.073242 | -0.091553 | -0.032044 | 0.059509 | -0.013733 | 0.059509 | 0.11902 | -0.15564 | 0.19226 | -0.082398
251.7812 | 122.5159 | -59.1751 | 63.1256 | -0.054245 | -0.0014356 | 0.066836 | 0.068064 | 0.0014675 | -4.7032e-05 | -0.001239 | -0.016571 | 0.0012411 | 0.065988 | 0.061269 | 0.00045928 | -0.0015529 | 0.00022852 | 0 | 63.089 | -59.2941 | 63.0066 | -59.0332 | -59.4498 | -59.4131 | 0.032043 | 63.0798 | -59.2072 | 63.1577 | -59.3033 | -59.3674 | -59.0515 | -0.11902 | -0.036621 | -0.11902 | -0.082397 | 0.14191 | 0.26093 | -0.23804 | -0.27466 | -0.11902 | -0.15564 | -0.032043 | -0.045776 | 0.032044 | 0.07782 | -0.12817 | -0.096131 | 0.1236 | -0.19226 | 0.15564 | -0.16022
252.7812 | 122.4701 | -59.1888 | 63.1348 | -0.10002 | -0.00036742 | 0.067264 | 0.069346 | 0.0014675 | -0.001115 | 0.00025641 | -0.044037 | 0.00060029 | 0.065347 | 0.068319 | -0.00060892 | -0.0021939 | 0.0023647 | -0.064087 | 63.0936 | -59.3262 | 62.9562 | -59.1019 | -59.5047 | -59.4635 | -0.07782 | 63.0981 | -59.2346 | 63.0524 | -59.2484 | -59.3948 | -59.0698 | -0.13733 | -0.041199 | -0.17853 | -0.13733 | 0.086975 | 0.2243 | -0.27466 | -0.31586 | -0.13733 | -0.17853 | -0.045776 | -0.036621 | -0.082398 | -0.045777 | -0.059509 | -0.013733 | 0.11902 | -0.20599 | 0.1648 | -0.16022
253.7812 | 122.4701 | -59.2072 | 63.1119 | -0.016098 | 0.001128 | 0.068545 | 0.068705 | 0.0025357 | -4.7032e-05 | 0.00068361 | -0.044037 | 0.0020957 | 0.064066 | 0.061056 | -0.00082252 | -0.0011257 | -0.0023349 | -0.041199 | 63.089 | -59.3353 | 62.9745 | -59.0973 | -59.4955 | -59.4315 | -0.054932 | 63.0798 | -59.2484 | 63.0707 | -59.2438 | -59.3765 | -59.079 | -0.12817 | -0.022888 | -0.13733 | -0.11444 | 0.10986 | 0.23804 | -0.22431 | -0.28839 | -0.096131 | -0.16022 | -0.041199 | -0.032043 | -0.041198 | -0.009155 | -0.036621 | 0.004578 | 0.12817 | -0.16937 | 0.16937 | -0.12817
254.7812 | 122.5296 | -59.198 | 63.1073 | -0.054245 | -0.00015382 | 0.066623 | 0.068278 | 0.00082671 | -0.00068783 | -0.0016662 | -0.022675 | 0.00038669 | 0.06492 | 0.063192 | -0.00039512 | -0.0030483 | 0.0015103 | -0.009155 | 63.089 | -59.3079 | 62.9883 | -59.0469 | -59.4635 | -59.4086 | 0.032043 | 63.0936 | -59.2209 | 63.121 | -59.1888 | -59.2896 | -58.9874 | -0.10986 | -0.018311 | -0.11902 | -0.10071 | 0.15106 | 0.26093 | -0.21057 | -0.2655 | -0.10071 | -0.15564 | -0.022888 | -0.013733 | 0.013733 | 0.027466 | 0.009155 | 0.032043 | 0.21057 | -0.091553 | 0.23346 | -0.068665
255.7812 | 122.4792 | -59.1705 | 63.1256 | -0.054245 | -0.001222 | 0.067904 | 0.066996 | 0.0012539 | 0.0029438 | -0.0031616 | -0.010467 | 0.0010275 | 0.065347 | 0.064901 | -0.00082252 | -0.0015529 | -0.0025485 | -0.027466 | 63.1073 | -59.2987 | 62.9745 | -59.0881 | -59.4498 | -59.3811 | 0.018311 | 63.0844 | -59.1934 | 63.1119 | -59.1614 | -59.3124 | -59.024 | -0.12817 | -0.01831 | -0.15106 | -0.13275 | 0.082397 | 0.21057 | -0.21057 | -0.27924 | -0.082398 | -0.15106 | -0.022888 | -0.041198 | -0.013733 | 0.027465 | 0.009155 | 0.032043 | 0.14648 | -0.14191 | 0.16937 | -0.11902
256.7812 | 122.5159 | -59.1797 | 63.1302 | -0.069504 | 0.00091438 | 0.06705 | 0.066996 | -0.0030185 | -0.00047423 | -0.0023072 | -0.0013123 | 0.00017309 | 0.065134 | 0.062764 | -0.0018905 | 0.00058326 | 0.0019375 | 0.004578 | 63.1027 | -59.3124 | 62.9883 | -59.0469 | -59.4406 | -59.404 | 0.045776 | 63.0981 | -59.23 | 63.1348 | -59.166 | -59.2896 | -59.0149 | -0.13275 | -0.027466 | -0.14191 | -0.11444 | 0.13275 | 0.2655 | -0.22431 | -0.26093 | -0.091553 | -0.12817 | -0.050355 | -0.032043 | 0.004578 | 0.036621 | 0.013732 | 0.064087 | 0.16479 | -0.10986 | 0.21515 | -0.059509
257.7812 | 122.4976 | -59.1843 | 63.1302 | -0.092392 | 0.0013416 | 0.067691 | 0.069133 | 0.0021085 | -0.001756 | -0.00059819 | -0.019623 | 0.00017309 | 0.064066 | 0.06426 | -0.0016769 | -0.0034757 | 0.0017239 | 0.009155 | 63.1027 | -59.3216 | 63.002 | -59.0469 | -59.4315 | -59.404 | 0.050354 | 63.0753 | -59.1843 | 63.1165 | -59.1843 | -59.3079 | -59.0332 | -0.13733 | -0.027466 | -0.12817 | -0.10071 | 0.13733 | 0.27466 | -0.21973 | -0.24719 | -0.082398 | -0.10986 | 0 | -0.054932 | -0.013733 | 0.041199 | 0 | 0 | 0.15106 | -0.1236 | 0.15106 | -0.1236
258.7812 | 122.4289 | -59.2117 | 63.1165 | -0.10765 | 0.0026234 | 0.069186 | 0.067637 | 0.0016813 | 0.00080757 | 0.0011108 | -0.056244 | 0.00060029 | 0.063852 | 0.060628 | 3.2078e-05 | -0.0015529 | 0.0021511 | -0.068665 | 63.0661 | -59.3536 | 62.9425 | -59.1339 | -59.523 | -59.4635 | -0.10986 | 63.0844 | -59.2392 | 63.0203 | -59.2804 | -59.4131 | -59.0836 | -0.14191 | -0.050354 | -0.17395 | -0.1236 | 0.07782 | 0.21973 | -0.25177 | -0.31128 | -0.10986 | -0.16937 | -0.027466 | -0.032043 | -0.09613 | -0.064087 | -0.068665 | -0.041199 | 0.12817 | -0.20142 | 0.15564 | -0.17395
259.7812 | 122.5067 | -59.1843 | 63.1165 | -0.084763 | -0.0010082 | 0.068118 | 0.067851 | 0.0016813 | 0.0010212 | -0.0016662 | -0.016571 | 0.0016685 | 0.064279 | 0.062764 | -0.00039512 | -0.0028347 | 0.0032193 | 0.013733 | 63.1119 | -59.3445 | 63.002 | -59.0378 | -59.4406 | -59.3994 | 0.027466 | 63.0753 | -59.2346 | 63.1119 | -59.1888 | -59.3079 | -59.0103 | -0.16022 | -0.004578 | -0.11444 | -0.10986 | 0.14648 | 0.3067 | -0.21515 | -0.25635 | -0.054932 | -0.096131 | -0.050354 | -0.041199 | -0.004578 | 0.036621 | -0.004578 | 0.045776 | 0.17395 | -0.1236 | 0.2243 | -0.073242
260.7812 | 122.5388 | -59.1431 | 63.1348 | -0.077133 | 0.001128 | 0.066409 | 0.068064 | 0.0021085 | -0.001115 | -0.0016662 | -0.0013123 | 0.00017309 | 0.065775 | 0.067891 | -0.0012497 | -0.0017667 | -0.00019868 | 0.004578 | 63.0981 | -59.3033 | 63.002 | -59.0424 | -59.4269 | -59.3903 | 0.045776 | 63.0936 | -59.2255 | 63.1119 | -59.1705 | -59.2987 | -59.0195 | -0.16022 | -0.036621 | -0.13275 | -0.096131 | 0.10071 | 0.26093 | -0.24719 | -0.28381 | -0.086975 | -0.1236 | -0.082398 | -0.041199 | -0.022889 | 0.01831 | -0.027466 | 0.054932 | 0.1236 | -0.15564 | 0.20599 | -0.073242
261.7812 | 122.4701 | -59.1888 | 63.1348 | -0.077133 | 5.9776e-05 | 0.067264 | 0.068278 | 0.0076627 | 0.0012348 | -0.0025208 | -0.0043641 | -0.00089511 | 0.065561 | 0.067037 | 0.00045928 | -0.0021939 | 0.0045011 | -0.009155 | 63.1165 | -59.2941 | 62.9974 | -59.0561 | -59.4269 | -59.3994 | 0.041199 | 63.0753 | -59.2117 | 63.1073 | -59.1843 | -59.3033 | -59.0332 | -0.10529 | -0.018311 | -0.13733 | -0.11902 | 0.13275 | 0.23804 | -0.21057 | -0.23804 | -0.10529 | -0.13275 | -0.022888 | -0.05951 | -0.027466 | 0.032044 | 0.004578 | 0.027466 | 0.15564 | -0.11444 | 0.17853 | -0.091553
262.7812 | 122.4792 | -59.198 | 63.1622 | 0.075455 | -0.0029308 | 0.066836 | 0.067637 | -0.0028049 | -0.00047423 | -0.00059819 | 0.032257 | 0.0012411 | 0.066202 | 0.063833 | 3.2078e-05 | -0.0039029 | 0.0030057 | -0.09613 | 63.1348 | -59.3216 | 62.9471 | -59.1202 | -59.4772 | -59.4131 | -0.018311 | 63.1027 | -59.1934 | 63.0661 | -59.2026 | -59.317 | -59.0286 | -0.1236 | -0.027465 | -0.21515 | -0.18768 | 0.07782 | 0.20142 | -0.21515 | -0.27924 | -0.091553 | -0.15564 | 0.004578 | -0.059509 | -0.09613 | -0.036621 | -0.004578 | -0.009156 | 0.16937 | -0.11902 | 0.1648 | -0.1236
263.7812 | 122.4655 | -59.1934 | 63.1622 | -0.023727 | -0.001222 | 0.069186 | 0.06721 | 0.00061311 | -0.0023968 | -0.00059819 | 0.047516 | 0.0023093 | 0.066415 | 0.062551 | -0.0012497 | -0.0028347 | 0.00086952 | -0.027466 | 63.1348 | -59.2987 | 63.0661 | -58.905 | -59.6786 | -59.5642 | -0.018311 | 63.1073 | -59.2163 | 63.0753 | -59.2209 | -59.317 | -59.0607 | -0.10529 | -0.027465 | -0.09613 | -0.068665 | 0.28839 | 0.39368 | -0.37079 | -0.48523 | -0.2655 | -0.37994 | -0.022889 | -0.054931 | -0.086975 | -0.032044 | -0.027466 | -0.004577 | 0.13275 | -0.1236 | 0.15564 | -0.10071
264.7812 | 122.5113 | -59.1751 | 63.1622 | -0.069504 | 0.00048698 | 0.066409 | 0.069133 | 0.0018949 | 0.0014484 | -0.0025208 | 0.013947 | 0.0031637 | 0.067911 | 0.065328 | 0.00024568 | -0.00069854 | -0.0061801 | -0.018311 | 63.1073 | -59.2987 | 62.9654 | -59.1064 | -59.4589 | -59.4223 | -0.013733 | 63.0936 | -59.2209 | 63.0798 | -59.1888 | -59.3216 | -59.0515 | -0.1236 | -0.054931 | -0.19684 | -0.14191 | 0.068665 | 0.19226 | -0.24719 | -0.28381 | -0.1236 | -0.16022 | -0.045776 | -0.068664 | -0.082397 | -0.013733 | -0.013733 | 0.032043 | 0.1236 | -0.14648 | 0.16937 | -0.10071
265.7812 | 122.4792 | -59.1934 | 63.1393 | -0.054245 | 0.0015552 | 0.066836 | 0.070201 | 0.00082671 | -0.0030376 | -0.00059819 | 0.016998 | 0.00017309 | 0.065561 | 0.067037 | 0.00045928 | -0.0021939 | -0.00062588 | -0.022888 | 63.1302 | -59.3399 | 62.9471 | -59.1202 | -59.4681 | -59.4086 | -0.009155 | 63.089 | -59.2072 | 63.089 | -59.2072 | -59.3308 | -59.0515 | -0.14648 | -0.009155 | -0.19226 | -0.1831 | 0.073242 | 0.21973 | -0.21515 | -0.27466 | -0.068664 | -0.12817 | -0.013733 | -0.050354 | -0.050354 | 0 | -0.013733 | 0 | 0.14191 | -0.13733 | 0.15564 | -0.1236
266.7812 | 122.4152 | -59.198 | 63.1439 | -0.016098 | 0.002837 | 0.069186 | 0.066783 | -0.0010959 | -0.00047423 | -0.0020936 | -0.047089 | 0.0014547 | 0.067056 | 0.063405 | -0.0014633 | -0.0017667 | -0.0027621 | 0.013733 | 63.1119 | -59.2941 | 63.0066 | -59.0378 | -59.4589 | -59.3948 | 0.036621 | 63.0707 | -59.2163 | 63.121 | -59.1751 | -59.3033 | -59.0012 | -0.09613 | -0.032044 | -0.13733 | -0.10529 | 0.16022 | 0.25635 | -0.19684 | -0.26093 | -0.10071 | -0.16479 | -0.018311 | -0.073242 | -0.022888 | 0.050354 | 0.022888 | 0.041199 | 0.19684 | -0.10529 | 0.21515 | -0.086975
267.7812 | 122.4884 | -59.1705 | 63.1577 | -0.061874 | -0.00036742 | 0.069613 | 0.068705 | -2.7692e-05 | -0.0013288 | -0.00059819 | -0.0074159 | -0.00025411 | 0.06492 | 0.064687 | 3.2078e-05 | -0.0013393 | -0.0031895 | -0.018311 | 63.1165 | -59.3124 | 62.9791 | -59.1019 | -59.4498 | -59.3994 | -0.004578 | 63.1165 | -59.2072 | 63.1073 | -59.1843 | -59.2896 | -59.0286 | -0.14191 | -0.041199 | -0.17853 | -0.13733 | 0.068664 | 0.21057 | -0.22888 | -0.27924 | -0.086975 | -0.13733 | -0.036621 | -0.041199 | -0.050354 | -0.009155 | -0.013733 | 0.022888 | 0.14191 | -0.11902 | 0.17853 | -0.082398
268.7812 | 122.5204 | -59.1751 | 63.1851 | -0.061874 | 5.9776e-05 | 0.067904 | 0.067637 | 0.0010403 | 0.0020892 | -0.0014526 | -0.010467 | 0.0010275 | 0.065347 | 0.057424 | -0.00018152 | -0.0030483 | 0.0040737 | -0.022888 | 63.1393 | -59.3216 | 62.9654 | -59.0927 | -59.4589 | -59.4315 | -0.004578 | 63.0981 | -59.23 | 63.0936 | -59.1797 | -59.3079 | -59.0515 | -0.14648 | -0.045777 | -0.21973 | -0.17395 | 0.082398 | 0.22888 | -0.25635 | -0.28381 | -0.10986 | -0.13733 | -0.054932 | -0.086975 | -0.091553 | -0.004578 | -0.004577 | 0.050355 | 0.1236 | -0.13275 | 0.17853 | -0.077819
269.7812 | 122.3969 | -59.1522 | 63.1485 | -0.031357 | -0.00036742 | 0.064914 | 0.066996 | 0.0012539 | -4.7032e-05 | -0.0023072 | 0.047516 | 0.0010275 | 0.066415 | 0.063405 | 0.00067288 | -0.0039029 | -0.0025485 | -0.036621 | 63.1348 | -59.3124 | 62.9608 | -59.0836 | -59.4635 | -59.404 | 0.013733 | 63.1302 | -59.2072 | 63.0112 | -58.992 | -59.2438 | -59.1385 | -0.16022 | -0.013733 | -0.18768 | -0.17395 | 0.068665 | 0.22888 | -0.25177 | -0.31128 | -0.091553 | -0.15106 | -0.054931 | -0.018311 | -0.13733 | -0.11902 | 0.16022 | 0.21515 | 0.013733 | -0.091552 | 0.068664 | -0.036621
270.7812 | 122.4838 | -59.1843 | 63.1485 | 0.14412 | 0.0013416 | 0.068118 | 0.069773 | -0.0019503 | 0.0027302 | -0.0042298 | 0.0017395 | 0.0018821 | 0.065988 | 0.068319 | -0.00060892 | -0.0026211 | 1.4916e-05 | -0.10986 | 63.1119 | -59.3216 | 62.9471 | -59.0927 | -59.4727 | -59.4177 | -0.004578 | 63.1027 | -59.2117 | 63.0753 | -59.2117 | -59.3262 | -59.0515 | -0.13733 | -0.036622 | -0.20142 | -0.16479 | 0.091553 | 0.22888 | -0.23346 | -0.28839 | -0.096131 | -0.15106 | -0.027466 | -0.045777 | -0.073243 | -0.027466 | -0.027466 | 0 | 0.13275 | -0.14191 | 0.16022 | -0.11444
271.7812 | 122.438 | -59.198 | 63.1393 | -0.0084687 | 0.00048698 | 0.067904 | 0.070414 | -0.00066869 | -0.001756 | -0.00038459 | -0.016571 | 0.0031637 | 0.062784 | 0.059987 | -0.00060892 | -0.0021939 | -0.0014805 | -0.013733 | 63.1165 | -59.3124 | 62.9242 | -59.2484 | -59.5322 | -59.4131 | -0.004578 | 63.0753 | -59.2163 | 63.0661 | -59.2117 | -59.3124 | -59.0378 | -0.11444 | -0.022888 | -0.21515 | -0.19226 | -0.050354 | 0.064087 | -0.21515 | -0.33417 | -0.10071 | -0.21973 | -0.018311 | -0.064087 | -0.073242 | -0.009155 | -0.013733 | 0.004578 | 0.16022 | -0.11444 | 0.17853 | -0.09613
272.7812 | 122.5067 | -59.1751 | 63.1622 | -0.046616 | -0.00036742 | 0.068545 | 0.067424 | -0.0025913 | 0.0012348 | -0.0027344 | -0.0013123 | 0.0016685 | 0.065347 | 0.060628 | -0.0021041 | -0.0026211 | -0.00041228 | -0.013733 | 63.1439 | -59.3033 | 62.9654 | -59.0927 | -59.4589 | -59.4131 | 0 | 63.0798 | -59.1934 | 63.0798 | -59.1797 | -59.2987 | -59.0332 | -0.12817 | -0.01831 | -0.19684 | -0.17853 | 0.082398 | 0.21057 | -0.23804 | -0.28381 | -0.10986 | -0.15564 | -0.01831 | -0.082397 | -0.082397 | 0 | -0.004577 | 0.013733 | 0.14191 | -0.1236 | 0.16022 | -0.10529
273.7812 | 122.438 | -59.2117 | 63.1165 | -0.11528 | 0.0015552 | 0.069613 | 0.070842 | 0.00039951 | 0.001662 | -0.00059819 | -0.059296 | 0.0023093 | 0.065988 | 0.06426 | -0.0014633 | -0.0021939 | 0.0010831 | -0.09613 | 63.0707 | -59.3674 | 62.8418 | -59.4269 | -59.4955 | -59.2987 | -0.11902 | 63.0615 | -59.2392 | 63.0478 | -59.2667 | -59.3719 | -59.079 | -0.15564 | -0.045776 | -0.27466 | -0.22888 | -0.21515 | -0.059509 | -0.086975 | -0.28381 | 0.068665 | -0.12817 | -0.027466 | -0.054932 | -0.068664 | -0.013732 | -0.054932 | -0.027466 | 0.13275 | -0.16022 | 0.16022 | -0.13275
274.7812 | 122.5113 | -59.166 | 63.1348 | -0.069504 | -0.00036742 | 0.066836 | 0.069773 | -0.0013095 | 0.0012348 | -0.0014526 | -0.0043641 | 0.0018821 | 0.065561 | 0.063833 | -0.00018152 | -0.0026211 | 0.0030057 | 0.13733 | 63.1897 | -59.2804 | 62.9379 | -59.0973 | -59.4727 | -59.4086 | -0.013733 | 63.1119 | -59.2163 | 63.0844 | -59.2072 | -59.3262 | -59.0698 | -0.11444 | 0.054931 | -0.19684 | -0.25177 | 0.068665 | 0.18311 | -0.24261 | -0.3067 | -0.12817 | -0.19226 | -0.050354 | -0.022889 | -0.050354 | -0.027465 | -0.041198 | 0.009156 | 0.096131 | -0.16022 | 0.14649 | -0.10986
275.7812 | 122.4609 | -59.166 | 63.1439 | -0.054245 | -0.0016492 | 0.066409 | 0.067637 | -0.00066869 | 0.00059397 | 0.00089721 | -0.047089 | 0.0020957 | 0.065347 | 0.056569 | 0.00024568 | -0.0013393 | 0.00022852 | -0.032043 | 63.1622 | -59.3262 | 62.9654 | -59.0836 | -59.3399 | -59.2117 | -0.11902 | 63.0936 | -59.2438 | 63.0661 | -59.198 | -59.3308 | -59.0607 | -0.16022 | 0.01831 | -0.17853 | -0.19684 | 0.082398 | 0.24262 | -0.045776 | -0.17395 | 0.11444 | -0.013733 | -0.077819 | -0.050354 | -0.07782 | -0.027466 | -0.032043 | 0.045776 | 0.10529 | -0.1648 | 0.18311 | -0.086976
276.7812 | 122.5021 | -59.1705 | 63.1439 | -0.046616 | 0.0013416 | 0.067477 | 0.065287 | 0.00082671 | 0.00059397 | 0.0011108 | 0.056671 | 0.00017309 | 0.063852 | 0.061483 | -0.0025315 | -0.0026211 | 0.0015103 | 0.027466 | 63.1027 | -59.2987 | 63.0386 | -58.992 | -59.3903 | -59.3536 | 0.068665 | 63.0798 | -59.1888 | 63.1073 | -59.1705 | -59.2987 | -59.0424 | -0.12817 | -0.041199 | -0.10529 | -0.064087 | 0.17853 | 0.3067 | -0.18311 | -0.21973 | -0.054932 | -0.091553 | -0.018311 | -0.064087 | -0.036621 | 0.027466 | 0 | 0.018311 | 0.12817 | -0.12817 | 0.14648 | -0.10986
277.7812 | 122.4609 | -59.1888 | 63.176 | -0.11528 | -0.0014356 | 0.067477 | 0.068492 | -0.0028049 | -0.0019696 | -0.00081179 | -0.0013123 | 0.00081389 | 0.065561 | 0.065114 | -0.00039512 | -0.0036893 | 0.0034329 | -0.004578 | 63.1256 | -59.3033 | 62.9791 | -59.0652 | -59.4543 | -59.3994 | 0.036621 | 63.089 | -59.1934 | 63.0661 | -59.2026 | -59.3124 | -59.0652 | -0.11444 | -0.050354 | -0.19684 | -0.14648 | 0.1236 | 0.23804 | -0.21057 | -0.2655 | -0.09613 | -0.15106 | -0.004577 | -0.086975 | -0.10986 | -0.022888 | -0.013733 | -0.009156 | 0.1236 | -0.1236 | 0.12817 | -0.11902
278.7812 | 122.4655 | -59.198 | 63.1393 | -0.092392 | -0.0010082 | 0.067904 | 0.070414 | 0.0014675 | -0.0015424 | 0.00047001 | -0.071503 | 0.00038669 | 0.062357 | 0.063405 | -0.0021041 | -0.0021939 | 0.0010831 | -0.036621 | 63.0844 | -59.3491 | 62.97 | -59.0927 | -59.491 | -59.4498 | -0.041199 | 63.0798 | -59.2255 | 63.0753 | -59.2346 | -59.3765 | -59.0744 | -0.15106 | -0.054931 | -0.16937 | -0.11444 | 0.10529 | 0.25635 | -0.25177 | -0.29297 | -0.10071 | -0.14191 | -0.027466 | -0.059509 | -0.064087 | -0.004578 | -0.036621 | -0.009155 | 0.1236 | -0.17853 | 0.15106 | -0.15106
279.7812 | 122.374 | -59.1705 | 63.1897 | -0.061874 | -0.00015382 | 0.066623 | 0.06956 | -0.0025913 | -0.001756 | -0.0025208 | 0.0078429 | -0.00046791 | 0.066629 | 0.059347 | -0.0010361 | -0.0032619 | 0.00044212 | -0.022888 | 63.1027 | -59.2987 | 62.9791 | -59.0836 | -59.4543 | -59.4131 | 0.013733 | 63.1119 | -59.23 | 63.0707 | -59.2026 | -59.3216 | -59.0698 | -0.12817 | -0.086975 | -0.21057 | -0.1236 | 0.086975 | 0.21515 | -0.24262 | -0.28381 | -0.11444 | -0.15564 | -0.05951 | -0.07782 | -0.11902 | -0.041198 | -0.032044 | 0.027466 | 0.10071 | -0.15106 | 0.16022 | -0.091552
280.7812 | 122.4472 | -59.1843 | 63.1668 | -0.13054 | -0.0010082 | 0.068759 | 0.068705 | -0.0043003 | 0.0012348 | -0.0016662 | 0.0078429 | -0.00046791 | 0.066415 | 0.062551 | -0.00060892 | -0.0039029 | -0.00062588 | 0.009155 | 63.089 | -59.2896 | 62.9883 | -59.0378 | -59.4406 | -59.3857 | 0.027466 | 63.1027 | -59.2026 | 63.0707 | -59.1843 | -59.3033 | -59.0515 | -0.10529 | -0.07782 | -0.17853 | -0.10071 | 0.14648 | 0.25177 | -0.20142 | -0.25635 | -0.09613 | -0.15106 | -0.018311 | -0.064087 | -0.09613 | -0.032043 | 0 | 0.018311 | 0.13275 | -0.11902 | 0.15106 | -0.10071
281.7812 | 122.4792 | -59.166 | 63.1439 | -0.046616 | 0.0015552 | 0.068759 | 0.064433 | 0.0010403 | -0.001115 | -0.0027344 | 0.032257 | 0.00038669 | 0.065988 | 0.06661 | -0.00018152 | -0.0026211 | 0.00086952 | 0.022888 | 63.0615 | -59.2941 | 63.0249 | -59.0149 | -59.3948 | -59.3491 | 0.068665 | 63.0936 | -59.2026 | 63.0936 | -59.1568 | -59.2667 | -59.0332 | -0.12817 | -0.082398 | -0.11902 | -0.036621 | 0.15106 | 0.27923 | -0.18311 | -0.22888 | -0.054932 | -0.10071 | -0.036621 | -0.050354 | -0.050354 | 0 | 0.009156 | 0.045777 | 0.13275 | -0.10071 | 0.16937 | -0.064087
282.7812 | 122.5342 | -59.1705 | 63.1439 | -0.084763 | 0.0024096 | 0.067264 | 0.066355 | -0.0017367 | 0.00080757 | -0.0023072 | 0.038361 | 0.00060029 | 0.064706 | 0.06191 | -0.00060892 | -0.0017667 | 0.0021511 | 0.036621 | 63.0615 | -59.3079 | 63.0615 | -58.9737 | -59.3719 | -59.317 | 0.11444 | 63.0844 | -59.1888 | 63.1302 | -59.1293 | -59.2575 | -58.992 | -0.13733 | -0.082398 | -0.082398 | 0 | 0.19684 | 0.33417 | -0.14648 | -0.20142 | -0.009156 | -0.064087 | -0.018311 | -0.059509 | -0.013733 | 0.045776 | 0.041199 | 0.05951 | 0.17853 | -0.086975 | 0.19684 | -0.068664
283.7812 | 122.644 | -59.0973 | 63.1531 | -0.077133 | 5.9776e-05 | 0.067904 | 0.066996 | -0.0030185 | -0.00068783 | -0.00017099 | 0.050568 | 0.00060029 | 0.062357 | 0.067464 | -0.0018905 | -0.0041165 | 0.00065592 | 0.036621 | 63.0524 | -59.2804 | 63.1073 | -58.9142 | -59.3216 | -59.2896 | 0.16022 | 63.0478 | -59.198 | 63.176 | -59.0698 | -59.2209 | -58.8547 | -0.18311 | -0.10071 | -0.045776 | 0.054932 | 0.1831 | 0.36621 | -0.19226 | -0.2243 | -0.009155 | -0.041198 | -0.10071 | -0.10529 | 0.022888 | 0.12817 | 0.027466 | 0.12817 | 0.24262 | -0.1236 | 0.34332 | -0.022888
284.7812 | 122.5754 | -59.1522 | 63.121 | -0.031357 | -0.0010082 | 0.066195 | 0.063792 | -0.0019503 | 0.00038017 | 0.00047001 | 0.0078429 | 0.0012411 | 0.064493 | 0.062337 | -0.0021041 | -0.0034757 | -0.00019868 | 0.022888 | 63.0936 | -59.2712 | 62.915 | -59.0424 | -59.2941 | -59.3399 | 0.14648 | 63.0478 | -59.2072 | 63.1805 | -59.0836 | -59.2072 | -58.9645 | -0.11902 | -0.027466 | -0.20599 | -0.17853 | 0.10986 | 0.22888 | -0.18768 | -0.14191 | -0.068665 | -0.022888 | -0.054931 | -0.073242 | 0.059509 | 0.13275 | 0.068665 | 0.1236 | 0.18768 | -0.054931 | 0.24261 | 0
285.7812 | 122.58 | -59.1522 | 63.089 | -0.031357 | 0.00027338 | 0.066195 | 0.068705 | -0.0036593 | 0.00059397 | -0.0023072 | 0.035309 | -0.00068151 | 0.065561 | 0.060415 | -0.0027451 | -0.0034757 | -0.00041228 | 0.045776 | 63.0844 | -59.2896 | 63.1119 | -58.9325 | -59.3262 | -59.2896 | 0.16022 | 63.0569 | -59.1705 | 63.1805 | -59.0698 | -59.2575 | -58.9966 | -0.13733 | -0.004577 | 0.022888 | 0.027465 | 0.21973 | 0.35706 | -0.13733 | -0.17395 | 0 | -0.036621 | -0.01831 | -0.032043 | 0.091553 | 0.1236 | 0.082398 | 0.10071 | 0.15564 | -0.10528 | 0.17395 | -0.086975
286.7812 | 122.4609 | -59.1614 | 63.1485 | -0.038986 | 0.00048698 | 0.067477 | 0.06721 | -0.0028049 | 0.00038017 | 0.0013244 | -0.010467 | 0.00017309 | 0.06492 | 0.065328 | -0.00060892 | -0.0032619 | -0.0031895 | 0.045776 | 63.0936 | -59.3124 | 63.0478 | -58.9874 | -59.4086 | -59.3262 | 0.064087 | 63.0936 | -59.198 | 63.0981 | -59.1614 | -59.285 | -59.0286 | -0.15106 | -0.054932 | -0.10071 | -0.045776 | 0.17395 | 0.32501 | -0.16479 | -0.24719 | -0.013733 | -0.09613 | -0.036621 | -0.054932 | -0.050354 | 0.004578 | 0 | 0.036621 | 0.13275 | -0.1236 | 0.16937 | -0.086975
287.7812 | 122.6028 | -59.166 | 63.0981 | -0.069504 | -0.00036742 | 0.068973 | 0.067851 | -0.0010959 | 0.0012348 | 0.00068361 | -0.034882 | 0.00017309 | 0.06257 | 0.068105 | -0.00039512 | -0.0021939 | -0.0038303 | 0.032043 | 63.0753 | -59.2896 | 63.0798 | -58.9371 | -59.2163 | -59.0698 | 0.013733 | 63.0386 | -59.1934 | 63.1851 | -59.0836 | -59.1888 | -58.9645 | -0.1236 | -0.022889 | -0.018311 | 0.004578 | 0.22888 | 0.35248 | 0.096131 | -0.050354 | 0.21973 | 0.073242 | -0.027465 | -0.05951 | 0.086975 | 0.14648 | 0.082398 | 0.10986 | 0.20142 | -0.022888 | 0.22888 | 0.004577
288.7812 | 122.493 | -59.1431 | 63.1393 | -0.038986 | -0.0010082 | 0.065768 | 0.068919 | 0.00039951 | -0.00068783 | -0.0018798 | -0.071503 | -0.00046791 | 0.065988 | 0.06661 | -0.00039512 | -0.0049709 | -0.0029757 | -0.009155 | 63.121 | -59.3399 | 62.9929 | -59.0927 | -59.4543 | -59.4223 | -0.004578 | 63.1165 | -59.2117 | 63.1073 | -59.166 | -59.3079 | -59.0515 | -0.19684 | -0.01831 | -0.14648 | -0.12817 | 0.050354 | 0.24719 | -0.27924 | -0.31128 | -0.082397 | -0.11444 | -0.068665 | -0.022888 | -0.032043 | -0.009155 | -0.022889 | 0.045776 | 0.091552 | -0.1648 | 0.16022 | -0.09613
289.7812 | 122.6166 | -59.1385 | 63.1027 | -0.046616 | 0.001128 | 0.065768 | 0.065715 | -0.00024129 | -0.0015424 | -0.0020936 | 0.041412 | 0.0012411 | 0.062143 | 0.062764 | 0.00024568 | -0.0032619 | 0.0025785 | 0.041199 | 63.0753 | -59.2896 | 63.089 | -58.9325 | -59.3262 | -59.2804 | 0.15106 | 63.0569 | -59.2026 | 63.208 | -59.0424 | -59.166 | -58.9371 | -0.15106 | -0.027466 | -0.013733 | 0.013733 | 0.20599 | 0.35706 | -0.14191 | -0.18768 | 0.009155 | -0.036621 | -0.064087 | -0.045776 | 0.10529 | 0.15106 | 0.096131 | 0.16022 | 0.20142 | -0.027466 | 0.2655 | 0.036621
290.7812 | 122.5571 | -59.166 | 63.1119 | -0.13817 | 0.00070078 | 0.068118 | 0.067851 | 0.00018591 | -0.001115 | 0.0011108 | 0.0047913 | 0.0012411 | 0.066202 | 0.06191 | -0.0014633 | -0.0047573 | -0.0014805 | -0.050354 | 63.0707 | -59.3308 | 63.0112 | -59.0378 | -59.4406 | -59.3582 | 0.009155 | 63.0524 | -59.2163 | 63.1302 | -59.1293 | -59.2438 | -58.9783 | -0.1648 | -0.041198 | -0.10071 | -0.05951 | 0.12817 | 0.29297 | -0.19226 | -0.27466 | -0.027465 | -0.10986 | -0.050354 | -0.059509 | 0.018311 | 0.07782 | 0.036622 | 0.086976 | 0.18768 | -0.077819 | 0.23804 | -0.027465
291.7812 | 122.5891 | -59.1476 | 63.0936 | -0.10002 | -0.0016492 | 0.068332 | 0.068492 | 0.00082671 | 0.00016657 | -0.00017099 | 0.0078429 | -4.0508e-05 | 0.065561 | 0.061696 | -0.00060892 | -0.0013393 | -0.0010531 | -0.036621 | 63.0524 | -59.3308 | 63.0386 | -59.0057 | -59.3994 | -59.3308 | 0.059509 | 63.0341 | -59.2209 | 63.1668 | -59.0881 | -59.2163 | -58.9508 | -0.18311 | -0.041199 | -0.054932 | -0.013733 | 0.14191 | 0.32501 | -0.18311 | -0.25177 | 0 | -0.068664 | -0.073242 | -0.059509 | 0.073242 | 0.13275 | 0.059509 | 0.13275 | 0.19684 | -0.068665 | 0.27008 | 0.004577
292.7812 | 122.5937 | -59.1614 | 63.1119 | -0.046616 | -0.0020764 | 0.067691 | 0.067851 | -0.0019503 | 0.0027302 | -0.0020936 | 0.013947 | -4.0508e-05 | 0.06257 | 0.064687 | -0.0031723 | -0.0024075 | -0.0012667 | -0.045776 | 63.0707 | -59.317 | 63.0157 | -59.0286 | -59.4177 | -59.3445 | 0.009155 | 63.0798 | -59.2117 | 63.1393 | -59.111 | -59.2346 | -58.9874 | -0.15564 | -0.041198 | -0.09613 | -0.054932 | 0.13275 | 0.28839 | -0.1831 | -0.25635 | -0.027465 | -0.10071 | -0.050354 | -0.032043 | 0.027466 | 0.059509 | 0.050354 | 0.10071 | 0.17395 | -0.073242 | 0.2243 | -0.022888
293.7812 | 122.4701 | -59.1751 | 63.1165 | -0.023727 | 0.00091438 | 0.068332 | 0.068492 | -0.00088229 | -0.0026104 | -0.001239 | -0.047089 | 0.0023093 | 0.065561 | 0.064901 | -0.0023177 | -0.0039029 | -0.0021213 | -0.027466 | 63.1119 | -59.3079 | 62.9837 | -59.0881 | -59.4498 | -59.4223 | 0.009155 | 63.1073 | -59.2346 | 63.1851 | -59.3536 | -59.4681 | -59.111 | -0.13275 | -0.004578 | -0.13275 | -0.12817 | 0.086975 | 0.21973 | -0.24719 | -0.27466 | -0.11444 | -0.14191 | -0.059509 | -0.009155 | 0.068665 | 0.07782 | -0.17853 | -0.11902 | 0.064087 | -0.29297 | 0.1236 | -0.23346
294.7812 | 122.58 | -59.1614 | 63.0798 | -0.13054 | -0.00015382 | 0.068118 | 0.066996 | -0.0043003 | 0.0018756 | -0.001239 | -0.0013123 | 0.00081389 | 0.065134 | 0.057638 | -0.0023177 | -0.0032619 | 0.0025785 | -0.054932 | 63.0844 | -59.3079 | 63.0295 | -59.0286 | -59.4177 | -59.3445 | 0.041199 | 63.0661 | -59.2209 | 63.1668 | -59.1019 | -59.2255 | -58.9645 | -0.14648 | 0.004578 | -0.050354 | -0.054932 | 0.13275 | 0.27924 | -0.1831 | -0.25635 | -0.036621 | -0.10986 | -0.059509 | -0.013733 | 0.086975 | 0.10071 | 0.059509 | 0.11902 | 0.19684 | -0.064087 | 0.25635 | -0.004578
295.7812 | 122.5388 | -59.166 | 63.0707 | -0.10765 | 0.0013416 | 0.068118 | 0.067851 | 0.00061311 | 0.001662 | -0.0040162 | -0.0043641 | -0.0011087 | 0.065561 | 0.058065 | -0.0016769 | -0.0021939 | -0.0025485 | -0.045776 | 63.089 | -59.2987 | 63.0203 | -59.0286 | -59.4131 | -59.3582 | 0 | 63.0478 | -59.198 | 63.1393 | -59.1019 | -59.2484 | -59.0057 | -0.13275 | 0.01831 | -0.050354 | -0.068664 | 0.13733 | 0.27008 | -0.19226 | -0.24719 | -0.059509 | -0.11444 | -0.032043 | -0.022888 | 0.068664 | 0.091552 | 0.064087 | 0.09613 | 0.16022 | -0.082397 | 0.19226 | -0.050354
296.7812 | 122.5479 | -59.166 | 63.0844 | -0.061874 | 0.00027338 | 0.069186 | 0.067637 | -0.00045509 | 0.0014484 | 0.00025641 | -0.028778 | 0.00017309 | 0.064066 | 0.063192 | -0.0016769 | -0.0034757 | 0.00022852 | -0.064087 | 63.0478 | -59.3124 | 63.0112 | -59.0378 | -59.317 | -59.3124 | -0.036621 | 63.0432 | -59.2209 | 63.121 | -59.1431 | -59.2667 | -59.024 | -0.14648 | -0.036621 | -0.073243 | -0.036622 | 0.12817 | 0.27466 | -0.14648 | -0.15106 | 0 | -0.004578 | -0.054931 | -0.041199 | 0.036621 | 0.07782 | 0.022889 | 0.07782 | 0.14191 | -0.10071 | 0.19684 | -0.045777
297.7812 | 122.5296 | -59.1934 | 63.0569 | -0.12291 | 0.0015552 | 0.067691 | 0.068705 | -0.00088229 | 0.00016657 | 0.00047001 | -0.034882 | -0.00025411 | 0.064706 | 0.063833 | 0.0013139 | -0.0039029 | -0.0021213 | -0.059509 | 63.0341 | -59.2987 | 63.0066 | -59.0378 | -59.4498 | -59.3811 | -0.009155 | 63.0615 | -59.23 | 63.1302 | -59.1614 | -59.2941 | -59.0332 | -0.10529 | -0.022888 | -0.050354 | -0.027466 | 0.15564 | 0.26093 | -0.18768 | -0.25635 | -0.082398 | -0.15106 | -0.036622 | 0.004577 | 0.073242 | 0.068665 | 0.032043 | 0.068665 | 0.16022 | -0.10071 | 0.19684 | -0.064086
298.7812 | 122.4701 | -59.1751 | 63.1119 | -0.0084687 | 0.00027338 | 0.069186 | 0.066355 | 0.0021085 | -0.00047423 | 0.00068361 | -0.047089 | 0.00038669 | 0.064279 | 0.062337 | -0.00039512 | -0.0041165 | -0.0025485 | 0.032043 | 63.0615 | -59.2987 | 63.0478 | -59.0012 | -59.3903 | -59.3491 | 0.073242 | 63.1027 | -59.198 | 63.1027 | -59.1705 | -59.2712 | -59.0286 | -0.1236 | -0.050354 | -0.064086 | -0.013732 | 0.17395 | 0.29755 | -0.17395 | -0.21515 | -0.050354 | -0.091553 | -0.022888 | -0.009155 | -0.009155 | 0 | 0.004578 | 0.027466 | 0.14648 | -0.09613 | 0.16937 | -0.073242
299.7812 | 122.493 | -59.1934 | 63.0981 | -0.1458 | 0.00048698 | 0.067477 | 0.068064 | -0.0047275 | 0.00059397 | 0.0023926 | -0.037933 | 0.00017309 | 0.06492 | 0.06191 | -0.00082252 | -0.0024075 | -0.0012667 | -0.064087 | 63.0432 | -59.317 | 63.0112 | -59.0424 | -59.4635 | -59.3719 | -0.036621 | 63.0386 | -59.2117 | 63.1302 | -59.1476 | -59.2804 | -59.0286 | -0.1236 | -0.054932 | -0.086976 | -0.032044 | 0.15106 | 0.27466 | -0.17853 | -0.27008 | -0.054931 | -0.14648 | -0.018311 | -0.05951 | 0.032043 | 0.091553 | 0.045776 | 0.064087 | 0.1648 | -0.086976 | 0.18311 | -0.068665
300.7812 | 122.5204 | -59.1797 | 63.0753 | -0.10765 | 0.00027338 | 0.067264 | 0.066569 | -0.00066869 | 0.00038017 | 0.003247 | -0.03183 | -0.0011087 | 0.063638 | 0.061696 | -0.00018152 | -0.0041165 | -0.0025485 | -0.07782 | 63.0295 | -59.3216 | 63.0249 | -59.0561 | -59.4589 | -59.3628 | -0.054932 | 63.0661 | -59.2209 | 63.1256 | -59.1522 | -59.2987 | -59.0195 | -0.14191 | -0.045776 | -0.050354 | -0.004578 | 0.1236 | 0.2655 | -0.18311 | -0.27924 | -0.041199 | -0.13733 | -0.041199 | -0.009155 | 0.050354 | 0.059509 | 0.027465 | 0.068664 | 0.16022 | -0.11902 | 0.20142 | -0.07782
301.7812 | 122.5204 | -59.1888 | 63.0844 | -0.069504 | 0.0015552 | 0.06705 | 0.068705 | 0.00039951 | 0.0018756 | -0.00017099 | -0.03183 | -0.00025411 | 0.063852 | 0.061696 | -0.0012497 | -0.0013393 | -0.00041228 | -0.082397 | 63.0478 | -59.3033 | 62.9974 | -59.0378 | -59.4727 | -59.3857 | -0.054932 | 63.0432 | -59.2163 | 63.1485 | -59.1568 | -59.285 | -59.024 | -0.11444 | -0.036621 | -0.086975 | -0.050354 | 0.15106 | 0.2655 | -0.19684 | -0.28381 | -0.082397 | -0.16937 | -0.027466 | -0.041199 | 0.064087 | 0.10529 | 0.032044 | 0.05951 | 0.16479 | -0.09613 | 0.19226 | -0.068664
302.7812 | 122.4976 | -59.1797 | 63.0981 | -0.092392 | 0.0017688 | 0.068332 | 0.068919 | -0.0025913 | 0.00059397 | 4.2808e-05 | -0.0074159 | 0.00038669 | 0.065561 | 0.06191 | -0.0033859 | -0.0036893 | 1.4916e-05 | -0.064087 | 63.0478 | -59.3216 | 63.0112 | -59.0378 | -59.436 | -59.3811 | -0.032043 | 63.0569 | -59.2438 | 63.089 | -59.1339 | -59.2484 | -59.0378 | -0.14191 | -0.050354 | -0.086976 | -0.036622 | 0.14191 | 0.28381 | -0.20142 | -0.25635 | -0.05951 | -0.11444 | -0.064087 | -0.041199 | -0.009156 | 0.032043 | 0.045776 | 0.10986 | 0.14191 | -0.068665 | 0.20599 | -0.004578
303.7969 | 122.4792 | -59.1797 | 63.121 | -0.054245 | -0.0014356 | 0.067477 | 0.071269 | 0.0025357 | 0.00080757 | 0.00068361 | -0.14474 | -4.0508e-05 | 0.065347 | 0.068319 | -0.0012497 | -5.7742e-05 | -0.0034031 | -0.036621 | 63.0753 | -59.3582 | 62.9883 | -59.079 | -59.5001 | -59.4452 | -0.041199 | 63.1027 | -59.2392 | 63.089 | -59.198 | -59.3628 | -59.0744 | -0.17853 | -0.045777 | -0.13275 | -0.086975 | 0.10071 | 0.27924 | -0.2655 | -0.32043 | -0.086975 | -0.14191 | -0.05951 | -0.018311 | -0.032044 | -0.013733 | -0.018311 | 0.041199 | 0.10529 | -0.18311 | 0.16479 | -0.1236
304.7812 | 122.5754 | -59.1568 | 63.0661 | -0.084763 | 0.0013416 | 0.070895 | 0.068064 | -0.0021641 | -0.00047423 | -0.0016662 | -0.059296 | -0.00089511 | 0.063425 | 0.062124 | -0.0010361 | -0.00069854 | 0.0019375 | -0.050354 | 63.0249 | -59.3353 | 63.0341 | -59.0057 | -59.436 | -59.3582 | 0.050354 | 63.0249 | -59.2667 | 63.1668 | -59.0973 | -59.2484 | -58.9874 | -0.17853 | -0.041199 | -0.032043 | 0.009156 | 0.15106 | 0.32959 | -0.20142 | -0.27924 | -0.022888 | -0.10071 | -0.10986 | -0.041199 | 0.10071 | 0.14191 | 0.059509 | 0.16937 | 0.16937 | -0.091553 | 0.27924 | 0.018311
305.7812 | 122.5479 | -59.1614 | 63.0615 | -0.20683 | 0.002837 | 0.070041 | 0.069773 | -0.00088229 | -0.0013288 | 0.00025641 | -0.028778 | 0.00017309 | 0.062143 | 0.06191 | 3.2078e-05 | -0.0034757 | 0.0010831 | -0.041199 | 63.0569 | -59.3079 | 63.0478 | -58.9828 | -59.404 | -59.3445 | 0.018311 | 63.0295 | -59.1888 | 63.1622 | -59.1339 | -59.2529 | -58.9783 | -0.14648 | -0.004577 | -0.013732 | -0.009155 | 0.17853 | 0.32501 | -0.1831 | -0.24262 | -0.036621 | -0.096131 | -0.027466 | -0.032043 | 0.10071 | 0.13275 | 0.027466 | 0.054932 | 0.18311 | -0.091553 | 0.21057 | -0.064087
306.7812 | 122.5571 | -59.1751 | 63.0569 | -0.16106 | 0.0015552 | 0.068759 | 0.068278 | 0.00018591 | -0.0013288 | 0.00089721 | -0.034882 | -0.0013223 | 0.062357 | 0.055715 | -0.0010361 | -0.0021939 | 0.0010831 | -0.050354 | 63.0112 | -59.3308 | 63.089 | -59.0332 | -59.3994 | -59.3262 | 0 | 63.0295 | -59.2255 | 63.1714 | -59.111 | -59.2667 | -58.9737 | -0.15564 | -0.045777 | 0.032043 | 0.07782 | 0.14191 | 0.29755 | -0.15106 | -0.2243 | 0.004578 | -0.068664 | -0.050354 | -0.027466 | 0.11444 | 0.14191 | 0.064087 | 0.11444 | 0.20142 | -0.091553 | 0.25177 | -0.041199
307.7812 | 122.5433 | -59.1705 | 63.0615 | -0.13054 | 0.00070078 | 0.0694 | 0.068705 | 0.0012539 | 0.0029438 | 0.00047001 | -0.059296 | 0.00017309 | 0.064279 | 0.06191 | -0.0014633 | -0.0036893 | 0.00086952 | -0.073242 | 63.0203 | -59.2987 | 63.0295 | -58.9966 | -59.4269 | -59.3491 | 0.018311 | 63.0524 | -59.2117 | 63.1393 | -59.1705 | -59.2667 | -58.9966 | -0.12817 | -0.041198 | -0.032043 | 0.009155 | 0.17395 | 0.30212 | -0.17853 | -0.25635 | -0.050354 | -0.12817 | -0.041199 | -0.009155 | 0.07782 | 0.086975 | 0 | 0.041199 | 0.17395 | -0.096131 | 0.21515 | -0.054932
308.7812 | 122.5204 | -59.1751 | 63.121 | -0.084763 | -0.00015382 | 0.067691 | 0.068064 | 0.0012539 | 0.00080757 | 0.0017518 | -0.025726 | -0.00089511 | 0.064066 | 0.063405 | -0.0010361 | -0.0030483 | -0.0036167 | 0.032043 | 63.1027 | -59.2941 | 63.0615 | -58.9737 | -59.3765 | -59.3445 | 0.07782 | 63.0798 | -59.198 | 63.1439 | -59.1202 | -59.2392 | -59.0012 | -0.11902 | -0.018311 | -0.05951 | -0.041199 | 0.20142 | 0.32043 | -0.16937 | -0.20142 | -0.050354 | -0.082398 | -0.022888 | -0.041199 | 0.022888 | 0.064087 | 0.054932 | 0.07782 | 0.17395 | -0.064087 | 0.19684 | -0.041199
309.7812 | 122.525 | -59.1843 | 63.0524 | -0.1992 | 0.0015552 | 0.066623 | 0.068064 | 0.0040311 | -4.7032e-05 | 0.0026062 | -0.044037 | 0.00081389 | 0.064493 | 0.067891 | -0.00018152 | -0.00048494 | -0.0019077 | 0.027466 | 63.0203 | -59.3491 | 63.0524 | -58.9828 | -59.4177 | -59.3536 | 0.013733 | 63.0386 | -59.2117 | 63.176 | -59.1293 | -59.2529 | -59.0012 | -0.16479 | -0.032043 | 0 | 0.032043 | 0.20142 | 0.36621 | -0.16937 | -0.23346 | -0.004578 | -0.068665 | -0.027466 | -0.013733 | 0.1236 | 0.13733 | 0.054932 | 0.082398 | 0.18311 | -0.068665 | 0.21057 | -0.041199
310.7812 | 122.5708 | -59.1568 | 63.0524 | -0.13817 | -0.0014356 | 0.067904 | 0.066996 | -0.0023777 | -4.7032e-05 | 0.0045288 | -0.077606 | 0.00038669 | 0.062357 | 0.066182 | -0.0014633 | -0.0011257 | 0.00065592 | -0.045776 | 63.0157 | -59.3262 | 62.9333 | -58.8684 | -59.2255 | -59.2896 | 0.022888 | 63.0432 | -59.1934 | 63.176 | -59.1156 | -59.2712 | -58.9737 | -0.16937 | -0.036621 | -0.11902 | -0.082397 | 0.28839 | 0.45776 | -0.13275 | -0.068665 | 0.036621 | 0.10071 | -0.036621 | -0.009155 | 0.1236 | 0.13275 | 0.041198 | 0.077819 | 0.1831 | -0.11444 | 0.21973 | -0.07782
311.7812 | 122.5388 | -59.1476 | 63.0936 | -0.1458 | -0.0010082 | 0.064914 | 0.067637 | -0.00045509 | 0.0014484 | -0.00038459 | 0.0047913 | -0.00025411 | 0.063852 | 0.059347 | -0.0021041 | -0.0024075 | 0.0021511 | 0.21057 | 63.1439 | -59.2438 | 63.0661 | -58.9462 | -59.3582 | -59.2987 | 0.11444 | 63.0615 | -59.1888 | 63.1805 | -59.0881 | -59.198 | -58.9508 | -0.09613 | 0.050354 | -0.027466 | -0.07782 | 0.20142 | 0.29755 | -0.15106 | -0.21057 | -0.054932 | -0.11444 | -0.041199 | -0.032044 | 0.086975 | 0.11902 | 0.059509 | 0.10071 | 0.19684 | -0.050354 | 0.23804 | -0.009155
312.7812 | 122.5845 | -59.1888 | 63.0844 | -0.11528 | -0.00015382 | 0.069827 | 0.069987 | -0.0010959 | -0.0013288 | 0.0011108 | -0.092865 | 0.00060029 | 0.062143 | 0.061269 | -0.0033859 | -0.0017667 | 0.0036465 | -0.091553 | 63.0478 | -59.3491 | 62.9974 | -59.0744 | -59.4818 | -59.4131 | -0.054932 | 63.0478 | -59.2209 | 63.1851 | -59.1476 | -59.2896 | -58.9783 | -0.16022 | -0.036621 | -0.086975 | -0.050354 | 0.11444 | 0.27466 | -0.2243 | -0.29297 | -0.064087 | -0.13275 | -0.032043 | -0.036621 | 0.10071 | 0.13733 | 0.041199 | 0.073242 | 0.21057 | -0.10071 | 0.24262 | -0.068665
313.7969 | 122.4747 | -59.1751 | 63.1027 | -0.038986 | -0.0014356 | 0.066409 | 0.066569 | 0.00039951 | 0.0014484 | 0.0011108 | -0.08371 | 0.00081389 | 0.064706 | 0.067251 | -0.00082252 | -0.0015529 | 0.0070645 | -0.013733 | 63.1393 | -59.3079 | 62.8647 | -59.0973 | -59.436 | -59.4406 | -0.032043 | 63.1256 | -59.2529 | 63.1027 | -59.166 | -59.3079 | -59.0607 | -0.13275 | 0.036621 | -0.23804 | -0.27466 | 0.07782 | 0.21057 | -0.2655 | -0.26093 | -0.13275 | -0.12817 | -0.07782 | 0.022888 | 0 | -0.022888 | 0.009155 | 0.086975 | 0.11444 | -0.13275 | 0.19226 | -0.054931
314.7812 | 122.5525 | -59.1568 | 63.0615 | -0.37468 | 0.0013416 | 0.067264 | 0.068919 | 0.00039951 | 0.001662 | -0.001239 | -0.053192 | 0.0016685 | 0.062357 | 0.065114 | -0.0014633 | -0.0011257 | 0.00065592 | 0 | 62.9837 | -59.2987 | 63.0524 | -58.9783 | -59.4086 | -59.3582 | 0 | 63.0432 | -59.2163 | 63.1668 | -59.1293 | -59.2484 | -58.9828 | -0.14191 | -0.077819 | -0.009155 | 0.068664 | 0.17853 | 0.32044 | -0.20142 | -0.25177 | -0.059509 | -0.10986 | -0.05951 | -0.01831 | 0.10529 | 0.1236 | 0.027466 | 0.086976 | 0.17395 | -0.091553 | 0.23346 | -0.032043
315.7812 | 122.5662 | -59.166 | 63.0707 | -0.11528 | -0.00015382 | 0.068118 | 0.066996 | -0.00066869 | 0.00016657 | 0.00025641 | -0.0013123 | -0.00089511 | 0.064706 | 0.064687 | -0.00082252 | -0.00027134 | -0.00041228 | -0.073242 | 63.0615 | -59.3124 | 63.0066 | -59.0424 | -59.4269 | -59.3491 | -0.036621 | 63.0524 | -59.2346 | 63.1256 | -59.1568 | -59.2529 | -58.992 | -0.14648 | -0.009156 | -0.064087 | -0.054931 | 0.1236 | 0.27008 | -0.18311 | -0.26093 | -0.036621 | -0.11444 | -0.068664 | -0.018311 | 0.054931 | 0.073242 | 0.009156 | 0.07782 | 0.17395 | -0.086975 | 0.24262 | -0.018311
316.7812 | 122.5754 | -59.1705 | 63.0569 | -0.077133 | 0.0015552 | 0.069613 | 0.069773 | 0.0023221 | 0.00016657 | -0.0014526 | -0.059296 | -0.00068151 | 0.065134 | 0.068319 | -0.0044541 | -5.7742e-05 | -0.0016941 | -0.045776 | 63.0386 | -59.3124 | 62.8967 | -59.1064 | -59.166 | -59.23 | 0.018311 | 63.002 | -59.2117 | 63.1714 | -59.1248 | -59.2438 | -58.9783 | -0.14191 | -0.018311 | -0.16022 | -0.14191 | 0.064087 | 0.20599 | -0.05951 | 0.004577 | 0.082397 | 0.14648 | -0.041199 | -0.054932 | 0.11444 | 0.16937 | 0.045776 | 0.086975 | 0.19226 | -0.073242 | 0.23346 | -0.032043
317.7812 | 122.5662 | -59.1888 | 63.0524 | -0.13054 | -0.00058102 | 0.068759 | 0.068919 | 0.0023221 | 0.00059397 | 0.0019654 | -0.044037 | -0.0023905 | 0.062357 | 0.067464 | -0.00082252 | -0.0011257 | -0.0034031 | -0.050354 | 63.0249 | -59.3079 | 63.0432 | -58.9966 | -59.5322 | -59.5413 | 0.13275 | 63.0249 | -59.2163 | 63.1989 | -59.1156 | -59.2438 | -58.9645 | -0.11902 | -0.027466 | -0.009155 | 0.018311 | 0.19226 | 0.31128 | -0.35248 | -0.34332 | -0.23346 | -0.2243 | -0.027466 | -0.027466 | 0.14648 | 0.17395 | 0.073242 | 0.10071 | 0.2243 | -0.054931 | 0.25177 | -0.027465
318.7812 | 122.5204 | -59.1568 | 63.0844 | -0.038986 | -0.00036742 | 0.067691 | 0.067424 | -0.00066869 | -0.00047423 | -0.0018798 | 0.013947 | -4.0508e-05 | 0.063211 | 0.064473 | -0.00018152 | -0.0032619 | 0.0012967 | -0.064087 | 63.0524 | -59.3124 | 63.089 | -59.111 | -59.3994 | -59.3308 | -0.032043 | 63.0707 | -59.2346 | 63.1393 | -59.1339 | -59.2712 | -59.0103 | -0.15564 | -0.032044 | 0.004577 | 0.036621 | 0.045776 | 0.20142 | -0.17395 | -0.24262 | -0.018311 | -0.086975 | -0.07782 | -0.013733 | 0.054931 | 0.068664 | 0.022888 | 0.10071 | 0.14648 | -0.11444 | 0.2243 | -0.036621
319.7812 | 122.525 | -59.1797 | 63.0798 | -0.10765 | 0.0013416 | 0.069827 | 0.069133 | -0.00066869 | 0.00016657 | -0.00059819 | -0.10812 | -0.0011087 | 0.063638 | 0.070028 | -0.0010361 | -0.0017667 | 0.0010831 | -0.10071 | 63.0707 | -59.3536 | 62.9929 | -59.0561 | -59.4681 | -59.4086 | -0.036621 | 63.0341 | -59.2255 | 63.1805 | -59.1385 | -59.3033 | -58.9874 | -0.17395 | -0.009155 | -0.086975 | -0.07782 | 0.1236 | 0.29755 | -0.22888 | -0.28839 | -0.054931 | -0.11444 | -0.045777 | -0.045776 | 0.10071 | 0.14648 | 0.041198 | 0.086975 | 0.19226 | -0.1236 | 0.23804 | -0.07782
320.7812 | 122.5616 | -59.166 | 63.0524 | -0.11528 | -0.00015382 | 0.068118 | 0.067851 | 0.00018591 | -0.00026063 | 0.0019654 | -0.047089 | 0.0014547 | 0.063638 | 0.061269 | -0.00018152 | 0.0020785 | 0.00044212 | -0.041199 | 63.0386 | -59.3262 | 63.0386 | -58.9966 | -59.4315 | -59.3582 | -0.004578 | 63.0157 | -59.2346 | 63.0569 | -58.9737 | -59.079 | -58.8867 | -0.16022 | -0.013733 | -0.013733 | 0 | 0.16937 | 0.32959 | -0.19226 | -0.2655 | -0.032043 | -0.10529 | -0.068664 | -0.036621 | 0.004578 | 0.041199 | 0.19226 | 0.26093 | 0.27924 | 0.086976 | 0.3479 | 0.15564
321.7812 | 122.4838 | -59.166 | 63.1119 | -0.092392 | 0.00091438 | 0.067691 | 0.066569 | -2.7692e-05 | -4.7032e-05 | 0.00047001 | -0.016571 | 0.00017309 | 0.063211 | 0.06661 | -0.0018905 | -0.0028347 | -0.00083948 | 0.036621 | 63.0981 | -59.2896 | 63.0524 | -58.9783 | -59.3903 | -59.3124 | 0.082397 | 63.0707 | -59.2117 | 63.1485 | -59.1248 | -59.2575 | -58.992 | -0.1236 | -0.013732 | -0.059509 | -0.045777 | 0.18768 | 0.31128 | -0.14648 | -0.2243 | -0.022888 | -0.10071 | -0.045776 | -0.041198 | 0.036622 | 0.07782 | 0.041199 | 0.086975 | 0.17395 | -0.091552 | 0.21973 | -0.045776
322.7812 | 122.5296 | -59.1614 | 63.0981 | -0.084763 | 0.001128 | 0.068973 | 0.068064 | -2.7692e-05 | 0.00038017 | -0.0010254 | -0.05014 | -0.00068151 | 0.06492 | 0.071096 | -0.00018152 | -5.7742e-05 | 0.0012967 | -0.050354 | 63.0386 | -59.3079 | 63.0615 | -58.9874 | -59.4223 | -59.3536 | 0.009155 | 63.0295 | -59.2117 | 63.1714 | -59.1156 | -59.2484 | -58.9828 | -0.14648 | -0.05951 | -0.036622 | 0.022888 | 0.17395 | 0.32043 | -0.19226 | -0.26093 | -0.045777 | -0.11444 | -0.050354 | -0.068665 | 0.073242 | 0.14191 | 0.045776 | 0.09613 | 0.17853 | -0.086975 | 0.22888 | -0.036621
323.7812 | 122.5845 | -59.1888 | 63.0615 | -0.061874 | 0.001128 | 0.068759 | 0.067637 | -0.00045509 | 0.00059397 | -0.00059819 | -0.03183 | -0.00025411 | 0.067056 | 0.064046 | 3.2078e-05 | -0.0019803 | 0.0027921 | -0.050354 | 63.0341 | -59.2941 | 63.0524 | -58.9737 | -59.4223 | -59.3399 | 0.009155 | 63.0295 | -59.23 | 63.1622 | -59.111 | -59.2346 | -58.9645 | -0.10529 | -0.027465 | -0.009155 | 0.01831 | 0.21515 | 0.32043 | -0.15106 | -0.23346 | -0.045777 | -0.12817 | -0.041199 | -0.032043 | 0.10071 | 0.13275 | 0.07782 | 0.11902 | 0.2243 | -0.045776 | 0.2655 | -0.004577
324.7812 | 122.5616 | -59.1797 | 63.0753 | -0.10002 | 0.0013416 | 0.068118 | 0.066569 | 0.0036039 | 0.00038017 | 0.00047001 | -0.028778 | 0.00038669 | 0.065134 | 0.062764 | 0.00045928 | -0.0013393 | -0.0023349 | -0.054932 | 63.0249 | -59.3124 | 63.0341 | -58.9783 | -59.4315 | -59.3536 | 0 | 63.0295 | -59.2117 | 63.1805 | -59.1248 | -59.2575 | -58.9966 | -0.13275 | -0.050354 | -0.041198 | 0.009156 | 0.20142 | 0.33417 | -0.17395 | -0.25177 | -0.041199 | -0.11902 | -0.032044 | -0.045776 | 0.10529 | 0.15106 | 0.054931 | 0.086975 | 0.18311 | -0.07782 | 0.21515 | -0.045776
325.7812 | 122.5937 | -59.1843 | 63.0478 | -0.11528 | 0.00091438 | 0.067691 | 0.068705 | -0.00088229 | 0.0025166 | -0.0014526 | -0.047089 | -0.0011087 | 0.063638 | 0.06661 | -0.0023177 | -0.00069854 | -0.0029757 | -0.041199 | 63.0661 | -59.3033 | 63.0478 | -59.0195 | -59.4315 | -59.3582 | 0 | 63.0112 | -59.2117 | 63.1805 | -59.0927 | -59.4498 | -59.0881 | -0.11902 | 0.01831 | 0 | -0.01831 | 0.1648 | 0.28381 | -0.17395 | -0.24719 | -0.054931 | -0.12817 | -0.027466 | -0.036622 | 0.13275 | 0.16937 | 0.091553 | 0.11902 | 0.09613 | -0.2655 | 0.1236 | -0.23804
326.7812 | 122.6486 | -59.1843 | 63.0615 | -0.11528 | -0.00015382 | 0.069186 | 0.068919 | -0.0010959 | 0.0020892 | 0.0026062 | -0.077606 | 0.00081389 | 0.063852 | 0.064901 | 0.0011001 | -0.0017667 | 0.0030057 | -0.032043 | 63.0203 | -59.3399 | 63.0386 | -58.9828 | -59.4223 | -59.3582 | -0.013733 | 63.0432 | -59.198 | 63.2538 | -59.317 | -59.3262 | -58.9325 | -0.15564 | -0.041198 | -0.022888 | 0.01831 | 0.20142 | 0.35706 | -0.17395 | -0.23804 | -0.01831 | -0.082397 | -0.013733 | -0.01831 | 0.19226 | 0.21057 | -0.13275 | -0.11902 | 0.25177 | -0.14191 | 0.2655 | -0.12817
327.7812 | 122.5204 | -59.1705 | 63.0478 | -0.10002 | 5.9776e-05 | 0.070041 | 0.068919 | -0.0019503 | 0.00016657 | 0.0028198 | -0.040985 | 0.00017309 | 0.064066 | 0.058492 | -0.00060892 | -0.0019803 | 0.0025785 | -0.073242 | 63.0112 | -59.3033 | 63.0386 | -58.992 | -59.4269 | -59.3353 | 0.004578 | 63.0249 | -59.2026 | 63.1577 | -59.1202 | -59.2484 | -58.992 | -0.13275 | -0.036622 | -0.009156 | 0.027466 | 0.17853 | 0.31128 | -0.16479 | -0.25635 | -0.032043 | -0.1236 | -0.032044 | -0.022889 | 0.10986 | 0.13275 | 0.050354 | 0.082398 | 0.17853 | -0.07782 | 0.21057 | -0.045776
328.7812 | 122.525 | -59.1614 | 63.0707 | -0.16106 | -0.001222 | 0.066195 | 0.069346 | -0.0051547 | -0.00068783 | -0.001239 | -0.03183 | 0.0014547 | 0.065347 | 0.062337 | 3.2078e-05 | -0.0019803 | -0.0023349 | -0.09613 | 63.0432 | -59.3262 | 63.0066 | -59.0378 | -59.4498 | -59.3765 | -0.073242 | 63.0249 | -59.2346 | 63.1348 | -59.166 | -59.2667 | -58.9966 | -0.16479 | -0.027466 | -0.064087 | -0.036621 | 0.1236 | 0.28839 | -0.21515 | -0.28839 | -0.050354 | -0.1236 | -0.073242 | -0.045777 | 0.064087 | 0.10986 | -0.004578 | 0.068664 | 0.1648 | -0.10529 | 0.23804 | -0.032044
329.7812 | 122.5525 | -59.1705 | 63.0707 | -0.084763 | 5.9776e-05 | 0.067477 | 0.067424 | -0.0013095 | 0.00038017 | 0.0011108 | -0.019623 | -0.00089511 | 0.061929 | 0.063192 | -0.0038131 | -0.0015529 | -0.00019868 | -0.091553 | 63.0524 | -59.3308 | 62.9791 | -59.0424 | -59.4635 | -59.3811 | -0.041199 | 63.0524 | -59.23 | 63.1256 | -59.1476 | -59.2484 | -58.9966 | -0.16022 | -0.018311 | -0.091553 | -0.073242 | 0.12817 | 0.28839 | -0.21057 | -0.29297 | -0.050354 | -0.13275 | -0.05951 | -0.018311 | 0.054931 | 0.073242 | 0.022888 | 0.082398 | 0.17395 | -0.07782 | 0.23346 | -0.01831
330.7812 | 122.525 | -59.1476 | 63.0798 | -0.084763 | -0.00015382 | 0.067691 | 0.069133 | 0.0023221 | 0.0014484 | -0.0023072 | -0.013519 | 0.0018821 | 0.06492 | 0.065969 | -0.00082252 | -0.0028347 | 0.0010831 | -0.073242 | 63.0569 | -59.3216 | 63.0157 | -59.0515 | -59.4727 | -59.3582 | -0.059509 | 63.0203 | -59.2209 | 63.121 | -59.1614 | -59.2575 | -58.9828 | -0.17395 | -0.022888 | -0.064087 | -0.041199 | 0.09613 | 0.27008 | -0.21057 | -0.32501 | -0.036621 | -0.15106 | -0.073242 | -0.059509 | 0.041199 | 0.10071 | -0.013733 | 0.059509 | 0.16479 | -0.10986 | 0.23804 | -0.036621
331.7812 | 122.5204 | -59.1751 | 63.1027 | -0.077133 | 0.00070078 | 0.067904 | 0.072337 | 0.00039951 | 0.0014484 | 0.00089721 | -0.12338 | 0.0010275 | 0.064493 | 0.067251 | -0.0018905 | -0.0021939 | -0.0036167 | -0.004578 | 63.0798 | -59.3262 | 63.0157 | -59.0424 | -59.4452 | -59.404 | 0 | 63.089 | -59.2163 | 63.121 | -59.2026 | -59.3262 | -59.0469 | -0.15106 | -0.022888 | -0.086975 | -0.064087 | 0.13275 | 0.28381 | -0.22888 | -0.27008 | -0.07782 | -0.11902 | -0.041199 | -0.013733 | 0.018311 | 0.032044 | -0.027466 | 0.013733 | 0.12817 | -0.15106 | 0.16937 | -0.10986
332.7812 | 122.5342 | -59.1843 | 63.0478 | -0.10002 | -0.001222 | 0.068118 | 0.067851 | 0.0014675 | 0.00059397 | -0.0010254 | -0.025726 | 0.00081389 | 0.064279 | 0.06191 | 0.0015275 | -0.0034757 | 0.00044212 | -0.045776 | 63.0112 | -59.3124 | 63.0386 | -58.992 | -59.4406 | -59.3582 | 0.009155 | 63.0066 | -59.2117 | 63.1531 | -59.1614 | -59.2896 | -59.0012 | -0.12817 | -0.036622 | -0.009156 | 0.027466 | 0.19226 | 0.32043 | -0.17395 | -0.25635 | -0.045776 | -0.12817 | -0.027466 | -0.041199 | 0.10529 | 0.14648 | 0.022888 | 0.050354 | 0.18311 | -0.10529 | 0.21057 | -0.07782
333.7812 | 122.7173 | -59.111 | 63.1485 | -0.092392 | 0.00091438 | 0.067691 | 0.066355 | -0.0028049 | -0.0013288 | 0.00025641 | 0.0047913 | -0.00089511 | 0.063425 | 0.063619 | -0.00060892 | -0.0032619 | -0.00019868 | 0.027466 | 63.0753 | -59.2758 | 63.1119 | -58.905 | -59.3216 | -59.2575 | 0.14191 | 63.0478 | -59.2072 | 63.1989 | -59.0652 | -59.1797 | -59.0103 | -0.16479 | -0.073243 | -0.036622 | 0.036621 | 0.20599 | 0.37079 | -0.14648 | -0.21057 | 0.018311 | -0.045776 | -0.09613 | -0.10071 | 0.050354 | 0.15106 | 0.045776 | 0.14191 | 0.10071 | -0.068664 | 0.19684 | 0.027466
334.7812 | 122.5662 | -59.1751 | 63.089 | 0.14412 | -0.001222 | 0.068545 | 0.069346 | 0.00061311 | 0.0018756 | 0.00025641 | 0.0078429 | -0.0013223 | 0.063425 | 0.069387 | -0.0012497 | -0.0030483 | 0.00086952 | -0.059509 | 63.0569 | -59.3124 | 63.0157 | -59.0286 | -59.4452 | -59.3308 | 0 | 63.0569 | -59.1843 | 63.1439 | -59.1293 | -59.2484 | -59.0149 | -0.13733 | -0.032043 | -0.073242 | -0.041199 | 0.14648 | 0.28381 | -0.15564 | -0.27008 | -0.018311 | -0.13275 | -0.009155 | -0.032043 | 0.054932 | 0.086975 | 0.045777 | 0.054932 | 0.16022 | -0.073242 | 0.16937 | -0.064087
335.7812 | 122.5662 | -59.1705 | 63.121 | -0.13817 | 0.00070078 | 0.067691 | 0.06721 | -0.0017367 | -0.00047423 | -0.0031616 | 0.032257 | -0.00089511 | 0.06492 | 0.058706 | -0.00018152 | -0.0026211 | -0.0046847 | -0.068665 | 63.0661 | -59.3216 | 63.002 | -59.0378 | -59.4086 | -59.3399 | -0.009155 | 63.0615 | -59.1843 | 63.1531 | -59.1293 | -59.2438 | -59.0057 | -0.15106 | -0.054932 | -0.11902 | -0.064087 | 0.13275 | 0.28381 | -0.16937 | -0.23804 | -0.018311 | -0.086975 | -0.013733 | -0.05951 | 0.032043 | 0.091553 | 0.041199 | 0.054932 | 0.16479 | -0.073242 | 0.17853 | -0.059509
336.7812 | 122.5571 | -59.1476 | 63.1165 | -0.1992 | 0.001128 | 0.069827 | 0.068278 | -0.0025913 | 0.00038017 | 4.2808e-05 | -0.019623 | -0.00025411 | 0.063425 | 0.062551 | -0.0025315 | -0.0015529 | 0.0010831 | -0.050354 | 63.0524 | -59.317 | 63.0157 | -59.024 | -59.4223 | -59.3399 | -0.004578 | 63.0798 | -59.2255 | 63.1531 | -59.111 | -59.2484 | -58.9874 | -0.16937 | -0.064087 | -0.10071 | -0.036621 | 0.1236 | 0.29297 | -0.19226 | -0.27466 | -0.022888 | -0.10529 | -0.07782 | -0.036621 | 0.036621 | 0.073242 | 0.036621 | 0.11444 | 0.16022 | -0.10071 | 0.23804 | -0.022888
337.7812 | 122.5525 | -59.1568 | 63.0707 | -0.11528 | 0.0015552 | 0.068545 | 0.066355 | -0.0047275 | -0.00047423 | 0.00047001 | -0.0043641 | 0.00017309 | 0.06257 | 0.067891 | 0.00045928 | -0.0026211 | 1.4916e-05 | -0.064087 | 63.0707 | -59.2941 | 63.0112 | -59.0332 | -59.4269 | -59.3536 | 0.004578 | 63.0524 | -59.2163 | 63.1622 | -59.1293 | -59.2438 | -58.992 | -0.13733 | 0 | -0.05951 | -0.05951 | 0.1236 | 0.26093 | -0.19684 | -0.27008 | -0.05951 | -0.13275 | -0.05951 | -0.018311 | 0.091552 | 0.10986 | 0.027466 | 0.086976 | 0.16479 | -0.086975 | 0.22431 | -0.027465
338.7812 | 122.5433 | -59.1431 | 63.0936 | -0.092392 | -0.00036742 | 0.067904 | 0.066783 | -0.0072909 | -4.7032e-05 | -0.00038459 | 0.010895 | 0.0020957 | 0.064706 | 0.063192 | -0.0012497 | -0.00027134 | -0.0042575 | 0.027466 | 63.0478 | -59.2804 | 63.0981 | -58.8959 | -59.317 | -59.2667 | 0.14648 | 63.0661 | -59.1888 | 63.2858 | -58.8409 | -59.4315 | -59.1248 | -0.13733 | -0.045776 | 0.004578 | 0.050354 | 0.24719 | 0.38452 | -0.1236 | -0.17395 | 0.013733 | -0.036621 | -0.045777 | -0.027466 | 0.19226 | 0.21973 | 0.30212 | 0.3479 | 0.01831 | -0.28839 | 0.064087 | -0.24262
339.7812 | 122.5479 | -59.1705 | 63.089 | -0.13054 | -0.00058102 | 0.067691 | 0.068492 | -0.00066869 | -0.00047423 | -0.0020936 | -0.0074159 | 0.00017309 | 0.063852 | 0.061696 | -0.0016769 | -0.0041165 | 0.00086952 | -0.059509 | 63.0661 | -59.3079 | 62.9974 | -59.0378 | -59.4223 | -59.3399 | -0.004578 | 63.0844 | -59.198 | 63.1577 | -59.1339 | -59.2529 | -59.0057 | -0.13733 | -0.022888 | -0.091552 | -0.068664 | 0.13275 | 0.27008 | -0.16937 | -0.25177 | -0.032044 | -0.11444 | -0.027466 | -0.004577 | 0.068665 | 0.073242 | 0.036621 | 0.064087 | 0.16479 | -0.082398 | 0.19226 | -0.054932
340.7812 | 122.525 | -59.1614 | 63.0844 | -0.069504 | 0.0019824 | 0.070468 | 0.066569 | -0.0032321 | -0.001115 | 0.00047001 | -0.019623 | -4.0508e-05 | 0.065347 | 0.06426 | 0.0013139 | -0.0026211 | -0.0019077 | -0.082397 | 63.0569 | -59.2941 | 63.0203 | -59.0652 | -59.4223 | -59.3399 | -0.013733 | 63.0478 | -59.2163 | 63.1485 | -59.1293 | -59.2438 | -59.0012 | -0.13275 | -0.027466 | -0.064087 | -0.036621 | 0.09613 | 0.22888 | -0.17853 | -0.26093 | -0.045777 | -0.12817 | -0.054932 | -0.036621 | 0.064087 | 0.10071 | 0.032044 | 0.086976 | 0.16022 | -0.082397 | 0.21515 | -0.027465
341.7969 | 122.4609 | -59.1751 | 63.1302 | -0.069504 | -0.0014356 | 0.068118 | 0.069133 | -2.7692e-05 | -4.7032e-05 | 4.2808e-05 | -0.14169 | 0.0031637 | 0.066202 | 0.068105 | -0.0038131 | 0.00036966 | -0.0036167 | -0.004578 | 63.0615 | -59.3262 | 63.0295 | -59.0424 | -59.4452 | -59.3948 | -0.13275 | 63.1073 | -59.2163 | 63.1348 | -59.1614 | -59.2941 | -59.0332 | -0.15106 | -0.068665 | -0.10071 | -0.032043 | 0.13275 | 0.28381 | -0.21973 | -0.27008 | -0.068664 | -0.11902 | -0.041199 | -0.022888 | 0.004578 | 0.027466 | 0.013733 | 0.054932 | 0.14191 | -0.11902 | 0.18311 | -0.077819
342.7812 | 122.5891 | -59.1339 | 63.0844 | -0.11528 | -0.00015382 | 0.068759 | 0.066783 | 0.0010403 | -0.0023968 | 0.00025641 | 0.016998 | -0.0011087 | 0.066843 | 0.060201 | -0.0023177 | -0.0026211 | -0.00062588 | 0 | 63.0753 | -59.2896 | 63.0981 | -58.9279 | -59.3445 | -59.2758 | 0.16479 | 63.0432 | -59.1705 | 63.2309 | -59.0469 | -59.1888 | -58.9371 | -0.15564 | -0.009156 | 0.013733 | 0.022889 | 0.20599 | 0.36163 | -0.14191 | -0.21057 | 0.013733 | -0.054931 | -0.036621 | -0.041199 | 0.14648 | 0.18768 | 0.086975 | 0.1236 | 0.19684 | -0.054932 | 0.23346 | -0.018311
343.7812 | 122.4518 | -59.1843 | 63.1393 | -0.12291 | 0.00091438 | 0.067691 | 0.068492 | -0.0043003 | -0.0013288 | -0.0016662 | -0.074554 | 0.00081389 | 0.064279 | 0.061056 | 0.00045928 | 0.00015586 | 0.0040737 | -0.11902 | 63.1165 | -59.3536 | 62.97 | -59.0881 | -59.4681 | -59.4177 | -0.009155 | 63.1073 | -59.2163 | 63.1165 | -59.1797 | -59.3079 | -59.0378 | -0.16937 | -0.022888 | -0.16937 | -0.14648 | 0.09613 | 0.2655 | -0.23346 | -0.28381 | -0.064087 | -0.11444 | -0.032044 | -0.032043 | -0.022888 | 0.009155 | 0.004578 | 0.036622 | 0.14648 | -0.1236 | 0.17853 | -0.091552
344.7812 | 122.5708 | -59.1522 | 63.121 | -0.10765 | -0.00058102 | 0.068118 | 0.066569 | -0.0013095 | 0.00080757 | 0.0011108 | 0.032257 | -0.0023905 | 0.064493 | 0.058919 | 0.00024568 | -0.0049709 | 1.4916e-05 | 0.009155 | 63.0798 | -59.2804 | 63.0753 | -58.9417 | -59.3628 | -59.2667 | 0.13275 | 63.0295 | -59.1934 | 63.2126 | -59.0424 | -59.1476 | -58.9371 | -0.12817 | -0.041199 | -0.045777 | -0.004578 | 0.21057 | 0.33875 | -0.11444 | -0.21057 | 0.013733 | -0.082397 | -0.041198 | -0.091553 | 0.091552 | 0.1831 | 0.10986 | 0.15106 | 0.21515 | 0.004578 | 0.25635 | 0.045776
345.7812 | 122.4838 | -59.166 | 63.0798 | -0.092392 | 5.9776e-05 | 0.068332 | 0.066783 | -0.0043003 | -0.0013288 | -0.00081179 | -0.0013123 | 0.0010275 | 0.065134 | 0.058492 | 0.00088648 | -0.0032619 | 0.00086952 | -0.091553 | 63.0478 | -59.2987 | 63.0203 | -59.0469 | -59.4498 | -59.3628 | -0.013733 | 63.0569 | -59.2026 | 63.1485 | -59.0332 | -59.2346 | -58.9737 | -0.13275 | -0.032043 | -0.059509 | -0.027466 | 0.11902 | 0.25177 | -0.19684 | -0.28381 | -0.064087 | -0.15106 | -0.036621 | -0.022888 | 0.068665 | 0.091553 | 0.13275 | 0.16937 | 0.19226 | -0.068664 | 0.22888 | -0.032043
346.7812 | 122.5525 | -59.1568 | 63.0981 | -0.10765 | 0.00048698 | 0.068118 | 0.068705 | -0.0017367 | 0.0020892 | -0.0023072 | -0.0043641 | -0.00025411 | 0.064279 | 0.065114 | 0.00067288 | -0.0039029 | 0.0010831 | -0.073242 | 63.0707 | -59.3124 | 62.9974 | -59.0332 | -59.436 | -59.3582 | -0.032043 | 63.0478 | -59.1888 | 63.1302 | -59.1568 | -59.2667 | -59.0057 | -0.15564 | -0.027466 | -0.10071 | -0.073242 | 0.1236 | 0.27924 | -0.20142 | -0.27924 | -0.045776 | -0.1236 | -0.032044 | -0.050354 | 0.032043 | 0.082397 | 0 | 0.032044 | 0.15106 | -0.10986 | 0.18311 | -0.07782
347.7812 | 122.5525 | -59.1522 | 63.0844 | -0.092392 | -0.0020764 | 0.067904 | 0.067424 | 0.0010403 | 0.0010212 | -0.00038459 | -0.019623 | -0.0015359 | 0.066629 | 0.058278 | 0.00045928 | -0.0013393 | 0.0017239 | -0.059509 | 63.1027 | -59.3262 | 63.0157 | -59.0469 | -59.4406 | -59.3811 | -0.018311 | 63.0341 | -59.2026 | 63.1439 | -59.1385 | -59.2255 | -58.9874 | -0.17395 | 0.01831 | -0.068665 | -0.086975 | 0.10529 | 0.27924 | -0.22888 | -0.28839 | -0.054932 | -0.11444 | -0.050354 | -0.050354 | 0.059509 | 0.10986 | 0.013733 | 0.064087 | 0.1648 | -0.073242 | 0.21515 | -0.022888
348.7812 | 122.4747 | -59.1476 | 63.1393 | -0.046616 | 0.00091438 | 0.067904 | 0.068064 | 0.0012539 | -0.00026063 | 0.0011108 | -0.025726 | -0.00046791 | 0.064279 | 0.067678 | -0.00060892 | -0.0028347 | 1.4916e-05 | 0.045776 | 63.1165 | -59.2804 | 63.0615 | -58.9737 | -59.3719 | -59.3308 | 0.086975 | 63.0753 | -59.2072 | 63.1393 | -59.1431 | -59.2529 | -59.0057 | -0.13275 | -0.022888 | -0.07782 | -0.054932 | 0.17395 | 0.3067 | -0.18311 | -0.2243 | -0.050354 | -0.091552 | -0.059509 | -0.064087 | 0 | 0.064087 | 0.004578 | 0.064087 | 0.14191 | -0.10529 | 0.20142 | -0.045777
349.7812 | 122.5662 | -59.1522 | 63.0798 | -0.092392 | 5.9776e-05 | 0.067477 | 0.068705 | -0.00088229 | 0.00059397 | 0.0026062 | 0.0017395 | 0.0012411 | 0.065347 | 0.067251 | -0.00018152 | -0.0024075 | 0.0012967 | -0.054932 | 63.0524 | -59.317 | 63.0249 | -59.0057 | -59.4131 | -59.3262 | 0.041199 | 63.0341 | -59.2117 | 63.1668 | -59.0927 | -59.23 | -58.9645 | -0.16479 | -0.027466 | -0.054932 | -0.027466 | 0.14648 | 0.31128 | -0.17395 | -0.26093 | -0.009155 | -0.09613 | -0.059509 | -0.045776 | 0.086975 | 0.13275 | 0.05951 | 0.11902 | 0.18768 | -0.07782 | 0.24719 | -0.018311
350.7812 | 122.5296 | -59.1614 | 63.0936 | -0.054245 | 5.9776e-05 | 0.067477 | 0.066569 | -0.0028049 | 0.0027302 | -0.0014526 | -0.0043641 | 0.00081389 | 0.064706 | 0.062978 | -0.00018152 | -0.0028347 | -0.0025485 | -0.064087 | 63.0844 | -59.3216 | 63.0432 | -59.0195 | -59.4177 | -59.3536 | 0.009155 | 63.0341 | -59.198 | 63.1531 | -59.0973 | -59.2209 | -58.9645 | -0.16022 | -0.009155 | -0.050354 | -0.041199 | 0.14191 | 0.30212 | -0.19226 | -0.25635 | -0.032044 | -0.096131 | -0.036621 | -0.059509 | 0.059509 | 0.11902 | 0.064087 | 0.10071 | 0.19684 | -0.059509 | 0.23346 | -0.022888
351.7812 | 122.5342 | -59.1522 | 63.0936 | -0.031357 | -0.00015382 | 0.068118 | 0.069346 | -0.0017367 | 0.0014484 | -0.001239 | -0.019623 | -0.00025411 | 0.063211 | 0.059987 | -0.0016769 | -0.0017667 | 0.0032193 | -0.091553 | 63.0615 | -59.2896 | 63.0112 | -59.0515 | -59.404 | -59.3262 | -0.004578 | 63.0386 | -59.2163 | 63.1348 | -59.1248 | -59.2392 | -59.0057 | -0.13733 | -0.032044 | -0.082398 | -0.050354 | 0.10071 | 0.23804 | -0.17395 | -0.25177 | -0.036621 | -0.11444 | -0.064087 | -0.054932 | 0.041199 | 0.096131 | 0.027466 | 0.091553 | 0.14648 | -0.086975 | 0.21057 | -0.022888
352.7812 | 122.5388 | -59.1751 | 63.0753 | -0.13054 | 0.00070078 | 0.065768 | 0.069133 | -0.0021641 | 0.0025166 | 0.0013244 | -0.03183 | 0.0012411 | 0.06257 | 0.066396 | -0.00039512 | -0.0028347 | -0.0034031 | -0.07782 | 63.0524 | -59.3079 | 63.0066 | -59.0469 | -59.4543 | -59.3628 | -0.041199 | 63.0432 | -59.2255 | 63.1302 | -59.1522 | -59.2621 | -58.992 | -0.13275 | -0.022888 | -0.068664 | -0.045776 | 0.12817 | 0.26093 | -0.18768 | -0.27924 | -0.054932 | -0.14648 | -0.050354 | -0.032043 | 0.054932 | 0.086975 | 0.022888 | 0.073242 | 0.18311 | -0.086975 | 0.23346 | -0.036621
353.7812 | 122.5571 | -59.1293 | 63.0981 | -0.10002 | -0.001222 | 0.066623 | 0.066783 | -0.00045509 | 0.0014484 | -0.0027344 | 0.035309 | 0.0014547 | 0.064066 | 0.065114 | -0.0014633 | -0.0034757 | 0.0021511 | 0.032043 | 63.0615 | -59.2896 | 63.2721 | -58.7631 | -59.7061 | -59.4223 | 0.07782 | 63.0615 | -59.1797 | 63.2034 | -59.0836 | -59.198 | -58.9462 | -0.16022 | -0.036622 | 0.17395 | 0.21057 | 0.36621 | 0.52643 | -0.29297 | -0.57678 | -0.13275 | -0.41656 | -0.050354 | -0.036622 | 0.10528 | 0.14191 | 0.045776 | 0.09613 | 0.18311 | -0.068665 | 0.23346 | -0.018311
354.7812 | 122.5113 | -59.1431 | 63.0753 | -0.10002 | 0.00027338 | 0.067691 | 0.069133 | 0.00082671 | 0.001662 | -0.0014526 | -0.013519 | 0.00060029 | 0.063852 | 0.068532 | -0.0040267 | -0.0024075 | -0.0034031 | -0.073242 | 63.0661 | -59.3124 | 63.0157 | -59.0561 | -59.4498 | -59.3765 | -0.041199 | 63.0524 | -59.2209 | 63.1073 | -59.0195 | -59.2667 | -58.9691 | -0.16937 | -0.009155 | -0.059509 | -0.050354 | 0.086975 | 0.25635 | -0.23346 | -0.3067 | -0.064087 | -0.13733 | -0.07782 | -0.022888 | 0.032044 | 0.054932 | 0.1236 | 0.20142 | 0.17395 | -0.1236 | 0.25177 | -0.045777
355.7812 | 122.5571 | -59.166 | 63.0936 | -0.18394 | 0.00070078 | 0.067264 | 0.067851 | -0.0021641 | 0.0020892 | -0.00038459 | -0.019623 | 0.00038669 | 0.063638 | 0.063405 | 0.00067288 | 0.00036966 | -0.0012667 | -0.032043 | 63.0753 | -59.3079 | 63.0203 | -59.0332 | -59.4177 | -59.3353 | 0 | 63.089 | -59.2072 | 63.1485 | -59.1156 | -59.2438 | -59.0012 | -0.14191 | -0.018311 | -0.073242 | -0.054931 | 0.13275 | 0.27466 | -0.16937 | -0.25177 | -0.027466 | -0.10986 | -0.041198 | -0.004578 | 0.054932 | 0.05951 | 0.050354 | 0.091552 | 0.16479 | -0.077819 | 0.20599 | -0.036621
356.7812 | 122.4792 | -59.1522 | 63.1119 | -0.016098 | 0.001128 | 0.06705 | 0.069773 | -0.00024129 | -0.0013288 | 0.0017518 | -0.12338 | 0.0025229 | 0.064066 | 0.062978 | -0.00018152 | 0.0012241 | -0.0016941 | 0.009155 | 63.1073 | -59.3216 | 63.0203 | -59.0515 | -59.4177 | -59.3903 | 0.054932 | 63.1073 | -59.1797 | 63.1256 | -59.1431 | -59.2941 | -59.0195 | -0.16937 | -0.004577 | -0.091552 | -0.086975 | 0.10071 | 0.27008 | -0.23804 | -0.2655 | -0.068665 | -0.096131 | -0.027465 | -0.004577 | 0.013733 | 0.01831 | 0.009156 | 0.036621 | 0.13275 | -0.14191 | 0.16022 | -0.11444
357.7812 | 122.5342 | -59.166 | 63.0936 | -0.038986 | 0.0024096 | 0.067691 | 0.067851 | 0.00018591 | 0.00080757 | -0.0023072 | -0.044037 | 0.00038669 | 0.065347 | 0.061269 | 0.0017411 | -0.00091214 | 0.00022852 | -0.050354 | 63.0524 | -59.3079 | 63.002 | -59.0424 | -59.4452 | -59.2896 | -0.027466 | 63.0569 | -59.2117 | 63.1302 | -59.1202 | -59.2438 | -58.992 | -0.14191 | -0.041199 | -0.091553 | -0.050354 | 0.1236 | 0.2655 | -0.1236 | -0.27923 | 0.01831 | -0.13733 | -0.045776 | -0.036621 | 0.036621 | 0.073242 | 0.045777 | 0.091553 | 0.17395 | -0.077819 | 0.21973 | -0.032043
358.7812 | 122.5525 | -59.1522 | 63.1027 | -0.084763 | 0.001128 | 0.068332 | 0.065501 | -0.00088229 | -0.00047423 | -0.0010254 | 0.0017395 | -4.0508e-05 | 0.063638 | 0.064687 | -0.00039512 | -0.0026211 | 0.00086952 | 0.041199 | 63.0707 | -59.3216 | 63.0798 | -58.96 | -59.3399 | -59.3079 | 0.091553 | 63.0844 | -59.2026 | 63.1577 | -59.111 | -59.2346 | -58.9966 | -0.16937 | -0.032043 | -0.022888 | 0.009155 | 0.19226 | 0.36163 | -0.15564 | -0.18768 | 0.013733 | -0.018311 | -0.050354 | -0.01831 | 0.054932 | 0.073242 | 0.041199 | 0.091553 | 0.15564 | -0.082397 | 0.20599 | -0.032043
359.7812 | 122.5159 | -59.1614 | 63.0569 | -0.084763 | 0.00070078 | 0.067264 | 0.066569 | -0.0040867 | -0.0015424 | -0.0027344 | -0.044037 | 0.00017309 | 0.064066 | 0.066182 | -0.0018905 | -0.00069854 | -0.0029757 | -0.059509 | 63.0341 | -59.3216 | 63.0249 | -58.9966 | -59.4498 | -59.3536 | 0.004578 | 63.0203 | -59.2072 | 63.1393 | -59.1385 | -59.2667 | -59.0057 | -0.16022 | -0.022888 | -0.032044 | -0.009156 | 0.1648 | 0.32501 | -0.19226 | -0.28839 | -0.032044 | -0.12817 | -0.045776 | -0.036621 | 0.082397 | 0.11902 | 0.022888 | 0.068664 | 0.15564 | -0.10529 | 0.20142 | -0.05951
360.7812 | 122.5021 | -59.1614 | 63.0478 | -0.077133 | 0.0026234 | 0.066409 | 0.066783 | -2.7692e-05 | -4.7032e-05 | -0.0027344 | -0.056244 | 0.00038669 | 0.06727 | 0.066396 | -0.00018152 | -0.0047573 | 0.0019375 | -0.064087 | 63.0341 | -59.3033 | 63.0386 | -58.9783 | -59.4223 | -59.3628 | 0.013733 | 63.0112 | -59.2255 | 63.1485 | -59.1568 | -59.2712 | -58.992 | -0.14191 | -0.013733 | -0.009156 | 0.004577 | 0.18311 | 0.32501 | -0.20142 | -0.26093 | -0.059509 | -0.11902 | -0.064087 | -0.036622 | 0.10071 | 0.13733 | 0.004578 | 0.068665 | 0.16937 | -0.10986 | 0.23346 | -0.045776
361.7812 | 122.5113 | -59.1614 | 63.0936 | -0.061874 | -0.0010082 | 0.067477 | 0.067424 | -0.0032321 | -0.00026063 | -0.00081179 | -0.040985 | 0.0012411 | 0.063425 | 0.06426 | -0.00018152 | -0.00069854 | -0.0025485 | -0.073242 | 63.0478 | -59.2804 | 63.0157 | -59.0332 | -59.4406 | -59.3719 | -0.050354 | 63.0341 | -59.2072 | 63.1485 | -59.1568 | -59.2667 | -59.0195 | -0.11902 | -0.045776 | -0.07782 | -0.032044 | 0.12817 | 0.24719 | -0.21057 | -0.27924 | -0.091552 | -0.16022 | -0.045776 | -0.059509 | 0.054932 | 0.11444 | 0.004578 | 0.050354 | 0.14191 | -0.10529 | 0.18768 | -0.05951
362.7812 | 122.5021 | -59.1751 | 63.0981 | -0.077133 | -0.00079462 | 0.068332 | 0.068705 | -0.0010959 | 0.0023028 | -0.0018798 | -0.03183 | 0.00038669 | 0.064493 | 0.063405 | 0.00088648 | -0.0019803 | -0.00083948 | -0.07782 | 63.0569 | -59.3033 | 63.0066 | -59.0424 | -59.4131 | -59.3536 | -0.086975 | 63.0524 | -59.198 | 63.1393 | -59.166 | -59.2758 | -59.0012 | -0.12817 | -0.041199 | -0.091553 | -0.050354 | 0.13275 | 0.26093 | -0.17853 | -0.23804 | -0.050354 | -0.10986 | -0.022888 | -0.045777 | 0.041198 | 0.086975 | 0.009155 | 0.032043 | 0.17395 | -0.10071 | 0.19684 | -0.07782
363.7812 | 122.58 | -59.1522 | 63.089 | -0.10002 | -0.00229 | 0.067264 | 0.066996 | -0.0038729 | -0.00068783 | 0.00068361 | -0.0074159 | 0.00060029 | 0.065775 | 0.06191 | -0.0023177 | -0.0021939 | -0.0036167 | -0.009155 | 63.0615 | -59.285 | 63.0615 | -58.9645 | -59.3536 | -59.2941 | 0.10529 | 63.0661 | -59.1888 | 63.1943 | -59.0561 | -59.1705 | -58.9325 | -0.13275 | -0.027466 | -0.027466 | 0 | 0.18768 | 0.32043 | -0.14191 | -0.20142 | -0.009155 | -0.068665 | -0.036621 | -0.022888 | 0.10529 | 0.12817 | 0.096131 | 0.13275 | 0.21973 | -0.01831 | 0.25635 | 0.018311
364.7812 | 122.4976 | -59.1522 | 63.1073 | -0.077133 | -0.00079462 | 0.069613 | 0.06721 | -0.0013095 | 0.00080757 | 0.00047001 | -0.044037 | -0.00089511 | 0.067911 | 0.06191 | 0.00067288 | -0.0028347 | 0.0023647 | -0.086975 | 63.0569 | -59.3033 | 62.9974 | -59.0561 | -59.4269 | -59.3536 | -0.041199 | 63.0661 | -59.2392 | 63.1348 | -59.1385 | -59.2575 | -59.0103 | -0.15106 | -0.050354 | -0.10986 | -0.059509 | 0.096131 | 0.24719 | -0.20142 | -0.27466 | -0.050354 | -0.1236 | -0.086975 | -0.041199 | 0.027466 | 0.068665 | 0.013733 | 0.10071 | 0.14191 | -0.10528 | 0.22888 | -0.01831
365.7812 | 122.5616 | -59.1705 | 63.0707 | -0.16106 | 0.00091438 | 0.067264 | 0.06956 | -0.00024129 | -0.0023968 | -0.00017099 | -0.044037 | -0.00025411 | 0.064706 | 0.066182 | -0.0010361 | -0.0028347 | -0.0031895 | -0.082397 | 63.0341 | -59.3262 | 63.0066 | -59.0286 | -59.4315 | -59.3628 | -0.032043 | 63.0432 | -59.2392 | 63.1302 | -59.1522 | -59.2575 | -59.0012 | -0.15564 | -0.036621 | -0.064087 | -0.027466 | 0.14191 | 0.29755 | -0.19226 | -0.26093 | -0.036621 | -0.10529 | -0.068665 | -0.027466 | 0.059509 | 0.086975 | 0.01831 | 0.086975 | 0.16937 | -0.086975 | 0.23804 | -0.01831
366.7969 | 122.4564 | -59.1568 | 63.1073 | -0.0084687 | -0.00058102 | 0.067904 | 0.07191 | 0.0050991 | 0.00016657 | 0.0028198 | -0.12949 | 0.0014547 | 0.065988 | 0.073873 | -0.0018905 | -0.0011257 | 0.00022852 | -0.036621 | 63.0661 | -59.3445 | 63.0203 | -59.0332 | -59.4498 | -59.4131 | 0.013733 | 63.0753 | -59.2255 | 63.1393 | -59.1751 | -59.317 | -59.024 | -0.18768 | -0.041199 | -0.086975 | -0.045776 | 0.1236 | 0.31128 | -0.25635 | -0.29297 | -0.068665 | -0.10529 | -0.068665 | -0.032044 | 0.032043 | 0.064087 | -0.018311 | 0.050354 | 0.13275 | -0.16022 | 0.20142 | -0.091553
367.7812 | 122.5479 | -59.166 | 63.0524 | -0.10002 | -0.00015382 | 0.067477 | 0.067424 | -0.00045509 | 0.00059397 | -0.001239 | -0.03183 | -0.00046791 | 0.063638 | 0.065328 | -0.0016769 | -0.0015529 | -0.00083948 | -0.009155 | 63.0112 | -59.3445 | 63.0157 | -58.9508 | -59.3948 | -59.3674 | 0.018311 | 63.0157 | -59.2117 | 63.1943 | -59.111 | -59.2346 | -58.9554 | -0.17853 | -0.041199 | -0.036621 | 0.004578 | 0.21515 | 0.39368 | -0.20142 | -0.22888 | -0.022889 | -0.050354 | -0.045776 | -0.036621 | 0.14191 | 0.17853 | 0.054932 | 0.10071 | 0.21057 | -0.068664 | 0.25635 | -0.022888
368.7812 | 122.5159 | -59.1568 | 63.0661 | -0.11528 | 0.00027338 | 0.068118 | 0.06721 | -0.0017367 | 0.0012348 | 0.0013244 | -0.028778 | 0.00017309 | 0.063638 | 0.06426 | 0.0019547 | -0.0039029 | 0.0021511 | -0.054932 | 63.0386 | -59.3079 | 63.0157 | -58.9783 | -59.4223 | -59.3536 | 0.009155 | 63.0295 | -59.2255 | 63.1577 | -59.1476 | -59.2621 | -58.9966 | -0.15106 | -0.027466 | -0.050354 | -0.022888 | 0.17853 | 0.32959 | -0.19684 | -0.2655 | -0.045777 | -0.11444 | -0.068665 | -0.036621 | 0.091553 | 0.12817 | 0.009155 | 0.07782 | 0.16022 | -0.10529 | 0.22888 | -0.036621
369.7812 | 122.5525 | -59.1843 | 63.0753 | -0.16869 | 0.0026234 | 0.068118 | 0.067851 | -0.00045509 | 0.0025166 | -0.00081179 | -0.022675 | -0.00089511 | 0.064706 | 0.062337 | -0.00039512 | -0.0043301 | -0.0014805 | -0.086975 | 63.0249 | -59.2987 | 63.002 | -59.024 | -59.4589 | -59.3719 | -0.036621 | 63.0249 | -59.2438 | 63.1119 | -59.1568 | -59.2575 | -58.992 | -0.11444 | -0.050354 | -0.073242 | -0.022888 | 0.16022 | 0.27466 | -0.18768 | -0.27466 | -0.073242 | -0.16022 | -0.059509 | -0.050354 | 0.036621 | 0.086975 | 0.027466 | 0.086975 | 0.19226 | -0.073242 | 0.25177 | -0.013733
370.7812 | 122.5067 | -59.1568 | 63.0661 | -0.10002 | 0.00048698 | 0.067904 | 0.069133 | -0.0034457 | 0.00038017 | 0.00047001 | -0.03183 | 0.0016685 | 0.063638 | 0.06426 | 0.00045928 | -0.0024075 | 0.0045011 | -0.082397 | 63.0341 | -59.3216 | 63.0157 | -59.0057 | -59.4406 | -59.3491 | -0.004578 | 63.0295 | -59.2346 | 63.1348 | -59.1522 | -59.2667 | -59.0057 | -0.16479 | -0.032043 | -0.050354 | -0.018311 | 0.15106 | 0.31586 | -0.19226 | -0.28381 | -0.027466 | -0.11902 | -0.07782 | -0.036621 | 0.068665 | 0.10529 | 0.004577 | 0.082397 | 0.15106 | -0.10986 | 0.22888 | -0.032044
371.7812 | 122.5479 | -59.1705 | 63.0569 | -0.12291 | 0.0013416 | 0.06705 | 0.06721 | 0.00061311 | 0.00080757 | 0.002179 | -0.016571 | 0.0010275 | 0.063425 | 0.059133 | -0.00082252 | -0.0028347 | 0.00086952 | -0.050354 | 63.0432 | -59.3124 | 63.0295 | -58.9737 | -59.4269 | -59.3399 | 0.009155 | 63.0157 | -59.2209 | 63.1577 | -59.1431 | -59.2575 | -58.9828 | -0.14191 | -0.013733 | -0.027466 | -0.013733 | 0.19684 | 0.33874 | -0.16937 | -0.25635 | -0.027466 | -0.11444 | -0.050354 | -0.041199 | 0.10071 | 0.14191 | 0.027466 | 0.07782 | 0.18768 | -0.086975 | 0.23804 | -0.036621
372.7812 | 122.525 | -59.1705 | 63.0524 | -0.10002 | 0.00070078 | 0.067264 | 0.068278 | -0.00066869 | 0.00059397 | 4.2808e-05 | -0.05014 | -4.0508e-05 | 0.062357 | 0.05956 | 0.0019547 | -0.00048494 | 0.0034329 | -0.032043 | 63.0569 | -59.2941 | 63.0432 | -58.9691 | -59.4131 | -59.3445 | 0 | 63.0157 | -59.2026 | 63.1805 | -59.1293 | -59.2438 | -58.9554 | -0.1236 | 0.004578 | -0.009155 | -0.013733 | 0.20142 | 0.32501 | -0.17395 | -0.24262 | -0.050354 | -0.11902 | -0.032044 | -0.036621 | 0.12817 | 0.1648 | 0.041199 | 0.073243 | 0.21515 | -0.073242 | 0.24719 | -0.041198
373.7812 | 122.58 | -59.1476 | 63.0478 | -0.22972 | 0.00070078 | 0.068332 | 0.068278 | 0.0018949 | 0.0040118 | -0.0010254 | -0.03183 | 0.0010275 | 0.065561 | 0.062551 | -0.00082252 | -0.0019803 | -0.0031895 | -0.050354 | 63.0249 | -59.3353 | 63.1393 | -58.8455 | -59.4543 | -59.3399 | 0.009155 | 63.0341 | -59.2117 | 63.1943 | -59.1064 | -59.2529 | -58.9691 | -0.18768 | -0.022889 | 0.091552 | 0.11444 | 0.30212 | 0.48981 | -0.19226 | -0.3067 | -0.004578 | -0.11902 | -0.064087 | -0.013733 | 0.14648 | 0.16022 | 0.041199 | 0.10529 | 0.17853 | -0.10529 | 0.24262 | -0.041199
374.7812 | 122.5525 | -59.1751 | 63.0661 | -0.084763 | 0.00048698 | 0.068332 | 0.068919 | -0.0017367 | 0.0027302 | -0.00059819 | -0.089813 | 0.0012411 | 0.062143 | 0.063192 | -0.00082252 | -0.0028347 | 0.0010831 | -0.10071 | 63.0478 | -59.3353 | 62.9883 | -59.0469 | -59.4727 | -59.3903 | -0.041199 | 63.0432 | -59.2255 | 63.1714 | -59.1431 | -59.2758 | -58.9554 | -0.16022 | -0.01831 | -0.07782 | -0.05951 | 0.12817 | 0.28839 | -0.21515 | -0.29755 | -0.054932 | -0.13733 | -0.050354 | -0.022888 | 0.10529 | 0.12817 | 0.032044 | 0.082398 | 0.21973 | -0.10071 | 0.27008 | -0.050354
375.7812 | 122.5571 | -59.1751 | 62.9929 | -0.092392 | 0.00048698 | 0.069827 | 0.06956 | -0.0017367 | 0.0010212 | -0.00081179 | -0.034882 | 0.00038669 | 0.065988 | 0.064046 | -0.00039512 | -0.0017667 | 0.0010831 | -0.064087 | 63.0524 | -59.3262 | 63.0295 | -58.9874 | -59.4223 | -59.3491 | -0.009155 | 63.0386 | -59.2026 | 63.1668 | -59.1248 | -59.2529 | -58.9645 | -0.15106 | 0.059509 | 0.036621 | -0.022888 | 0.18768 | 0.33875 | -0.17395 | -0.24719 | -0.022888 | -0.09613 | -0.027466 | 0.045776 | 0.17395 | 0.12817 | 0.050354 | 0.07782 | 0.21057 | -0.07782 | 0.23804 | -0.050354
376.7812 | 122.5296 | -59.1064 | 63.1119 | -0.023727 | 0.0030506 | 0.066409 | 0.065715 | -0.0015231 | 0.0012348 | 0.002179 | -0.0043641 | -0.0011087 | 0.065561 | 0.064901 | 0.00024568 | -0.0032619 | -0.0027621 | 0.027466 | 63.1027 | -59.2804 | 63.0798 | -58.9554 | -59.3536 | -59.2804 | 0.10986 | 63.0753 | -59.1843 | 63.1714 | -59.0927 | -59.1934 | -58.9462 | -0.17395 | -0.009155 | -0.032043 | -0.022888 | 0.15106 | 0.32501 | -0.17395 | -0.24719 | 0 | -0.073242 | -0.07782 | -0.036621 | 0.05951 | 0.096131 | 0.013733 | 0.091553 | 0.16022 | -0.086975 | 0.23804 | -0.009155
377.7812 | 122.5616 | -59.1705 | 63.0753 | -0.084763 | -0.00058102 | 0.067691 | 0.066996 | -0.0034457 | 0.0029438 | 0.00089721 | -0.034882 | -0.0013223 | 0.062784 | 0.065328 | -0.0031723 | -0.0013393 | 0.0019375 | -0.068665 | 63.0432 | -59.317 | 62.97 | -59.1888 | -59.3536 | -59.1705 | -0.004578 | 63.0066 | -59.1888 | 63.176 | -59.1293 | -59.2255 | -58.9737 | -0.14648 | -0.032043 | -0.10529 | -0.073242 | -0.018311 | 0.12817 | 0 | -0.18311 | 0.14648 | -0.036621 | -0.018311 | -0.068664 | 0.10071 | 0.16937 | 0.041199 | 0.05951 | 0.19684 | -0.054932 | 0.21515 | -0.036621
378.7812 | 122.5662 | -59.1431 | 63.0615 | -0.069504 | -0.00036742 | 0.067264 | 0.06721 | 0.0010403 | 0.0012348 | -0.00038459 | -0.065399 | 0.00038669 | 0.06492 | 0.064687 | -0.00060892 | -0.0011257 | -0.0034031 | -0.041199 | 63.0112 | -59.317 | 63.0341 | -58.9874 | -59.4315 | -59.3491 | -0.022888 | 62.9837 | -59.2484 | 63.1668 | -59.079 | -59.2484 | -58.9737 | -0.17395 | -0.050354 | -0.027465 | 0.022889 | 0.15564 | 0.32959 | -0.20599 | -0.28839 | -0.032043 | -0.11444 | -0.10529 | -0.077819 | 0.10529 | 0.1831 | 0.064087 | 0.16937 | 0.16937 | -0.10529 | 0.27466 | 0
379.7812 | 122.5845 | -59.1431 | 63.089 | -0.10765 | 0.00091438 | 0.0647 | 0.066996 | -0.0040867 | -0.0021832 | -0.0033752 | 0.068878 | -0.00025411 | 0.065134 | 0.061483 | 0.00067288 | -0.0017667 | -0.0012667 | 0.027466 | 63.0844 | -59.2529 | 63.2675 | -58.7952 | -59.7198 | -59.404 | 0.059509 | 63.0569 | -59.1705 | 63.1943 | -59.0561 | -59.1934 | -58.9554 | -0.10986 | -0.004577 | 0.17853 | 0.1831 | 0.3479 | 0.45776 | -0.26093 | -0.57678 | -0.15106 | -0.46692 | -0.027466 | -0.032043 | 0.10529 | 0.13733 | 0.086975 | 0.11444 | 0.18768 | -0.050354 | 0.21515 | -0.022888
380.7812 | 122.4838 | -59.1797 | 63.0661 | -0.092392 | -0.0014356 | 0.06705 | 0.068919 | 0.00039951 | 0.001662 | -0.00081179 | -0.10202 | 0.00060029 | 0.064706 | 0.068532 | 0.00024568 | -0.00048494 | -0.0029757 | -0.11444 | 63.0661 | -59.3674 | 62.9608 | -59.0927 | -59.5047 | -59.436 | -0.082397 | 63.0569 | -59.2484 | 63.1256 | -59.1568 | -59.2896 | -59.0057 | -0.18768 | 0 | -0.10529 | -0.10529 | 0.086975 | 0.27466 | -0.25635 | -0.32501 | -0.068664 | -0.13733 | -0.068665 | -0.009155 | 0.059509 | 0.068664 | 0.022888 | 0.091553 | 0.17395 | -0.10986 | 0.24261 | -0.041199
381.7812 | 122.5342 | -59.166 | 63.0615 | -0.13054 | 0.00027338 | 0.069186 | 0.070842 | -0.0010959 | 0.00059397 | 0.0011108 | -0.068451 | 0.00017309 | 0.065134 | 0.064901 | -0.00060892 | -0.0017667 | 1.4916e-05 | -0.10986 | 63.089 | -59.3445 | 62.9425 | -59.1064 | -59.5184 | -59.4315 | -0.10071 | 63.0798 | -59.2529 | 63.1302 | -59.1934 | -59.317 | -59.024 | -0.17853 | 0.027466 | -0.11902 | -0.14648 | 0.05951 | 0.23804 | -0.2655 | -0.35248 | -0.086976 | -0.17395 | -0.086975 | 0.018311 | 0.068665 | 0.050354 | -0.027465 | 0.05951 | 0.14191 | -0.15106 | 0.22888 | -0.064087
382.7812 | 122.5662 | -59.1568 | 63.0432 | -0.10765 | 0.00091438 | 0.066836 | 0.069773 | 0.0031765 | -0.001756 | -0.0014526 | -0.047089 | 0.0010275 | 0.064066 | 0.062337 | 3.2078e-05 | -0.0011257 | 0.00044212 | -0.086975 | 63.0432 | -59.317 | 63.002 | -59.0469 | -59.436 | -59.3582 | -0.045776 | 63.0066 | -59.2209 | 63.1714 | -59.1385 | -59.2575 | -58.9371 | -0.16022 | 0 | -0.041199 | -0.041199 | 0.10986 | 0.27008 | -0.20142 | -0.27924 | -0.041198 | -0.11902 | -0.064087 | -0.036621 | 0.12817 | 0.1648 | 0.01831 | 0.082397 | 0.21973 | -0.10071 | 0.28381 | -0.036621
383.7812 | 122.5708 | -59.1248 | 63.089 | -0.11528 | 0.0015552 | 0.065341 | 0.065501 | 0.0016813 | -4.7032e-05 | 0.00068361 | 0.035309 | 0.0018821 | 0.06257 | 0.063192 | -0.0014633 | -0.0013393 | 0.00022852 | 0.018311 | 63.0707 | -59.2667 | 63.0798 | -58.9142 | -59.3308 | -59.2621 | 0.15106 | 63.0615 | -59.1751 | 63.2172 | -59.0378 | -59.1614 | -58.9233 | -0.14191 | -0.01831 | -0.009155 | 0.009155 | 0.21057 | 0.35248 | -0.13733 | -0.20599 | 0.004578 | -0.064087 | -0.050354 | -0.027466 | 0.12817 | 0.15564 | 0.086975 | 0.13733 | 0.20142 | -0.036621 | 0.25177 | 0.013733
384.7969 | 122.493 | -59.1751 | 63.2126 | -0.092392 | 5.9776e-05 | 0.068118 | 0.068278 | -0.0013095 | -0.001115 | -0.00017099 | -0.013519 | 0.00060029 | 0.064066 | 0.066182 | -0.00082252 | -0.0019803 | -0.00041228 | -0.041199 | 63.1073 | -59.3124 | 62.9791 | -59.0561 | -59.4498 | -59.3903 | -0.009155 | 63.0936 | -59.2117 | 63.0844 | -59.1797 | -59.3079 | -59.0332 | -0.13733 | -0.10528 | -0.23346 | -0.12817 | 0.11902 | 0.25635 | -0.21515 | -0.27466 | -0.07782 | -0.13733 | -0.036621 | -0.11902 | -0.12817 | -0.009155 | -0.004577 | 0.032044 | 0.14191 | -0.13275 | 0.17853 | -0.09613
385.7812 | 122.4976 | -59.1614 | 63.089 | -0.13054 | 0.00091438 | 0.067691 | 0.069346 | -0.00024129 | 0.00016657 | -0.0042298 | -0.0013123 | 0.0010275 | 0.06492 | 0.06426 | -0.0014633 | -0.0013393 | -0.0034031 | -0.086975 | 63.0478 | -59.3079 | 62.9929 | -59.024 | -59.436 | -59.3491 | -0.045776 | 63.0341 | -59.1934 | 63.121 | -59.1568 | -59.2484 | -58.992 | -0.14648 | -0.041198 | -0.09613 | -0.054932 | 0.13733 | 0.28381 | -0.18768 | -0.27466 | -0.041199 | -0.12817 | -0.032043 | -0.054931 | 0.032044 | 0.086975 | 0.004578 | 0.036621 | 0.16937 | -0.086975 | 0.20142 | -0.054932
386.7812 | 122.5571 | -59.1431 | 63.1073 | -0.077133 | -0.00036742 | 0.066623 | 0.067851 | -0.00066869 | 0.00080757 | -0.00081179 | 0.013947 | -4.0508e-05 | 0.066843 | 0.064473 | -0.0023177 | -0.0019803 | 0.00022852 | 0 | 63.0936 | -59.2941 | 63.0661 | -58.9508 | -59.3491 | -59.2804 | 0.14648 | 63.0432 | -59.1843 | 63.2126 | -59.0332 | -59.166 | -58.9188 | -0.15106 | -0.013733 | -0.041199 | -0.027466 | 0.19226 | 0.34332 | -0.13733 | -0.20599 | 0.013732 | -0.054932 | -0.041199 | -0.064087 | 0.10528 | 0.16937 | 0.10986 | 0.15106 | 0.2243 | -0.022889 | 0.2655 | 0.01831
387.7812 | 122.5433 | -59.1522 | 63.0524 | -0.19157 | -0.001222 | 0.068118 | 0.067637 | -2.7692e-05 | -0.001756 | 0.0013244 | -0.025726 | 0.00017309 | 0.066202 | 0.068319 | -0.0021041 | -0.0021939 | 0.00022852 | 0.086975 | 63.0341 | -59.285 | 63.0569 | -58.9874 | -59.4086 | -59.3445 | 0.013733 | 63.0295 | -59.2117 | 63.1668 | -59.1202 | -59.2346 | -58.9828 | -0.13275 | -0.01831 | 0.004578 | 0.022888 | 0.1648 | 0.29755 | -0.19226 | -0.25635 | -0.059509 | -0.1236 | -0.059509 | -0.022888 | 0.11444 | 0.13733 | 0.032044 | 0.091553 | 0.16937 | -0.082397 | 0.22888 | -0.022888
388.7812 | 122.5479 | -59.1431 | 63.0753 | -0.10765 | -0.0027172 | 0.065982 | 0.069773 | -2.7692e-05 | -0.00047423 | -0.0023072 | 0.023102 | 0.00081389 | 0.065347 | 0.058065 | -0.00060892 | -0.0026211 | 0.00044212 | -0.022888 | 63.0707 | -59.2896 | 63.0386 | -58.9645 | -59.3674 | -59.2987 | 0.082397 | 63.0478 | -59.2072 | 63.1805 | -59.0652 | -59.1705 | -58.9325 | -0.14649 | -0.004577 | -0.036621 | -0.032044 | 0.17853 | 0.32501 | -0.15564 | -0.22431 | -0.009155 | -0.07782 | -0.064087 | -0.027465 | 0.10529 | 0.13275 | 0.077819 | 0.14191 | 0.21057 | -0.027466 | 0.27466 | 0.036621
389.7812 | 122.5754 | -59.1293 | 63.0615 | -0.092392 | -0.0016492 | 0.066409 | 0.066355 | -0.0030185 | 0.0018756 | -0.0016662 | 0.023102 | 0.0010275 | 0.064706 | 0.061269 | -0.0023177 | -0.0039029 | -0.0012667 | -0.018311 | 63.0478 | -59.317 | 63.0569 | -58.9874 | -59.3765 | -59.3124 | 0.082397 | 63.0707 | -59.1797 | 63.176 | -59.0652 | -59.1751 | -58.9325 | -0.18768 | -0.013732 | -0.004577 | 0.009155 | 0.14191 | 0.32959 | -0.18311 | -0.24719 | 0.004578 | -0.059509 | -0.050354 | 0.009156 | 0.11444 | 0.10529 | 0.064086 | 0.11444 | 0.19684 | -0.045777 | 0.24719 | 0.004577
390.7812 | 122.5479 | -59.166 | 63.0615 | -0.11528 | 5.9776e-05 | 0.0694 | 0.069133 | 0.0014675 | -0.00047423 | -0.00081179 | -0.028778 | 0.0014547 | 0.066843 | 0.062124 | -0.0018905 | -0.0028347 | 0.0015103 | -0.059509 | 63.0386 | -59.3079 | 63.0295 | -58.9828 | -59.4269 | -59.3491 | 0.009155 | 63.0295 | -59.2163 | 63.1439 | -59.1476 | -59.2667 | -58.9874 | -0.14191 | -0.022888 | -0.032043 | -0.009155 | 0.18311 | 0.32501 | -0.18311 | -0.26093 | -0.041199 | -0.11902 | -0.050354 | -0.032043 | 0.082398 | 0.11444 | 0.018311 | 0.068665 | 0.17853 | -0.10071 | 0.22888 | -0.050354
391.7812 | 122.5113 | -59.166 | 63.0432 | -0.10765 | 0.0013416 | 0.067691 | 0.069133 | 0.0014675 | -0.0015424 | -0.00017099 | -0.047089 | 0.0027365 | 0.063638 | 0.066396 | -0.0027451 | -0.0015529 | -0.0010531 | -0.050354 | 63.0341 | -59.3308 | 63.0432 | -58.9874 | -59.4177 | -59.3308 | -0.004578 | 63.0569 | -59.2255 | 63.1393 | -59.1568 | -59.2621 | -58.9874 | -0.1648 | -0.009155 | 0 | 0.009155 | 0.17853 | 0.34332 | -0.1648 | -0.25177 | 0 | -0.086975 | -0.059509 | 0.013733 | 0.09613 | 0.082397 | 0.009156 | 0.068665 | 0.17853 | -0.09613 | 0.23804 | -0.036621
392.7812 | 122.5388 | -59.1843 | 63.0707 | -0.092392 | 0.00048698 | 0.068332 | 0.068278 | 0.0010403 | -0.00026063 | 0.00047001 | -0.068451 | -0.00068151 | 0.062997 | 0.067678 | -0.0027451 | -0.0034757 | 0.0030057 | -0.054932 | 63.0386 | -59.3262 | 63.0432 | -58.9828 | -59.4223 | -59.3399 | 0.018311 | 63.0524 | -59.198 | 63.1531 | -59.1431 | -59.2575 | -58.96 | -0.14191 | -0.032044 | -0.027466 | 0.004578 | 0.20142 | 0.34332 | -0.15564 | -0.23804 | -0.013733 | -0.09613 | -0.013733 | -0.018311 | 0.082397 | 0.10071 | 0.041199 | 0.054932 | 0.2243 | -0.073242 | 0.23804 | -0.059509
393.7812 | 122.5571 | -59.1522 | 63.1027 | -0.12291 | 0.0015552 | 0.06705 | 0.067637 | 0.00018591 | 0.0014484 | 0.00047001 | -0.0043641 | 0.00017309 | 0.068124 | 0.05956 | 0.00045928 | -0.0026211 | -0.0031895 | -0.050354 | 63.0661 | -59.3079 | 63.0341 | -59.0012 | -59.3994 | -59.3353 | 0.027466 | 63.0386 | -59.1888 | 63.1668 | -59.0881 | -59.2026 | -58.9462 | -0.15564 | -0.036621 | -0.068664 | -0.032043 | 0.15106 | 0.3067 | -0.1831 | -0.24719 | -0.027466 | -0.091553 | -0.036621 | -0.064087 | 0.064087 | 0.12817 | 0.064087 | 0.10071 | 0.20599 | -0.050354 | 0.24262 | -0.013733
394.7812 | 122.4792 | -59.166 | 63.1348 | -0.038986 | 0.0015552 | 0.065768 | 0.069133 | -0.0028049 | -0.00026063 | 0.0043152 | -0.068451 | 0.00060029 | 0.063425 | 0.065114 | -0.0014633 | -0.0019803 | -0.0027621 | -0.004578 | 63.121 | -59.3216 | 62.9883 | -59.0744 | -59.4543 | -59.3994 | 0.004578 | 63.1256 | -59.1888 | 63.1302 | -59.1568 | -59.2987 | -59.0057 | -0.15564 | -0.013733 | -0.14648 | -0.13275 | 0.091553 | 0.24719 | -0.23346 | -0.28839 | -0.07782 | -0.13275 | -0.022888 | -0.009156 | -0.004578 | 0.004578 | 0.009156 | 0.032044 | 0.16022 | -0.13275 | 0.18311 | -0.10986
395.7812 | 122.5204 | -59.1705 | 63.0615 | -0.016098 | 0.00070078 | 0.067477 | 0.068492 | 0.00082671 | 0.0018756 | -0.00017099 | -0.022675 | 0.0010275 | 0.065134 | 0.0696 | 3.2078e-05 | -0.0034757 | 0.00044212 | 0 | 63.0478 | -59.3079 | 63.0707 | -59.0012 | -59.404 | -59.3216 | 0.027466 | 63.0203 | -59.2026 | 63.1714 | -59.1339 | -59.2438 | -58.9828 | -0.13733 | -0.013732 | 0.009156 | 0.022888 | 0.16937 | 0.3067 | -0.15106 | -0.23346 | -0.013733 | -0.096131 | -0.032044 | -0.041198 | 0.10986 | 0.15106 | 0.036621 | 0.068665 | 0.18768 | -0.073242 | 0.21973 | -0.041198
396.7812 | 122.493 | -59.1293 | 63.1027 | -0.11528 | -0.0025036 | 0.064914 | 0.067637 | -0.0028049 | -0.001115 | 0.00025641 | 0.0017395 | -4.0508e-05 | 0.064279 | 0.063833 | -0.00060892 | -0.0041165 | -0.0023349 | 0.032043 | 63.0569 | -59.285 | 63.0615 | -58.9554 | -59.3582 | -59.2941 | 0.1236 | 63.0569 | -59.1797 | 63.1622 | -59.0744 | -59.198 | -58.96 | -0.15564 | -0.045776 | -0.041199 | 0.004577 | 0.17395 | 0.32959 | -0.16479 | -0.22888 | -0.009155 | -0.073242 | -0.050354 | -0.045776 | 0.059509 | 0.10528 | 0.054931 | 0.10529 | 0.16937 | -0.068665 | 0.21973 | -0.018311
397.7812 | 122.5067 | -59.1568 | 63.0615 | -0.11528 | 0.0017688 | 0.068332 | 0.06721 | 0.00082671 | 0.00080757 | 4.2808e-05 | -0.044037 | -4.0508e-05 | 0.065347 | 0.063192 | -0.0016769 | -0.0011257 | 1.4916e-05 | -0.059509 | 63.0386 | -59.2941 | 63.0157 | -58.992 | -59.4269 | -59.3491 | -0.018311 | 63.0066 | -59.2346 | 63.1439 | -59.1385 | -59.2667 | -59.0057 | -0.13733 | -0.022888 | -0.045776 | -0.022888 | 0.16479 | 0.30212 | -0.19226 | -0.27008 | -0.054932 | -0.13275 | -0.07782 | -0.054931 | 0.082398 | 0.13733 | 0.01831 | 0.09613 | 0.15106 | -0.10986 | 0.22888 | -0.032044
398.7812 | 122.5342 | -59.1202 | 63.121 | -0.031357 | -0.00015382 | 0.064914 | 0.067424 | -0.0017367 | 0.00016657 | 4.2808e-05 | -0.0043641 | 0.0010275 | 0.063425 | 0.061269 | 0.00024568 | -0.0026211 | 0.00086952 | 0.009155 | 63.0661 | -59.2804 | 63.1027 | -58.8776 | -59.3124 | -59.2392 | 0.1236 | 63.0707 | -59.1751 | 63.1943 | -59.0515 | -59.1751 | -58.9371 | -0.16022 | -0.054932 | -0.018311 | 0.036621 | 0.24262 | 0.40283 | -0.11902 | -0.19226 | 0.041199 | -0.032043 | -0.054932 | -0.050354 | 0.073242 | 0.1236 | 0.068664 | 0.1236 | 0.18311 | -0.054932 | 0.23804 | 0
399.7812 | 122.5159 | -59.1431 | 63.0753 | -0.061874 | 0.0013416 | 0.069827 | 0.069346 | 0.0014675 | -0.00026063 | -0.0023072 | -0.034882 | 0.00038669 | 0.065561 | 0.05721 | -0.00039512 | -0.0021939 | -0.0010531 | -0.050354 | 63.0341 | -59.3033 | 63.0295 | -59.0057 | -59.4086 | -59.3582 | -0.009155 | 63.0203 | -59.198 | 63.1393 | -59.1431 | -59.2667 | -58.9828 | -0.16022 | -0.041198 | -0.045776 | -0.004578 | 0.13733 | 0.29755 | -0.21515 | -0.2655 | -0.054931 | -0.10529 | -0.054932 | -0.054931 | 0.064087 | 0.11902 | 0 | 0.054932 | 0.16022 | -0.1236 | 0.21515 | -0.068665
400.7812 | 122.5616 | -59.1705 | 63.0615 | -0.13817 | 0.00048698 | 0.068545 | 0.07191 | -0.0010959 | 0.0020892 | -0.0025208 | -0.040985 | -0.0013223 | 0.062784 | 0.0696 | -0.00039512 | -0.0015529 | 0.0012967 | -0.059509 | 63.0066 | -59.2941 | 63.0295 | -58.9828 | -59.4131 | -59.3353 | 0.013733 | 63.0295 | -59.1843 | 63.1714 | -59.1385 | -59.2529 | -58.9874 | -0.1236 | -0.054931 | -0.032043 | 0.022888 | 0.18768 | 0.31128 | -0.16479 | -0.24262 | -0.041199 | -0.11902 | -0.013733 | -0.032043 | 0.10986 | 0.14191 | 0.032043 | 0.045776 | 0.18311 | -0.082398 | 0.19684 | -0.068665
401.7812 | 122.5662 | -59.1202 | 63.089 | -0.10765 | 0.00070078 | 0.06705 | 0.068705 | -0.0036593 | 0.00059397 | -0.0020936 | -0.0043641 | 0.00038669 | 0.065347 | 0.059774 | 0.0011001 | -0.0053983 | 0.00065592 | -0.054932 | 63.0798 | -59.2941 | 63.0249 | -59.0149 | -59.3948 | -59.317 | 0.027466 | 63.0524 | -59.2117 | 63.1714 | -59.0927 | -59.1934 | -58.96 | -0.17395 | -0.009155 | -0.064087 | -0.054932 | 0.10529 | 0.27923 | -0.19684 | -0.27466 | -0.022889 | -0.10071 | -0.091553 | -0.036621 | 0.082398 | 0.11902 | 0.027466 | 0.11902 | 0.16022 | -0.073242 | 0.25177 | 0.018311
402.7812 | 122.5433 | -59.1568 | 63.0844 | -0.00083923 | -0.0010082 | 0.0694 | 0.070201 | 0.00061311 | -4.7032e-05 | -0.0048706 | -0.044037 | 0.00017309 | 0.066415 | 0.063833 | -0.00039512 | -0.0030483 | -0.0044711 | -0.091553 | 63.0249 | -59.3079 | 63.0386 | -58.9966 | -59.4086 | -59.317 | 0.018311 | 63.0432 | -59.2072 | 63.1714 | -59.1248 | -59.2575 | -58.9783 | -0.15106 | -0.05951 | -0.045777 | 0.013733 | 0.16022 | 0.31128 | -0.16022 | -0.25177 | -0.009156 | -0.10071 | -0.050354 | -0.041199 | 0.086975 | 0.12817 | 0.032043 | 0.082397 | 0.17853 | -0.10071 | 0.22888 | -0.050354
403.7812 | 122.5983 | -59.1614 | 63.0753 | -0.069504 | 0.00091438 | 0.06705 | 0.067424 | 0.00082671 | 0.0014484 | 0.002179 | -0.0074159 | 0.00060029 | 0.065988 | 0.065328 | 0.0023819 | -0.00069854 | 0.0010831 | 0.036621 | 63.089 | -59.2896 | 63.089 | -58.9188 | -59.3124 | -59.2712 | 0.14191 | 63.0753 | -59.1568 | 63.1622 | -58.9874 | -59.2529 | -59.0012 | -0.12817 | 0.013733 | 0.013733 | 0 | 0.24262 | 0.37079 | -0.10986 | -0.15106 | 0.018311 | -0.022888 | 0.004578 | 0 | 0.086975 | 0.086975 | 0.17395 | 0.16937 | 0.16022 | -0.091553 | 0.15564 | -0.096131
404.7812 | 122.5433 | -59.1568 | 63.0798 | -0.16106 | 0.001128 | 0.067264 | 0.067424 | 0.0010403 | -0.0023968 | -0.00059819 | -0.03183 | -0.00046791 | 0.063425 | 0.064046 | -0.0018905 | -0.0032619 | -0.00041228 | -0.068665 | 63.0249 | -59.2987 | 63.0157 | -59.0103 | -59.4406 | -59.3628 | -0.018311 | 63.0341 | -59.2163 | 63.1256 | -59.1385 | -59.2758 | -59.0012 | -0.14191 | -0.054932 | -0.064087 | -0.009155 | 0.14648 | 0.28839 | -0.20599 | -0.28381 | -0.064087 | -0.14191 | -0.05951 | -0.045776 | 0.045776 | 0.091552 | 0.01831 | 0.07782 | 0.15564 | -0.11902 | 0.21515 | -0.059509
405.7812 | 122.5067 | -59.1568 | 63.0798 | -0.10765 | 0.0013416 | 0.0694 | 0.065074 | -0.0030185 | -0.001115 | 0.0017518 | -0.03183 | 0.00060029 | 0.064706 | 0.065114 | 0.0019547 | -0.0028347 | 0.0010831 | -0.054932 | 63.0249 | -59.3124 | 63.0249 | -58.9966 | -59.4269 | -59.3491 | -0.013733 | 63.0066 | -59.2346 | 63.1485 | -59.1522 | -59.2758 | -59.0012 | -0.15564 | -0.054932 | -0.054932 | 0 | 0.16022 | 0.31586 | -0.19226 | -0.27008 | -0.036621 | -0.11444 | -0.07782 | -0.073242 | 0.068665 | 0.14191 | 0.004577 | 0.082397 | 0.15564 | -0.11902 | 0.23346 | -0.041199
406.7812 | 122.438 | -59.166 | 63.1393 | -0.061874 | -0.00079462 | 0.067264 | 0.06956 | -0.0021641 | 0.0010212 | 0.00068361 | -0.047089 | 0.00060029 | 0.064706 | 0.064473 | 0.00045928 | -0.0021939 | 0.0021511 | 0.036621 | 63.089 | -59.2987 | 63.0386 | -58.9966 | -59.404 | -59.3491 | 0.041199 | 63.089 | -59.1843 | 63.0936 | -59.1614 | -59.2621 | -59.0515 | -0.13275 | -0.050354 | -0.10071 | -0.050354 | 0.16937 | 0.30212 | -0.18311 | -0.23804 | -0.050354 | -0.10529 | -0.01831 | -0.050354 | -0.045776 | 0.004578 | 0.004578 | 0.022888 | 0.11444 | -0.09613 | 0.13275 | -0.07782
407.7812 | 122.5159 | -59.1431 | 63.0844 | -0.10765 | 0.00070078 | 0.067691 | 0.069133 | -0.00066869 | -0.0015424 | -0.0031616 | -0.016571 | 0.00017309 | 0.063852 | 0.061483 | 3.2078e-05 | -0.0028347 | 0.00086952 | -0.073242 | 63.0386 | -59.285 | 63.002 | -59.0378 | -59.436 | -59.3674 | -0.022888 | 63.0569 | -59.2072 | 63.1393 | -59.1385 | -59.2804 | -58.9874 | -0.14191 | -0.045777 | -0.082398 | -0.036621 | 0.10528 | 0.24719 | -0.22431 | -0.29297 | -0.082398 | -0.15106 | -0.064087 | -0.027466 | 0.054931 | 0.082397 | 0.004577 | 0.068664 | 0.15564 | -0.13733 | 0.21973 | -0.073243
408.7969 | 122.4655 | -59.166 | 63.1027 | -0.031357 | 5.9776e-05 | 0.069613 | 0.072337 | 0.0018949 | -0.00047423 | -0.0016662 | -0.080658 | -0.00025411 | 0.062997 | 0.071523 | -0.0016769 | -0.0017667 | -0.0027621 | 0.004578 | 63.0753 | -59.3216 | 63.0157 | -59.0057 | -59.5139 | -59.4589 | 0.004578 | 63.089 | -59.2026 | 63.1531 | -59.1705 | -59.317 | -59.0286 | -0.15564 | -0.027466 | -0.086975 | -0.059509 | 0.16022 | 0.31586 | -0.29297 | -0.3479 | -0.13733 | -0.19226 | -0.036621 | -0.013733 | 0.050354 | 0.064087 | -0.004577 | 0.032044 | 0.13733 | -0.15106 | 0.17395 | -0.11444
409.7812 | 122.5708 | -59.1705 | 63.0432 | -0.11528 | 0.00027338 | 0.068545 | 0.066355 | -2.7692e-05 | 0.0020892 | 0.0041016 | -0.047089 | 0.00060029 | 0.065134 | 0.065328 | -0.00018152 | -0.0011257 | 0.00044212 | -0.054932 | 63.0295 | -59.3124 | 62.9883 | -59.1888 | -59.3308 | -59.1843 | -0.013733 | 63.0524 | -59.198 | 63.1668 | -59.1156 | -59.2117 | -58.9691 | -0.14191 | -0.013733 | -0.054932 | -0.041199 | -0.018311 | 0.1236 | -0.013733 | -0.16022 | 0.12817 | -0.018311 | -0.027466 | 0.009155 | 0.1236 | 0.11444 | 0.054931 | 0.082397 | 0.20142 | -0.041199 | 0.22888 | -0.013733
410.7812 | 122.5204 | -59.1614 | 63.0981 | -0.18394 | 0.00091438 | 0.069186 | 0.067637 | -2.7692e-05 | 0.0010212 | 0.00047001 | -0.010467 | -0.00068151 | 0.064493 | 0.061696 | -0.0029587 | -0.0013393 | -0.00062588 | -0.054932 | 63.0386 | -59.3124 | 63.0112 | -58.9645 | -59.4223 | -59.3399 | 0 | 63.0386 | -59.2255 | 63.1439 | -59.1293 | -59.2667 | -59.0012 | -0.15106 | -0.05951 | -0.086976 | -0.027466 | 0.19684 | 0.3479 | -0.17853 | -0.26093 | -0.027466 | -0.10986 | -0.064087 | -0.05951 | 0.045776 | 0.10529 | 0.032044 | 0.096131 | 0.16022 | -0.10529 | 0.2243 | -0.041199
411.7969 | 122.4701 | -59.1568 | 63.1256 | -0.038986 | 0.0013416 | 0.068118 | 0.068919 | 0.0036039 | 0.00080757 | 0.00047001 | -0.086761 | 0.0023093 | 0.063211 | 0.064687 | -0.0010361 | -0.00048494 | 0.0017239 | 0.009155 | 63.0981 | -59.317 | 62.9974 | -59.0286 | -59.4086 | -59.3811 | 0.032043 | 63.0981 | -59.1797 | 63.1439 | -59.1476 | -59.3079 | -59.0195 | -0.16022 | -0.027465 | -0.12817 | -0.10071 | 0.12817 | 0.28839 | -0.22431 | -0.25177 | -0.064087 | -0.091552 | -0.022888 | -0.027465 | 0.018311 | 0.045776 | 0.009155 | 0.032043 | 0.13733 | -0.15106 | 0.16022 | -0.12817
412.7812 | 122.5479 | -59.1522 | 63.089 | -0.10002 | -0.00015382 | 0.067264 | 0.06721 | -0.0030185 | 0.00080757 | -0.001239 | 0.029205 | 0.00017309 | 0.065988 | 0.067891 | 0.00088648 | -0.0026211 | -0.0053257 | -0.064087 | 63.0798 | -59.317 | 63.002 | -59.024 | -59.4177 | -59.3674 | -0.032043 | 63.0478 | -59.2072 | 63.1302 | -59.111 | -59.2484 | -58.9645 | -0.16479 | -0.009155 | -0.086975 | -0.07782 | 0.12817 | 0.29297 | -0.21515 | -0.2655 | -0.050354 | -0.10071 | -0.054931 | -0.041198 | 0.041199 | 0.082397 | 0.041199 | 0.09613 | 0.18768 | -0.09613 | 0.24261 | -0.041199
413.7812 | 122.5479 | -59.1339 | 63.0844 | -0.12291 | 0.00091438 | 0.068545 | 0.065715 | -0.0075045 | 0.0025166 | -0.0042298 | 0.02005 | 0.0012411 | 0.062997 | 0.06661 | -0.00018152 | -0.0015529 | -0.00083948 | -0.036621 | 63.0615 | -59.2758 | 63.0341 | -58.9737 | -59.3811 | -59.3033 | 0.054932 | 63.0661 | -59.1888 | 63.2034 | -59.0881 | -59.1843 | -58.9554 | -0.14191 | -0.022889 | -0.050354 | -0.027465 | 0.16022 | 0.30212 | -0.16937 | -0.24719 | -0.027466 | -0.10529 | -0.054932 | -0.018311 | 0.11902 | 0.13733 | 0.045776 | 0.10071 | 0.17853 | -0.050354 | 0.23346 | 0.004578
414.7812 | 122.5296 | -59.1431 | 63.1073 | -0.084763 | 0.0015552 | 0.069613 | 0.068492 | -0.00088229 | 0.00080757 | -0.0027344 | -0.022675 | 0.0010275 | 0.066415 | 0.062337 | -0.00039512 | -0.0028347 | -0.0029757 | -0.086975 | 63.0432 | -59.2987 | 62.9883 | -59.0286 | -59.436 | -59.3445 | -0.054932 | 63.0386 | -59.2026 | 63.1256 | -59.1339 | -59.2484 | -59.0012 | -0.15564 | -0.064087 | -0.11902 | -0.054932 | 0.11444 | 0.27008 | -0.20142 | -0.29297 | -0.045776 | -0.13733 | -0.05951 | -0.068665 | 0.01831 | 0.086975 | 0.009155 | 0.068665 | 0.14191 | -0.10529 | 0.20142 | -0.045776
415.7812 | 122.5342 | -59.1568 | 63.0798 | -0.15343 | 0.0015552 | 0.069186 | 0.069987 | -0.00088229 | 0.0023028 | 0.001538 | -0.028778 | -0.0023905 | 0.065134 | 0.067891 | 0.0011001 | -0.0028347 | 0.00086952 | -0.041199 | 63.0295 | -59.2896 | 63.0203 | -58.9828 | -59.4223 | -59.3399 | 0.018311 | 63.0203 | -59.2163 | 63.1668 | -59.1293 | -59.2667 | -58.9828 | -0.13275 | -0.050354 | -0.059509 | -0.009155 | 0.17395 | 0.3067 | -0.18311 | -0.2655 | -0.050354 | -0.13275 | -0.05951 | -0.059509 | 0.086975 | 0.14648 | 0.027466 | 0.086976 | 0.17395 | -0.10986 | 0.23346 | -0.050354
416.7969 | 122.5342 | -59.1431 | 63.0569 | -0.092392 | 0.00070078 | 0.068973 | 0.070414 | 0.00082671 | 0.00080757 | -0.0020936 | -0.095917 | 0.0029501 | 0.065347 | 0.064473 | -0.0027451 | -0.0017667 | -0.0027621 | -0.082397 | 63.0341 | -59.3445 | 63.0112 | -59.0149 | -59.4131 | -59.3582 | -0.032043 | 63.0249 | -59.2255 | 63.1714 | -59.1248 | -59.2484 | -58.9508 | -0.20142 | -0.022888 | -0.045777 | -0.022889 | 0.12817 | 0.32959 | -0.21515 | -0.27008 | -0.013733 | -0.068665 | -0.082398 | -0.032044 | 0.11444 | 0.14649 | 0.01831 | 0.10071 | 0.19226 | -0.10529 | 0.27466 | -0.022888
417.7812 | 122.5159 | -59.1522 | 63.0707 | -0.10765 | -0.00015382 | 0.068332 | 0.068278 | -0.00088229 | 0.00080757 | 0.0013244 | 0.0047913 | 0.00081389 | 0.06492 | 0.067251 | 0.00088648 | -0.0039029 | 0.0017239 | -0.073242 | 63.0661 | -59.2941 | 63.0066 | -58.9096 | -59.4223 | -59.3308 | -0.045776 | 63.0615 | -59.2026 | 63.1302 | -59.1339 | -59.2484 | -59.0195 | -0.14191 | -0.004578 | -0.064087 | -0.059509 | 0.24262 | 0.38452 | -0.17853 | -0.27008 | -0.036622 | -0.12817 | -0.050354 | -0.009156 | 0.059509 | 0.068665 | 0.018311 | 0.068665 | 0.13275 | -0.09613 | 0.18311 | -0.045776
418.7812 | 122.4518 | -59.1843 | 63.1302 | -0.069504 | 0.00048698 | 0.068332 | 0.069133 | -0.0021641 | -0.00090143 | 0.0013244 | -0.03183 | -0.0011087 | 0.066415 | 0.066396 | -0.0021041 | -0.0026211 | -0.0027621 | 0.036621 | 63.0661 | -59.2758 | 63.0524 | -58.9737 | -59.3582 | -59.317 | 0.068665 | 63.0981 | -59.1614 | 63.1302 | -59.1202 | -59.2484 | -58.9966 | -0.091553 | -0.064087 | -0.07782 | -0.013733 | 0.21057 | 0.30212 | -0.13275 | -0.17395 | -0.041199 | -0.082397 | 0.022888 | -0.032043 | 0 | 0.032043 | 0.064087 | 0.041199 | 0.18768 | -0.064087 | 0.1648 | -0.086975
419.7812 | 122.5113 | -59.1431 | 63.0707 | -0.084763 | -0.00079462 | 0.067904 | 0.067424 | -0.0023777 | 0.001662 | -0.0020936 | -0.047089 | 0.00017309 | 0.065775 | 0.062978 | -0.00060892 | -0.0017667 | -0.00019868 | -0.068665 | 63.0478 | -59.317 | 63.0203 | -59.1064 | -59.2804 | -59.3033 | 0.018311 | 63.0432 | -59.1888 | 63.1348 | -59.1431 | -59.2484 | -59.0012 | -0.17395 | -0.022888 | -0.050354 | -0.027466 | 0.036621 | 0.21057 | -0.16022 | -0.13733 | 0.013733 | 0.036621 | -0.045777 | -0.027466 | 0.064087 | 0.091553 | 0 | 0.045777 | 0.14191 | -0.10529 | 0.18768 | -0.059509
420.7812 | 122.5433 | -59.1522 | 63.1027 | -0.10002 | -0.00229 | 0.068973 | 0.067424 | 0.00018591 | -4.7032e-05 | 0.0017518 | -0.0074159 | 0.0010275 | 0.065347 | 0.06426 | -0.00082252 | -0.0030483 | 0.0032193 | -0.068665 | 63.0753 | -59.2896 | 63.002 | -59.0286 | -59.4177 | -59.2484 | -0.036621 | 63.0707 | -59.1888 | 63.1165 | -59.1202 | -59.2392 | -58.9874 | -0.13733 | -0.027466 | -0.10071 | -0.073242 | 0.1236 | 0.26093 | -0.09613 | -0.2655 | 0.041199 | -0.12817 | -0.036621 | -0.032043 | 0.013733 | 0.045776 | 0.032044 | 0.068665 | 0.1648 | -0.086975 | 0.20142 | -0.050354
421.7969 | 122.4838 | -59.0973 | 63.1531 | -0.10765 | 0.00027338 | 0.066836 | 0.068919 | -0.0019503 | -4.7032e-05 | -0.0018798 | 0.029205 | 0.00060029 | 0.064706 | 0.062764 | -0.0016769 | -0.0030483 | 0.00065592 | 0.032043 | 63.0707 | -59.2758 | 63.1027 | -58.905 | -59.317 | -59.2621 | 0.15106 | 63.0569 | -59.1797 | 63.1943 | -59.0469 | -59.1614 | -58.9096 | -0.17853 | -0.082397 | -0.050354 | 0.032043 | 0.19226 | 0.37079 | -0.16479 | -0.21973 | 0.013733 | -0.041199 | -0.082397 | -0.09613 | 0.041199 | 0.13733 | 0.050354 | 0.13275 | 0.18768 | -0.064087 | 0.27008 | 0.01831
422.7812 | 122.5433 | -59.1522 | 63.0936 | -0.069504 | -0.00079462 | 0.066836 | 0.067851 | -0.0019503 | 0.0020892 | -0.0033752 | -0.019623 | -0.0013223 | 0.061716 | 0.062337 | -0.00082252 | -0.0017667 | -0.00083948 | -0.059509 | 63.0478 | -59.3033 | 63.0112 | -59.0378 | -59.4269 | -59.3491 | 0.10529 | 63.0569 | -59.1797 | 63.121 | -59.111 | -59.2667 | -58.9828 | -0.15106 | -0.045776 | -0.082398 | -0.036622 | 0.11444 | 0.2655 | -0.19684 | -0.27466 | -0.045776 | -0.1236 | -0.027465 | -0.036621 | 0.027466 | 0.064087 | 0.041199 | 0.068664 | 0.16937 | -0.11444 | 0.19684 | -0.086976
423.7812 | 122.4518 | -59.1934 | 63.1393 | -0.12291 | 0.00091438 | 0.067477 | 0.070628 | -2.7692e-05 | 0.00080757 | -0.002948 | -0.08371 | 0.0012411 | 0.063425 | 0.068746 | -0.0025315 | -0.0021939 | 0.00044212 | -0.022888 | 63.121 | -59.2804 | 63.0066 | -59.079 | -59.5047 | -59.3811 | -0.022888 | 63.1119 | -59.2117 | 63.0936 | -59.1888 | -59.3033 | -59.0378 | -0.086976 | -0.01831 | -0.13275 | -0.11444 | 0.11444 | 0.20142 | -0.18768 | -0.31128 | -0.10071 | -0.2243 | -0.018311 | -0.027466 | -0.045776 | -0.01831 | 0.004577 | 0.022888 | 0.15564 | -0.10986 | 0.17395 | -0.091553
424.7812 | 122.5296 | -59.166 | 63.0981 | -0.11528 | 0.00048698 | 0.068332 | 0.068705 | 0.0010403 | 0.00016657 | 0.0026062 | -0.0013123 | 0.00017309 | 0.065775 | 0.067037 | 0.00045928 | -0.00069854 | 0.00086952 | -0.064087 | 63.0707 | -59.2987 | 63.0066 | -59.0332 | -59.4223 | -59.3536 | -0.027466 | 63.0569 | -59.2072 | 63.1485 | -59.1522 | -59.2392 | -58.992 | -0.13275 | -0.027466 | -0.091553 | -0.064087 | 0.13275 | 0.2655 | -0.18768 | -0.25635 | -0.054932 | -0.1236 | -0.041198 | -0.041199 | 0.050354 | 0.091553 | 0.013733 | 0.054931 | 0.17395 | -0.073242 | 0.21515 | -0.032044
425.7812 | 122.4976 | -59.1614 | 63.0707 | -0.10765 | -0.00036742 | 0.068118 | 0.068919 | -0.0023777 | -0.00090143 | -0.0031616 | 0.0017395 | -0.0015359 | 0.063852 | 0.061483 | 3.2078e-05 | -0.0013393 | -0.0016941 | -0.054932 | 63.0386 | -59.3079 | 63.0249 | -59.0057 | -59.4086 | -59.3308 | 0.013733 | 63.0478 | -59.1934 | 63.1439 | -59.1019 | -59.2163 | -58.9874 | -0.14648 | -0.032044 | -0.045777 | -0.013733 | 0.15564 | 0.30212 | -0.16937 | -0.24719 | -0.022889 | -0.10071 | -0.032043 | -0.022888 | 0.073242 | 0.09613 | 0.059509 | 0.091552 | 0.17395 | -0.054932 | 0.20599 | -0.022889
426.8125 | 122.5433 | -59.1522 | 63.0524 | -0.13054 | 0.0017688 | 0.068759 | 0.068919 | -0.00066869 | 0.0012348 | 0.00089721 | -0.11728 | -0.0019631 | 0.06257 | 0.072164 | -0.00018152 | -0.00027134 | -0.0042575 | -0.041199 | 63.0203 | -59.3308 | 63.0295 | -58.9645 | -59.4315 | -59.5093 | 0.18768 | 63.0432 | -59.1934 | 63.1714 | -59.0973 | -59.2209 | -58.9508 | -0.17853 | -0.032043 | -0.022888 | 0.009155 | 0.18768 | 0.36621 | -0.35705 | -0.27924 | -0.17853 | -0.10071 | -0.041198 | -0.009155 | 0.11902 | 0.12817 | 0.054932 | 0.09613 | 0.20142 | -0.068664 | 0.24261 | -0.027466
427.7812 | 122.5433 | -59.1568 | 63.0798 | -0.1992 | 0.0015552 | 0.067477 | 0.068492 | -0.0057955 | 0.0031574 | 0.00025641 | 0.0017395 | -4.0508e-05 | 0.066202 | 0.061483 | -0.0014633 | -0.00091214 | -0.00041228 | -0.036621 | 63.0661 | -59.285 | 63.002 | -59.0195 | -59.3994 | -59.3262 | 0.004578 | 63.0615 | -59.2117 | 63.1485 | -59.1019 | -59.2209 | -58.9874 | -0.12817 | -0.013733 | -0.07782 | -0.064087 | 0.13733 | 0.2655 | -0.16937 | -0.24262 | -0.041199 | -0.11444 | -0.054932 | -0.018311 | 0.068665 | 0.086976 | 0.054931 | 0.10986 | 0.16937 | -0.064087 | 0.2243 | -0.009155
428.7812 | 122.5937 | -59.1202 | 63.0844 | -0.12291 | 0.00091438 | 0.069186 | 0.066569 | 0.0027493 | 0.0027302 | 0.0019654 | 0.0078429 | -0.0023905 | 0.063852 | 0.061056 | 3.2078e-05 | -0.0043301 | -0.0016941 | -0.027466 | 63.0844 | -59.2621 | 63.0569 | -58.9691 | -59.3719 | -59.3216 | 0.082397 | 63.0524 | -59.1751 | 63.1897 | -59.079 | -59.1843 | -58.9417 | -0.14191 | 0 | -0.027466 | -0.027466 | 0.15106 | 0.29297 | -0.20142 | -0.25177 | -0.059509 | -0.10986 | -0.054932 | -0.032044 | 0.10529 | 0.13733 | 0.041199 | 0.096131 | 0.17853 | -0.064087 | 0.23346 | -0.009155
429.7969 | 122.5525 | -59.1614 | 63.089 | -0.12291 | -0.0016492 | 0.070041 | 0.066996 | 0.0014675 | 0.001662 | -0.0016662 | -0.065399 | -0.00046791 | 0.063638 | 0.068532 | -0.0010361 | -0.0032619 | -0.0010531 | -0.20599 | 62.9929 | -59.3491 | 63.0844 | -58.9783 | -59.4269 | -59.3308 | 0.018311 | 63.0157 | -59.2209 | 63.1714 | -59.0927 | -59.2392 | -58.9645 | -0.18768 | -0.09613 | -0.004577 | 0.091553 | 0.18311 | 0.37079 | -0.16937 | -0.2655 | 0.01831 | -0.07782 | -0.059509 | -0.073242 | 0.082398 | 0.15564 | 0.068665 | 0.12817 | 0.19684 | -0.07782 | 0.25635 | -0.018311
430.7812 | 122.5571 | -59.1293 | 63.0981 | -0.061874 | 5.9776e-05 | 0.066409 | 0.067637 | -0.0028049 | -0.0015424 | -0.0020936 | -0.022675 | -0.0011087 | 0.063638 | 0.064687 | -0.0012497 | -0.0024075 | 0.0021511 | -0.027466 | 63.0615 | -59.2621 | 63.0478 | -58.9783 | -59.3719 | -59.2941 | 0.054932 | 63.0615 | -59.1888 | 63.1851 | -59.0698 | -59.1888 | -58.9325 | -0.13275 | -0.036622 | -0.050354 | -0.013732 | 0.15106 | 0.28381 | -0.16479 | -0.24262 | -0.032043 | -0.10986 | -0.05951 | -0.036622 | 0.086975 | 0.1236 | 0.059509 | 0.11902 | 0.19684 | -0.05951 | 0.25635 | 0
431.7969 | 122.4747 | -59.1614 | 63.1348 | -0.061874 | -0.00015382 | 0.06705 | 0.069987 | -0.00066869 | 0.0025166 | -0.001239 | -0.10812 | 0.0020957 | 0.064493 | 0.06896 | -0.0012497 | -0.00069854 | -0.0010531 | -0.018311 | 63.0981 | -59.3079 | 63.0295 | -59.0332 | -59.404 | -59.3857 | 0.045776 | 63.1302 | -59.2209 | 63.1485 | -59.1248 | -59.2804 | -59.0057 | -0.14648 | -0.036621 | -0.10529 | -0.068665 | 0.12817 | 0.27466 | -0.2243 | -0.24262 | -0.07782 | -0.096131 | -0.059509 | -0.004578 | 0.013733 | 0.018311 | 0.036621 | 0.09613 | 0.15564 | -0.11902 | 0.21515 | -0.05951
432.7812 | 122.5937 | -59.1843 | 63.0981 | -0.069504 | -0.00015382 | 0.067691 | 0.066996 | -0.0013095 | 0.0014484 | -0.00017099 | -0.019623 | 0.0016685 | 0.065347 | 0.06426 | -0.0012497 | -0.0013393 | -0.0010531 | -0.041199 | 63.089 | -59.2804 | 63.0112 | -58.9783 | -59.3811 | -59.3033 | 0.045776 | 63.0615 | -59.198 | 63.176 | -59.079 | -59.1934 | -58.9554 | -0.096131 | -0.009156 | -0.086976 | -0.07782 | 0.20599 | 0.30213 | -0.11902 | -0.19684 | -0.022888 | -0.10071 | -0.013733 | -0.036622 | 0.077819 | 0.11444 | 0.10529 | 0.11902 | 0.22888 | -0.009155 | 0.24262 | 0.004578
433.7812 | 122.5571 | -59.1431 | 63.1165 | -0.069504 | -0.00015382 | 0.064486 | 0.068919 | -0.0049411 | 0.00016657 | -0.0014526 | -0.025726 | 0.00038669 | 0.062143 | 0.071523 | 0.00024568 | -0.0024075 | -0.0034031 | 0.25635 | 63.1165 | -59.2209 | 63.0341 | -58.96 | -59.2804 | -59.3033 | 0.14191 | 63.0569 | -59.1888 | 63.1851 | -59.0607 | -59.1843 | -58.9371 | -0.07782 | 0 | -0.082397 | -0.082397 | 0.1831 | 0.26093 | -0.16022 | -0.13733 | -0.082398 | -0.05951 | -0.045777 | -0.059509 | 0.068665 | 0.12817 | 0.082397 | 0.12817 | 0.20599 | -0.041199 | 0.25177 | 0.004578
434.7969 | 122.5616 | -59.1705 | 63.1119 | -0.061874 | -0.00079462 | 0.066195 | 0.068064 | -0.0015231 | 0.0010212 | -0.00059819 | -0.016571 | -4.0508e-05 | 0.066202 | 0.063405 | -0.0021041 | -0.0026211 | 0.0030057 | -0.027466 | 63.0661 | -59.2712 | 63.0066 | -59.0103 | -59.4131 | -59.3399 | 0 | 63.0432 | -59.1797 | 63.1577 | -59.1019 | -59.2346 | -58.992 | -0.10071 | -0.045776 | -0.10529 | -0.059509 | 0.16022 | 0.26093 | -0.16937 | -0.24262 | -0.068665 | -0.14191 | -0.009155 | -0.068664 | 0.045777 | 0.11444 | 0.068664 | 0.077819 | 0.17853 | -0.064087 | 0.18768 | -0.054932
435.7812 | 122.5342 | -59.1248 | 63.1073 | -0.092392 | 0.0017688 | 0.06705 | 0.068064 | -0.0015231 | 0.0012348 | 0.00089721 | 0.010895 | 0.0020957 | 0.065561 | 0.065328 | -0.0012497 | -0.0036893 | 0.0038601 | -0.027466 | 63.0524 | -59.3033 | 63.0478 | -58.9828 | -59.3628 | -59.3216 | 0.073242 | 63.0661 | -59.1843 | 63.1805 | -59.0607 | -59.1843 | -58.9371 | -0.17853 | -0.054932 | -0.059509 | -0.004577 | 0.14191 | 0.32043 | -0.19684 | -0.23804 | -0.01831 | -0.059509 | -0.059509 | -0.041199 | 0.073242 | 0.11444 | 0.064087 | 0.1236 | 0.18768 | -0.059509 | 0.24719 | 0
436.7812 | 122.5937 | -59.1476 | 63.1119 | -0.092392 | -0.00036742 | 0.065555 | 0.068492 | -0.0023777 | -0.001115 | -0.0027344 | 0.0047913 | 0.00060029 | 0.065775 | 0.056783 | 0.00067288 | -0.0026211 | 0.00065592 | -0.022888 | 63.0615 | -59.2941 | 63.0432 | -58.9737 | -59.3674 | -59.2987 | 0.091553 | 63.0707 | -59.1888 | 63.176 | -59.0607 | -59.1843 | -58.9462 | -0.14648 | -0.050354 | -0.068664 | -0.01831 | 0.17395 | 0.32043 | -0.15106 | -0.21973 | -0.004578 | -0.073243 | -0.041199 | -0.041198 | 0.064087 | 0.10529 | 0.086975 | 0.12817 | 0.20142 | -0.036621 | 0.24262 | 0.004578
437.7812 | 122.5113 | -59.1476 | 63.1073 | 0.0067903 | 0.0013416 | 0.06705 | 0.068705 | 0.0042447 | 0.0010212 | -0.0018798 | -0.016571 | -0.0013223 | 0.064493 | 0.061483 | -0.00060892 | -0.0045437 | 0.0010831 | -0.059509 | 63.0386 | -59.317 | 63.002 | -59.024 | -59.4131 | -59.317 | -0.004578 | 63.0524 | -59.1797 | 63.1439 | -59.1064 | -59.2392 | -58.992 | -0.16937 | -0.068665 | -0.10529 | -0.036621 | 0.1236 | 0.29297 | -0.16937 | -0.2655 | 0 | -0.09613 | -0.032043 | -0.054932 | 0.036621 | 0.091553 | 0.041199 | 0.073242 | 0.15564 | -0.091553 | 0.18768 | -0.05951
438.7812 | 122.4976 | -59.1248 | 63.1119 | -0.10765 | 0.00048698 | 0.068759 | 0.067424 | 0.0029629 | 0.00059397 | 0.001538 | -0.03183 | -0.00025411 | 0.064493 | 0.059774 | -0.0018905 | -0.0021939 | -0.0016941 | 0.045776 | 63.0753 | -59.2712 | 63.0753 | -58.9417 | -59.3491 | -59.2941 | 0.11444 | 63.0615 | -59.1751 | 63.176 | -59.0744 | -59.1934 | -58.96 | -0.14648 | -0.036621 | -0.036621 | 0 | 0.18311 | 0.32959 | -0.16937 | -0.2243 | -0.022888 | -0.07782 | -0.050354 | -0.050354 | 0.064087 | 0.11444 | 0.050354 | 0.10071 | 0.16479 | -0.068664 | 0.21515 | -0.01831
439.7812 | 122.4884 | -59.166 | 63.1073 | -0.11528 | 0.001128 | 0.066836 | 0.069133 | 0.0023221 | -0.001756 | 0.0030334 | -0.13559 | 0.00060029 | 0.065988 | 0.061269 | -0.0010361 | -0.0015529 | 0.0012967 | 0.013733 | 63.0981 | -59.3124 | 63.0157 | -59.024 | -59.4177 | -59.3811 | 0.036621 | 63.0798 | -59.1797 | 63.1393 | -59.1568 | -59.2941 | -59.0149 | -0.14648 | -0.009155 | -0.091553 | -0.082398 | 0.14191 | 0.28839 | -0.21515 | -0.25177 | -0.068665 | -0.10529 | -0.013732 | -0.027466 | 0.032043 | 0.059509 | 0.009156 | 0.022888 | 0.15106 | -0.12817 | 0.16479 | -0.11444
440.7812 | 122.5388 | -59.1385 | 63.0844 | -0.092392 | -0.00058102 | 0.064486 | 0.067637 | -0.0017367 | 0.0012348 | 4.2808e-05 | 0.026154 | -0.0028177 | 0.063852 | 0.062337 | -0.0016769 | -0.0034757 | -0.00041228 | 0.050354 | 63.0707 | -59.2529 | 63.0798 | -58.8913 | -59.3216 | -59.2758 | 0.43488 | 63.1119 | -59.1385 | 63.1622 | -59.079 | -59.1248 | -58.905 | -0.11444 | -0.013733 | -0.004578 | 0.009155 | 0.24719 | 0.36163 | -0.13733 | -0.1831 | -0.022888 | -0.068664 | 0 | 0.027465 | 0.077819 | 0.050354 | 0.05951 | 0.05951 | 0.23346 | 0.013733 | 0.23346 | 0.013733
441.7812 | 122.4747 | -59.1431 | 63.1073 | -0.046616 | -0.0016492 | 0.070254 | 0.07426 | 0.0012539 | -0.0015424 | 0.0013244 | -0.092865 | 0.0020957 | 0.066202 | 0.067251 | 3.2078e-05 | -0.00091214 | -0.0025485 | 0.036621 | 63.0844 | -59.3216 | 63.0157 | -59.0286 | -59.4269 | -59.3857 | 0.045776 | 63.1027 | -59.198 | 63.1485 | -59.1476 | -59.2712 | -58.992 | -0.17853 | -0.022888 | -0.091553 | -0.068665 | 0.11444 | 0.29297 | -0.24262 | -0.28381 | -0.064087 | -0.10529 | -0.054932 | -0.004578 | 0.041199 | 0.045777 | -0.004578 | 0.050354 | 0.15106 | -0.12817 | 0.20599 | -0.073242
442.7812 | 122.493 | -59.1339 | 63.1302 | -0.054245 | 0.00091438 | 0.066623 | 0.067637 | -0.0013095 | 0.0014484 | -0.00038459 | 0.010895 | -0.00046791 | 0.065134 | 0.06661 | 0.00067288 | -0.0019803 | 0.0017239 | 0.041199 | 63.0707 | -59.2804 | 63.0478 | -58.9462 | -59.3353 | -59.2987 | 0.10071 | 63.0844 | -59.198 | 63.1393 | -59.1064 | -59.2117 | -58.9737 | -0.14649 | -0.059509 | -0.082397 | -0.022888 | 0.18768 | 0.33417 | -0.1648 | -0.20142 | -0.01831 | -0.054931 | -0.064087 | -0.045776 | 0.009155 | 0.054931 | 0.027466 | 0.091553 | 0.16022 | -0.07782 | 0.2243 | -0.013733
443.7812 | 122.5616 | -59.1248 | 63.1027 | -0.11528 | 0.00027338 | 0.065555 | 0.067851 | 0.00061311 | 0.00038017 | 0.0019654 | 0.038361 | 0.0012411 | 0.064279 | 0.059133 | 0.00024568 | -0.0036893 | -0.0014805 | 0 | 63.0707 | -59.2804 | 63.0661 | -58.9142 | -59.3262 | -59.2712 | 0.14191 | 63.0524 | -59.1888 | 63.2172 | -59.024 | -59.1385 | -58.9005 | -0.15564 | -0.032043 | -0.036621 | -0.004578 | 0.21057 | 0.36621 | -0.14648 | -0.20142 | 0.009156 | -0.045776 | -0.064087 | -0.050354 | 0.11444 | 0.16479 | 0.10071 | 0.16479 | 0.2243 | -0.013733 | 0.28839 | 0.050354
444.7812 | 122.5113 | -59.1797 | 63.1165 | -0.077133 | 0.0015552 | 0.070041 | 0.070842 | 0.0025357 | 0.00059397 | -0.001239 | -0.10812 | 0.0010275 | 0.064493 | 0.0696 | 3.2078e-05 | -0.00091214 | 0.0051419 | 0 | 63.089 | -59.3124 | 63.0066 | -59.0149 | -59.4223 | -59.3857 | 0.036621 | 63.1119 | -59.2163 | 63.1302 | -59.166 | -59.3491 | -59.1476 | -0.13275 | -0.027466 | -0.10986 | -0.082397 | 0.16479 | 0.29755 | -0.20599 | -0.24262 | -0.073242 | -0.10986 | -0.036622 | -0.004578 | 0.013733 | 0.018311 | 0.013732 | 0.050354 | 0.032043 | -0.16937 | 0.068665 | -0.13275
445.7812 | 122.525 | -59.1385 | 63.1348 | -0.061874 | 5.9776e-05 | 0.068759 | 0.067851 | -0.0019503 | 0.00080757 | -0.00017099 | 0.013947 | 0.0010275 | 0.062784 | 0.06426 | -0.00060892 | -0.0030483 | -0.00062588 | 0.045776 | 63.0844 | -59.2667 | 63.0478 | -58.9508 | -59.3445 | -59.3079 | 0.09613 | 63.0707 | -59.1431 | 63.121 | -59.1248 | -59.2392 | -58.9874 | -0.12817 | -0.050354 | -0.086975 | -0.036621 | 0.18768 | 0.31586 | -0.16937 | -0.20599 | -0.041198 | -0.077819 | -0.004577 | -0.064087 | -0.013733 | 0.050354 | 0.013733 | 0.01831 | 0.15106 | -0.10071 | 0.15564 | -0.096131
446.7812 | 122.3328 | -59.1431 | 63.1622 | -0.054245 | 5.9776e-05 | 0.06705 | 0.066355 | -0.0034457 | 0.00016657 | -0.0052978 | -0.022675 | 0.00017309 | 0.06492 | 0.06896 | 3.2078e-05 | -0.0043301 | 0.0032193 | -0.027466 | 63.121 | -59.2941 | 62.97 | -59.0424 | -59.4223 | -59.3811 | 0.009155 | 63.1073 | -59.2072 | 63.0753 | -59.1751 | -59.3033 | -59.0515 | -0.15106 | -0.041198 | -0.19226 | -0.15106 | 0.10071 | 0.25177 | -0.23804 | -0.27924 | -0.086976 | -0.12817 | -0.064087 | -0.054931 | -0.086975 | -0.032044 | -0.032044 | 0.032043 | 0.091552 | -0.16022 | 0.15564 | -0.096131
447.7812 | 122.4564 | -59.1568 | 63.1302 | -0.12291 | -0.00058102 | 0.068973 | 0.071482 | -0.0015231 | -0.00047423 | -0.00081179 | -0.092865 | 0.00081389 | 0.066202 | 0.067891 | -0.0010361 | -0.0011257 | 0.00065592 | -0.013733 | 63.0798 | -59.3124 | 62.9929 | -59.0469 | -59.4452 | -59.3994 | -0.013733 | 63.0936 | -59.2072 | 63.1027 | -59.198 | -59.3353 | -59.0469 | -0.15564 | -0.050354 | -0.13733 | -0.086975 | 0.10986 | 0.2655 | -0.24262 | -0.28839 | -0.086975 | -0.13275 | -0.050354 | -0.036621 | -0.027466 | 0.009155 | -0.041199 | 0.009155 | 0.10986 | -0.17853 | 0.16022 | -0.12817
448.7812 | 122.4609 | -59.1248 | 63.1165 | -0.10002 | 0.0015552 | 0.067904 | 0.065287 | 0.0027493 | 0.0018756 | -0.0020936 | -0.0043641 | 0.0014547 | 0.065134 | 0.058492 | -0.0014633 | -0.0019803 | -0.0034031 | 0.045776 | 63.0661 | -59.2667 | 63.0478 | -58.9645 | -59.3536 | -59.3033 | 0.091553 | 63.0844 | -59.1568 | 63.1256 | -59.1156 | -59.2438 | -58.992 | -0.14191 | -0.050354 | -0.068664 | -0.01831 | 0.16022 | 0.30212 | -0.17853 | -0.22888 | -0.036621 | -0.086975 | -0.032043 | -0.032043 | 0.009155 | 0.041198 | 0.009155 | 0.041198 | 0.13275 | -0.11902 | 0.16479 | -0.086975
449.7812 | 122.4701 | -59.1522 | 63.1393 | -0.038986 | 0.00048698 | 0.066623 | 0.068064 | -0.0015231 | 0.00080757 | -0.0014526 | 0.0078429 | 0.0010275 | 0.065775 | 0.064046 | -0.0010361 | -0.00091214 | 0.0025785 | 0.032043 | 63.0798 | -59.2896 | 63.0432 | -58.9645 | -59.3445 | -59.317 | 0.082397 | 63.0844 | -59.1934 | 63.1348 | -59.1385 | -59.2392 | -58.9966 | -0.13733 | -0.059509 | -0.09613 | -0.036621 | 0.18768 | 0.32501 | -0.16479 | -0.19226 | -0.027466 | -0.054931 | -0.041198 | -0.054931 | -0.004577 | 0.050354 | 0.013733 | 0.054931 | 0.15564 | -0.086975 | 0.19684 | -0.045777
450.7812 | 122.4289 | -59.1888 | 63.1165 | -0.10765 | 0.00048698 | 0.071109 | 0.072123 | 0.0014675 | -4.7032e-05 | -0.00017099 | -0.10202 | 0.00038669 | 0.065561 | 0.068532 | -0.0010361 | -0.0026211 | -0.0014805 | -0.068665 | 63.0753 | -59.3216 | 62.9425 | -59.0836 | -59.5001 | -59.4543 | -0.1236 | 63.0844 | -59.2209 | 63.0615 | -59.2438 | -59.3811 | -59.0698 | -0.13275 | -0.041199 | -0.17395 | -0.13275 | 0.10529 | 0.23804 | -0.2655 | -0.31128 | -0.13275 | -0.17853 | -0.032043 | -0.032043 | -0.054932 | -0.022889 | -0.054931 | -0.022888 | 0.11902 | -0.19226 | 0.15106 | -0.16022
451.7812 | 122.5021 | -59.1293 | 63.1348 | -0.069504 | 0.00091438 | 0.067477 | 0.067424 | 0.00039951 | 0.001662 | 0.0013244 | 0.0017395 | -4.0508e-05 | 0.065775 | 0.058919 | 3.2078e-05 | -0.0017667 | -0.0012667 | 0.027466 | 63.0936 | -59.2758 | 63.0432 | -58.9645 | -59.3628 | -59.317 | 0.091553 | 63.0661 | -59.1614 | 63.1165 | -59.111 | -59.2392 | -58.9783 | -0.14648 | -0.041199 | -0.091553 | -0.050354 | 0.16479 | 0.31128 | -0.18768 | -0.23346 | -0.041199 | -0.086975 | -0.032044 | -0.068665 | -0.018311 | 0.050354 | 0.01831 | 0.050354 | 0.15106 | -0.10986 | 0.18311 | -0.07782
452.7812 | 122.438 | -59.1705 | 63.1485 | -0.031357 | -0.00058102 | 0.069613 | 0.068064 | 0.0038175 | -0.00026063 | 0.0023926 | 0.0047913 | -4.0508e-05 | 0.066415 | 0.063192 | -0.00060892 | -0.0021939 | -0.0021213 | -0.018311 | 63.1302 | -59.285 | 62.9562 | -59.0424 | -59.436 | -59.3857 | 0.004578 | 63.1073 | -59.1843 | 63.0707 | -59.166 | -59.285 | -59.0286 | -0.11444 | -0.018311 | -0.19226 | -0.17395 | 0.12817 | 0.24262 | -0.21515 | -0.2655 | -0.10071 | -0.15106 | -0.013733 | -0.041199 | -0.07782 | -0.036621 | 0.004577 | 0.01831 | 0.14191 | -0.11444 | 0.15564 | -0.10071
453.7812 | 122.4792 | -59.1293 | 63.1073 | -0.069504 | 0.00048698 | 0.066836 | 0.070201 | 0.0036039 | 0.0023028 | -0.00017099 | -0.08371 | 0.0020957 | 0.064279 | 0.067037 | 3.2078e-05 | -0.00091214 | -0.00083948 | 0.022888 | 63.1165 | -59.3033 | 63.0157 | -59.0012 | -59.4177 | -59.3903 | 0.036621 | 63.1073 | -59.2072 | 63.1302 | -59.1614 | -59.285 | -58.9828 | -0.17395 | 0.009155 | -0.091553 | -0.10071 | 0.12817 | 0.30212 | -0.26093 | -0.28839 | -0.086975 | -0.11444 | -0.07782 | 0 | 0.022888 | 0.022888 | -0.032044 | 0.045776 | 0.14648 | -0.15564 | 0.2243 | -0.07782
454.7812 | 122.4289 | -59.1705 | 63.1577 | -0.023727 | -0.0016492 | 0.065555 | 0.068919 | -0.0013095 | 0.0012348 | -0.0010254 | -0.037933 | 0.0010275 | 0.064706 | 0.065542 | 0.00045928 | 0.00036966 | -0.00062588 | -0.07782 | 63.1393 | -59.2758 | 62.9929 | -59.0973 | -59.4543 | -59.3628 | -0.004578 | 63.0707 | -59.1843 | 63.0981 | -59.1705 | -59.2987 | -59.0515 | -0.10529 | -0.018311 | -0.16479 | -0.14648 | 0.073242 | 0.17853 | -0.19226 | -0.28381 | -0.086975 | -0.17853 | -0.013733 | -0.086975 | -0.059509 | 0.027466 | 0 | 0.013733 | 0.11902 | -0.12817 | 0.13275 | -0.11444
455.7812 | 122.4564 | -59.1339 | 63.1393 | -0.061874 | 0.002837 | 0.069827 | 0.068492 | -0.0021641 | 0.00080757 | 0.00089721 | -0.025726 | 0.0016685 | 0.066843 | 0.065969 | -0.00018152 | -0.0011257 | 0.00044212 | -0.009155 | 63.0981 | -59.2987 | 62.97 | -59.0469 | -59.4315 | -59.3765 | 0.018311 | 63.0981 | -59.2163 | 63.1165 | -59.1339 | -59.2941 | -58.992 | -0.1648 | -0.041198 | -0.16937 | -0.12817 | 0.086975 | 0.25177 | -0.24262 | -0.29755 | -0.07782 | -0.13275 | -0.082398 | -0.041198 | -0.022888 | 0.01831 | 0 | 0.082398 | 0.14191 | -0.16022 | 0.22431 | -0.077819
456.7812 | 122.5067 | -59.1705 | 63.1439 | -0.10765 | 0.0019824 | 0.065341 | 0.068919 | 0.0038175 | 0.0012348 | -0.0018798 | -0.074554 | 0.0023093 | 0.064493 | 0.064901 | -0.0010361 | -0.0015529 | -0.0029757 | -0.15564 | 63.0615 | -59.3262 | 62.9654 | -59.0195 | -59.4177 | -59.3811 | 0.022888 | 63.0753 | -59.198 | 63.1073 | -59.1614 | -59.3033 | -59.0103 | -0.15564 | -0.082398 | -0.17853 | -0.09613 | 0.15106 | 0.3067 | -0.21057 | -0.24719 | -0.054932 | -0.091553 | -0.027466 | -0.068665 | -0.036621 | 0.032044 | 0.009155 | 0.036621 | 0.16022 | -0.13275 | 0.18768 | -0.10529
457.7812 | 122.4747 | -59.1431 | 63.1714 | -0.069504 | 0.002196 | 0.066623 | 0.066996 | 0.0018949 | 0.001662 | -0.0016662 | 0.0047913 | -0.0019631 | 0.064279 | 0.056356 | 0.0019547 | -0.0041165 | 0.0021511 | 0.018311 | 63.0981 | -59.2941 | 62.9837 | -59.024 | -59.4131 | -59.3536 | 0.064087 | 63.1119 | -59.1797 | 63.0753 | -59.166 | -59.285 | -59.0286 | -0.15106 | -0.073242 | -0.18768 | -0.11444 | 0.11902 | 0.27008 | -0.21057 | -0.27008 | -0.05951 | -0.11902 | -0.036621 | -0.05951 | -0.096131 | -0.036621 | -0.022889 | 0.013732 | 0.11444 | -0.14191 | 0.15106 | -0.10529
458.7812 | 122.4564 | -59.1385 | 63.1622 | 0.0067903 | 0.00048698 | 0.067691 | 0.068064 | -0.0010959 | -0.00047423 | 0.00047001 | 0.1116 | 0.0012411 | 0.066415 | 0.062764 | 0.0011001 | -0.0026211 | 0.0025785 | -0.036621 | 63.1165 | -59.3079 | 62.9562 | -59.0515 | -59.436 | -59.3308 | -0.12817 | 63.121 | -59.1934 | 63.0798 | -59.2026 | -59.3124 | -59.0469 | -0.16937 | -0.045776 | -0.20599 | -0.16022 | 0.086975 | 0.25635 | -0.19226 | -0.29755 | -0.022889 | -0.12817 | -0.054931 | -0.041198 | -0.082397 | -0.041199 | -0.064087 | -0.009156 | 0.091553 | -0.17395 | 0.14648 | -0.11902
459.7812 | 122.4976 | -59.1705 | 63.1256 | -0.046616 | 0.001128 | 0.068545 | 0.068278 | 0.0023221 | 0.0012348 | 0.00047001 | -0.056244 | 0.00081389 | 0.067056 | 0.064687 | -0.00060892 | -0.0030483 | -0.0012667 | 0.022888 | 63.121 | -59.2941 | 63.0066 | -59.0103 | -59.3948 | -59.3674 | 0.073242 | 63.0936 | -59.2026 | 63.121 | -59.1476 | -59.2896 | -58.9874 | -0.1236 | -0.004577 | -0.11902 | -0.11444 | 0.16022 | 0.28381 | -0.19684 | -0.2243 | -0.073243 | -0.10071 | -0.032044 | -0.032043 | -0.004577 | 0.027466 | 0.022888 | 0.054932 | 0.18311 | -0.11902 | 0.21515 | -0.086975
460.7969 | 122.525 | -59.1705 | 63.0798 | -0.16106 | -0.00058102 | 0.068118 | 0.067851 | 0.0029629 | 0.00038017 | 0.0013244 | -0.037933 | 0.0020957 | 0.066629 | 0.068319 | -0.00082252 | 0.00036966 | -0.00019868 | -0.10529 | 63.0753 | -59.3262 | 62.9517 | -59.0652 | -59.4772 | -59.4177 | -0.082397 | 63.0569 | -59.2392 | 63.0981 | -59.166 | -59.285 | -58.9783 | -0.15564 | -0.004578 | -0.12817 | -0.1236 | 0.10529 | 0.26093 | -0.24719 | -0.3067 | -0.091553 | -0.15106 | -0.068665 | -0.022888 | 0.018311 | 0.041199 | 0.004577 | 0.073242 | 0.19226 | -0.11444 | 0.26093 | -0.045776
461.7812 | 122.4792 | -59.1614 | 63.1485 | -0.061874 | -0.00015382 | 0.068545 | 0.06956 | 0.0025357 | 0.00016657 | -0.00059819 | -0.016571 | 0.0012411 | 0.065561 | 0.065114 | -0.0016769 | -0.0028347 | 0.0017239 | -0.018311 | 63.1165 | -59.3079 | 62.97 | -59.0652 | -59.4543 | -59.404 | -0.004578 | 63.1073 | -59.1888 | 63.089 | -59.1614 | -59.285 | -59.0149 | -0.14648 | -0.032044 | -0.17853 | -0.14648 | 0.09613 | 0.24261 | -0.24262 | -0.29297 | -0.096131 | -0.14648 | -0.027466 | -0.041199 | -0.05951 | -0.018311 | 0 | 0.027466 | 0.14648 | -0.1236 | 0.17395 | -0.09613
462.7812 | 122.4335 | -59.1934 | 63.1073 | -0.061874 | 0.00048698 | 0.06705 | 0.069346 | 0.0055265 | 0.001662 | -0.00059819 | -0.013519 | 0.0018821 | 0.066415 | 0.064901 | -0.0014633 | -0.0028347 | -0.0025485 | -0.10986 | 63.0661 | -59.3491 | 63.0798 | -58.8776 | -59.5688 | -59.4727 | -0.13275 | 63.0615 | -59.2163 | 63.0661 | -59.23 | -59.3674 | -59.0744 | -0.15564 | -0.041199 | -0.027466 | 0.013733 | 0.31586 | 0.4715 | -0.27924 | -0.37537 | -0.1236 | -0.21973 | -0.022889 | -0.045777 | -0.041199 | 0.004578 | -0.036622 | -0.013733 | 0.11902 | -0.17395 | 0.14191 | -0.15106
463.7812 | 122.4609 | -59.166 | 63.121 | -0.092392 | 5.9776e-05 | 0.066195 | 0.068064 | -0.00088229 | 0.0057208 | 0.0011108 | -0.034882 | 0.0016685 | 0.06492 | 0.06661 | -0.0012497 | -0.0030483 | 0.0015103 | -0.004578 | 63.1073 | -59.3262 | 63.002 | -59.0149 | -59.4223 | -59.3536 | 0.027466 | 63.089 | -59.2072 | 63.1165 | -59.1431 | -59.2346 | -58.9645 | -0.16022 | -0.013733 | -0.11902 | -0.10529 | 0.15106 | 0.31128 | -0.18768 | -0.25635 | -0.027466 | -0.09613 | -0.041198 | -0.032044 | -0.004578 | 0.027466 | 0.022889 | 0.064087 | 0.20142 | -0.068664 | 0.24261 | -0.027466
464.7812 | 122.4243 | -59.1522 | 63.1439 | -0.069504 | -0.0018628 | 0.067477 | 0.068064 | -0.00088229 | 0.00080757 | -0.0025208 | 0.023102 | 0.00038669 | 0.06727 | 0.063833 | 0.00088648 | -0.0030483 | 0.00065592 | -0.045776 | 63.121 | -59.3079 | 62.9562 | -59.0515 | -59.436 | -59.3811 | 0 | 63.1073 | -59.2072 | 63.089 | -59.1797 | -59.2941 | -59.0469 | -0.15564 | -0.022888 | -0.18768 | -0.16479 | 0.10071 | 0.25635 | -0.22888 | -0.28381 | -0.073243 | -0.12817 | -0.054931 | -0.036621 | -0.054932 | -0.018311 | -0.027465 | 0.027466 | 0.10529 | -0.14191 | 0.16022 | -0.086975
465.7969 | 122.5296 | -59.1431 | 63.0615 | -0.069504 | 0.00091438 | 0.068973 | 0.067424 | 0.0025357 | 0.00016657 | -0.00081179 | -0.074554 | -0.0019631 | 0.065347 | 0.062978 | -0.00060892 | -0.0034757 | -0.0025485 | -0.032043 | 63.0295 | -59.2941 | 63.0203 | -58.9828 | -59.404 | -59.4315 | 0.15106 | 62.9929 | -59.198 | 63.176 | -59.1156 | -59.23 | -58.9371 | -0.15106 | -0.032043 | -0.041198 | -0.009155 | 0.16022 | 0.31128 | -0.28839 | -0.26093 | -0.13733 | -0.10986 | -0.054932 | -0.068664 | 0.11444 | 0.1831 | 0.027465 | 0.082397 | 0.20599 | -0.086976 | 0.26093 | -0.032044
466.7812 | 122.4655 | -59.1568 | 63.1393 | -0.031357 | -0.0014356 | 0.066195 | 0.067637 | 0.0010403 | 0.00059397 | -0.0031616 | 0.016998 | 0.00060029 | 0.066415 | 0.066823 | 0.00088648 | -0.0026211 | 0.00065592 | -0.018311 | 63.1256 | -59.3124 | 63.1256 | -58.9966 | -59.7107 | -59.5184 | -0.041199 | 63.1119 | -59.2072 | 63.0844 | -59.1797 | -59.3079 | -59.024 | -0.15564 | -0.013733 | -0.013733 | 0 | 0.16022 | 0.31586 | -0.36163 | -0.55389 | -0.20599 | -0.39825 | -0.050354 | -0.027466 | -0.054931 | -0.027465 | -0.022888 | 0.027466 | 0.13275 | -0.15106 | 0.1831 | -0.10071
467.7812 | 122.4426 | -59.1476 | 63.1577 | -0.10765 | 5.9776e-05 | 0.065555 | 0.069346 | -0.0021641 | 0.0031574 | 0.002179 | -0.022675 | 0.00017309 | 0.065347 | 0.061056 | -0.0014633 | -0.0017667 | -0.0014805 | -0.018311 | 63.1073 | -59.3124 | 62.9471 | -59.0607 | -59.4452 | -59.3948 | -0.013733 | 63.1073 | -59.198 | 63.1165 | -59.1476 | -59.2987 | -59.0103 | -0.16479 | -0.050354 | -0.21057 | -0.16022 | 0.086975 | 0.25177 | -0.24719 | -0.29755 | -0.082397 | -0.13275 | -0.050354 | -0.050354 | -0.041199 | 0.009155 | 0 | 0.050354 | 0.13733 | -0.15106 | 0.18768 | -0.10071
468.7969 | 122.5113 | -59.1705 | 63.089 | -0.16869 | -0.00015382 | 0.068759 | 0.067424 | 0.0010403 | 0.00016657 | -0.0014526 | 0.038361 | -0.0011087 | 0.067056 | 0.058919 | -0.0018905 | -0.0034757 | 0.0023647 | -0.068665 | 63.0341 | -59.317 | 63.002 | -59.0195 | -59.4269 | -59.3445 | -0.022888 | 63.0386 | -59.2026 | 63.1165 | -59.1568 | -59.2804 | -58.992 | -0.14648 | -0.054931 | -0.086975 | -0.032044 | 0.15106 | 0.29755 | -0.17395 | -0.25635 | -0.027465 | -0.10986 | -0.032044 | -0.050354 | 0.027466 | 0.07782 | 0.013733 | 0.045777 | 0.17853 | -0.10986 | 0.21057 | -0.07782
469.7812 | 122.4472 | -59.1843 | 63.1668 | -0.18394 | 0.00048698 | 0.06705 | 0.069773 | -0.00024129 | 0.00080757 | 4.2808e-05 | -0.025726 | 0.0025229 | 0.065561 | 0.061056 | -0.00082252 | -0.0017667 | -0.00019868 | -0.018311 | 63.1348 | -59.3262 | 62.9608 | -59.0561 | -59.4681 | -59.4177 | 0.004578 | 63.0981 | -59.1934 | 63.0981 | -59.1614 | -59.2987 | -59.0103 | -0.14191 | -0.032043 | -0.20599 | -0.17395 | 0.12817 | 0.27008 | -0.23346 | -0.28381 | -0.091553 | -0.14191 | -0.009155 | -0.068664 | -0.068664 | 0 | 0.022888 | 0.032043 | 0.17395 | -0.11444 | 0.18311 | -0.10529
470.7969 | 122.5983 | -59.1568 | 63.0661 | -0.084763 | 0.0015552 | 0.069186 | 0.068278 | 0.0014675 | 0.0020892 | -0.00059819 | 0.0017395 | -0.00025411 | 0.065347 | 0.062337 | -0.0014633 | -0.0021939 | 0.0010831 | -0.050354 | 63.0295 | -59.3033 | 63.0295 | -58.992 | -59.404 | -59.3353 | 0.036621 | 62.9974 | -59.2484 | 63.1531 | -59.0744 | -59.2346 | -58.9645 | -0.14648 | -0.036621 | -0.036621 | 0 | 0.16479 | 0.31128 | -0.17853 | -0.24719 | -0.032043 | -0.10071 | -0.091553 | -0.068664 | 0.086975 | 0.15564 | 0.082397 | 0.17395 | 0.19226 | -0.07782 | 0.28381 | 0.013733
471.7969 | 122.4792 | -59.1614 | 63.1027 | -0.092392 | 0.0017688 | 0.069827 | 0.070201 | 0.0018949 | 0.00016657 | -0.0016662 | -0.047089 | 0.0025229 | 0.065775 | 0.072164 | -0.00018152 | -0.0013393 | 1.4916e-05 | -0.004578 | 63.089 | -59.3124 | 62.9791 | -59.0332 | -59.4406 | -59.3994 | 0.013733 | 63.0753 | -59.2026 | 63.089 | -59.1843 | -59.3079 | -59.024 | -0.15106 | -0.013733 | -0.1236 | -0.10986 | 0.12817 | 0.27924 | -0.23804 | -0.27924 | -0.086975 | -0.12817 | -0.041199 | -0.027466 | -0.013733 | 0.013733 | -0.022888 | 0.018311 | 0.13733 | -0.14648 | 0.17853 | -0.10529
472.7812 | 122.4701 | -59.1522 | 63.1393 | -0.046616 | 5.9776e-05 | 0.066836 | 0.067637 | -0.0015231 | 0.00038017 | -0.0010254 | 0.041412 | 0.00081389 | 0.064493 | 0.064046 | -0.00039512 | -0.0036893 | -0.0048983 | 0.027466 | 63.089 | -59.285 | 62.9974 | -58.9783 | -59.3719 | -59.3079 | 0.07782 | 63.089 | -59.1705 | 63.1073 | -59.1476 | -59.2621 | -59.0332 | -0.13275 | -0.050354 | -0.14191 | -0.091552 | 0.17395 | 0.3067 | -0.15564 | -0.21973 | -0.022888 | -0.086975 | -0.01831 | -0.050354 | -0.032043 | 0.018311 | 0.004578 | 0.022888 | 0.11902 | -0.10986 | 0.13733 | -0.091553
473.7812 | 122.3831 | -59.1431 | 63.1577 | -0.046616 | 0.0017688 | 0.066836 | 0.067424 | -0.0025913 | 0.0025166 | -0.00017099 | 0.0047913 | -0.0011087 | 0.064279 | 0.061269 | -0.0010361 | -0.0034757 | 0.0027921 | -0.018311 | 63.121 | -59.2804 | 62.9654 | -59.0378 | -59.4269 | -59.3765 | 0.004578 | 63.089 | -59.198 | 63.0798 | -59.166 | -59.285 | -59.0469 | -0.13733 | -0.036621 | -0.19226 | -0.15564 | 0.10528 | 0.24262 | -0.23346 | -0.28381 | -0.09613 | -0.14648 | -0.054932 | -0.068665 | -0.07782 | -0.009155 | -0.022889 | 0.032043 | 0.09613 | -0.14191 | 0.15106 | -0.086975
474.7812 | 122.4426 | -59.1843 | 63.1348 | -0.10002 | 0.00091438 | 0.067477 | 0.069987 | 0.0033901 | 0.0023028 | -0.0023072 | -0.059296 | 0.0018821 | 0.06492 | 0.068105 | -0.00039512 | -0.00027134 | 0.0015103 | -0.045776 | 63.0661 | -59.3353 | 62.9608 | -59.0652 | -59.4772 | -59.4269 | -0.036621 | 63.0936 | -59.2255 | 63.0844 | -59.2072 | -59.3445 | -59.0561 | -0.15106 | -0.068665 | -0.17395 | -0.10529 | 0.11902 | 0.27008 | -0.24261 | -0.29297 | -0.091553 | -0.14191 | -0.041199 | -0.041199 | -0.050354 | -0.009155 | -0.022888 | 0.018311 | 0.12817 | -0.16022 | 0.16937 | -0.11902
475.7812 | 122.406 | -59.2255 | 63.1119 | -0.038986 | 0.002837 | 0.071322 | 0.068919 | 0.0046719 | 0.0012348 | 4.2808e-05 | -0.062347 | 0.0010275 | 0.065561 | 0.069173 | 0.0013139 | -0.0017667 | 0.00065592 | -0.073242 | 63.0524 | -59.3582 | 62.9196 | -59.1293 | -59.5184 | -59.4635 | -0.13733 | 63.0707 | -59.2484 | 63.0478 | -59.2529 | -59.3857 | -59.0744 | -0.13275 | -0.059509 | -0.19226 | -0.13275 | 0.096131 | 0.22888 | -0.23804 | -0.29297 | -0.10529 | -0.16022 | -0.022888 | -0.041198 | -0.064086 | -0.022888 | -0.027466 | -0.004578 | 0.15106 | -0.16022 | 0.17395 | -0.13733
476.7812 | 122.4976 | -59.1431 | 63.1165 | -0.054245 | 0.002196 | 0.068118 | 0.065287 | -0.00045509 | 0.00038017 | 0.0023926 | 0.050568 | -0.00046791 | 0.06492 | 0.058706 | 0.00067288 | -0.0032619 | -0.0023349 | 0.032043 | 63.0753 | -59.2712 | 63.0386 | -58.9645 | -59.3491 | -59.3079 | 0.09613 | 63.0569 | -59.198 | 63.1256 | -59.1064 | -59.2255 | -58.9737 | -0.12817 | -0.041199 | -0.07782 | -0.036621 | 0.17853 | 0.3067 | -0.1648 | -0.20599 | -0.036621 | -0.07782 | -0.054932 | -0.059509 | 0.009155 | 0.068664 | 0.036621 | 0.091553 | 0.16937 | -0.082398 | 0.2243 | -0.027466
477.7812 | 122.4884 | -59.166 | 63.1256 | -0.084763 | 0.0019824 | 0.066836 | 0.068919 | 0.0027493 | 0.0018756 | -0.00081179 | -0.047089 | 0.0016685 | 0.067056 | 0.063833 | -0.0014633 | -0.0013393 | 0.0010831 | -0.009155 | 63.1119 | -59.317 | 63.0249 | -58.9966 | -59.4681 | -59.404 | 0.032043 | 63.089 | -59.23 | 63.0936 | -59.166 | -59.2941 | -59.0103 | -0.15106 | -0.013733 | -0.10071 | -0.086975 | 0.16937 | 0.32044 | -0.23804 | -0.30212 | -0.086975 | -0.15106 | -0.064087 | -0.036621 | -0.032043 | 0.004578 | 0 | 0.064087 | 0.15564 | -0.12817 | 0.21973 | -0.064086
478.7812 | 122.5616 | -59.1476 | 63.0844 | 0.16701 | -0.0014356 | 0.072818 | 0.065715 | -0.00024129 | -0.0013288 | 0.0017518 | 0.087189 | -0.0013223 | 0.06492 | 0.062551 | -0.00039512 | -0.0036893 | -0.0038303 | -0.11444 | 63.0844 | -59.2438 | 63.089 | -58.9371 | -59.3445 | -59.2621 | 0.10529 | 63.0798 | -59.166 | 63.1531 | -59.0698 | -59.1614 | -58.9279 | -0.09613 | 0 | 0.004577 | 0.004577 | 0.21057 | 0.3067 | -0.11444 | -0.19684 | -0.018311 | -0.10071 | -0.018311 | -0.004578 | 0.068664 | 0.073242 | 0.07782 | 0.096131 | 0.21973 | -0.013733 | 0.23804 | 0.004578
479.7812 | 122.4976 | -59.1156 | 63.1348 | -0.11528 | 0.0015552 | 0.068118 | 0.066996 | 0.00082671 | 0.0029438 | -0.00081179 | 0.023102 | -0.00068151 | 0.06492 | 0.058278 | 0.00024568 | -0.0026211 | -0.00041228 | 0.027466 | 63.0753 | -59.2712 | 63.0524 | -58.9462 | -59.3445 | -59.2987 | 0.10986 | 63.0753 | -59.1568 | 63.1073 | -59.111 | -59.2346 | -58.9783 | -0.15564 | -0.05951 | -0.082398 | -0.022888 | 0.16937 | 0.32501 | -0.18311 | -0.22888 | -0.027466 | -0.073242 | -0.041198 | -0.05951 | -0.027466 | 0.032044 | 0.004578 | 0.045776 | 0.13733 | -0.11902 | 0.17853 | -0.07782
480.7812 | 122.4976 | -59.166 | 63.1393 | -0.038986 | 0.0019824 | 0.068545 | 0.071269 | -0.00024129 | 0.00038017 | 0.00047001 | -0.040985 | 0.0012411 | 0.065775 | 0.070882 | -0.0014633 | -0.00091214 | -0.0019077 | 0.009155 | 63.1119 | -59.3079 | 63.0157 | -59.0057 | -59.4269 | -59.3811 | 0.041199 | 63.089 | -59.1751 | 63.1302 | -59.1476 | -59.2896 | -59.0057 | -0.14191 | -0.027466 | -0.1236 | -0.09613 | 0.16022 | 0.30212 | -0.21515 | -0.26093 | -0.073243 | -0.11902 | -0.009155 | -0.050354 | -0.009155 | 0.041199 | 0.018311 | 0.027466 | 0.16022 | -0.1236 | 0.16937 | -0.11444
481.7812 | 122.58 | -59.1293 | 63.1027 | -0.054245 | -0.0035718 | 0.064486 | 0.065287 | -0.0017367 | -4.7032e-05 | -0.00081179 | 0.035309 | -0.0015359 | 0.066202 | 0.064687 | -0.00060892 | -0.0015529 | 0.00022852 | 0.29297 | 63.121 | -59.2117 | 63.002 | -58.9508 | -59.2896 | -59.2941 | 0.13733 | 63.0524 | -59.1431 | 63.1393 | -59.0652 | -59.1843 | -58.9462 | -0.082398 | 0.018311 | -0.10071 | -0.11902 | 0.17853 | 0.26093 | -0.16479 | -0.16022 | -0.082397 | -0.07782 | -0.013733 | -0.050354 | 0.036621 | 0.086975 | 0.064086 | 0.077819 | 0.18311 | -0.054932 | 0.19684 | -0.041199
482.7812 | 122.5159 | -59.1202 | 63.1393 | -0.023727 | 0.00070078 | 0.0694 | 0.067637 | -0.0021641 | -0.00068783 | -0.0020936 | 0.035309 | -0.0019631 | 0.064493 | 0.057638 | 0.0013139 | -0.0013393 | 0.0012967 | 0.013733 | 63.0615 | -59.2529 | 63.0524 | -58.9554 | -59.3445 | -59.2987 | 0.10986 | 63.0844 | -59.1522 | 63.1348 | -59.079 | -59.2163 | -58.9783 | -0.13275 | -0.07782 | -0.086975 | -0.009155 | 0.1648 | 0.29755 | -0.17853 | -0.2243 | -0.045776 | -0.091552 | -0.032044 | -0.054931 | -0.004577 | 0.050354 | 0.041199 | 0.073243 | 0.14191 | -0.096131 | 0.17395 | -0.064087
483.7812 | 122.4426 | -59.1385 | 63.121 | -0.12291 | 0.0015552 | 0.068332 | 0.068705 | 0.00082671 | 0.00016657 | 0.00089721 | -0.040985 | 0.00081389 | 0.066629 | 0.066182 | -0.00060892 | -0.0028347 | -0.0021213 | -0.009155 | 63.0753 | -59.3399 | 62.9929 | -59.0103 | -59.5734 | -59.4315 | 0.018311 | 63.0753 | -59.1934 | 63.121 | -59.166 | -59.3033 | -58.9874 | -0.20142 | -0.045777 | -0.12817 | -0.082397 | 0.12817 | 0.32959 | -0.29297 | -0.43487 | -0.091553 | -0.23346 | -0.054931 | -0.045777 | 0 | 0.045777 | -0.027466 | 0.027465 | 0.15106 | -0.16479 | 0.20599 | -0.10986
484.7812 | 122.5708 | -59.1293 | 63.0753 | -0.11528 | -0.00036742 | 0.068332 | 0.066783 | -0.00066869 | 0.00059397 | -0.002948 | 0.065826 | -0.00089511 | 0.064706 | 0.060628 | 0.0011001 | -0.0021939 | 0.0012967 | -0.036621 | 63.0798 | -59.2758 | 63.0203 | -58.9874 | -59.3811 | -59.3124 | 0.036621 | 63.0478 | -59.1888 | 63.1622 | -59.0927 | -59.2026 | -58.9325 | -0.14648 | 0.004578 | -0.054931 | -0.059509 | 0.14191 | 0.28839 | -0.18311 | -0.25177 | -0.036621 | -0.10529 | -0.05951 | -0.027465 | 0.086975 | 0.11444 | 0.036621 | 0.096131 | 0.19684 | -0.073243 | 0.25635 | -0.013733
485.7812 | 122.5525 | -59.1385 | 63.0753 | -0.13054 | 0.00070078 | 0.068545 | 0.066569 | 0.0010403 | 0.0031574 | -0.0014526 | 0.038361 | -0.00068151 | 0.064279 | 0.058278 | -0.0014633 | -0.0036893 | 0.0017239 | -0.068665 | 63.0661 | -59.2896 | 63.0112 | -59.0012 | -59.4086 | -59.3353 | 0.004578 | 63.0295 | -59.1934 | 63.1485 | -59.1064 | -59.2163 | -58.9554 | -0.15106 | -0.009155 | -0.064087 | -0.054932 | 0.13733 | 0.28839 | -0.19684 | -0.27008 | -0.045776 | -0.11902 | -0.054931 | -0.045776 | 0.073243 | 0.11902 | 0.032044 | 0.086975 | 0.18311 | -0.07782 | 0.23804 | -0.022889
486.7812 | 122.4792 | -59.1385 | 63.1531 | -0.077133 | 0.00070078 | 0.067264 | 0.065501 | 0.0012539 | -0.00047423 | 0.00025641 | -0.0043641 | -0.00068151 | 0.065134 | 0.061696 | -0.00082252 | -0.0028347 | 0.00044212 | 0.036621 | 63.0707 | -59.2667 | 63.0295 | -58.9737 | -59.3857 | -59.3216 | 0.073242 | 63.0981 | -59.1843 | 63.1073 | -59.1522 | -59.2758 | -59.024 | -0.12817 | -0.082397 | -0.1236 | -0.041199 | 0.16479 | 0.29297 | -0.1831 | -0.24719 | -0.054931 | -0.11902 | -0.045776 | -0.054931 | -0.045776 | 0.009155 | -0.013733 | 0.032043 | 0.11444 | -0.13733 | 0.16022 | -0.091553
487.7812 | 122.5525 | -59.1019 | 63.0981 | -0.061874 | 0.00070078 | 0.068118 | 0.063365 | -0.0021641 | 0.0031574 | -0.0025208 | 0.17874 | 0.00060029 | 0.064066 | 0.062978 | -0.00060892 | -0.0017667 | -0.0034031 | 0.022888 | 63.0569 | -59.2804 | 63.0753 | -58.9096 | -59.317 | -59.2712 | 0.054932 | 63.0936 | -59.1797 | 63.1851 | -59.0469 | -59.1843 | -58.9371 | -0.17853 | -0.041199 | -0.022889 | 0.01831 | 0.19226 | 0.37079 | -0.16937 | -0.21515 | 0.009156 | -0.036621 | -0.077819 | -0.004578 | 0.086975 | 0.091553 | 0.054932 | 0.13275 | 0.1648 | -0.082397 | 0.24261 | -0.004578
488.7812 | 122.5708 | -59.0973 | 63.1073 | 0.052567 | 0.0015552 | 0.068332 | 0.067424 | 0.00018591 | 0.00059397 | -0.0027344 | 0.0047913 | 0.00017309 | 0.064493 | 0.063192 | 3.2078e-05 | -0.0039029 | 0.00086952 | -0.009155 | 63.0478 | -59.2667 | 63.0844 | -58.9325 | -59.3445 | -59.2896 | 0.12817 | 63.0524 | -59.1614 | 63.1577 | -59.0561 | -59.1797 | -58.9279 | -0.16937 | -0.059509 | -0.022888 | 0.036621 | 0.16479 | 0.33417 | -0.19226 | -0.24719 | -0.022888 | -0.077819 | -0.064087 | -0.054932 | 0.050354 | 0.10529 | 0.041199 | 0.10529 | 0.16937 | -0.082397 | 0.23346 | -0.01831
489.7812 | 122.5113 | -59.1019 | 63.1256 | -0.10002 | 0.00027338 | 0.068545 | 0.066569 | -2.7692e-05 | 0.0014484 | 0.001538 | 0.0017395 | -0.00046791 | 0.065134 | 0.067037 | 0.00024568 | -0.0036893 | 0.00022852 | 0.022888 | 63.0707 | -59.2804 | 63.0615 | -58.96 | -59.3536 | -59.2987 | 0.10986 | 63.0844 | -59.1751 | 63.1393 | -59.1019 | -59.2438 | -58.9874 | -0.17853 | -0.054931 | -0.064087 | -0.009156 | 0.14191 | 0.32044 | -0.19684 | -0.25177 | -0.01831 | -0.073242 | -0.073242 | -0.041198 | 0.013733 | 0.054931 | 0 | 0.073242 | 0.11444 | -0.14191 | 0.18768 | -0.068664
490.7812 | 122.5388 | -59.1339 | 63.089 | -0.1458 | 0.00027338 | 0.068118 | 0.069987 | -0.00088229 | 0.0029438 | -0.00081179 | 0.016998 | -0.0017495 | 0.062997 | 0.063405 | -0.0016769 | -0.0028347 | -0.00019868 | -0.045776 | 63.0478 | -59.3079 | 63.002 | -59.0195 | -59.4131 | -59.3491 | 0 | 63.0615 | -59.1934 | 63.1302 | -59.0973 | -59.23 | -58.9645 | -0.17395 | -0.041198 | -0.086975 | -0.045777 | 0.11444 | 0.28839 | -0.21515 | -0.27924 | -0.041199 | -0.10529 | -0.059509 | -0.027466 | 0.041199 | 0.068665 | 0.036621 | 0.09613 | 0.16937 | -0.096131 | 0.22888 | -0.036622
491.7812 | 122.3831 | -59.1339 | 63.1897 | -0.031357 | 0.0019824 | 0.065768 | 0.069133 | -0.0019503 | -0.0019696 | -0.00017099 | -0.037933 | 0.0029501 | 0.064706 | 0.068105 | -0.00018152 | -0.0021939 | -0.00062588 | -0.009155 | 63.089 | -59.3033 | 62.9745 | -59.0424 | -59.436 | -59.3903 | 0.009155 | 63.1439 | -59.2026 | 63.0753 | -59.166 | -59.2941 | -59.0378 | -0.16937 | -0.10071 | -0.21515 | -0.11444 | 0.091553 | 0.26093 | -0.25635 | -0.30212 | -0.086975 | -0.13275 | -0.068665 | -0.045776 | -0.11444 | -0.068665 | -0.032044 | 0.036621 | 0.09613 | -0.16022 | 0.16479 | -0.091552
492.7812 | 122.525 | -59.1339 | 63.0707 | -0.13817 | 0.0015552 | 0.069827 | 0.068705 | 0.00039951 | 0.0010212 | -0.00017099 | -0.028778 | -0.00089511 | 0.062357 | 0.063619 | 3.2078e-05 | -0.0021939 | -0.00083948 | -0.064087 | 63.0157 | -59.3033 | 63.002 | -59.0012 | -59.4406 | -59.3628 | -0.022888 | 63.0386 | -59.1888 | 63.1119 | -59.1248 | -59.2575 | -58.9966 | -0.16937 | -0.054932 | -0.068665 | -0.013733 | 0.13275 | 0.30212 | -0.22888 | -0.3067 | -0.059509 | -0.13733 | -0.054932 | -0.032044 | 0.041198 | 0.073242 | 0.009155 | 0.064087 | 0.13733 | -0.1236 | 0.19226 | -0.068664
493.7812 | 122.5296 | -59.1431 | 63.089 | -0.13817 | 0.00048698 | 0.067904 | 0.068492 | 0.00018591 | 0.0014484 | -0.00038459 | 0.0017395 | 0.00038669 | 0.063211 | 0.063405 | 0.00067288 | -0.0011257 | -0.0021213 | -0.091553 | 63.0386 | -59.3308 | 62.9974 | -59.0424 | -59.4315 | -59.3445 | -0.027466 | 63.0386 | -59.1934 | 63.1073 | -59.1385 | -59.2667 | -58.9966 | -0.18768 | -0.050354 | -0.091552 | -0.041198 | 0.10071 | 0.28839 | -0.20142 | -0.28839 | -0.013732 | -0.10071 | -0.050354 | -0.050354 | 0.018311 | 0.068665 | 0.004577 | 0.054931 | 0.14648 | -0.1236 | 0.19684 | -0.073243
494.7812 | 122.5754 | -59.1385 | 63.0707 | -0.084763 | 0.00027338 | 0.065982 | 0.066996 | 0.0010403 | 0.0025166 | -0.0014526 | 0.0017395 | -0.0015359 | 0.065347 | 0.06426 | 0.00067288 | -0.0017667 | 0.0040737 | -0.050354 | 63.0524 | -59.2804 | 63.0203 | -58.9874 | -59.3765 | -59.3216 | 0.022888 | 63.0432 | -59.198 | 63.1714 | -59.0927 | -59.2117 | -58.9691 | -0.14191 | -0.018311 | -0.050354 | -0.032043 | 0.15106 | 0.29297 | -0.1831 | -0.23804 | -0.041198 | -0.09613 | -0.059509 | -0.027466 | 0.10071 | 0.12817 | 0.045777 | 0.10529 | 0.16937 | -0.073242 | 0.22888 | -0.013733
495.7812 | 122.5159 | -59.1339 | 63.0615 | -0.10002 | -0.0016492 | 0.070041 | 0.066783 | -0.0015231 | 0.00016657 | -0.0027344 | 0.013947 | -4.0508e-05 | 0.064279 | 0.061056 | 0.00045928 | -0.0017667 | -0.0012667 | -0.045776 | 63.0524 | -59.2804 | 63.0157 | -59.0103 | -59.404 | -59.3216 | 0.013733 | 63.0341 | -59.198 | 63.1577 | -59.0973 | -59.2163 | -58.9645 | -0.14649 | -0.009155 | -0.045776 | -0.036621 | 0.1236 | 0.27008 | -0.18768 | -0.27008 | -0.041198 | -0.1236 | -0.064087 | -0.027465 | 0.096131 | 0.1236 | 0.036621 | 0.10071 | 0.16937 | -0.082398 | 0.23346 | -0.018311
496.7969 | 122.4518 | -59.166 | 63.1485 | -0.031357 | -0.0010082 | 0.067691 | 0.069133 | 0.0012539 | -0.00090143 | 0.00025641 | -0.053192 | 0.0014547 | 0.063852 | 0.064687 | -0.00060892 | -0.0024075 | -0.0021213 | -0.009155 | 63.1256 | -59.317 | 62.9837 | -59.0607 | -59.436 | -59.3948 | -0.018311 | 63.1165 | -59.2026 | 63.1027 | -59.1476 | -59.2987 | -59.024 | -0.15106 | -0.022889 | -0.16479 | -0.14191 | 0.10529 | 0.25635 | -0.22888 | -0.27008 | -0.077819 | -0.11902 | -0.036621 | -0.032044 | -0.045777 | -0.013733 | 0.018311 | 0.054932 | 0.14191 | -0.13275 | 0.17853 | -0.09613
497.7812 | 122.5021 | -59.166 | 63.0936 | -0.16869 | -0.0016492 | 0.066623 | 0.068705 | 0.00039951 | 0.0012348 | 0.00089721 | -0.074554 | 0.0010275 | 0.062143 | 0.06426 | -0.00082252 | -0.0017667 | 0.0010831 | -0.059509 | 63.0661 | -59.285 | 62.9974 | -59.0424 | -59.2941 | -59.2987 | 0 | 63.0615 | -59.2072 | 63.1165 | -59.1064 | -59.2163 | -58.9966 | -0.11902 | -0.027466 | -0.09613 | -0.068664 | 0.1236 | 0.24262 | -0.13275 | -0.12817 | -0.013733 | -0.009155 | -0.041198 | -0.032044 | 0.022888 | 0.054932 | 0.05951 | 0.10071 | 0.16937 | -0.050354 | 0.21057 | -0.009156
498.7812 | 122.5296 | -59.1476 | 63.0569 | -0.16106 | 0.0024096 | 0.066623 | 0.069773 | -0.0019503 | 0.0020892 | 0.0013244 | -0.040985 | -0.0017495 | 0.064066 | 0.06896 | -0.0021041 | -0.00048494 | 0.00044212 | -0.022888 | 63.0386 | -59.3124 | 63.0341 | -58.9645 | -59.3994 | -59.3399 | 0.009155 | 63.0203 | -59.1888 | 63.1531 | -59.1064 | -59.2484 | -58.9554 | -0.16479 | -0.018311 | -0.022888 | -0.004577 | 0.1831 | 0.3479 | -0.19226 | -0.25177 | -0.027466 | -0.086975 | -0.041199 | -0.036621 | 0.09613 | 0.13275 | 0.041199 | 0.082398 | 0.19226 | -0.10071 | 0.23346 | -0.059509
499.7812 | 122.5937 | -59.1064 | 63.0798 | -0.077133 | 0.00048698 | 0.065127 | 0.068492 | -0.0034457 | 0.0010212 | -0.0014526 | 0.029205 | -0.00025411 | 0.064706 | 0.063405 | 0.00045928 | -0.0047573 | 0.00065592 | 0.045776 | 63.0753 | -59.2575 | 63.1119 | -58.905 | -59.2987 | -59.2804 | 0.1236 | 63.0524 | -59.1614 | 63.3682 | -59.3216 | -59.317 | -58.9874 | -0.15106 | -0.004578 | 0.032043 | 0.036621 | 0.20142 | 0.35248 | -0.17395 | -0.19226 | -0.022889 | -0.041199 | -0.054932 | -0.027466 | 0.28839 | 0.31586 | -0.21515 | -0.16022 | 0.11902 | -0.21057 | 0.17395 | -0.15564
500.7812 | 122.5113 | -59.1385 | 63.1119 | -0.10765 | -0.0010082 | 0.066409 | 0.067637 | 0.00039951 | 0.00080757 | -0.00081179 | -0.040985 | -0.00068151 | 0.064706 | 0.064687 | -0.00018152 | -0.0039029 | -0.00041228 | -0.1236 | 63.0661 | -59.2896 | 62.9974 | -59.024 | -59.4086 | -59.3582 | 0 | 63.0615 | -59.198 | 63.1348 | -59.111 | -59.2346 | -58.9783 | -0.15106 | -0.045776 | -0.11444 | -0.068664 | 0.11444 | 0.2655 | -0.21973 | -0.27008 | -0.068664 | -0.11902 | -0.059509 | -0.050354 | 0.022889 | 0.073243 | 0.027466 | 0.086975 | 0.16022 | -0.09613 | 0.21973 | -0.036621
501.7812 | 122.5983 | -59.1385 | 63.1073 | -0.023727 | -0.00058102 | 0.067264 | 0.067637 | -0.0025913 | -0.00068783 | 0.0017518 | -0.010467 | 0.00017309 | 0.064706 | 0.062551 | -0.0016769 | -0.0034757 | -0.0021213 | 0 | 63.0524 | -59.2758 | 63.0798 | -58.9188 | -59.3353 | -59.2529 | 0.15564 | 63.0661 | -59.1522 | 63.2263 | -59.0332 | -59.1339 | -58.9279 | -0.13733 | -0.054932 | -0.027466 | 0.027466 | 0.21973 | 0.35706 | -0.11444 | -0.19684 | 0.022888 | -0.059509 | -0.013733 | -0.041199 | 0.11902 | 0.16022 | 0.10529 | 0.11902 | 0.21057 | 0.004578 | 0.22431 | 0.018311
502.7812 | 122.5113 | -59.1705 | 63.0661 | -0.10002 | 0.00048698 | 0.068332 | 0.068705 | 0.0014675 | -0.00047423 | 0.00025641 | -0.047089 | -0.0011087 | 0.064066 | 0.06426 | -0.00060892 | -0.0021939 | 0.00065592 | -0.059509 | 63.0341 | -59.3079 | 62.9883 | -59.0057 | -59.4406 | -59.3674 | -0.032043 | 63.0524 | -59.1888 | 63.1302 | -59.1339 | -59.2712 | -59.0012 | -0.13733 | -0.032043 | -0.07782 | -0.045777 | 0.16479 | 0.30212 | -0.19684 | -0.27008 | -0.05951 | -0.13275 | -0.018311 | -0.013733 | 0.064087 | 0.07782 | 0.036621 | 0.054932 | 0.16937 | -0.10071 | 0.18768 | -0.082397
503.7812 | 122.6028 | -59.1431 | 63.0707 | -0.092392 | -0.00079462 | 0.065127 | 0.070201 | -0.0032321 | -0.001115 | -0.0010254 | -0.016571 | 0.00081389 | 0.065561 | 0.062124 | -0.0010361 | -0.0028347 | 0.0023647 | -0.086975 | 63.0432 | -59.2941 | 63.0066 | -59.0286 | -59.4269 | -59.3491 | -0.036621 | 63.0341 | -59.2255 | 63.1348 | -59.1202 | -59.3353 | -58.9966 | -0.15106 | -0.027466 | -0.064087 | -0.036621 | 0.11444 | 0.2655 | -0.20599 | -0.28381 | -0.054932 | -0.13275 | -0.082398 | -0.036621 | 0.064087 | 0.10071 | 0.022888 | 0.10529 | 0.14648 | -0.19226 | 0.22888 | -0.10986
504.7812 | 122.5296 | -59.1339 | 63.0707 | -0.084763 | -0.001222 | 0.068545 | 0.067851 | -0.0051547 | 0.0014484 | -0.00038459 | -0.016571 | -0.00025411 | 0.064279 | 0.068532 | 0.00045928 | -0.00091214 | -0.00062588 | -0.032043 | 63.0432 | -59.2804 | 63.0615 | -58.9645 | -59.3399 | -59.2896 | 0.050354 | 63.0478 | -59.1797 | 63.1851 | -59.0607 | -59.1705 | -58.9325 | -0.14649 | -0.027466 | -0.009156 | 0.01831 | 0.16937 | 0.31586 | -0.15564 | -0.20599 | -0.009155 | -0.059509 | -0.045776 | -0.022888 | 0.11444 | 0.13733 | 0.073242 | 0.11902 | 0.20142 | -0.036621 | 0.24719 | 0.009155
505.7812 | 122.5204 | -59.1431 | 63.0249 | -0.11528 | -0.00015382 | 0.067904 | 0.069773 | -0.0019503 | 0.00016657 | 0.0019654 | -0.08371 | 0.00038669 | 0.064066 | 0.068319 | -0.0018905 | -5.7742e-05 | -0.0061801 | -0.045776 | 63.0203 | -59.2896 | 63.0432 | -58.9691 | -59.4177 | -59.3399 | 0.022888 | 63.002 | -59.2438 | 63.1485 | -59.0652 | -59.2392 | -58.9691 | -0.14649 | -0.004577 | 0.018311 | 0.022888 | 0.17395 | 0.32044 | -0.19684 | -0.27466 | -0.050354 | -0.12817 | -0.10071 | -0.022888 | 0.1236 | 0.14648 | 0.077819 | 0.17853 | 0.17395 | -0.096131 | 0.27466 | 0.004577
506.7812 | 122.5479 | -59.1476 | 63.0386 | -0.12291 | 0.00070078 | 0.068332 | 0.068064 | 0.00061311 | 0.00080757 | 0.001538 | -0.071503 | 0.00017309 | 0.06492 | 0.064473 | -0.0014633 | -0.0028347 | 0.00044212 | -0.045776 | 63.0249 | -59.285 | 63.0432 | -58.9691 | -59.404 | -59.3353 | 0.009155 | 63.0203 | -59.198 | 63.1714 | -59.1064 | -59.2438 | -58.9691 | -0.13733 | -0.013733 | 0.004578 | 0.018311 | 0.17853 | 0.31586 | -0.18768 | -0.25635 | -0.050354 | -0.11902 | -0.050354 | -0.01831 | 0.13275 | 0.15106 | 0.041199 | 0.091553 | 0.17853 | -0.09613 | 0.22888 | -0.045776
507.7812 | 122.5113 | -59.1385 | 63.0707 | -0.1458 | -0.00036742 | 0.067264 | 0.069346 | 0.0012539 | 0.0018756 | -0.0014526 | -0.028778 | -0.00046791 | 0.065347 | 0.060628 | -0.00060892 | -0.00091214 | -0.0016941 | -0.068665 | 63.0341 | -59.2987 | 62.9974 | -58.992 | -59.4269 | -59.3628 | -0.027466 | 63.0295 | -59.198 | 63.1302 | -59.1385 | -59.2529 | -58.992 | -0.16022 | -0.036621 | -0.073242 | -0.036621 | 0.14648 | 0.3067 | -0.2243 | -0.28839 | -0.064087 | -0.12817 | -0.059509 | -0.041199 | 0.059509 | 0.10071 | 0 | 0.059509 | 0.14648 | -0.11444 | 0.20599 | -0.054932
508.7812 | 122.4976 | -59.1843 | 63.0936 | -0.15343 | 0.0019824 | 0.067691 | 0.068919 | -0.00045509 | 0.0020892 | 0.00047001 | -0.037933 | -0.00068151 | 0.063211 | 0.06661 | 3.2078e-05 | -0.0024075 | 0.0034329 | -0.036621 | 63.0432 | -59.2804 | 62.9883 | -59.0195 | -59.3994 | -59.3124 | -0.013733 | 63.0524 | -59.1797 | 63.1348 | -59.1019 | -59.23 | -58.9645 | -0.096131 | -0.050354 | -0.10529 | -0.054932 | 0.1648 | 0.26093 | -0.12817 | -0.21515 | -0.032043 | -0.11902 | 0.004578 | -0.041199 | 0.041199 | 0.082398 | 0.082397 | 0.077819 | 0.21973 | -0.045777 | 0.21515 | -0.050355
509.7812 | 122.4335 | -59.1431 | 63.1348 | -0.046616 | 0.0017688 | 0.068545 | 0.067424 | -0.0025913 | -0.00068783 | -0.0010254 | -0.047089 | -0.00068151 | 0.063211 | 0.062764 | 0.00045928 | -0.0028347 | 0.0019375 | 0.022888 | 63.0753 | -59.2758 | 63.0432 | -58.9828 | -59.3674 | -59.3216 | 0.059509 | 63.0661 | -59.166 | 63.1027 | -59.1385 | -59.2667 | -59.0286 | -0.13275 | -0.05951 | -0.091553 | -0.032043 | 0.16022 | 0.29297 | -0.17853 | -0.22431 | -0.045776 | -0.091553 | -0.022889 | -0.068665 | -0.032044 | 0.036621 | 0.004577 | 0.027466 | 0.11444 | -0.1236 | 0.13733 | -0.10071
510.7812 | 122.5159 | -59.1614 | 63.0798 | -0.1458 | -0.00079462 | 0.068973 | 0.067424 | 0.0031765 | 0.00038017 | 0.00047001 | -0.05014 | -0.00025411 | 0.064493 | 0.065114 | -0.0021041 | -0.0015529 | -0.0012667 | -0.054932 | 63.0295 | -59.317 | 63.0341 | -58.992 | -59.3994 | -59.3262 | 0.004578 | 63.0386 | -59.198 | 63.1485 | -59.1339 | -59.2529 | -58.9691 | -0.15564 | -0.050354 | -0.045776 | 0.004578 | 0.16937 | 0.32501 | -0.16479 | -0.23804 | -0.009155 | -0.082397 | -0.036621 | -0.041199 | 0.068665 | 0.10986 | 0.027466 | 0.064087 | 0.19226 | -0.091553 | 0.22888 | -0.054932
511.7812 | 122.4609 | -59.1431 | 63.1165 | -0.077133 | 0.001128 | 0.066409 | 0.068705 | 0.00018591 | 0.00016657 | -0.0035888 | -0.010467 | 0.00081389 | 0.062997 | 0.064687 | -0.00082252 | 0.00036966 | 0.00044212 | 0.022888 | 63.0753 | -59.2712 | 63.0615 | -58.9508 | -59.3536 | -59.2941 | 0.086975 | 63.0936 | -59.1522 | 63.1531 | -59.0973 | -59.2072 | -58.9828 | -0.12817 | -0.041199 | -0.054932 | -0.013733 | 0.19226 | 0.32043 | -0.15106 | -0.21057 | -0.022888 | -0.082398 | -0.009156 | -0.022888 | 0.036621 | 0.059509 | 0.045776 | 0.054932 | 0.16022 | -0.064087 | 0.16937 | -0.054931
512.7812 | 122.4884 | -59.166 | 63.0569 | -0.10002 | 0.002837 | 0.070254 | 0.069987 | -0.0045139 | 0.001662 | -0.00038459 | -0.040985 | 0.0010275 | 0.067056 | 0.062124 | 0.00088648 | -0.00091214 | 0.00022852 | -0.064087 | 63.0341 | -59.2987 | 62.9974 | -59.0012 | -59.4315 | -59.3582 | -0.041199 | 63.0432 | -59.2117 | 63.1165 | -59.1385 | -59.2392 | -58.9874 | -0.13275 | -0.022888 | -0.059509 | -0.036621 | 0.16479 | 0.29755 | -0.19226 | -0.2655 | -0.059509 | -0.13275 | -0.045776 | -0.013733 | 0.059509 | 0.073242 | 0.027466 | 0.073242 | 0.17853 | -0.073242 | 0.2243 | -0.027466
513.7812 | 122.5662 | -59.1476 | 63.0386 | -0.13817 | -0.0018628 | 0.069613 | 0.069346 | -0.0013095 | 0.003371 | 0.003247 | -0.098968 | 0.0010275 | 0.062357 | 0.065114 | -0.00082252 | -0.0019803 | 0.0021511 | -0.054932 | 63.0295 | -59.3216 | 63.0295 | -58.9691 | -59.404 | -59.3445 | -0.022888 | 63.0341 | -59.2117 | 63.1714 | -59.1019 | -59.2392 | -58.9417 | -0.17395 | -0.009155 | -0.009155 | 0 | 0.17853 | 0.35248 | -0.19684 | -0.25635 | -0.022888 | -0.082398 | -0.064087 | -0.004577 | 0.13275 | 0.13733 | 0.045776 | 0.10986 | 0.20599 | -0.091553 | 0.27008 | -0.027466
514.7812 | 122.5845 | -59.1156 | 63.0844 | -0.069504 | 0.001128 | 0.064273 | 0.066569 | -0.0064365 | -4.7032e-05 | 0.00047001 | -0.0013123 | -0.00068151 | 0.065347 | 0.064046 | -0.00082252 | -0.0019803 | -0.0019077 | -0.027466 | 63.0661 | -59.2575 | 63.0478 | -58.9645 | -59.3582 | -59.2712 | 0.07782 | 63.0569 | -59.2026 | 63.1805 | -59.0515 | -59.1568 | -58.9233 | -0.14191 | -0.018311 | -0.036621 | -0.01831 | 0.15106 | 0.29297 | -0.15564 | -0.24261 | -0.013733 | -0.10071 | -0.086975 | -0.027466 | 0.09613 | 0.1236 | 0.064087 | 0.15106 | 0.19226 | -0.041198 | 0.27924 | 0.045777
515.7812 | 122.5204 | -59.166 | 63.0615 | -0.11528 | 0.001128 | 0.06705 | 0.066783 | 0.00018591 | -0.00068783 | -0.00059819 | -0.037933 | 0.0018821 | 0.065561 | 0.062337 | -0.00039512 | -0.0030483 | -0.00062588 | -0.068665 | 63.0203 | -59.3033 | 63.0341 | -58.9783 | -59.404 | -59.3308 | -0.009155 | 63.0157 | -59.198 | 63.1531 | -59.1202 | -59.2392 | -58.9874 | -0.13733 | -0.041198 | -0.027465 | 0.013733 | 0.18768 | 0.32501 | -0.1648 | -0.23804 | -0.027466 | -0.10071 | -0.032043 | -0.045776 | 0.091553 | 0.13733 | 0.045777 | 0.07782 | 0.17853 | -0.073242 | 0.21057 | -0.041199
516.7812 | 122.5388 | -59.1248 | 63.1027 | -0.054245 | -0.005708 | 0.064273 | 0.066142 | 0.0010403 | -4.7032e-05 | 0.00068361 | -0.0074159 | -0.00089511 | 0.063638 | 0.064901 | -0.0023177 | -0.0021939 | 1.4916e-05 | 0.26093 | 63.1485 | -59.2255 | 63.0295 | -58.9279 | -59.3079 | -59.2758 | 0.10986 | 63.0798 | -59.1568 | 63.1805 | -59.0607 | -59.1843 | -58.9279 | -0.10071 | 0.045777 | -0.073242 | -0.11902 | 0.19684 | 0.29755 | -0.15106 | -0.18311 | -0.050354 | -0.082397 | -0.032043 | -0.022888 | 0.07782 | 0.10071 | 0.064087 | 0.09613 | 0.19684 | -0.059509 | 0.22888 | -0.027466
518.0469 | 122.5571 | -59.1431 | 63.0524 | -0.084763 | 0.001128 | 0.068545 | 0.067637 | 0.00061311 | 0.0012348 | -0.00081179 | 0.029205 | 0.0012411 | 0.064279 | 0.066182 | -0.00018152 | -0.00091214 | 0.00044212 | -0.032043 | 63.0157 | -59.3353 | 63.0707 | -59.1293 | -59.5963 | -59.1843 | -0.036621 | 63.0295 | -59.1888 | 63.1805 | -59.1202 | -59.2255 | -58.9645 | -0.19226 | -0.036621 | 0.018311 | 0.054932 | 0.013733 | 0.20599 | -0.041199 | -0.45319 | 0.15106 | -0.26093 | -0.045777 | -0.022888 | 0.12817 | 0.15106 | 0.022888 | 0.068665 | 0.17853 | -0.082398 | 0.2243 | -0.036621
519.2344 | 122.5113 | -59.1339 | 63.0798 | -0.10002 | -0.00036742 | 0.067691 | 0.068919 | -0.0028049 | 0.00080757 | 0.0011108 | -0.0074159 | 0.00081389 | 0.063852 | 0.05721 | -0.00060892 | -0.0036893 | 0.0023647 | -0.045776 | 63.0798 | -59.2941 | 63.0249 | -59.0012 | -59.3903 | -59.3262 | 0.013733 | 63.0569 | -59.1934 | 63.1531 | -59.0836 | -59.2117 | -58.9508 | -0.16022 | 0 | -0.054932 | -0.054932 | 0.13275 | 0.29297 | -0.19226 | -0.25635 | -0.032044 | -0.096131 | -0.059509 | -0.022888 | 0.073242 | 0.09613 | 0.050354 | 0.10986 | 0.1831 | -0.07782 | 0.24261 | -0.018311
520.3281 | 122.4518 | -59.1385 | 63.1348 | -0.061874 | -0.00079462 | 0.067264 | 0.065928 | -0.0010959 | -0.00047423 | 0.00068361 | -0.047089 | 0.00060029 | 0.063425 | 0.065755 | -0.00060892 | -0.0032619 | 0.00065592 | 0.009155 | 63.0569 | -59.2941 | 63.0203 | -58.9828 | -59.3674 | -59.3262 | 0.07782 | 63.1073 | -59.1843 | 63.1165 | -59.1431 | -59.2529 | -58.992 | -0.15564 | -0.07782 | -0.11444 | -0.036621 | 0.15564 | 0.31128 | -0.18768 | -0.22888 | -0.032044 | -0.073243 | -0.045776 | -0.027466 | -0.018311 | 0.009155 | -0.004577 | 0.041199 | 0.14648 | -0.11444 | 0.19226 | -0.068665
521.7656 | 122.4564 | -59.1431 | 63.1165 | -0.061874 | 0.00048698 | 0.068332 | 0.068919 | 0.0018949 | 0.00016657 | 0.00047001 | -0.12949 | 0.0014547 | 0.064706 | 0.064901 | -0.00060892 | -0.0039029 | -0.0025485 | 0.004578 | 63.1027 | -59.3033 | 62.9883 | -59.0378 | -59.4177 | -59.3765 | 0.032043 | 63.1073 | -59.1751 | 63.1439 | -59.1293 | -59.2758 | -58.9966 | -0.16022 | -0.013733 | -0.12817 | -0.11444 | 0.10528 | 0.2655 | -0.23346 | -0.27466 | -0.073242 | -0.11444 | -0.032044 | -0.009155 | 0.027466 | 0.036621 | 0.013733 | 0.045777 | 0.14648 | -0.13275 | 0.17853 | -0.10071
522.3281 | 122.4518 | -59.166 | 63.1348 | -0.046616 | 0.00048698 | 0.068973 | 0.065928 | 0.0010403 | 0.0012348 | 0.001538 | -0.044037 | 0.0016685 | 0.064493 | 0.05956 | -0.0012497 | -0.0039029 | 0.0017239 | 0 | 63.1165 | -59.2758 | 62.9745 | -59.0561 | -59.436 | -59.3765 | 0 | 63.1027 | -59.2072 | 63.0707 | -59.1568 | -59.2758 | -59.0378 | -0.10986 | -0.018311 | -0.16022 | -0.14191 | 0.10986 | 0.21973 | -0.21057 | -0.27008 | -0.10071 | -0.16022 | -0.041198 | -0.032044 | -0.064087 | -0.032043 | 0.009156 | 0.050354 | 0.12817 | -0.10986 | 0.16937 | -0.068665
524.4531 | 122.3831 | -59.198 | 63.0936 | -0.15343 | 0.00091438 | 0.069186 | 0.070414 | 0.0023221 | 0.00016657 | -0.00038459 | -0.092865 | 0.00081389 | 0.067056 | 0.071523 | -0.00018152 | -0.0028347 | -0.0025485 | -0.064087 | 63.1073 | -59.3308 | 62.9288 | -59.1156 | -59.5139 | -59.4589 | -0.10986 | 63.0981 | -59.2438 | 63.0615 | -59.2575 | -59.3948 | -59.0698 | -0.13275 | 0.013733 | -0.16479 | -0.17853 | 0.082397 | 0.21515 | -0.26093 | -0.31586 | -0.12817 | -0.1831 | -0.045776 | 0.004578 | -0.032044 | -0.036622 | -0.059509 | -0.013733 | 0.12817 | -0.19684 | 0.17395 | -0.15106
525.4531 | 122.3877 | -59.1476 | 63.1714 | -0.046616 | -0.00015382 | 0.067691 | 0.068492 | -0.0032321 | 0.0012348 | -0.00038459 | -0.019623 | 0.0016685 | 0.065988 | 0.062124 | 0.0015275 | -0.0030483 | -0.0021213 | -0.041199 | 63.121 | -59.2758 | 62.9654 | -59.0378 | -59.4315 | -59.3857 | 0.004578 | 63.121 | -59.1843 | 63.0753 | -59.1431 | -59.2987 | -59.0149 | -0.12817 | -0.050354 | -0.20599 | -0.15564 | 0.10986 | 0.23804 | -0.23804 | -0.28381 | -0.10986 | -0.15564 | -0.036621 | -0.050354 | -0.096131 | -0.045777 | 0.004578 | 0.041199 | 0.13275 | -0.15106 | 0.16937 | -0.11444
526.4844 | 122.5021 | -59.1385 | 63.1256 | -0.077133 | -0.0010082 | 0.068545 | 0.066142 | -0.0040867 | 0.0018756 | -0.00038459 | 0.0078429 | 0.0010275 | 0.063852 | 0.057851 | -0.00060892 | -0.0026211 | 0.00065592 | 0.009155 | 63.089 | -59.2941 | 63.0478 | -58.96 | -59.3536 | -59.3033 | 0.07782 | 63.0936 | -59.166 | 63.1302 | -59.1293 | -59.2209 | -58.9828 | -0.15564 | -0.036621 | -0.077819 | -0.041198 | 0.17853 | 0.33417 | -0.16479 | -0.21515 | -0.009156 | -0.05951 | -0.027466 | -0.032043 | 0.004578 | 0.036621 | 0.009156 | 0.036622 | 0.15564 | -0.082397 | 0.18311 | -0.054931
530.8281 | 122.6028 | -59.1156 | 63.0615 | -0.11528 | 0.00027338 | 0.067691 | 0.065287 | -0.0019503 | -0.00047423 | -0.0020936 | 0.050568 | 0.00038669 | 0.064493 | 0.05721 | -0.00018152 | -0.0060391 | 0.0017239 | -0.013733 | 63.0844 | -59.2758 | 63.0478 | -58.9508 | -59.3445 | -59.285 | 0.10529 | 63.0249 | -59.1934 | 63.1851 | -59.0424 | -59.1431 | -58.9188 | -0.16022 | 0.022889 | -0.013732 | -0.036621 | 0.16479 | 0.32501 | -0.16937 | -0.22888 | -0.009155 | -0.068664 | -0.077819 | -0.036621 | 0.1236 | 0.16022 | 0.073243 | 0.15106 | 0.19684 | -0.027465 | 0.27466 | 0.050354
531.4062 | 122.5662 | -59.1568 | 63.0661 | -0.069504 | 0.0013416 | 0.06705 | 0.067424 | 0.00039951 | 0.0023028 | 0.00047001 | 0.0017395 | 0.00017309 | 0.062784 | 0.063192 | -0.0012497 | -0.0030483 | -0.00062588 | -0.045776 | 63.0569 | -59.3124 | 63.0112 | -58.9874 | -59.3903 | -59.3262 | 0.032043 | 63.0661 | -59.1843 | 63.1577 | -59.0698 | -59.1934 | -58.9371 | -0.15564 | -0.009155 | -0.054932 | -0.045777 | 0.16937 | 0.32501 | -0.16937 | -0.23346 | -0.013733 | -0.07782 | -0.027466 | 0 | 0.091553 | 0.091553 | 0.086975 | 0.11444 | 0.21973 | -0.036621 | 0.24719 | -0.009155
531.9688 | 122.4518 | -59.1293 | 63.1302 | -0.092392 | 0.0017688 | 0.070254 | 0.067637 | -0.0023777 | 0.0014484 | -0.0023072 | -0.0013123 | 0.0018821 | 0.065561 | 0.066396 | -0.00060892 | -0.0028347 | -0.0036167 | 0.036621 | 63.0753 | -59.2712 | 63.0432 | -58.9828 | -59.3628 | -59.2987 | 0.07782 | 63.0936 | -59.1614 | 63.1073 | -59.1431 | -59.2575 | -58.992 | -0.14191 | -0.054932 | -0.086975 | -0.032043 | 0.14648 | 0.28839 | -0.16937 | -0.23346 | -0.027466 | -0.091553 | -0.032044 | -0.036621 | -0.022888 | 0.013733 | -0.013733 | 0.018311 | 0.13733 | -0.12817 | 0.16937 | -0.09613
532.8125 | 122.4426 | -59.1888 | 63.0981 | -0.069504 | -0.00058102 | 0.068118 | 0.071482 | 0.0014675 | 0.00080757 | 0.00089721 | -0.10812 | 0.0014547 | 0.066629 | 0.066182 | -0.0016769 | -0.00069854 | 0.00086952 | -0.064087 | 63.0707 | -59.3491 | 62.9425 | -59.0927 | -59.5093 | -59.4223 | -0.091553 | 63.0661 | -59.2117 | 63.0707 | -59.2346 | -59.3719 | -59.0744 | -0.16022 | -0.027466 | -0.15564 | -0.12817 | 0.096131 | 0.25635 | -0.23346 | -0.32043 | -0.073242 | -0.16022 | -0.022888 | -0.032044 | -0.027466 | 0.004578 | -0.045776 | -0.022888 | 0.11444 | -0.18311 | 0.13733 | -0.16022
533.7812 | 122.5571 | -59.1431 | 63.0386 | -0.13817 | -0.00036742 | 0.068332 | 0.068064 | 0.0044583 | 0.0027302 | 0.0013244 | -0.047089 | -0.00089511 | 0.064706 | 0.063192 | -0.00039512 | 0.00015586 | -0.0036167 | -0.050354 | 63.0524 | -59.3079 | 63.0203 | -59.0012 | -59.404 | -59.3445 | -0.018311 | 63.0203 | -59.198 | 63.1622 | -59.0973 | -59.2438 | -58.9325 | -0.1648 | 0.013733 | -0.01831 | -0.032043 | 0.14191 | 0.3067 | -0.20142 | -0.26093 | -0.036621 | -0.096131 | -0.054932 | -0.01831 | 0.1236 | 0.14191 | 0.045776 | 0.10071 | 0.21057 | -0.10071 | 0.2655 | -0.045776
534.7812 | 122.5662 | -59.1156 | 63.0661 | -0.092392 | 0.00027338 | 0.065341 | 0.065501 | -0.0023777 | 0.0018756 | 0.0011108 | 0.041412 | 0.00038669 | 0.063852 | 0.066396 | -0.00060892 | -0.0049709 | -0.0016941 | 0.036621 | 63.0615 | -59.2804 | 63.0844 | -58.8776 | -59.3079 | -59.2667 | 0.13733 | 63.0707 | -59.1797 | 63.2584 | -58.8089 | -59.404 | -59.1064 | -0.1648 | -0.004578 | 0.018311 | 0.022889 | 0.23804 | 0.40283 | -0.15106 | -0.19226 | 0.013733 | -0.027465 | -0.064086 | 0.004578 | 0.19226 | 0.18768 | 0.3067 | 0.37079 | 0.009156 | -0.28839 | 0.073242 | -0.22431
535.7812 | 122.493 | -59.1248 | 63.0844 | -0.12291 | 0.00091438 | 0.068118 | 0.068278 | -0.0034457 | 0.00016657 | -0.00081179 | 0.02005 | -0.00089511 | 0.065347 | 0.065328 | 0.00024568 | -0.0019803 | 0.00044212 | -0.073242 | 63.0386 | -59.2896 | 62.9883 | -59.0195 | -59.4131 | -59.3491 | -0.036621 | 63.0478 | -59.1888 | 63.1256 | -59.1293 | -59.2621 | -58.9783 | -0.1648 | -0.045777 | -0.096131 | -0.050354 | 0.10529 | 0.27008 | -0.2243 | -0.28839 | -0.059509 | -0.1236 | -0.064087 | -0.036621 | 0.041198 | 0.077819 | -0.004577 | 0.05951 | 0.14648 | -0.13733 | 0.21057 | -0.073242
536.7812 | 122.4838 | -59.1568 | 63.0798 | -0.11528 | 0.00027338 | 0.069613 | 0.06956 | 0.00061311 | -0.00068783 | -0.001239 | -0.019623 | 0.00017309 | 0.061929 | 0.063192 | -0.0012497 | -0.0030483 | -0.00083948 | -0.059509 | 63.0432 | -59.2758 | 63.002 | -58.992 | -59.4315 | -59.3399 | -0.009155 | 63.0203 | -59.1888 | 63.1256 | -59.1156 | -59.2346 | -58.992 | -0.11902 | -0.036621 | -0.07782 | -0.041199 | 0.16479 | 0.28381 | -0.18311 | -0.27466 | -0.064087 | -0.15564 | -0.032044 | -0.059509 | 0.045776 | 0.10529 | 0.041198 | 0.073242 | 0.16479 | -0.07782 | 0.19684 | -0.045776
537.7812 | 122.5388 | -59.1293 | 63.1073 | -0.16106 | -0.00015382 | 0.067264 | 0.068492 | 0.0048855 | -0.00090143 | 0.0019654 | 0.0078429 | 0.00060029 | 0.063638 | 0.064046 | -0.00060892 | -0.0028347 | 0.00022852 | 0.027466 | 63.0569 | -59.2758 | 63.0798 | -58.9279 | -59.3216 | -59.2804 | 0.10529 | 63.0615 | -59.1797 | 63.1668 | -59.0698 | -59.2026 | -58.9279 | -0.14648 | -0.050354 | -0.027466 | 0.022888 | 0.20142 | 0.3479 | -0.15106 | -0.19226 | -0.004578 | -0.045776 | -0.050354 | -0.045777 | 0.059509 | 0.10529 | 0.059509 | 0.10986 | 0.20142 | -0.073243 | 0.25177 | -0.022889
538.7812 | 122.4838 | -59.1476 | 63.0615 | -0.084763 | 5.9776e-05 | 0.070468 | 0.065715 | -0.00088229 | 0.00038017 | 0.00068361 | -0.040985 | 0.0025229 | 0.063638 | 0.06661 | 0.00067288 | -0.0030483 | 0.0021511 | -0.050354 | 63.0432 | -59.317 | 63.0112 | -58.9828 | -59.404 | -59.3308 | -0.004578 | 63.0341 | -59.198 | 63.1393 | -59.1431 | -59.2392 | -58.9783 | -0.16937 | -0.01831 | -0.050354 | -0.032044 | 0.16479 | 0.33417 | -0.18311 | -0.25635 | -0.013733 | -0.086975 | -0.050354 | -0.027465 | 0.07782 | 0.10528 | 0.004578 | 0.054932 | 0.16937 | -0.091553 | 0.21973 | -0.041199
539.7812 | 122.5296 | -59.1431 | 63.0661 | -0.10765 | -0.00015382 | 0.06705 | 0.067424 | 0.0023221 | 0.00038017 | -0.0016662 | 0.02005 | 0.00060029 | 0.065134 | 0.061269 | 0.00045928 | -0.0036893 | 0.0015103 | -0.064087 | 63.0707 | -59.2712 | 63.0157 | -59.0012 | -59.4543 | -59.3536 | -0.032043 | 63.0386 | -59.2026 | 63.1165 | -59.1202 | -59.2163 | -58.9462 | -0.12817 | 0.004578 | -0.050354 | -0.054932 | 0.14191 | 0.27008 | -0.21057 | -0.31128 | -0.082398 | -0.18311 | -0.05951 | -0.027466 | 0.050354 | 0.07782 | 0.022888 | 0.082398 | 0.19684 | -0.073243 | 0.25635 | -0.013733
540.7812 | 122.5479 | -59.1156 | 63.0341 | -0.16869 | 0.001128 | 0.068759 | 0.06956 | -0.00066869 | -4.7032e-05 | -0.00081179 | -0.062347 | 0.0012411 | 0.063852 | 0.070241 | -0.00060892 | -0.0021939 | 0.00086952 | -0.045776 | 63.0203 | -59.2941 | 63.0569 | -58.9874 | -59.3994 | -59.3262 | -0.21057 | 62.9929 | -59.23 | 63.1805 | -59.0836 | -59.2392 | -58.9462 | -0.17853 | -0.013733 | 0.022888 | 0.036621 | 0.12817 | 0.3067 | -0.21057 | -0.28381 | -0.032044 | -0.10529 | -0.11444 | -0.041199 | 0.14648 | 0.18768 | 0.032044 | 0.14648 | 0.16937 | -0.1236 | 0.28381 | -0.009155
541.7812 | 122.58 | -59.1339 | 63.0615 | -0.12291 | -0.00036742 | 0.069827 | 0.069133 | -0.0017367 | 0.0012348 | 0.001538 | -0.047089 | -4.0508e-05 | 0.064493 | 0.067251 | 0.0013139 | -0.0043301 | -0.0025485 | -0.036621 | 63.0386 | -59.2987 | 63.0432 | -58.9737 | -59.3903 | -59.3399 | -0.009155 | 63.0112 | -59.1843 | 63.295 | -59.1339 | -59.2575 | -58.9645 | -0.1648 | -0.022888 | -0.01831 | 0.004578 | 0.16022 | 0.32501 | -0.20599 | -0.25635 | -0.041199 | -0.091553 | -0.050354 | -0.050354 | 0.23346 | 0.28381 | 0 | 0.050354 | 0.16937 | -0.1236 | 0.21973 | -0.073242
542.7812 | 122.5662 | -59.1339 | 63.0844 | -0.10765 | -0.00079462 | 0.065341 | 0.066142 | 0.00061311 | 0.0027302 | 4.2808e-05 | 0.026154 | -0.0015359 | 0.063211 | 0.06426 | -0.0010361 | -0.0036893 | -0.0021213 | -0.059509 | 63.0524 | -59.285 | 63.0249 | -58.9828 | -59.3903 | -59.3262 | 0.018311 | 63.0341 | -59.1888 | 63.1714 | -59.0881 | -59.2072 | -58.9417 | -0.15106 | -0.032044 | -0.05951 | -0.027466 | 0.15106 | 0.30212 | -0.19226 | -0.25635 | -0.041199 | -0.10529 | -0.054932 | -0.050354 | 0.086975 | 0.13733 | 0.045776 | 0.10071 | 0.19226 | -0.073242 | 0.24719 | -0.01831
543.7812 | 122.4976 | -59.1476 | 63.0753 | -0.069504 | 0.00027338 | 0.068759 | 0.068064 | 0.0023221 | 0.0035846 | -0.0014526 | 0.032257 | 0.00017309 | 0.062143 | 0.070241 | -0.0010361 | -0.00069854 | -0.0016941 | -0.059509 | 63.0707 | -59.3079 | 62.9837 | -59.0195 | -59.4269 | -59.3445 | -0.073242 | 63.0341 | -59.2026 | 63.1256 | -59.1156 | -59.2438 | -58.9828 | -0.16022 | -0.004577 | -0.091552 | -0.086975 | 0.12817 | 0.28839 | -0.19684 | -0.27924 | -0.036621 | -0.11902 | -0.054932 | -0.041198 | 0.050354 | 0.091552 | 0.032043 | 0.086975 | 0.16479 | -0.09613 | 0.21973 | -0.041198
544.7812 | 122.5113 | -59.1431 | 63.0707 | -0.092392 | -0.0010082 | 0.069186 | 0.066355 | -0.0034457 | 0.00059397 | -0.00038459 | 0.010895 | -0.0013223 | 0.065561 | 0.065755 | -0.0040267 | -0.0026211 | -0.0023349 | -0.068665 | 63.0478 | -59.2987 | 63.0066 | -59.0057 | -59.4131 | -59.3399 | -0.032043 | 63.0524 | -59.1888 | 63.0981 | -59.1339 | -59.2163 | -59.0195 | -0.15564 | -0.022888 | -0.064087 | -0.041199 | 0.13733 | 0.29297 | -0.19684 | -0.27008 | -0.041199 | -0.11444 | -0.045777 | -0.018311 | 0.027466 | 0.045777 | 0.009155 | 0.054932 | 0.1236 | -0.073243 | 0.16937 | -0.027466
545.7812 | 122.493 | -59.1476 | 63.0661 | -0.1458 | 0.0013416 | 0.068759 | 0.068064 | -0.00066869 | 0.00059397 | 0.00025641 | -0.0074159 | 0.0010275 | 0.065775 | 0.068746 | 3.2078e-05 | -0.0024075 | 0.0025785 | -0.07782 | 63.0432 | -59.2987 | 63.002 | -59.0012 | -59.4315 | -59.3308 | -0.045776 | 63.0432 | -59.2255 | 63.1256 | -59.1431 | -59.2621 | -58.9737 | -0.15106 | -0.022888 | -0.064087 | -0.041199 | 0.14648 | 0.29755 | -0.18311 | -0.28381 | -0.032044 | -0.13275 | -0.07782 | -0.022888 | 0.059509 | 0.082397 | 0.004578 | 0.082398 | 0.17395 | -0.11444 | 0.25177 | -0.036621
546.7812 | 122.5342 | -59.1293 | 63.0661 | -0.12291 | 0.00048698 | 0.067904 | 0.068705 | -0.0038729 | 0.001662 | 0.0013244 | -0.025726 | 0.0010275 | 0.065134 | 0.070241 | -0.0014633 | -0.0015529 | -0.00041228 | -0.07782 | 63.0569 | -59.2987 | 63.002 | -59.0195 | -59.4223 | -59.3399 | -0.036621 | 63.0569 | -59.2026 | 63.1393 | -59.1385 | -59.166 | -58.8272 | -0.16937 | -0.009155 | -0.064087 | -0.054932 | 0.10986 | 0.27924 | -0.21057 | -0.29297 | -0.041199 | -0.1236 | -0.073243 | -0.009155 | 0.073242 | 0.082397 | -0.009156 | 0.064087 | 0.30212 | -0.036622 | 0.37537 | 0.036621
547.7812 | 122.6257 | -59.1339 | 63.0798 | -0.061874 | -0.0020764 | 0.067264 | 0.065928 | -0.00024129 | -0.00090143 | -0.0020936 | -0.0043641 | 0.00038669 | 0.064706 | 0.065542 | -0.0010361 | -0.0079617 | -0.0012667 | 0.032043 | 63.0936 | -59.2529 | 63.089 | -58.9005 | -59.3079 | -59.2529 | 0.13733 | 63.0707 | -59.1614 | 63.3911 | -59.2575 | -59.3628 | -59.0332 | -0.11902 | 0.013733 | 0.009155 | -0.004578 | 0.23346 | 0.35248 | -0.11902 | -0.17395 | 0 | -0.054931 | -0.027466 | -0.009155 | 0.31128 | 0.32043 | -0.1236 | -0.09613 | 0.10071 | -0.22888 | 0.12817 | -0.20142
548.7812 | 122.5296 | -59.1522 | 63.0707 | -0.13054 | -0.00036742 | 0.06705 | 0.069133 | -0.00088229 | -0.00026063 | -0.00017099 | -0.047089 | 0.0012411 | 0.062997 | 0.065542 | 0.00088648 | -0.0034757 | 0.00044212 | -0.068665 | 63.0341 | -59.3124 | 63.0249 | -58.9828 | -59.4131 | -59.3399 | 0.032043 | 63.0524 | -59.198 | 63.1531 | -59.1339 | -59.2621 | -58.9828 | -0.16022 | -0.036621 | -0.045777 | -0.009156 | 0.16937 | 0.32959 | -0.18768 | -0.26093 | -0.027466 | -0.10071 | -0.045776 | -0.018311 | 0.082397 | 0.10071 | 0.018311 | 0.064087 | 0.16937 | -0.10986 | 0.21515 | -0.064087
549.7812 | 122.4792 | -59.1339 | 63.0524 | -0.054245 | 5.9776e-05 | 0.068545 | 0.065074 | 0.0021085 | 0.001662 | -0.0010254 | -0.03183 | 0.00060029 | 0.064279 | 0.068319 | -0.0023177 | -0.0036893 | 0.00086952 | -0.064087 | 63.0524 | -59.2896 | 63.0066 | -58.9966 | -59.4543 | -59.3582 | -0.022888 | 63.0341 | -59.2163 | 63.1485 | -59.1064 | -59.2392 | -58.9645 | -0.15564 | 0 | -0.045776 | -0.045776 | 0.13733 | 0.29297 | -0.2243 | -0.32044 | -0.068664 | -0.16479 | -0.082398 | -0.01831 | 0.096131 | 0.11444 | 0.027466 | 0.10986 | 0.16937 | -0.10529 | 0.25177 | -0.022888
550.7812 | 122.5159 | -59.1339 | 63.0844 | -0.054245 | -0.001222 | 0.068118 | 0.068919 | -0.0062229 | 0.001662 | -0.0016662 | -0.037933 | -0.0011087 | 0.066843 | 0.068105 | -0.00082252 | -0.00048494 | -0.0023349 | -0.091553 | 63.0569 | -59.3308 | 63.002 | -59.0103 | -59.4315 | -59.3536 | -0.059509 | 63.0386 | -59.2163 | 63.121 | -59.1202 | -59.2529 | -58.9874 | -0.19684 | -0.027466 | -0.082398 | -0.054932 | 0.1236 | 0.32044 | -0.21973 | -0.29755 | -0.022888 | -0.10071 | -0.082398 | -0.045777 | 0.036621 | 0.082398 | 0.013733 | 0.096131 | 0.14648 | -0.11902 | 0.22888 | -0.036621
551.7812 | 122.5021 | -59.1431 | 63.0524 | -0.10765 | -0.00229 | 0.065982 | 0.070414 | -0.00045509 | 0.0020892 | 0.0013244 | 0.010895 | -0.00089511 | 0.06022 | 0.064473 | 0.0011001 | -0.0030483 | -0.00083948 | -0.073242 | 63.0661 | -59.2987 | 63.0112 | -59.0195 | -59.4269 | -59.3353 | 0.050354 | 63.0569 | -59.166 | 63.1027 | -59.0927 | -59.2575 | -58.9691 | -0.15564 | 0.013733 | -0.041199 | -0.054932 | 0.1236 | 0.27924 | -0.19226 | -0.28381 | -0.036621 | -0.12817 | -0.022889 | 0.004578 | 0.050354 | 0.045776 | 0.050354 | 0.073243 | 0.17395 | -0.11444 | 0.19684 | -0.091552
552.7812 | 122.5571 | -59.1156 | 63.0798 | -0.031357 | 0.00048698 | 0.067904 | 0.068278 | -2.7692e-05 | 0.0012348 | -0.0014526 | -0.016571 | -0.0015359 | 0.063852 | 0.065328 | -0.00082252 | -0.0045437 | 1.4916e-05 | -0.009155 | 63.0661 | -59.2575 | 63.089 | -58.9371 | -59.3262 | -59.2667 | 0.10986 | 63.0432 | -59.1522 | 63.2034 | -59.0149 | -59.1522 | -58.8959 | -0.14191 | -0.013733 | 0.009155 | 0.022888 | 0.17853 | 0.32043 | -0.15106 | -0.21057 | -0.009156 | -0.068665 | -0.036621 | -0.036621 | 0.1236 | 0.16022 | 0.10071 | 0.13733 | 0.21973 | -0.036621 | 0.25635 | 0
553.7812 | 122.4701 | -59.1339 | 63.0753 | -0.10765 | -0.00058102 | 0.06705 | 0.068705 | -0.00088229 | 0.0023028 | 0.00047001 | -0.053192 | -0.00025411 | 0.064493 | 0.069387 | -0.00060892 | -0.0017667 | 0.0034329 | -0.064087 | 63.0112 | -59.3033 | 63.0295 | -58.9874 | -59.4223 | -59.3445 | -0.013733 | 63.0249 | -59.2209 | 63.1302 | -59.1385 | -59.2621 | -58.992 | -0.16937 | -0.064087 | -0.045776 | 0.018311 | 0.14648 | 0.31586 | -0.21057 | -0.28839 | -0.041198 | -0.11902 | -0.086975 | -0.050354 | 0.054932 | 0.10529 | -0.004578 | 0.082397 | 0.14191 | -0.12817 | 0.22888 | -0.041199
554.7812 | 122.5204 | -59.166 | 63.089 | -0.077133 | -0.00015382 | 0.06705 | 0.069346 | -0.0013095 | 0.0025166 | -0.0016662 | -0.019623 | -0.0021769 | 0.063425 | 0.05956 | 0.0013139 | -0.0026211 | 0.00065592 | 0.068665 | 63.0936 | -59.2896 | 62.9929 | -58.9874 | -59.436 | -59.3353 | -0.027466 | 63.0432 | -59.2163 | 63.1485 | -59.0881 | -59.2117 | -58.9645 | -0.1236 | 0.004578 | -0.09613 | -0.10071 | 0.17853 | 0.30212 | -0.16937 | -0.27008 | -0.045776 | -0.14648 | -0.050354 | -0.045776 | 0.05951 | 0.10529 | 0.07782 | 0.12817 | 0.20142 | -0.045776 | 0.25177 | 0.004578
555.7812 | 122.4838 | -59.1431 | 63.0661 | -0.11528 | 0.00091438 | 0.068759 | 0.067637 | -0.00066869 | -0.0013288 | 0.0017518 | -0.040985 | 0.00038669 | 0.064066 | 0.068532 | -0.0012497 | -0.0030483 | 0.00065592 | -0.082397 | 63.0478 | -59.3033 | 62.9974 | -59.0332 | -59.4452 | -59.3582 | -0.050354 | 63.0249 | -59.2163 | 63.1439 | -59.1248 | -59.2575 | -58.9828 | -0.16022 | -0.01831 | -0.068664 | -0.050354 | 0.10986 | 0.27008 | -0.21515 | -0.30212 | -0.054931 | -0.14191 | -0.073243 | -0.041199 | 0.07782 | 0.11902 | 0.01831 | 0.091553 | 0.16022 | -0.11444 | 0.23346 | -0.041198
556.7812 | 122.5113 | -59.1293 | 63.0798 | -0.046616 | 0.00091438 | 0.065555 | 0.067637 | -0.0025913 | 0.00080757 | 0.00068361 | 0.0047913 | 0.00017309 | 0.064066 | 0.062551 | -0.00082252 | -0.0017667 | -0.0014805 | -0.045776 | 63.0478 | -59.2896 | 63.0249 | -58.992 | -59.3903 | -59.3124 | 0.018311 | 63.0432 | -59.166 | 63.1805 | -59.0698 | -59.1888 | -58.9508 | -0.16022 | -0.032043 | -0.054932 | -0.022889 | 0.13733 | 0.29755 | -0.18311 | -0.26093 | -0.022888 | -0.10071 | -0.036622 | -0.036621 | 0.10071 | 0.13733 | 0.059509 | 0.096131 | 0.17853 | -0.05951 | 0.21515 | -0.022888
557.7812 | 122.5571 | -59.1522 | 63.0981 | -0.092392 | -0.00079462 | 0.068332 | 0.066783 | -0.0015231 | -0.0015424 | -0.0014526 | -0.0074159 | 0.00038669 | 0.063638 | 0.068532 | 0.00024568 | -0.0017667 | 0.0017239 | -0.013733 | 63.0707 | -59.2896 | 63.0707 | -58.9371 | -59.3308 | -59.2758 | 0.11444 | 63.0661 | -59.1843 | 63.2126 | -59.0378 | -59.1476 | -58.9142 | -0.13733 | -0.027466 | -0.027466 | 0 | 0.21515 | 0.35248 | -0.1236 | -0.17853 | 0.013733 | -0.041199 | -0.032043 | -0.032044 | 0.11444 | 0.14648 | 0.11444 | 0.14648 | 0.23804 | 0.004578 | 0.27008 | 0.036621
558.7812 | 122.5021 | -59.1522 | 63.0661 | -0.11528 | -0.00058102 | 0.068973 | 0.067851 | -0.00066869 | 0.0023028 | -0.00081179 | -0.059296 | -0.0017495 | 0.065561 | 0.073659 | 0.0011001 | -0.0021939 | 0.00086952 | -0.050354 | 63.0386 | -59.2896 | 63.0295 | -58.992 | -59.4269 | -59.3353 | 0 | 63.0295 | -59.2026 | 63.1302 | -59.1293 | -59.2575 | -58.9828 | -0.13733 | -0.027466 | -0.036621 | -0.009155 | 0.16022 | 0.29755 | -0.1831 | -0.27466 | -0.045776 | -0.13733 | -0.050354 | -0.036621 | 0.064087 | 0.10071 | 0.022889 | 0.073243 | 0.16937 | -0.10528 | 0.21973 | -0.054931
559.7812 | 122.5342 | -59.1339 | 63.1073 | -0.061874 | 0.00091438 | 0.069827 | 0.068064 | 0.00039951 | -0.00090143 | -0.0010254 | -0.010467 | -0.0011087 | 0.065134 | 0.068319 | 0.00045928 | -0.0024075 | -0.0010531 | -0.073242 | 63.0707 | -59.2804 | 62.9929 | -59.0057 | -59.3994 | -59.3491 | -0.004578 | 63.0249 | -59.1888 | 63.1622 | -59.0973 | -59.2255 | -58.9554 | -0.14649 | -0.036621 | -0.11444 | -0.07782 | 0.12817 | 0.27466 | -0.21515 | -0.2655 | -0.068664 | -0.11902 | -0.054932 | -0.082398 | 0.054931 | 0.13733 | 0.036621 | 0.091553 | 0.17853 | -0.091553 | 0.23346 | -0.036621
560.7812 | 122.5525 | -59.1385 | 63.0386 | -0.10765 | -0.00015382 | 0.068759 | 0.068492 | 0.0012539 | 0.0012348 | 0.001538 | -0.08371 | 0.0027365 | 0.065347 | 0.071523 | -0.0048813 | -0.00069854 | 0.00086952 | -0.041199 | 63.0295 | -59.3262 | 63.0478 | -58.9691 | -59.3994 | -59.3124 | -0.12817 | 63.002 | -59.2209 | 63.208 | -59.0927 | -59.2117 | -58.9142 | -0.18768 | -0.009155 | 0.009156 | 0.018311 | 0.16937 | 0.35706 | -0.17395 | -0.26093 | 0.013733 | -0.073242 | -0.082397 | -0.036621 | 0.16937 | 0.20599 | 0.045777 | 0.12817 | 0.2243 | -0.073242 | 0.3067 | 0.009155
561.7812 | 122.5388 | -59.1385 | 63.0249 | -0.16106 | -0.0010082 | 0.070041 | 0.069133 | 0.0036039 | 0.0014484 | 0.001538 | -0.092865 | 0.00081389 | 0.062784 | 0.065328 | 0.0011001 | -0.0017667 | 0.00086952 | -0.027466 | 63.0249 | -59.3124 | 63.0295 | -58.9737 | -59.3994 | -59.3308 | 0.004578 | 63.0249 | -59.2117 | 63.2034 | -58.9645 | -59.2392 | -58.9966 | -0.17395 | 0 | 0.004578 | 0.004578 | 0.16479 | 0.33874 | -0.19226 | -0.26093 | -0.018311 | -0.086975 | -0.073242 | 0 | 0.17853 | 0.17853 | 0.17395 | 0.24719 | 0.14191 | -0.10071 | 0.21515 | -0.027466
562.7812 | 122.4747 | -59.1522 | 63.0661 | -0.11528 | 0.00027338 | 0.069613 | 0.067424 | -0.0023777 | 0.00080757 | 0.0013244 | -0.022675 | -0.00046791 | 0.064279 | 0.066823 | -0.00082252 | -0.0026211 | -0.00041228 | -0.10071 | 63.0386 | -59.2987 | 63.0066 | -59.0057 | -59.4452 | -59.3445 | -0.045776 | 63.0432 | -59.2072 | 63.1348 | -59.1385 | -59.2346 | -58.9645 | -0.14648 | -0.027466 | -0.059509 | -0.032043 | 0.14648 | 0.29297 | -0.19226 | -0.29297 | -0.045776 | -0.14648 | -0.054931 | -0.022888 | 0.068665 | 0.091553 | 0.013733 | 0.068664 | 0.18768 | -0.082397 | 0.24261 | -0.027466
563.7812 | 122.5159 | -59.1385 | 63.0798 | -0.10765 | 0.00048698 | 0.068973 | 0.066569 | -0.0051547 | 0.0025166 | 0.00047001 | -0.028778 | -0.0013223 | 0.065561 | 0.068105 | -0.0027451 | -0.0015529 | -0.0019077 | -0.064087 | 63.0798 | -59.2804 | 62.9929 | -59.0149 | -59.4223 | -59.3308 | -0.027466 | 63.0478 | -59.1934 | 63.1714 | -59.1431 | -59.2072 | -58.9645 | -0.14191 | 0 | -0.086975 | -0.086975 | 0.1236 | 0.2655 | -0.19226 | -0.28381 | -0.050354 | -0.14191 | -0.054931 | -0.032043 | 0.091553 | 0.1236 | -0.004577 | 0.050354 | 0.17395 | -0.068664 | 0.22888 | -0.013733
564.7812 | 122.5754 | -59.1156 | 63.0844 | -0.092392 | 0.00070078 | 0.066836 | 0.069773 | -0.0040867 | -4.7032e-05 | -0.0025208 | -0.016571 | 0.00060029 | 0.067056 | 0.067037 | -0.0010361 | -0.0015529 | 1.4916e-05 | -0.009155 | 63.0661 | -59.2667 | 63.0753 | -58.9279 | -59.3124 | -59.285 | 0.11902 | 63.0569 | -59.1614 | 63.2034 | -59.0286 | -59.1431 | -58.9142 | -0.15106 | -0.018311 | -0.009156 | 0.009155 | 0.18768 | 0.33875 | -0.16937 | -0.19684 | -0.01831 | -0.045776 | -0.045776 | -0.027466 | 0.11902 | 0.14648 | 0.086976 | 0.13275 | 0.20142 | -0.027465 | 0.24719 | 0.018311
565.7812 | 122.5067 | -59.1522 | 63.0615 | -0.069504 | 0.0015552 | 0.0694 | 0.067637 | 0.00061311 | 0.001662 | -0.00081179 | -0.062347 | 0.00017309 | 0.065134 | 0.06426 | -0.0021041 | -0.0024075 | 0.00065592 | -0.064087 | 63.0386 | -59.3033 | 63.0249 | -58.9874 | -59.4315 | -59.3399 | -0.022888 | 63.0249 | -59.1751 | 63.1439 | -59.1156 | -59.2346 | -58.9783 | -0.15106 | -0.022888 | -0.036621 | -0.013733 | 0.1648 | 0.31586 | -0.18768 | -0.27924 | -0.036621 | -0.12817 | -0.022888 | -0.036621 | 0.082398 | 0.11902 | 0.036621 | 0.059509 | 0.17395 | -0.082397 | 0.19684 | -0.059509
566.7812 | 122.4792 | -59.1568 | 63.0524 | -0.13054 | 0.001128 | 0.067691 | 0.069773 | -2.7692e-05 | -4.7032e-05 | -0.0014526 | -0.044037 | -4.0508e-05 | 0.065775 | 0.066182 | 0.00045928 | -0.0019803 | -0.00019868 | -0.064087 | 63.0249 | -59.2941 | 63.0295 | -58.9737 | -59.4131 | -59.3216 | -0.004578 | 63.0341 | -59.2026 | 63.1439 | -59.111 | -59.2346 | -58.9874 | -0.13733 | -0.027466 | -0.022888 | 0.004578 | 0.1831 | 0.32043 | -0.16479 | -0.25635 | -0.027466 | -0.11902 | -0.045777 | -0.01831 | 0.091553 | 0.10986 | 0.045776 | 0.091553 | 0.16937 | -0.07782 | 0.21515 | -0.032043
567.7812 | 122.5479 | -59.1293 | 63.0707 | -0.038986 | -0.0020764 | 0.065127 | 0.067851 | -0.0021641 | 0.003371 | -0.00038459 | 0.0078429 | -0.00025411 | 0.06257 | 0.059774 | -0.0018905 | -0.0019803 | 0.00022852 | -0.022888 | 63.0707 | -59.2941 | 63.0432 | -58.9645 | -59.3445 | -59.2941 | 0.050354 | 63.0524 | -59.1476 | 63.1897 | -59.0515 | -59.166 | -58.9188 | -0.16479 | 0 | -0.027466 | -0.027466 | 0.16479 | 0.32959 | -0.16479 | -0.21515 | 0 | -0.050354 | -0.018311 | -0.018311 | 0.11902 | 0.13733 | 0.077819 | 0.09613 | 0.21057 | -0.036622 | 0.22888 | -0.018311
568.7812 | 122.5342 | -59.1156 | 63.0478 | -0.18394 | 0.002837 | 0.070468 | 0.070842 | -0.0030185 | 0.0018756 | 0.00068361 | -0.18137 | 0.0044455 | 0.062357 | 0.06661 | -0.0016769 | -0.0013393 | 0.0015103 | -0.036621 | 63.0341 | -59.2987 | 63.0341 | -58.9691 | -59.3903 | -59.3308 | -0.059509 | 62.97 | -59.2209 | 63.1989 | -59.0881 | -59.2072 | -58.9142 | -0.18311 | -0.013733 | -0.013733 | 0 | 0.14648 | 0.32959 | -0.21515 | -0.27466 | -0.032044 | -0.091553 | -0.10529 | -0.07782 | 0.15106 | 0.22888 | 0.027466 | 0.13275 | 0.20142 | -0.091552 | 0.3067 | 0.013733
569.7812 | 122.5342 | -59.1248 | 63.0753 | -0.12291 | -0.00079462 | 0.066409 | 0.06721 | -0.0017367 | 0.00038017 | -0.0016662 | -0.037933 | 0.0012411 | 0.06492 | 0.06191 | 0.0021683 | -0.0017667 | -0.0016941 | -0.082397 | 63.0615 | -59.2941 | 62.9883 | -59.0057 | -59.4131 | -59.3491 | -0.064087 | 63.0432 | -59.2117 | 63.1393 | -59.1339 | -59.2484 | -58.96 | -0.16937 | -0.013733 | -0.086975 | -0.073242 | 0.11902 | 0.28839 | -0.2243 | -0.28839 | -0.054932 | -0.11902 | -0.086975 | -0.032043 | 0.064087 | 0.09613 | -0.009155 | 0.07782 | 0.16479 | -0.1236 | 0.25177 | -0.036621
570.7812 | 122.4792 | -59.1568 | 63.0707 | -0.092392 | -0.00079462 | 0.070468 | 0.069133 | -0.0036593 | 0.0020892 | -0.0014526 | -0.059296 | 0.0016685 | 0.063425 | 0.066396 | 0.00024568 | -0.0015529 | -0.0031895 | -0.07782 | 63.0295 | -59.317 | 63.0066 | -58.992 | -59.4315 | -59.3399 | -0.050354 | 63.0753 | -59.198 | 63.1439 | -59.1202 | -59.2392 | -58.9874 | -0.16022 | -0.041199 | -0.064087 | -0.022888 | 0.16479 | 0.32501 | -0.18311 | -0.27466 | -0.022888 | -0.11444 | -0.041199 | 0.004577 | 0.073242 | 0.068665 | 0.036621 | 0.07782 | 0.16937 | -0.082398 | 0.21057 | -0.041199
571.7812 | 122.5159 | -59.1522 | 63.0753 | -0.10002 | 0.0030506 | 0.067264 | 0.066569 | -0.00088229 | 0.0014484 | 0.0017518 | -0.028778 | -0.00089511 | 0.064706 | 0.063192 | -0.0016769 | -0.0019803 | 0.0027921 | -0.091553 | 63.0569 | -59.3033 | 63.0066 | -59.0149 | -59.3994 | -59.3399 | -0.032043 | 63.0432 | -59.1843 | 63.1439 | -59.1156 | -59.1797 | -58.873 | -0.15106 | -0.01831 | -0.068664 | -0.050354 | 0.13733 | 0.28839 | -0.18768 | -0.24719 | -0.036621 | -0.09613 | -0.032043 | -0.032043 | 0.068665 | 0.10071 | 0.036621 | 0.068664 | 0.27924 | -0.027465 | 0.31128 | 0.004578
572.7812 | 122.5067 | -59.1156 | 63.121 | -0.038986 | 0.00048698 | 0.067477 | 0.068278 | 0.00082671 | 0.001662 | 0.00089721 | -0.028778 | 0.00060029 | 0.06492 | 0.064687 | -0.00082252 | -0.0036893 | 0.0036465 | 0.004578 | 63.0524 | -59.2804 | 63.0707 | -58.9279 | -59.3262 | -59.2712 | 0.10986 | 63.0753 | -59.166 | 63.1851 | -59.0652 | -59.1797 | -58.9279 | -0.1648 | -0.068665 | -0.050354 | 0.018311 | 0.18768 | 0.35248 | -0.15564 | -0.21057 | 0.009156 | -0.045776 | -0.050354 | -0.045777 | 0.064087 | 0.10986 | 0.050354 | 0.10071 | 0.18768 | -0.064086 | 0.23804 | -0.013732
573.7812 | 122.493 | -59.1339 | 63.0569 | -0.13817 | 0.001128 | 0.068332 | 0.069346 | 0.0027493 | 0.0023028 | -0.00081179 | -0.034882 | -0.0013223 | 0.064066 | 0.057851 | -0.0010361 | -0.0026211 | -0.0021213 | -0.050354 | 63.0341 | -59.2987 | 63.0386 | -58.96 | -59.4131 | -59.3308 | 0.022888 | 63.0524 | -59.1934 | 63.1668 | -59.1339 | -59.2346 | -58.9691 | -0.1648 | -0.022888 | -0.018311 | 0.004577 | 0.17395 | 0.33875 | -0.19684 | -0.27924 | -0.032044 | -0.11444 | -0.059509 | -0.004578 | 0.10986 | 0.11444 | 0 | 0.059509 | 0.16479 | -0.10071 | 0.2243 | -0.041199
574.7812 | 122.5433 | -59.1248 | 63.0615 | -0.054245 | 0.0026234 | 0.068118 | 0.069346 | 0.00082671 | 0.0035846 | 0.00025641 | -0.016571 | 0.00060029 | 0.064493 | 0.06426 | -0.00082252 | -0.0041165 | 0.0038601 | -0.041199 | 63.0478 | -59.2804 | 63.0203 | -58.9737 | -59.3857 | -59.3033 | 0.022888 | 63.0524 | -59.2163 | 63.1851 | -59.0652 | -59.1797 | -58.9371 | -0.15564 | -0.013732 | -0.041198 | -0.027466 | 0.15106 | 0.3067 | -0.17853 | -0.26093 | -0.022888 | -0.10528 | -0.091553 | -0.009155 | 0.1236 | 0.13275 | 0.059509 | 0.15106 | 0.18768 | -0.054931 | 0.27924 | 0.036622
575.7812 | 122.5113 | -59.1248 | 63.0478 | 0.060196 | -0.00036742 | 0.068759 | 0.070628 | 0.0018949 | -0.00026063 | -0.0014526 | -0.071503 | -4.0508e-05 | 0.063638 | 0.064687 | -0.0010361 | -0.0028347 | 0.0015103 | 0.009155 | 63.0203 | -59.317 | 63.0341 | -58.9737 | -59.3811 | -59.3079 | 0.022888 | 63.0203 | -59.2346 | 63.1622 | -59.1019 | -59.2255 | -58.9645 | -0.19226 | -0.027466 | -0.013733 | 0.013733 | 0.15106 | 0.34332 | -0.18311 | -0.25635 | 0.009156 | -0.064087 | -0.10986 | -0.027466 | 0.11444 | 0.14191 | 0.022888 | 0.13275 | 0.16022 | -0.10071 | 0.27008 | 0.009155
576.7812 | 122.5571 | -59.1385 | 63.0569 | -0.1458 | 0.0017688 | 0.067691 | 0.067424 | 0.00039951 | 0.0010212 | -0.0023072 | -0.053192 | 0.00060029 | 0.063852 | 0.065114 | -0.00039512 | -0.0017667 | 0.00065592 | -0.059509 | 63.0386 | -59.3079 | 63.0066 | -58.9783 | -59.4269 | -59.3536 | -0.009155 | 63.0295 | -59.1888 | 63.1439 | -59.1293 | -59.2575 | -58.9645 | -0.16937 | -0.018311 | -0.050354 | -0.032043 | 0.16022 | 0.32959 | -0.21515 | -0.28839 | -0.045777 | -0.11902 | -0.050354 | -0.027466 | 0.086975 | 0.11444 | 0.009156 | 0.05951 | 0.17395 | -0.11902 | 0.2243 | -0.068664
577.7812 | 122.5388 | -59.1064 | 63.089 | -0.10002 | 0.002837 | 0.066195 | 0.066996 | -0.0043003 | -0.00068783 | -0.00017099 | 0.02005 | 0.00038669 | 0.062784 | 0.062337 | -0.0018905 | -0.00069854 | -0.0021213 | 0.027466 | 63.0981 | -59.2529 | 63.089 | -58.873 | -59.3033 | -59.2346 | 0.14648 | 63.0615 | -59.1751 | 63.2126 | -59.0332 | -59.1476 | -58.9188 | -0.14648 | 0.009156 | 0 | -0.009156 | 0.23346 | 0.37994 | -0.12817 | -0.19684 | 0.018311 | -0.050354 | -0.068665 | -0.027466 | 0.1236 | 0.15106 | 0.073242 | 0.14191 | 0.18768 | -0.041199 | 0.25635 | 0.027466
578.7812 | 122.5159 | -59.1293 | 63.0524 | -0.10002 | -0.001222 | 0.067264 | 0.067637 | -0.0010959 | -4.7032e-05 | 0.00025641 | -0.047089 | -0.0017495 | 0.062997 | 0.064901 | -0.0018905 | -0.0021939 | -0.00083948 | -0.054932 | 63.0341 | -59.317 | 63.0295 | -58.9554 | -59.3903 | -59.3124 | 0.009155 | 63.0341 | -59.1888 | 63.1714 | -59.1064 | -59.2255 | -58.9462 | -0.18768 | -0.01831 | -0.022888 | -0.004578 | 0.17395 | 0.36163 | -0.18311 | -0.26093 | 0.004578 | -0.073242 | -0.05951 | -0.01831 | 0.11902 | 0.13733 | 0.022888 | 0.082398 | 0.18311 | -0.096131 | 0.24262 | -0.036621
579.7812 | 122.5204 | -59.1385 | 63.0707 | -0.11528 | -0.00079462 | 0.068973 | 0.068492 | 0.0027493 | -0.00068783 | 0.00025641 | -0.037933 | -0.00046791 | 0.06257 | 0.056997 | -0.0010361 | -0.0019803 | 0.0010831 | -0.07782 | 63.0615 | -59.3033 | 63.0157 | -58.9737 | -59.4269 | -59.3491 | -0.027466 | 63.0203 | -59.1797 | 63.1531 | -59.1202 | -59.2484 | -58.9737 | -0.16479 | -0.009156 | -0.054932 | -0.045776 | 0.16479 | 0.32959 | -0.21057 | -0.28839 | -0.045776 | -0.1236 | -0.041198 | -0.050354 | 0.082397 | 0.13275 | 0.018311 | 0.059509 | 0.16479 | -0.10986 | 0.20599 | -0.068665
580.7812 | 122.4747 | -59.1431 | 63.0707 | -0.084763 | 0.00091438 | 0.067477 | 0.068492 | -0.00088229 | 0.0014484 | 0.00089721 | -0.037933 | -0.00089511 | 0.065347 | 0.058492 | -0.00060892 | -0.0017667 | -0.00083948 | -0.09613 | 63.0341 | -59.285 | 62.9929 | -59.0149 | -59.404 | -59.3353 | -0.045776 | 63.0386 | -59.1888 | 63.1302 | -59.1156 | -59.2621 | -58.9783 | -0.14191 | -0.036621 | -0.07782 | -0.041199 | 0.12817 | 0.27008 | -0.19226 | -0.26093 | -0.050354 | -0.11902 | -0.045777 | -0.032044 | 0.059509 | 0.091553 | 0.027465 | 0.073242 | 0.16479 | -0.11902 | 0.21057 | -0.073242
581.7812 | 122.5113 | -59.1522 | 63.0569 | -0.10765 | 5.9776e-05 | 0.067477 | 0.070628 | -0.00066869 | 0.0018756 | -0.00038459 | -0.10202 | -0.00089511 | 0.064279 | 0.065328 | -0.0023177 | -0.0019803 | -0.00062588 | -0.10071 | 63.0341 | -59.3353 | 62.9791 | -59.0424 | -59.4635 | -59.3903 | -0.073242 | 63.0432 | -59.2255 | 63.1805 | -59.111 | -59.285 | -58.9737 | -0.1831 | -0.022888 | -0.07782 | -0.054932 | 0.10986 | 0.29297 | -0.23804 | -0.31128 | -0.054932 | -0.12817 | -0.073242 | -0.013733 | 0.1236 | 0.13733 | 0.041199 | 0.11444 | 0.17853 | -0.13275 | 0.25177 | -0.059509
582.7812 | 122.5204 | -59.1385 | 63.0798 | -0.10765 | 5.9776e-05 | 0.066409 | 0.067637 | -0.0030185 | 0.0012348 | 0.0023926 | 0.0017395 | 0.0010275 | 0.068765 | 0.065755 | 0.00067288 | -0.0024075 | -0.00019868 | -0.082397 | 63.0753 | -59.285 | 63.0066 | -59.0057 | -59.4131 | -59.3216 | -0.018311 | 63.0615 | -59.1888 | 63.1531 | -59.079 | -59.2072 | -58.9462 | -0.14648 | -0.004578 | -0.073242 | -0.068664 | 0.13275 | 0.27924 | -0.1831 | -0.27466 | -0.036621 | -0.12817 | -0.050354 | -0.018311 | 0.073242 | 0.091553 | 0.05951 | 0.10986 | 0.19226 | -0.068664 | 0.24262 | -0.01831
583.7812 | 122.5479 | -59.1431 | 63.0524 | -0.084763 | 0.0013416 | 0.067691 | 0.068705 | 0.0018949 | 0.0010212 | 0.0011108 | -0.068451 | 0.0018821 | 0.064493 | 0.062764 | -0.0010361 | -0.0019803 | -0.0019077 | -0.050354 | 63.0157 | -59.3216 | 63.0432 | -58.9737 | -59.3994 | -59.3216 | -0.1236 | 62.9929 | -59.2072 | 63.2034 | -59.0973 | -59.2209 | -58.9233 | -0.17853 | -0.036621 | -0.009155 | 0.027466 | 0.16937 | 0.3479 | -0.17853 | -0.25635 | 0 | -0.07782 | -0.064087 | -0.059509 | 0.15106 | 0.21057 | 0.045776 | 0.10986 | 0.21973 | -0.07782 | 0.28381 | -0.013733
584.7812 | 122.4884 | -59.1476 | 63.0798 | -0.023727 | -0.00058102 | 0.068759 | 0.068064 | -0.00088229 | 0.001662 | -0.0014526 | -0.0074159 | 0.00017309 | 0.063425 | 0.063619 | 0.00024568 | -0.0045437 | -0.0031895 | -0.068665 | 63.0524 | -59.2712 | 63.0157 | -59.0103 | -59.4131 | -59.3262 | -0.045776 | 63.0341 | -59.2072 | 63.1348 | -59.111 | -59.2758 | -59.0149 | -0.1236 | -0.027466 | -0.064087 | -0.036621 | 0.13733 | 0.26093 | -0.17853 | -0.2655 | -0.054932 | -0.14191 | -0.059509 | -0.045776 | 0.054932 | 0.10071 | 0.036621 | 0.09613 | 0.13275 | -0.12817 | 0.19226 | -0.068665
585.7812 | 122.5433 | -59.1339 | 63.0478 | -0.11528 | 0.00048698 | 0.0694 | 0.068492 | 0.0012539 | 0.0035846 | 0.00025641 | -0.092865 | -0.00025411 | 0.06492 | 0.062764 | -0.0012497 | -0.0013393 | 0.0036465 | -0.059509 | 63.0295 | -59.3033 | 63.0157 | -58.9828 | -59.3948 | -59.3399 | -0.004578 | 63.0341 | -59.198 | 63.1805 | -59.0881 | -59.2209 | -58.9508 | -0.16937 | -0.018311 | -0.032044 | -0.013733 | 0.15106 | 0.32043 | -0.20599 | -0.26093 | -0.036621 | -0.091552 | -0.064087 | -0.013733 | 0.13275 | 0.14648 | 0.045776 | 0.10986 | 0.1831 | -0.086975 | 0.24719 | -0.022888
586.7812 | 122.5204 | -59.1522 | 63.0478 | -0.1458 | 0.00091438 | 0.067691 | 0.069133 | -0.0023777 | -4.7032e-05 | 4.2808e-05 | -0.095917 | 0.0012411 | 0.063425 | 0.069814 | -0.0012497 | -0.0026211 | -0.00062588 | -0.10071 | 63.0661 | -59.3262 | 62.9883 | -59.0378 | -59.4452 | -59.3811 | -0.045776 | 63.0157 | -59.1888 | 63.1897 | -59.1202 | -59.2346 | -58.9279 | -0.17395 | 0.01831 | -0.05951 | -0.07782 | 0.11444 | 0.28839 | -0.22888 | -0.29297 | -0.054932 | -0.11902 | -0.036621 | -0.032044 | 0.14191 | 0.17395 | 0.032044 | 0.068665 | 0.22431 | -0.082397 | 0.26093 | -0.045776
587.7812 | 122.525 | -59.1522 | 63.0386 | -0.11528 | -0.00036742 | 0.068759 | 0.068919 | -0.0013095 | 0.00038017 | -0.0016662 | -0.10202 | 0.00017309 | 0.062143 | 0.065328 | -0.00039512 | 0.00036966 | -0.0010531 | -0.11444 | 63.0524 | -59.317 | 62.9654 | -59.0515 | -59.4681 | -59.404 | -0.086975 | 63.0386 | -59.2346 | 63.1531 | -59.1293 | -59.2712 | -58.9783 | -0.16479 | 0.013733 | -0.073242 | -0.086975 | 0.10071 | 0.2655 | -0.25177 | -0.31586 | -0.086975 | -0.15106 | -0.082397 | 0 | 0.11444 | 0.11444 | 0.022889 | 0.10529 | 0.17395 | -0.11902 | 0.25635 | -0.036621
588.7812 | 122.5388 | -59.1339 | 63.0569 | -0.13817 | -0.00079462 | 0.069186 | 0.065074 | -0.00066869 | 0.0010212 | 0.0017518 | -0.062347 | 0.0010275 | 0.064706 | 0.065969 | -0.0012497 | -0.0039029 | 0.0015103 | -0.054932 | 63.0249 | -59.2941 | 63.0203 | -58.9737 | -59.3308 | -59.3124 | 0.009155 | 63.0341 | -59.1797 | 63.1668 | -59.0927 | -59.2163 | -58.9279 | -0.16022 | -0.032044 | -0.036621 | -0.004577 | 0.16022 | 0.32043 | -0.17853 | -0.19684 | -0.018311 | -0.036622 | -0.045776 | -0.022888 | 0.10986 | 0.13275 | 0.041199 | 0.086975 | 0.20599 | -0.082398 | 0.25177 | -0.036622
589.7812 | 122.4655 | -59.1705 | 63.089 | -0.11528 | -0.00058102 | 0.069613 | 0.069133 | -0.0010959 | 0.0020892 | 0.00047001 | -0.08371 | 0.0012411 | 0.06492 | 0.06426 | 0.00024568 | -0.0019803 | -0.00083948 | -0.086975 | 63.0707 | -59.3216 | 62.9379 | -59.0652 | -59.4955 | -59.436 | -0.11902 | 63.0661 | -59.2117 | 63.1027 | -59.1934 | -59.3124 | -59.0103 | -0.15106 | -0.01831 | -0.15106 | -0.13275 | 0.10529 | 0.25635 | -0.2655 | -0.32501 | -0.11444 | -0.17395 | -0.041199 | -0.022888 | 0.013733 | 0.036621 | -0.022888 | 0.018311 | 0.16022 | -0.14191 | 0.20142 | -0.10071
590.7812 | 122.4792 | -59.1751 | 63.0524 | -0.046616 | -0.00015382 | 0.065982 | 0.06956 | 0.0033901 | 0.0012348 | -0.00038459 | -0.08371 | 0.00081389 | 0.063852 | 0.071309 | 3.2078e-05 | -0.0013393 | 0.0017239 | -0.11444 | 63.0707 | -59.3353 | 62.9608 | -59.0561 | -59.4772 | -59.3994 | -0.086975 | 63.0432 | -59.2255 | 63.1393 | -59.1293 | -59.285 | -58.9828 | -0.16022 | 0.018311 | -0.091553 | -0.10986 | 0.11902 | 0.27924 | -0.2243 | -0.30212 | -0.064087 | -0.14191 | -0.050354 | -0.009155 | 0.086975 | 0.09613 | 0.045777 | 0.096131 | 0.19226 | -0.10986 | 0.24262 | -0.059509
591.7812 | 122.4609 | -59.166 | 63.0844 | -0.084763 | 0.00027338 | 0.06705 | 0.068705 | -0.0015231 | 0.001662 | 0.0013244 | -0.089813 | 0.00060029 | 0.064493 | 0.061483 | -0.00082252 | -0.0024075 | 0.0019375 | -0.10071 | 63.0844 | -59.3308 | 62.9288 | -59.0744 | -59.4772 | -59.4131 | -0.10986 | 63.0661 | -59.2255 | 63.1165 | -59.1751 | -59.3124 | -59.0057 | -0.1648 | 0 | -0.15564 | -0.15564 | 0.091553 | 0.25635 | -0.24719 | -0.31128 | -0.082397 | -0.14648 | -0.059509 | -0.018311 | 0.032043 | 0.050354 | -0.009155 | 0.050354 | 0.16022 | -0.14648 | 0.21973 | -0.086975
592.7812 | 122.5342 | -59.1568 | 63.0661 | -0.13054 | 0.00048698 | 0.067264 | 0.066996 | 0.0029629 | 0.00016657 | 0.0017518 | -0.047089 | 0.00017309 | 0.065988 | 0.066823 | 3.2078e-05 | -0.0024075 | 0.00065592 | -0.1236 | 63.0707 | -59.2987 | 62.9425 | -59.0561 | -59.4635 | -59.3857 | -0.086975 | 63.0157 | -59.23 | 63.1622 | -59.1202 | -59.2438 | -58.9691 | -0.14191 | 0.004578 | -0.1236 | -0.12817 | 0.10071 | 0.24262 | -0.22888 | -0.3067 | -0.086975 | -0.16479 | -0.073243 | -0.050354 | 0.09613 | 0.14648 | 0.036621 | 0.10986 | 0.18768 | -0.086975 | 0.26093 | -0.013732
593.7812 | 122.406 | -59.1568 | 63.0936 | -0.084763 | -0.00036742 | 0.067904 | 0.070628 | -0.0017367 | 0.0018756 | 0.00068361 | 0.013947 | 0.0029501 | 0.064066 | 0.064901 | 0.0013139 | 0.0010105 | 0.00065592 | -0.11902 | 63.089 | -59.3353 | 62.9105 | -59.1156 | -59.7885 | -59.6558 | -0.027466 | 63.0981 | -59.2026 | 63.0936 | -59.2163 | -59.3536 | -59.0515 | -0.17853 | -0.004578 | -0.18311 | -0.17853 | 0.041198 | 0.21973 | -0.49896 | -0.63171 | -0.32044 | -0.45319 | -0.045777 | 0.004578 | 0 | -0.004578 | -0.05951 | -0.013733 | 0.10529 | -0.19684 | 0.15106 | -0.15106
594.7812 | 122.5433 | -59.1385 | 63.0798 | -0.061874 | 0.0013416 | 0.066623 | 0.070201 | -0.00088229 | 0.00038017 | 0.0034606 | -0.071503 | -0.00089511 | 0.065561 | 0.060415 | -0.00039512 | -0.0019803 | 0.0040737 | -0.10071 | 63.0753 | -59.3308 | 62.9791 | -59.0149 | -59.4452 | -59.3765 | -0.064087 | 63.0661 | -59.1705 | 63.1393 | -59.1248 | -59.23 | -58.9371 | -0.19226 | -0.004578 | -0.10071 | -0.09613 | 0.1236 | 0.31586 | -0.23804 | -0.3067 | -0.045776 | -0.11444 | -0.032043 | -0.013733 | 0.059509 | 0.073242 | 0.013733 | 0.045776 | 0.20142 | -0.091553 | 0.23346 | -0.05951
595.7812 | 122.3969 | -59.1614 | 63.1302 | -0.10765 | 0.00070078 | 0.0694 | 0.067851 | -0.0032321 | -0.00026063 | 0.002179 | -0.10202 | 0.0035911 | 0.067484 | 0.057638 | -0.00039512 | -0.0015529 | 0.0027921 | -0.068665 | 63.0707 | -59.317 | 62.9333 | -59.0881 | -59.5001 | -59.436 | -0.13275 | 63.0707 | -59.2026 | 63.0386 | -59.23 | -59.3674 | -59.0652 | -0.15564 | -0.059509 | -0.19684 | -0.13733 | 0.073242 | 0.22888 | -0.27466 | -0.33874 | -0.11902 | -0.1831 | -0.041199 | -0.059509 | -0.091553 | -0.032044 | -0.068665 | -0.027466 | 0.09613 | -0.20599 | 0.13733 | -0.16479
596.7812 | 122.4838 | -59.1843 | 63.0981 | -0.084763 | -0.00036742 | 0.068759 | 0.069346 | 0.0038175 | 0.00080757 | -0.00038459 | -0.098968 | 0.0018821 | 0.065561 | 0.072805 | 3.2078e-05 | -0.00069854 | -0.0046847 | -0.10986 | 63.0936 | -59.3262 | 62.9059 | -59.1293 | -59.523 | -59.4681 | -0.13275 | 63.0615 | -59.2163 | 63.0707 | -59.2209 | -59.3628 | -59.0378 | -0.14191 | -0.004578 | -0.19226 | -0.18768 | 0.054932 | 0.19684 | -0.28381 | -0.33874 | -0.14191 | -0.19684 | -0.032044 | -0.036622 | -0.027466 | 0.009156 | -0.036621 | -0.004577 | 0.14648 | -0.17853 | 0.17853 | -0.14648
597.7812 | 122.5067 | -59.1522 | 63.0707 | -0.13817 | 0.001128 | 0.064914 | 0.064647 | 0.00039951 | 0.00038017 | -0.0014526 | 0.016998 | -4.0508e-05 | 0.065561 | 0.066182 | 0.00024568 | -0.0021939 | 0.0034329 | -0.068665 | 63.0707 | -59.2621 | 63.0157 | -58.96 | -59.3903 | -59.3308 | -0.018311 | 63.0157 | -59.1843 | 63.1348 | -59.1431 | -59.2484 | -58.9691 | -0.10986 | 0 | -0.054932 | -0.054932 | 0.19226 | 0.30212 | -0.17853 | -0.23804 | -0.068665 | -0.12817 | -0.032043 | -0.054932 | 0.064087 | 0.11902 | 0.009156 | 0.041199 | 0.18311 | -0.09613 | 0.21515 | -0.064087
598.7812 | 122.4152 | -59.1522 | 63.1119 | -0.077133 | 0.00091438 | 0.068118 | 0.069773 | 0.00082671 | 0.00059397 | 0.00047001 | -0.08371 | 0.0012411 | 0.064066 | 0.06426 | -0.00018152 | -0.0015529 | 0.00086952 | -0.041199 | 63.089 | -59.285 | 62.9517 | -59.0515 | -59.4635 | -59.404 | -0.073242 | 63.089 | -59.1843 | 63.0753 | -59.198 | -59.3399 | -59.0332 | -0.13275 | -0.022888 | -0.16022 | -0.13733 | 0.10071 | 0.23346 | -0.25177 | -0.31128 | -0.11902 | -0.17853 | -0.032043 | -0.022888 | -0.036621 | -0.013733 | -0.045776 | -0.013733 | 0.11902 | -0.18768 | 0.15106 | -0.15564
599.7812 | 122.4518 | -59.166 | 63.1302 | -0.077133 | -0.0014356 | 0.068332 | 0.070628 | 0.0012539 | 0.0012348 | 0.001538 | -0.068451 | 0.0070089 | 0.067697 | 0.067251 | -0.0016769 | -0.0021939 | 0.0025785 | -0.10071 | 63.089 | -59.317 | 62.915 | -59.111 | -59.5093 | -59.4498 | -0.4303 | 63.002 | -59.2484 | 63.1165 | -59.2209 | -59.3765 | -59.0469 | -0.15106 | -0.041199 | -0.21515 | -0.17395 | 0.054932 | 0.20599 | -0.28381 | -0.34332 | -0.13275 | -0.19226 | -0.082397 | -0.12817 | -0.013733 | 0.11444 | -0.054931 | 0.027466 | 0.11902 | -0.21057 | 0.20142 | -0.12817
600.7812 | 122.4243 | -59.1431 | 63.1302 | -0.10002 | -0.0010082 | 0.068545 | 0.069346 | -2.7692e-05 | 0.0010212 | 4.2808e-05 | -0.059296 | -0.00025411 | 0.067484 | 0.067037 | -0.00082252 | -0.0024075 | -0.00083948 | -0.036621 | 63.0981 | -59.3124 | 62.9517 | -59.0469 | -59.4498 | -59.404 | -0.054932 | 63.0661 | -59.1797 | 63.089 | -59.1934 | -59.3399 | -59.024 | -0.16937 | -0.032043 | -0.17853 | -0.14649 | 0.09613 | 0.2655 | -0.26093 | -0.3067 | -0.091553 | -0.13733 | -0.036621 | -0.064087 | -0.041199 | 0.022888 | -0.050354 | -0.013733 | 0.11902 | -0.19684 | 0.15564 | -0.16022
601.7812 | 122.4747 | -59.1705 | 63.121 | -0.046616 | 0.00070078 | 0.065341 | 0.068705 | 0.00039951 | 0.0031574 | 0.00047001 | -0.047089 | 0.0027365 | 0.065347 | 0.064687 | -0.00060892 | -0.0026211 | 0.0015103 | 0.004578 | 63.1073 | -59.2941 | 62.9883 | -59.0012 | -59.4086 | -59.3628 | 0.022888 | 63.0798 | -59.1522 | 63.1622 | -59.2667 | -59.285 | -58.992 | -0.1236 | -0.013733 | -0.13275 | -0.11902 | 0.16937 | 0.29297 | -0.19226 | -0.23804 | -0.068665 | -0.11444 | 0.01831 | -0.041199 | 0.041198 | 0.082397 | -0.096131 | -0.11444 | 0.17853 | -0.11444 | 0.16022 | -0.13275
602.7812 | 122.4518 | -59.1248 | 63.1577 | -0.054245 | -0.0020764 | 0.067477 | 0.067637 | 0.0016813 | 0.00059397 | 0.00047001 | -0.016571 | 0.00017309 | 0.06727 | 0.060201 | -0.00039512 | -0.0024075 | 0.0019375 | -0.027466 | 63.1393 | -59.285 | 62.9654 | -59.0378 | -59.4177 | -59.3628 | -0.009155 | 63.0936 | -59.166 | 63.1165 | -59.111 | -59.2438 | -58.9783 | -0.16022 | -0.018311 | -0.19226 | -0.17395 | 0.086975 | 0.24719 | -0.23804 | -0.29297 | -0.07782 | -0.13275 | -0.041199 | -0.064087 | -0.041199 | 0.022888 | 0.013733 | 0.054932 | 0.14648 | -0.11902 | 0.18768 | -0.077819
603.7812 | 122.4609 | -59.1705 | 63.0753 | -0.092392 | -0.0020764 | 0.064486 | 0.067851 | 0.00082671 | -0.00047423 | -0.0018798 | -0.034882 | 0.0020957 | 0.065134 | 0.063833 | 0.00067288 | -0.0030483 | -0.0012667 | -0.11902 | 63.0569 | -59.317 | 62.9379 | -59.0652 | -59.4772 | -59.3994 | -0.09613 | 63.0432 | -59.1843 | 63.0981 | -59.1614 | -59.2987 | -59.0012 | -0.14648 | -0.01831 | -0.13733 | -0.11902 | 0.10529 | 0.25177 | -0.22888 | -0.3067 | -0.082397 | -0.16022 | -0.013733 | -0.032043 | 0.022889 | 0.054932 | 0.009155 | 0.022888 | 0.16937 | -0.12817 | 0.18311 | -0.11444
604.7812 | 122.4243 | -59.1522 | 63.1348 | -0.061874 | 0.00027338 | 0.069186 | 0.068064 | -0.0043003 | 0.0010212 | -0.0018798 | -0.086761 | 0.00060029 | 0.067056 | 0.065328 | -0.00039512 | -0.00069854 | 0.00022852 | -0.054932 | 63.0844 | -59.3262 | 62.9517 | -59.0378 | -59.4635 | -59.4131 | -0.073242 | 63.1073 | -59.2072 | 63.0753 | -59.1934 | -59.3262 | -59.0286 | -0.17395 | -0.050354 | -0.18311 | -0.13275 | 0.11444 | 0.28839 | -0.26093 | -0.31128 | -0.086975 | -0.13733 | -0.054931 | -0.027466 | -0.05951 | -0.032044 | -0.041198 | 0.013733 | 0.1236 | -0.17395 | 0.17853 | -0.11902
605.7812 | 122.4426 | -59.1476 | 63.1119 | -0.12291 | 0.00070078 | 0.066836 | 0.068492 | 0.0036039 | -0.00068783 | -0.002948 | -0.056244 | -0.00025411 | 0.063638 | 0.063833 | -0.0016769 | -0.0021939 | -0.0027621 | -0.027466 | 63.1073 | -59.2941 | 62.9837 | -59.0103 | -59.4269 | -59.3765 | -0.022888 | 63.0661 | -59.1934 | 63.089 | -59.1751 | -59.3033 | -59.0195 | -0.14648 | -0.004577 | -0.12817 | -0.1236 | 0.13733 | 0.28381 | -0.22888 | -0.27924 | -0.082398 | -0.13275 | -0.045776 | -0.045776 | -0.022888 | 0.022888 | -0.027466 | 0.01831 | 0.12817 | -0.15564 | 0.17395 | -0.10986
606.7812 | 122.4564 | -59.1248 | 63.121 | -0.054245 | 0.0015552 | 0.068118 | 0.067637 | -0.0021641 | -0.00026063 | -0.0014526 | 0.0047913 | 0.00060029 | 0.06492 | 0.063833 | -0.00039512 | -0.0015529 | 0.00044212 | -0.018311 | 63.121 | -59.2484 | 62.9745 | -59.0012 | -59.3811 | -59.3491 | 0.018311 | 63.1027 | -59.1934 | 63.1119 | -59.1248 | -59.2438 | -58.992 | -0.1236 | 0 | -0.14648 | -0.14648 | 0.1236 | 0.24719 | -0.2243 | -0.25635 | -0.10071 | -0.13275 | -0.068664 | -0.018311 | -0.009156 | 0.009155 | 0 | 0.068664 | 0.13275 | -0.11902 | 0.20142 | -0.050354
607.7812 | 122.3969 | -59.1705 | 63.1119 | -0.092392 | 0.0030506 | 0.067477 | 0.068492 | -0.0013095 | 0.0023028 | 0.0017518 | -0.028778 | 0.0014547 | 0.064706 | 0.067891 | 0.00088648 | -0.0017667 | 1.4916e-05 | -0.09613 | 63.0661 | -59.2941 | 62.9288 | -59.1064 | -59.5047 | -59.4543 | -0.15106 | 63.0798 | -59.2117 | 63.0341 | -59.2346 | -59.3765 | -59.0652 | -0.1236 | -0.045776 | -0.1831 | -0.13733 | 0.064087 | 0.18768 | -0.28381 | -0.33417 | -0.16022 | -0.21057 | -0.041199 | -0.032043 | -0.077819 | -0.045776 | -0.064087 | -0.022888 | 0.10529 | -0.20599 | 0.14648 | -0.16479
608.7812 | 122.4472 | -59.1431 | 63.1165 | -0.069504 | 0.00048698 | 0.068332 | 0.06721 | 0.00061311 | 0.0031574 | -0.0033752 | -0.040985 | 0.0012411 | 0.065561 | 0.065542 | 0.00045928 | -0.0024075 | -0.00062588 | -0.10071 | 63.0844 | -59.3124 | 62.9288 | -59.0652 | -59.4635 | -59.4086 | -0.11902 | 63.0798 | -59.2072 | 63.1073 | -59.1797 | -59.3124 | -59.0057 | -0.16937 | -0.032043 | -0.18768 | -0.15564 | 0.077819 | 0.24719 | -0.2655 | -0.32044 | -0.09613 | -0.15106 | -0.064087 | -0.036621 | -0.009155 | 0.027466 | -0.036621 | 0.027466 | 0.13733 | -0.16937 | 0.20142 | -0.10529
609.7812 | 122.4426 | -59.1568 | 63.1622 | -0.061874 | -0.00015382 | 0.067904 | 0.064647 | -0.0043003 | -4.7032e-05 | -0.0010254 | 0.02005 | 0.0020957 | 0.064279 | 0.064901 | -0.0018905 | -0.00048494 | 0.0012967 | -0.036621 | 63.1577 | -59.2484 | 62.9608 | -59.0332 | -59.4131 | -59.3719 | 0.21057 | 63.1622 | -59.1385 | 63.0569 | -59.1614 | -59.2758 | -59.024 | -0.091553 | -0.004577 | -0.20142 | -0.19684 | 0.1236 | 0.21515 | -0.21515 | -0.25635 | -0.1236 | -0.1648 | 0.01831 | 0 | -0.10528 | -0.10528 | -0.004578 | -0.022888 | 0.13275 | -0.11902 | 0.11444 | -0.13733
610.7812 | 122.4838 | -59.1339 | 63.1485 | -0.031357 | 0.00048698 | 0.066409 | 0.070201 | -2.7692e-05 | -0.00090143 | 0.0026062 | -0.025726 | 0.0027365 | 0.065988 | 0.06426 | -0.00060892 | -0.0032619 | -0.00062588 | -0.022888 | 63.1165 | -59.2712 | 63.0112 | -59.0149 | -59.3811 | -59.3445 | 0.036621 | 63.1027 | -59.166 | 63.1256 | -59.1156 | -59.2529 | -58.9691 | -0.13733 | -0.032044 | -0.13733 | -0.10529 | 0.11902 | 0.25635 | -0.21057 | -0.24719 | -0.073242 | -0.10986 | -0.032044 | -0.045777 | -0.022889 | 0.022888 | 0.01831 | 0.050354 | 0.16479 | -0.11902 | 0.19684 | -0.086975
611.7812 | 122.4335 | -59.1476 | 63.1622 | -0.069504 | -0.00079462 | 0.067477 | 0.066142 | 0.00039951 | -0.0021832 | -0.0025208 | 0.038361 | 0.0018821 | 0.065347 | 0.058706 | 0.00045928 | -0.0026211 | -0.0012667 | -0.041199 | 63.121 | -59.2712 | 63.0936 | -59.1614 | -59.6466 | -59.4315 | -0.041199 | 63.1073 | -59.1705 | 63.0844 | -59.1522 | -59.2712 | -58.9828 | -0.1236 | -0.041198 | -0.068664 | -0.027466 | -0.013733 | 0.10986 | -0.28381 | -0.49896 | -0.16022 | -0.37537 | -0.022888 | -0.054931 | -0.077819 | -0.022888 | -0.004578 | 0.01831 | 0.16479 | -0.1236 | 0.18768 | -0.10071
612.7812 | 122.438 | -59.1568 | 63.1302 | -0.092392 | 0.0019824 | 0.067264 | 0.071055 | 0.0010403 | 0.0025166 | -0.00081179 | -0.05014 | 0.0012411 | 0.067911 | 0.068319 | -0.0016769 | -0.0017667 | 0.0023647 | -0.036621 | 63.0981 | -59.2896 | 62.9791 | -59.0195 | -59.4452 | -59.3765 | -0.022888 | 63.089 | -59.1751 | 63.0798 | -59.1751 | -59.3124 | -59.0057 | -0.13275 | -0.032043 | -0.15106 | -0.11902 | 0.13733 | 0.27008 | -0.21973 | -0.28839 | -0.086975 | -0.15564 | -0.018311 | -0.041199 | -0.050354 | -0.009155 | -0.018311 | 0 | 0.15106 | -0.15564 | 0.16937 | -0.13733
613.7969 | 122.4747 | -59.1156 | 63.0707 | -0.13817 | -0.00058102 | 0.069186 | 0.066569 | -0.0013095 | 0.0010212 | 0.0013244 | 0.050568 | 0.00060029 | 0.06727 | 0.055715 | -0.0010361 | -0.0024075 | -0.00019868 | -0.054932 | 63.0615 | -59.2667 | 63.0295 | -58.9417 | -59.3857 | -59.3216 | -0.004578 | 62.9929 | -59.198 | 63.1439 | -59.1156 | -59.2438 | -58.9462 | -0.15106 | -0.009156 | -0.041199 | -0.032043 | 0.17395 | 0.32501 | -0.20599 | -0.27008 | -0.054931 | -0.11902 | -0.082397 | -0.07782 | 0.073242 | 0.15106 | 0 | 0.082397 | 0.16937 | -0.12817 | 0.25177 | -0.045776
614.7812 | 122.4518 | -59.1293 | 63.1531 | -0.092392 | 5.9776e-05 | 0.066409 | 0.069133 | 0.00061311 | 0.0023028 | -0.0027344 | 0.041412 | 0.0023093 | 0.066415 | 0.06426 | -0.0021041 | -0.0032619 | 0.0032193 | -0.036621 | 63.1439 | -59.2667 | 63.0798 | -58.8364 | -59.6329 | -59.5139 | -0.004578 | 63.1119 | -59.1843 | 63.089 | -59.1614 | -59.285 | -59.0057 | -0.13733 | -0.009155 | -0.073242 | -0.064087 | 0.29297 | 0.4303 | -0.38452 | -0.50354 | -0.24719 | -0.36621 | -0.054932 | -0.041199 | -0.064087 | -0.022888 | -0.032044 | 0.022888 | 0.1236 | -0.15564 | 0.17853 | -0.10071
615.7812 | 122.4014 | -59.1202 | 63.1531 | -0.023727 | -0.00036742 | 0.068332 | 0.065928 | -0.0025913 | 0.00080757 | 0.0013244 | 0.041412 | 0.00038669 | 0.066202 | 0.067464 | -0.0012497 | -0.0026211 | -0.0034031 | 0.022888 | 63.121 | -59.2667 | 63.0203 | -58.9508 | -59.3628 | -59.317 | 0.054932 | 63.0844 | -59.1568 | 63.1027 | -59.1064 | -59.2575 | -59.0012 | -0.14648 | -0.032043 | -0.13275 | -0.10071 | 0.16937 | 0.31586 | -0.19684 | -0.24262 | -0.050354 | -0.09613 | -0.036621 | -0.068664 | -0.050354 | 0.01831 | 0.013733 | 0.050354 | 0.11902 | -0.13733 | 0.15564 | -0.10071
616.7812 | 122.4838 | -59.111 | 63.1668 | -0.077133 | 0.00048698 | 0.065555 | 0.066996 | -0.0021641 | 0.00059397 | 0.0019654 | 0.038361 | -0.00046791 | 0.063852 | 0.062978 | -0.0014633 | -0.00027134 | -0.0019077 | 0.045776 | 63.0936 | -59.2255 | 63.0524 | -58.9325 | -59.3124 | -59.2712 | 0.09613 | 63.0936 | -59.1476 | 63.1348 | -59.111 | -59.2163 | -58.9508 | -0.11444 | -0.073242 | -0.11444 | -0.041199 | 0.17853 | 0.29297 | -0.16022 | -0.20142 | -0.045776 | -0.086975 | -0.036621 | -0.073242 | -0.032043 | 0.041199 | 0 | 0.036621 | 0.16022 | -0.10529 | 0.19684 | -0.068665
617.7812 | 122.4747 | -59.1293 | 63.1531 | -0.25261 | -0.0025036 | 0.067691 | 0.071269 | 0.00061311 | 0.001662 | -0.0010254 | -0.056244 | 0.00081389 | 0.063638 | 0.067464 | -0.0014633 | -0.0015529 | -0.0010531 | -0.009155 | 63.0753 | -59.3079 | 63.0203 | -58.992 | -59.4177 | -59.3719 | 0.009155 | 63.1165 | -59.1751 | 63.1119 | -59.1522 | -59.285 | -58.9783 | -0.17853 | -0.07782 | -0.13275 | -0.054931 | 0.13733 | 0.31586 | -0.24262 | -0.28839 | -0.064087 | -0.10986 | -0.045777 | -0.036621 | -0.041199 | -0.004578 | -0.022889 | 0.022888 | 0.15106 | -0.15564 | 0.19684 | -0.10986
618.7969 | 122.5021 | -59.0927 | 63.0936 | -0.1458 | 0.0013416 | 0.0694 | 0.067851 | -2.7692e-05 | 0.0029438 | -0.0023072 | 0.016998 | -4.0508e-05 | 0.065347 | 0.061056 | 0.00045928 | -0.0034757 | 0.00065592 | -0.050354 | 63.0661 | -59.2438 | 63.0341 | -58.9462 | -59.3582 | -59.2941 | 0.032043 | 63.0478 | -59.1888 | 63.1577 | -59.0927 | -59.2026 | -58.9417 | -0.15106 | -0.027466 | -0.059509 | -0.032043 | 0.14648 | 0.29755 | -0.20142 | -0.2655 | -0.050354 | -0.11444 | -0.096131 | -0.045776 | 0.064087 | 0.10986 | 0 | 0.096131 | 0.15106 | -0.10986 | 0.24719 | -0.013733
619.7812 | 122.4655 | -59.1248 | 63.1165 | -0.023727 | 0.0032642 | 0.068545 | 0.069987 | -0.00045509 | 0.00038017 | -0.0010254 | -0.040985 | 0.00038669 | 0.063638 | 0.06896 | -0.0010361 | -0.0026211 | -0.0021213 | -0.018311 | 63.1119 | -59.2758 | 63.0295 | -58.9783 | -59.3811 | -59.3491 | 0.036621 | 63.1256 | -59.1705 | 63.1531 | -59.1431 | -59.2163 | -58.9233 | -0.15106 | -0.004578 | -0.086975 | -0.082397 | 0.14648 | 0.29755 | -0.2243 | -0.25635 | -0.073242 | -0.10529 | -0.045776 | 0.009155 | 0.036621 | 0.027466 | -0.01831 | 0.027466 | 0.20142 | -0.091553 | 0.24719 | -0.045777
620.7812 | 122.4564 | -59.0881 | 63.121 | -0.069504 | 0.0026234 | 0.068118 | 0.066783 | -0.00045509 | 0.0014484 | -0.00017099 | 0.056671 | 0.00060029 | 0.064279 | 0.064473 | -0.00018152 | -0.0039029 | -0.0038303 | 0.041199 | 63.1119 | -59.2575 | 63.0478 | -58.9188 | -59.3216 | -59.2667 | 0.10071 | 63.1073 | -59.1431 | 63.1393 | -59.0607 | -59.2117 | -58.9462 | -0.16937 | -0.009156 | -0.073242 | -0.064086 | 0.16937 | 0.33874 | -0.17853 | -0.23346 | -0.009156 | -0.064087 | -0.054931 | -0.013733 | 0.01831 | 0.032043 | 0.027466 | 0.082397 | 0.14191 | -0.1236 | 0.19684 | -0.068665
621.7812 | 122.438 | -59.1248 | 63.1714 | -0.092392 | -0.0010082 | 0.065768 | 0.068064 | -2.7692e-05 | -4.7032e-05 | -0.00059819 | 0.093292 | 0.0027365 | 0.068338 | 0.070882 | -0.0025315 | -0.0017667 | -0.0031895 | -0.032043 | 63.1302 | -59.285 | 62.97 | -59.0332 | -59.4131 | -59.3536 | -0.09613 | 63.1485 | -59.1751 | 63.0753 | -59.1385 | -59.3079 | -59.0286 | -0.16022 | -0.041199 | -0.20142 | -0.16022 | 0.091553 | 0.25177 | -0.22888 | -0.28839 | -0.068665 | -0.12817 | -0.050354 | -0.022888 | -0.096131 | -0.073243 | -0.013733 | 0.036621 | 0.096131 | -0.18311 | 0.14648 | -0.13275
622.7812 | 122.406 | -59.1705 | 63.1348 | -0.10765 | -0.00058102 | 0.070254 | 0.070628 | 0.0033901 | 0.00038017 | -0.0020936 | -0.080658 | 0.0016685 | 0.063211 | 0.067251 | -0.00018152 | -0.0019803 | -0.00019868 | -0.07782 | 63.1165 | -59.317 | 62.9425 | -59.1064 | -59.491 | -59.4223 | -0.14191 | 63.0936 | -59.2072 | 63.0524 | -59.2163 | -59.3628 | -59.0515 | -0.14648 | -0.018311 | -0.19226 | -0.17395 | 0.064087 | 0.21057 | -0.25177 | -0.32043 | -0.10529 | -0.17395 | -0.036621 | -0.041199 | -0.082398 | -0.041199 | -0.045777 | -0.009156 | 0.11902 | -0.19226 | 0.15564 | -0.15564
623.7812 | 122.5891 | -59.1064 | 63.0936 | -0.084763 | 0.001128 | 0.067904 | 0.06721 | -0.0036593 | 0.0012348 | -0.0050842 | 0.081085 | -4.0508e-05 | 0.063211 | 0.057638 | 0.0013139 | -0.0047573 | 1.4916e-05 | 0.004578 | 63.1073 | -59.2392 | 63.1073 | -58.9005 | -59.2987 | -59.23 | 0.14648 | 63.0707 | -59.1385 | 63.2126 | -58.992 | -59.1064 | -58.8776 | -0.13275 | 0.013733 | 0.013733 | 0 | 0.20599 | 0.33874 | -0.1236 | -0.19226 | 0.009155 | -0.059509 | -0.032044 | -0.022888 | 0.11902 | 0.14191 | 0.11444 | 0.14648 | 0.22888 | 0 | 0.26093 | 0.032044
624.7812 | 122.4976 | -59.1019 | 63.1485 | -0.023727 | 0.00091438 | 0.067691 | 0.066569 | -0.0015231 | 0.0010212 | -0.0027344 | 0.065826 | -0.00068151 | 0.063852 | 0.062551 | -0.0014633 | -0.0032619 | -0.00083948 | 0.027466 | 63.1073 | -59.2667 | 63.0798 | -58.9142 | -59.3033 | -59.2575 | 0.11902 | 63.0753 | -59.1293 | 63.1668 | -59.0607 | -59.166 | -58.9096 | -0.16479 | -0.041199 | -0.068665 | -0.027466 | 0.18768 | 0.35248 | -0.15564 | -0.20142 | 0.009156 | -0.036621 | -0.027465 | -0.073243 | 0.01831 | 0.091553 | 0.041199 | 0.068664 | 0.19226 | -0.064087 | 0.21973 | -0.036622
625.7812 | 122.4426 | -59.0927 | 63.176 | -0.092392 | -0.00015382 | 0.068118 | 0.066569 | 0.0012539 | 0.00080757 | -0.00038459 | 0.016998 | 0.00017309 | 0.067697 | 0.059774 | 0.00088648 | -0.0026211 | 0.0019375 | 0.027466 | 63.1302 | -59.2575 | 63.0386 | -58.9417 | -59.3445 | -59.2987 | 0.054932 | 63.0981 | -59.1522 | 63.1073 | -59.0927 | -59.2209 | -58.9966 | -0.16479 | -0.045776 | -0.13733 | -0.091553 | 0.15106 | 0.31586 | -0.20599 | -0.25177 | -0.041199 | -0.086975 | -0.05951 | -0.077819 | -0.068664 | 0.009155 | 0 | 0.05951 | 0.09613 | -0.12817 | 0.15564 | -0.068664
626.7812 | 122.5296 | -59.0744 | 63.1256 | -0.069504 | 0.0013416 | 0.068545 | 0.067637 | -0.0010959 | -0.0036786 | -0.0048706 | 0.068878 | -0.00025411 | 0.065775 | 0.063192 | -0.0010361 | -0.0019803 | -0.00041228 | 0.036621 | 63.1027 | -59.2346 | 63.2996 | -58.9462 | -59.6375 | -59.317 | 0.064087 | 63.0615 | -59.1568 | 63.208 | -59.0286 | -59.1293 | -58.8821 | -0.16022 | -0.022888 | 0.17395 | 0.19684 | 0.12817 | 0.28839 | -0.24262 | -0.56305 | -0.082398 | -0.40283 | -0.082397 | -0.064087 | 0.082398 | 0.14648 | 0.045777 | 0.12817 | 0.19226 | -0.054931 | 0.27466 | 0.027466
627.7812 | 122.4518 | -59.1156 | 63.1622 | -0.054245 | -0.00058102 | 0.065555 | 0.07191 | -0.0015231 | 0.0031574 | 4.2808e-05 | -0.037933 | 0.0012411 | 0.065775 | 0.065755 | -0.0014633 | -0.0019803 | 0.0017239 | -0.013733 | 63.1714 | -59.2712 | 62.9837 | -59.0149 | -59.3948 | -59.3628 | 0.022888 | 63.1256 | -59.1705 | 63.1165 | -59.1064 | -59.2529 | -58.9645 | -0.15564 | 0.009156 | -0.17853 | -0.18768 | 0.10071 | 0.25635 | -0.24719 | -0.27923 | -0.091553 | -0.1236 | -0.054931 | -0.036621 | -0.045776 | -0.009155 | 0.009156 | 0.064087 | 0.15106 | -0.13733 | 0.20599 | -0.082398
628.7812 | 122.5571 | -59.0698 | 63.121 | -0.031357 | -0.00058102 | 0.067477 | 0.065287 | -0.0010959 | 0.0012348 | -0.00017099 | 0.056671 | -0.00025411 | 0.065775 | 0.062124 | -0.0021041 | -0.0043301 | 0.00065592 | 0.022888 | 63.0936 | -59.2255 | 63.1256 | -58.8547 | -59.2621 | -59.2117 | 0.12817 | 63.0753 | -59.1385 | 63.1989 | -58.9783 | -59.1248 | -58.8776 | -0.15564 | -0.027466 | 0.004577 | 0.032043 | 0.21515 | 0.37079 | -0.14191 | -0.19226 | 0.013733 | -0.036621 | -0.068665 | -0.045777 | 0.07782 | 0.1236 | 0.091553 | 0.16022 | 0.19226 | -0.054932 | 0.26093 | 0.013733
629.7812 | 122.4518 | -59.1339 | 63.1439 | -0.054245 | 0.0015552 | 0.067477 | 0.068919 | 0.0044583 | -0.001115 | 0.00047001 | -0.065399 | 0.0027365 | 0.065988 | 0.068746 | -0.0014633 | -0.0024075 | 0.0032193 | 0 | 63.1668 | -59.2941 | 63.0295 | -58.992 | -59.3719 | -59.3445 | 0.045776 | 63.121 | -59.1705 | 63.1485 | -59.0927 | -59.2163 | -58.9462 | -0.16022 | 0.022888 | -0.11444 | -0.13733 | 0.14191 | 0.30212 | -0.21057 | -0.23804 | -0.050354 | -0.07782 | -0.036621 | -0.022888 | 0.004578 | 0.027466 | 0.041199 | 0.07782 | 0.18768 | -0.082398 | 0.2243 | -0.045777
630.7812 | 122.4838 | -59.1202 | 63.1027 | -0.046616 | -0.00015382 | 0.067904 | 0.066783 | 0.0012539 | -0.00090143 | -0.0014526 | 0.032257 | 0.0010275 | 0.066843 | 0.062978 | 0.00088648 | -0.0019803 | 0.0017239 | -0.054932 | 63.0798 | -59.2758 | 63.0249 | -58.96 | -59.3536 | -59.2896 | 0.018311 | 63.089 | -59.1751 | 63.1714 | -59.0515 | -59.1751 | -58.9279 | -0.15564 | -0.022888 | -0.07782 | -0.054932 | 0.16022 | 0.31586 | -0.16937 | -0.23346 | -0.013733 | -0.07782 | -0.054932 | -0.013733 | 0.068665 | 0.082398 | 0.068664 | 0.1236 | 0.19226 | -0.054932 | 0.24719 | 0
631.7812 | 122.6028 | -59.0973 | 63.1348 | -0.10002 | 0.00070078 | 0.06705 | 0.065715 | -2.7692e-05 | 0.0012348 | 0.0019654 | 0.013947 | -0.00046791 | 0.064279 | 0.063619 | -0.0029587 | -0.0019803 | 0.0032193 | -0.018311 | 63.1119 | -59.2484 | 63.0661 | -58.9142 | -59.3079 | -59.2667 | 0.07782 | 63.0524 | -59.1568 | 63.1989 | -59.0103 | -59.1385 | -58.8913 | -0.15106 | -0.022889 | -0.068665 | -0.045776 | 0.1831 | 0.33417 | -0.16937 | -0.21057 | -0.018311 | -0.059509 | -0.059509 | -0.082398 | 0.064087 | 0.14648 | 0.086975 | 0.14648 | 0.20599 | -0.041199 | 0.2655 | 0.01831
632.7812 | 122.4609 | -59.1248 | 63.1577 | -0.00083923 | -0.00229 | 0.066623 | 0.068492 | 0.0046719 | -0.001115 | -0.00017099 | -0.019623 | 0.00081389 | 0.06492 | 0.06661 | -0.00060892 | -0.0026211 | -0.0031895 | -0.022888 | 63.1577 | -59.2758 | 62.9791 | -59.0378 | -59.3994 | -59.3491 | 0.013733 | 63.1393 | -59.1568 | 63.1897 | -59.0286 | -59.4589 | -59.1705 | -0.15106 | 0 | -0.17853 | -0.17853 | 0.086975 | 0.23804 | -0.2243 | -0.27466 | -0.073242 | -0.1236 | -0.032043 | -0.018311 | 0.032043 | 0.050354 | 0.096131 | 0.12817 | -0.045776 | -0.33417 | -0.013733 | -0.30212
633.7812 | 122.5342 | -59.1064 | 63.1348 | -0.061874 | 0.0017688 | 0.068973 | 0.067637 | -0.00024129 | 0.0046528 | 4.2808e-05 | 0.0078429 | 0.0014547 | 0.063852 | 0.063405 | 0.00088648 | -0.0026211 | 0.0030057 | 0.027466 | 63.1027 | -59.2117 | 63.0478 | -58.8043 | -59.3536 | -59.3033 | 0.11902 | 63.0707 | -59.1431 | 63.208 | -59.0103 | -59.1568 | -58.8913 | -0.10529 | -0.032044 | -0.086975 | -0.054931 | 0.30212 | 0.40741 | -0.19684 | -0.24719 | -0.091553 | -0.14191 | -0.036621 | -0.064087 | 0.073242 | 0.13733 | 0.09613 | 0.13275 | 0.21515 | -0.050354 | 0.25177 | -0.013733
634.7812 | 122.4747 | -59.111 | 63.1439 | -0.031357 | -0.00015382 | 0.067264 | 0.072551 | 0.0016813 | 0.00038017 | -0.0020936 | -0.092865 | 0.00038669 | 0.064706 | 0.07195 | -0.0023177 | -0.0026211 | 0.00044212 | 0.022888 | 63.1348 | -59.2575 | 63.0341 | -58.96 | -59.3811 | -59.3399 | 0.041199 | 63.1165 | -59.1522 | 63.1531 | -59.1156 | -59.2438 | -58.9554 | -0.14648 | -0.009155 | -0.10986 | -0.10071 | 0.15106 | 0.29755 | -0.22888 | -0.27008 | -0.082398 | -0.1236 | -0.041199 | -0.027466 | 0.009155 | 0.036621 | -0.004578 | 0.036621 | 0.15564 | -0.13275 | 0.19684 | -0.091552
635.7812 | 122.5067 | -59.0607 | 63.1027 | -0.046616 | 0.00048698 | 0.068759 | 0.065074 | 0.0012539 | 0.00016657 | -0.00081179 | 0.035309 | 0.00017309 | 0.063211 | 0.059347 | -0.0012497 | -0.0030483 | 0.0012967 | 0.036621 | 63.0981 | -59.2255 | 63.1027 | -58.8593 | -59.285 | -59.23 | 0.16022 | 63.0661 | -59.1522 | 63.121 | -59.1248 | -59.0698 | -58.8455 | -0.1648 | -0.004577 | 0 | 0.004577 | 0.20142 | 0.36621 | -0.16937 | -0.2243 | -0.004578 | -0.059509 | -0.091553 | -0.036621 | 0.018311 | 0.054932 | -0.064087 | 0.027466 | 0.21515 | -0.009155 | 0.3067 | 0.082398
636.7812 | 122.5159 | -59.0836 | 63.1989 | -0.077133 | -0.0010082 | 0.066836 | 0.067637 | 0.00039951 | -0.00026063 | 0.00089721 | 0.053619 | -0.00046791 | 0.065134 | 0.058492 | 0.00067288 | -0.00091214 | 0.00044212 | 0.041199 | 63.1302 | -59.2163 | 63.1256 | -58.8455 | -59.2529 | -59.2072 | 0.16479 | 63.0936 | -59.1064 | 63.2034 | -58.9828 | -59.1248 | -58.8776 | -0.13275 | -0.068665 | -0.073243 | -0.004578 | 0.23804 | 0.37079 | -0.1236 | -0.16937 | 0.009156 | -0.036621 | -0.022888 | -0.10529 | 0.004577 | 0.10986 | 0.10071 | 0.1236 | 0.20599 | -0.041199 | 0.22888 | -0.018311
637.7812 | 122.4792 | -59.1019 | 63.1531 | -0.046616 | -0.0010082 | 0.066623 | 0.069987 | -0.00066869 | 0.001662 | 0.0013244 | -0.071503 | 0.00060029 | 0.063425 | 0.075155 | -0.0038131 | -5.7742e-05 | -0.00019868 | 0.009155 | 63.1439 | -59.2712 | 63.0341 | -58.9874 | -59.3536 | -59.3262 | 0.041199 | 63.1073 | -59.1431 | 63.1714 | -59.0927 | -59.2392 | -58.96 | -0.16937 | -0.009155 | -0.11902 | -0.10986 | 0.11444 | 0.28381 | -0.2243 | -0.25177 | -0.054932 | -0.082398 | -0.041198 | -0.045776 | 0.018311 | 0.064087 | 0.009156 | 0.050354 | 0.14191 | -0.13733 | 0.1831 | -0.096131
638.7812 | 122.5525 | -59.0744 | 63.1439 | -0.046616 | 5.9776e-05 | 0.067904 | 0.066355 | -0.00088229 | 0.0010212 | -0.00059819 | -0.10202 | -4.0508e-05 | 0.064066 | 0.062551 | -0.0010361 | -0.0051845 | 0.0019375 | 0.032043 | 63.0981 | -59.2209 | 63.1256 | -58.8867 | -58.8776 | -59.1202 | 0.17395 | 63.0936 | -59.1064 | 63.1805 | -58.9966 | -59.1248 | -58.9096 | -0.14648 | -0.045776 | -0.018311 | 0.027465 | 0.18768 | 0.33417 | -0.045776 | 0.19684 | 0.10071 | 0.34332 | -0.032043 | -0.050354 | 0.036621 | 0.086975 | 0.07782 | 0.10986 | 0.16479 | -0.050354 | 0.19684 | -0.018311
639.7812 | 122.4289 | -59.1248 | 63.1805 | -0.038986 | -0.0010082 | 0.068118 | 0.065501 | -0.0034457 | -4.7032e-05 | -0.0016662 | -0.022675 | -4.0508e-05 | 0.065134 | 0.060415 | -0.0010361 | -0.0041165 | -0.00041228 | 0.013733 | 63.1348 | -59.2529 | 63.0157 | -58.9828 | -59.3719 | -59.3033 | 0.041199 | 63.1256 | -59.1339 | 63.1073 | -59.1064 | -59.2484 | -58.9874 | -0.12817 | -0.045776 | -0.1648 | -0.11902 | 0.14191 | 0.27008 | -0.17853 | -0.24719 | -0.050354 | -0.11902 | -0.009155 | -0.054932 | -0.073242 | -0.01831 | 0.018311 | 0.027466 | 0.13733 | -0.1236 | 0.14648 | -0.11444
640.7812 | 122.6532 | -59.0332 | 63.1897 | -0.10765 | -0.00058102 | 0.067477 | 0.066783 | -0.0015231 | 0.00059397 | -0.0025208 | 0.053619 | -0.00068151 | 0.06257 | 0.060201 | -0.0023177 | -0.0017667 | 0.00044212 | 0.032043 | 63.0844 | -59.2072 | 63.1302 | -58.8318 | -59.2575 | -59.1934 | 0.15106 | 63.0844 | -59.1248 | 63.2126 | -58.9828 | -59.1293 | -59.0332 | -0.17395 | -0.10529 | -0.059509 | 0.045776 | 0.20142 | 0.37537 | -0.16022 | -0.2243 | 0.013733 | -0.050354 | -0.091553 | -0.10529 | 0.022888 | 0.12817 | 0.050354 | 0.14191 | 0 | -0.09613 | 0.091553 | -0.004577
641.7812 | 122.5525 | -59.0881 | 63.1256 | -0.084763 | 0.00070078 | 0.066409 | 0.066783 | 0.00061311 | 0.00038017 | 0.001538 | 0.02005 | -0.0013223 | 0.064706 | 0.060842 | -0.0021041 | -0.0043301 | 0.00044212 | 0.013733 | 63.1027 | -59.2255 | 63.121 | -58.8547 | -59.2712 | -59.2346 | 0.16022 | 63.0936 | -59.1202 | 63.2401 | -58.9966 | -59.111 | -58.8684 | -0.13733 | -0.022888 | -0.004577 | 0.018311 | 0.23346 | 0.37079 | -0.14648 | -0.1831 | -0.009155 | -0.045776 | -0.032043 | -0.032043 | 0.11444 | 0.14648 | 0.091553 | 0.1236 | 0.21973 | -0.022888 | 0.25177 | 0.009155
642.7812 | 122.5388 | -59.0652 | 63.1622 | -0.10002 | 0.001128 | 0.066836 | 0.069346 | -0.0025913 | 0.00059397 | -0.0020936 | 0.1116 | -0.0011087 | 0.06492 | 0.065969 | -0.0016769 | -0.0039029 | -0.0046847 | 0.036621 | 63.1073 | -59.2117 | 63.121 | -58.8776 | -59.2804 | -59.2255 | 0.35248 | 63.1439 | -59.1156 | 63.1805 | -58.9645 | -59.1156 | -58.873 | -0.14648 | -0.054931 | -0.041198 | 0.013733 | 0.18768 | 0.33417 | -0.16022 | -0.21515 | -0.013733 | -0.068665 | -0.050354 | -0.01831 | 0.018311 | 0.036621 | 0.10071 | 0.15106 | 0.19226 | -0.050354 | 0.24262 | 0
643.7812 | 122.5296 | -59.079 | 63.1256 | -0.092392 | 5.9776e-05 | 0.067477 | 0.068492 | -0.0010959 | 0.0018756 | 0.00047001 | 0.010895 | -0.00068151 | 0.063211 | 0.061056 | -0.0010361 | -0.0026211 | -0.0014805 | -0.009155 | 63.0936 | -59.2575 | 63.0798 | -58.9142 | -59.3079 | -59.2346 | 0.068665 | 63.089 | -59.1248 | 63.2309 | -59.0057 | -59.1476 | -58.8821 | -0.17853 | -0.032043 | -0.045776 | -0.013733 | 0.16479 | 0.34332 | -0.15564 | -0.22888 | 0.022888 | -0.050354 | -0.045777 | -0.036621 | 0.10529 | 0.14191 | 0.073242 | 0.11902 | 0.19684 | -0.068665 | 0.24262 | -0.022888
644.7812 | 122.4243 | -59.0973 | 63.1668 | -0.17632 | -0.00015382 | 0.067477 | 0.071482 | 0.0012539 | -4.7032e-05 | 0.00025641 | -0.071503 | 0.0029501 | 0.063638 | 0.067464 | -0.0016769 | -0.0032619 | -0.0012667 | -0.013733 | 63.1302 | -59.2758 | 63.0157 | -58.9966 | -59.4131 | -59.3628 | -0.009155 | 63.1577 | -59.1293 | 63.1393 | -59.0973 | -59.2392 | -58.9645 | -0.17853 | -0.036621 | -0.15106 | -0.11444 | 0.10071 | 0.27924 | -0.2655 | -0.31586 | -0.086975 | -0.13733 | -0.032043 | -0.009155 | -0.027466 | -0.018311 | 0 | 0.032043 | 0.13275 | -0.14191 | 0.16479 | -0.10986
645.7812 | 122.5021 | -59.1064 | 63.1302 | -0.054245 | -0.001222 | 0.065982 | 0.066569 | 0.00018591 | -0.00090143 | 0.00089721 | 0.029205 | -0.0019631 | 0.065561 | 0.064901 | -0.0010361 | -0.0045437 | 0.00065592 | -0.068665 | 63.1027 | -59.2392 | 63.0569 | -58.9508 | -59.3399 | -59.2712 | 0.022888 | 63.0798 | -59.1385 | 63.1943 | -59.0332 | -59.1614 | -58.9005 | -0.13275 | -0.027466 | -0.073242 | -0.045776 | 0.15564 | 0.28839 | -0.16479 | -0.23346 | -0.032043 | -0.10071 | -0.032044 | -0.050354 | 0.064087 | 0.11444 | 0.073242 | 0.10529 | 0.20599 | -0.054932 | 0.23804 | -0.022888
646.7812 | 122.5662 | -59.0698 | 63.1165 | -0.11528 | -0.00058102 | 0.068332 | 0.065501 | 0.0010403 | 0.0029438 | -0.0025208 | 0.02005 | 0.00017309 | 0.065561 | 0.063619 | -0.0014633 | -0.0049709 | -0.0019077 | 0.032043 | 63.1165 | -59.2346 | 63.1393 | -58.8226 | -59.2529 | -59.1934 | 0.13275 | 63.0936 | -59.1202 | 63.2172 | -58.992 | -59.0927 | -58.8638 | -0.1648 | 0 | 0.022888 | 0.022888 | 0.24719 | 0.41199 | -0.1236 | -0.18311 | 0.041199 | -0.018311 | -0.050354 | -0.022888 | 0.10071 | 0.1236 | 0.07782 | 0.12817 | 0.20599 | -0.022888 | 0.25635 | 0.027466
647.7812 | 122.5479 | -59.0515 | 63.1119 | -0.10765 | 5.9776e-05 | 0.068118 | 0.068705 | 0.00039951 | 0.0014484 | -0.0025208 | -0.034882 | -0.00046791 | 0.063211 | 0.066396 | -0.00039512 | -0.0041165 | 0.00065592 | 0.045776 | 63.0981 | -59.2117 | 63.1165 | -58.5526 | -59.111 | -59.4086 | 0.19226 | 63.089 | -59.1339 | 63.2126 | -58.9828 | -59.1476 | -58.8959 | -0.16022 | -0.013732 | 0.004578 | 0.01831 | 0.49896 | 0.65918 | -0.35706 | -0.059509 | -0.19684 | 0.10071 | -0.082397 | -0.022888 | 0.10071 | 0.1236 | 0.068665 | 0.15106 | 0.15564 | -0.09613 | 0.23804 | -0.013733
648.7812 | 122.5204 | -59.079 | 63.1119 | -0.1458 | -0.0025036 | 0.068332 | 0.068705 | 0.0010403 | 0.00016657 | -0.00017099 | -0.010467 | -0.0013223 | 0.064066 | 0.070028 | -0.0014633 | -0.00091214 | -0.0025485 | -0.054932 | 63.0981 | -59.2575 | 63.0386 | -58.96 | -59.3536 | -59.2667 | -0.009155 | 63.089 | -59.1522 | 63.1577 | -59.0424 | -59.1888 | -58.9142 | -0.17853 | -0.013732 | -0.073242 | -0.05951 | 0.11902 | 0.29755 | -0.18768 | -0.27466 | -0.009156 | -0.096131 | -0.073243 | -0.022888 | 0.045777 | 0.068665 | 0.036621 | 0.10986 | 0.16479 | -0.10986 | 0.23804 | -0.036621
649.7812 | 122.5662 | -59.0561 | 63.121 | -0.077133 | -0.0014356 | 0.068332 | 0.069987 | -0.0036593 | 0.0027302 | 0.0011108 | 0.0047913 | 0.00017309 | 0.062357 | 0.066823 | -0.0029587 | -0.0030483 | -0.0016941 | 0 | 63.1027 | -59.2209 | 63.1302 | -58.8638 | -59.2758 | -59.2026 | 0.13733 | 63.0936 | -59.1248 | 63.2446 | -58.9645 | -59.0973 | -58.8455 | -0.16479 | -0.018311 | 0.009155 | 0.027466 | 0.19226 | 0.35706 | -0.14648 | -0.21973 | 0.01831 | -0.054932 | -0.068665 | -0.027466 | 0.1236 | 0.15106 | 0.091552 | 0.16022 | 0.21057 | -0.041199 | 0.27924 | 0.027466
650.7812 | 122.5067 | -59.0927 | 63.1119 | -0.13054 | 5.9776e-05 | 0.067477 | 0.067637 | -0.0043003 | 0.00038017 | -0.0027344 | 0.0017395 | -0.00089511 | 0.062997 | 0.069814 | -0.00060892 | -0.0015529 | 0.00022852 | -0.022888 | 63.0981 | -59.2346 | 63.0524 | -58.9554 | -59.3582 | -59.2575 | -0.009155 | 63.0798 | -59.1339 | 63.1668 | -59.0378 | -59.1934 | -58.9142 | -0.14191 | -0.013732 | -0.059509 | -0.045777 | 0.13733 | 0.27924 | -0.16479 | -0.2655 | -0.022888 | -0.1236 | -0.041199 | -0.032043 | 0.054932 | 0.086975 | 0.054931 | 0.09613 | 0.17853 | -0.10071 | 0.21973 | -0.059509
651.7812 | 122.4792 | -59.0973 | 63.1073 | -0.10765 | 0.001128 | 0.070041 | 0.067424 | -0.0015231 | -0.00047423 | 0.00089721 | -0.077606 | 0.0031637 | 0.064279 | 0.063833 | -0.0023177 | -0.0021939 | -0.0012667 | -0.036621 | 63.0753 | -59.2575 | 63.0798 | -58.905 | -59.3399 | -59.2896 | 0.022888 | 63.0569 | -59.1339 | 63.2034 | -59.0561 | -59.1934 | -58.905 | -0.16022 | -0.032044 | -0.027466 | 0.004578 | 0.19226 | 0.35248 | -0.19226 | -0.24262 | -0.032044 | -0.082398 | -0.036621 | -0.050354 | 0.09613 | 0.14648 | 0.041199 | 0.07782 | 0.19226 | -0.09613 | 0.22888 | -0.059509
652.7812 | 122.5067 | -59.0973 | 63.1073 | -0.12291 | 5.9776e-05 | 0.069186 | 0.066142 | 0.00039951 | 0.00016657 | 0.002179 | -0.056244 | 0.0010275 | 0.06492 | 0.065542 | -0.00082252 | -0.0026211 | -0.0025485 | -0.054932 | 63.0661 | -59.2255 | 63.0844 | -58.905 | -59.3857 | -59.2896 | 0 | 63.0798 | -59.1614 | 63.1851 | -59.0607 | -59.2026 | -58.9279 | -0.12817 | -0.041199 | -0.022888 | 0.018311 | 0.19226 | 0.32044 | -0.19226 | -0.28839 | -0.064087 | -0.16022 | -0.064087 | -0.027466 | 0.07782 | 0.10529 | 0.036621 | 0.10071 | 0.16937 | -0.10529 | 0.23346 | -0.041199
653.7812 | 122.5113 | -59.0881 | 63.1073 | -0.12291 | 5.9776e-05 | 0.068545 | 0.068919 | -0.0010959 | 0.0014484 | -0.00059819 | -0.03183 | 0.00017309 | 0.062997 | 0.061483 | -0.00060892 | -0.0030483 | -0.0044711 | -0.07782 | 63.0569 | -59.2392 | 63.0478 | -58.9417 | -59.3719 | -59.2804 | -0.045776 | 63.089 | -59.1339 | 63.176 | -59.0652 | -59.1843 | -58.9279 | -0.15106 | -0.050354 | -0.059509 | -0.009155 | 0.14648 | 0.29755 | -0.19226 | -0.28381 | -0.041199 | -0.13275 | -0.045776 | -0.018311 | 0.068664 | 0.086975 | 0.022888 | 0.068664 | 0.16022 | -0.09613 | 0.20599 | -0.050354
654.7812 | 122.4838 | -59.079 | 63.1439 | -0.092392 | -0.00229 | 0.069186 | 0.065501 | -0.0025913 | 0.0018756 | 0.00089721 | -0.019623 | -0.0011087 | 0.062784 | 0.059987 | 0.00088648 | -0.0034757 | -0.0023349 | -0.073242 | 63.1073 | -59.2438 | 63.0157 | -58.9645 | -59.3399 | -59.285 | -0.041199 | 63.1073 | -59.1522 | 63.1805 | -59.0469 | -59.1888 | -58.9096 | -0.1648 | -0.036621 | -0.12817 | -0.091553 | 0.11444 | 0.27923 | -0.20599 | -0.26093 | -0.041199 | -0.096131 | -0.073243 | -0.036621 | 0.036621 | 0.073242 | 0.032043 | 0.10529 | 0.16937 | -0.10986 | 0.24262 | -0.036621
655.7812 | 122.5342 | -59.1019 | 63.1073 | -0.092392 | -0.00079462 | 0.066409 | 0.068492 | -0.0030185 | -0.00026063 | 0.0011108 | -0.0043641 | -0.00046791 | 0.063211 | 0.06191 | 0.0011001 | -0.0058255 | -0.0023349 | -0.07782 | 63.1073 | -59.2575 | 63.0432 | -58.9554 | -59.3719 | -59.2941 | -0.041199 | 63.1119 | -59.1385 | 63.1805 | -59.0469 | -59.198 | -58.9233 | -0.15564 | 0 | -0.064087 | -0.064087 | 0.14649 | 0.30212 | -0.19226 | -0.27008 | -0.036621 | -0.11444 | -0.036621 | 0.004577 | 0.073242 | 0.068665 | 0.054932 | 0.091553 | 0.17853 | -0.09613 | 0.21515 | -0.059509
656.7812 | 122.6074 | -59.1019 | 63.0936 | -0.12291 | -0.00079462 | 0.06705 | 0.068705 | 0.00039951 | 0.0025166 | -0.00038459 | -0.086761 | 0.0014547 | 0.06257 | 0.070241 | -0.0031723 | -0.00048494 | -0.0016941 | -0.059509 | 63.0386 | -59.2484 | 63.0981 | -58.9233 | -59.3628 | -59.285 | 0.004578 | 63.089 | -59.1156 | 63.2309 | -59.2392 | -59.1522 | -58.7906 | -0.14648 | -0.054932 | 0.004578 | 0.05951 | 0.17853 | 0.32501 | -0.1831 | -0.26093 | -0.036621 | -0.11444 | -0.013733 | -0.004578 | 0.13733 | 0.14191 | -0.13733 | -0.1236 | 0.31128 | -0.050354 | 0.32501 | -0.036621
657.7812 | 122.5525 | -59.079 | 63.0798 | -0.10002 | 0.0013416 | 0.069613 | 0.068278 | 0.0012539 | 0.0012348 | 0.0019654 | -0.098968 | -4.0508e-05 | 0.063211 | 0.059987 | -0.0016769 | -0.0021939 | -0.0023349 | -0.045776 | 63.0661 | -59.2712 | 63.0798 | -58.9096 | -59.3445 | -59.4406 | 0.17395 | 63.0936 | -59.1293 | 63.2172 | -59.0103 | -59.166 | -58.8867 | -0.19226 | -0.013733 | 0 | 0.013733 | 0.16937 | 0.36163 | -0.36163 | -0.2655 | -0.16937 | -0.073242 | -0.050354 | 0.013733 | 0.13733 | 0.1236 | 0.068664 | 0.11902 | 0.19226 | -0.086976 | 0.24261 | -0.036622
658.7812 | 122.5159 | -59.0881 | 63.0981 | -0.069504 | -0.0018628 | 0.070041 | 0.068064 | 0.0012539 | 0.00059397 | 0.0017518 | -0.068451 | 0.00038669 | 0.065561 | 0.064901 | -0.0016769 | -0.0019803 | 0.0010831 | -0.073242 | 63.089 | -59.2712 | 63.0753 | -58.9188 | -59.3491 | -59.2758 | -0.022888 | 63.0478 | -59.1385 | 63.2401 | -59.0195 | -59.198 | -58.8776 | -0.1831 | -0.009156 | -0.022889 | -0.013733 | 0.16937 | 0.35248 | -0.18768 | -0.26093 | -0.004578 | -0.07782 | -0.050354 | -0.050354 | 0.14191 | 0.19226 | 0.068665 | 0.11902 | 0.21057 | -0.10986 | 0.26093 | -0.059509
659.7812 | 122.4747 | -59.1019 | 63.1027 | -0.13817 | -0.00036742 | 0.067691 | 0.067851 | -0.00066869 | -0.0015424 | -0.002948 | -0.05014 | 0.0023093 | 0.063852 | 0.064901 | -0.00082252 | -0.0036893 | -0.0034031 | -0.09613 | 63.0569 | -59.2621 | 63.0478 | -58.9417 | -59.3628 | -59.2987 | -0.068665 | 63.0981 | -59.1614 | 63.1714 | -59.0607 | -59.2117 | -58.9233 | -0.16022 | -0.045776 | -0.054931 | -0.009155 | 0.16022 | 0.32043 | -0.19684 | -0.26093 | -0.036621 | -0.10071 | -0.059509 | -0.004577 | 0.068665 | 0.073242 | 0.041199 | 0.10071 | 0.17853 | -0.10986 | 0.23804 | -0.050354
660.7812 | 122.4197 | -59.0973 | 63.1119 | -0.084763 | -0.00036742 | 0.068545 | 0.070201 | 0.0023221 | 0.0012348 | 0.0034606 | -0.12643 | 0.0010275 | 0.065134 | 0.068105 | -0.0016769 | -0.0030483 | 0.0012967 | -0.09613 | 63.1256 | -59.2758 | 62.9929 | -59.0286 | -59.436 | -59.3719 | -0.1236 | 63.1165 | -59.1614 | 63.1577 | -59.1156 | -59.2896 | -58.9737 | -0.17853 | 0.013733 | -0.11902 | -0.13275 | 0.068665 | 0.24719 | -0.27466 | -0.33874 | -0.09613 | -0.16022 | -0.064087 | 0.004578 | 0.045777 | 0.041199 | -0.018311 | 0.045776 | 0.1236 | -0.19226 | 0.18768 | -0.12817
661.7812 | 122.4518 | -59.1156 | 63.1302 | -0.10002 | 0.0017688 | 0.067691 | 0.070414 | 0.0038175 | 0.0014484 | 0.0011108 | -0.074554 | 0.0016685 | 0.062784 | 0.063619 | 0.00024568 | -0.0017667 | -0.0010531 | -0.1236 | 63.1256 | -59.2896 | 62.9837 | -58.7769 | -59.3719 | -59.4269 | -0.1236 | 63.1348 | -59.1476 | 63.1348 | -59.1385 | -59.3079 | -58.9966 | -0.17395 | -0.004578 | -0.14648 | -0.14191 | 0.33875 | 0.5127 | -0.31128 | -0.25635 | -0.13733 | -0.082397 | -0.032043 | 0.004578 | 0.004578 | 0 | -0.022888 | 0.009155 | 0.11902 | -0.19226 | 0.15106 | -0.16022
662.7812 | 122.4884 | -59.079 | 63.1165 | -0.077133 | 0.00027338 | 0.068545 | 0.067851 | -0.0028049 | -0.00047423 | 0.00068361 | -0.0043641 | 0.00038669 | 0.064706 | 0.058706 | 0.00067288 | -0.0028347 | -0.0036167 | -0.091553 | 63.0798 | -59.2163 | 63.0249 | -58.9371 | -59.3536 | -59.285 | -0.068665 | 63.0524 | -59.1476 | 63.1851 | -59.0607 | -59.1843 | -58.905 | -0.13733 | -0.036621 | -0.091553 | -0.054932 | 0.14191 | 0.27924 | -0.20599 | -0.27466 | -0.068664 | -0.13733 | -0.068665 | -0.064087 | 0.068665 | 0.13275 | 0.01831 | 0.086975 | 0.17395 | -0.10529 | 0.24262 | -0.036621
663.7812 | 122.4197 | -59.0881 | 63.1622 | -0.077133 | 0.002837 | 0.067691 | 0.071269 | 0.00039951 | 0.0012348 | 0.0013244 | -0.12338 | 0.0012411 | 0.062784 | 0.067251 | -0.0014633 | -0.00091214 | 0.0040737 | -0.073242 | 63.1485 | -59.2712 | 62.9883 | -59.0195 | -59.4452 | -59.3765 | -0.12817 | 63.1165 | -59.1797 | 63.1119 | -59.1705 | -59.2987 | -58.9966 | -0.1831 | -0.013732 | -0.17395 | -0.16022 | 0.068665 | 0.25177 | -0.28839 | -0.35705 | -0.10529 | -0.17395 | -0.091552 | -0.045776 | -0.050354 | -0.004578 | -0.082397 | 0.009155 | 0.091553 | -0.21057 | 0.18311 | -0.11902
664.7812 | 122.4564 | -59.0836 | 63.1302 | -0.26024 | -0.0031446 | 0.065127 | 0.068278 | -0.0025913 | 0.0027302 | 0.0028198 | -0.08371 | 0.00060029 | 0.065134 | 0.066396 | 0.0013139 | -0.0030483 | -0.0040439 | 0.045776 | 63.1119 | -59.2712 | 62.9745 | -59.0195 | -59.436 | -59.3903 | -0.11444 | 63.1256 | -59.1797 | 63.1485 | -59.1339 | -59.2758 | -58.9645 | -0.18768 | -0.018311 | -0.15564 | -0.13733 | 0.064087 | 0.25177 | -0.3067 | -0.35248 | -0.11902 | -0.16479 | -0.09613 | -0.004578 | 0.018311 | 0.022889 | -0.050354 | 0.045776 | 0.11902 | -0.19226 | 0.21515 | -0.096131
665.7812 | 122.4701 | -59.0744 | 63.1393 | -0.084763 | 0.00048698 | 0.068332 | 0.068278 | 0.0036039 | -0.001115 | 0.00025641 | -0.089813 | 0.0016685 | 0.065775 | 0.068746 | -0.00060892 | 0.0016513 | -0.00019868 | -0.013733 | 63.1577 | -59.2392 | 63.0386 | -58.9233 | -59.3536 | -59.3079 | 0.036621 | 63.1348 | -59.1431 | 63.0615 | -58.9508 | -59.2255 | -58.9417 | -0.16479 | 0.018311 | -0.10071 | -0.11902 | 0.15106 | 0.31586 | -0.23346 | -0.27924 | -0.068664 | -0.11444 | -0.068664 | -0.004577 | -0.07782 | -0.073243 | 0.1236 | 0.19226 | 0.13275 | -0.15106 | 0.20142 | -0.082398
666.7812 | 122.4426 | -59.1064 | 63.1531 | -0.046616 | 0.00027338 | 0.071109 | 0.068919 | 0.0014675 | 0.00016657 | 4.2808e-05 | -0.089813 | 0.0010275 | 0.064279 | 0.063619 | 0.0011001 | -0.0021939 | -0.00019868 | -0.004578 | 63.1485 | -59.2438 | 63.0569 | -58.9325 | -59.3399 | -59.2941 | 0.018311 | 63.1439 | -59.1248 | 63.1668 | -59.0744 | -59.2346 | -58.9325 | -0.13733 | -0.004577 | -0.09613 | -0.091553 | 0.17395 | 0.31128 | -0.18768 | -0.23346 | -0.050354 | -0.096131 | -0.018311 | -0.009155 | 0.013733 | 0.022888 | 0.032043 | 0.050354 | 0.17395 | -0.12817 | 0.19226 | -0.10986
667.7812 | 122.3785 | -59.1156 | 63.1577 | -0.11528 | -0.001222 | 0.068759 | 0.071269 | 0.0023221 | 0.00059397 | 4.2808e-05 | -0.086761 | 0.0016685 | 0.066415 | 0.060201 | 0.00067288 | -0.00091214 | -0.0051121 | -0.07782 | 63.1622 | -59.2667 | 62.9791 | -59.024 | -59.436 | -59.3857 | -0.14648 | 63.1302 | -59.1522 | 63.1073 | -59.1568 | -59.3079 | -59.0057 | -0.15106 | 0.004577 | -0.17853 | -0.1831 | 0.091553 | 0.24262 | -0.27008 | -0.32043 | -0.11902 | -0.16937 | -0.036621 | -0.027466 | -0.050354 | -0.022888 | -0.041198 | -0.004577 | 0.10986 | -0.19226 | 0.14648 | -0.15564
668.7812 | 122.4335 | -59.0927 | 63.1622 | -0.10765 | 0.00070078 | 0.069186 | 0.071055 | 0.0012539 | 0.00059397 | 0.00047001 | -0.086761 | 0.00081389 | 0.065561 | 0.064046 | 0.0011001 | -0.0026211 | -0.00041228 | -0.022888 | 63.1485 | -59.2346 | 63.0569 | -58.9233 | -59.3674 | -59.3216 | -0.004578 | 63.1165 | -59.1156 | 63.1714 | -59.0836 | -59.2209 | -58.9325 | -0.14191 | -0.013732 | -0.10528 | -0.091553 | 0.16937 | 0.31128 | -0.22888 | -0.27466 | -0.086975 | -0.13275 | -0.022889 | -0.045776 | 0.009156 | 0.054932 | 0.009155 | 0.032044 | 0.16022 | -0.12817 | 0.18311 | -0.10529
669.7812 | 122.4014 | -59.1019 | 63.1897 | -0.12291 | 0.00048698 | 0.067691 | 0.068064 | -0.0021641 | 0.00080757 | -0.00017099 | -0.062347 | -0.00046791 | 0.065134 | 0.069814 | 0.00045928 | -0.0013393 | -0.0010531 | -0.050354 | 63.1622 | -59.2392 | 63.0203 | -58.9691 | -59.3857 | -59.3445 | -0.054932 | 63.1256 | -59.1248 | 63.1256 | -59.1202 | -59.2621 | -58.9508 | -0.13733 | -0.027466 | -0.16937 | -0.14191 | 0.13275 | 0.27008 | -0.24261 | -0.28381 | -0.10529 | -0.14648 | -0.022888 | -0.064087 | -0.064087 | 0 | -0.01831 | 0.004578 | 0.15106 | -0.16022 | 0.17395 | -0.13733
670.7812 | 122.4518 | -59.0515 | 63.1531 | -0.046616 | 0.00048698 | 0.067691 | 0.068278 | 0.0021085 | 0.00080757 | -0.00017099 | -0.068451 | 0.0018821 | 0.067056 | 0.064687 | 0.0013139 | 0.0025059 | 0.0023647 | -0.009155 | 63.1805 | -59.23 | 63.0569 | -58.905 | -59.3445 | -59.2941 | 0.018311 | 63.1348 | -59.111 | 63.0936 | -59.0378 | -59.2163 | -58.905 | -0.17853 | 0.027466 | -0.09613 | -0.1236 | 0.14648 | 0.32501 | -0.24261 | -0.29297 | -0.064086 | -0.11444 | -0.059509 | -0.01831 | -0.059509 | -0.041199 | 0.013733 | 0.073242 | 0.14648 | -0.1648 | 0.20599 | -0.10529
671.7812 | 122.4197 | -59.079 | 63.176 | -0.18394 | 0.0036914 | 0.066195 | 0.067637 | 0.0012539 | 0.00059397 | -0.00038459 | -0.074554 | 0.0016685 | 0.064493 | 0.063405 | 0.00045928 | -0.00027134 | 0.00086952 | -0.009155 | 63.1439 | -59.2209 | 63.0478 | -58.9188 | -59.3582 | -59.3124 | 0.018311 | 63.1439 | -59.111 | 63.1805 | -59.0836 | -59.2117 | -58.8959 | -0.14191 | -0.032043 | -0.12817 | -0.09613 | 0.16022 | 0.30212 | -0.23346 | -0.27924 | -0.091553 | -0.13733 | -0.032044 | -0.032043 | 0.004578 | 0.036621 | -0.004578 | 0.027466 | 0.1831 | -0.13275 | 0.21515 | -0.10071
672.7812 | 122.4289 | -59.079 | 63.1851 | -0.10002 | 0.0015552 | 0.068973 | 0.069773 | 0.0021085 | 0.0027302 | 0.0017518 | -0.047089 | 0.0029501 | 0.06492 | 0.058278 | 0.0011001 | -0.0024075 | 0.0032193 | -0.032043 | 63.176 | -59.2117 | 63.0203 | -58.9645 | -59.3903 | -59.3399 | -0.045776 | 63.1485 | -59.1385 | 63.1485 | -59.111 | -59.2621 | -58.9462 | -0.13275 | -0.009156 | -0.16479 | -0.15564 | 0.11444 | 0.24719 | -0.26093 | -0.31128 | -0.12817 | -0.17853 | -0.05951 | -0.036621 | -0.036621 | 0 | -0.032044 | 0.027466 | 0.13275 | -0.18311 | 0.19226 | -0.1236
673.7812 | 122.4472 | -59.0836 | 63.1851 | -0.023727 | 0.00048698 | 0.066623 | 0.068919 | -2.7692e-05 | 0.0014484 | -0.00038459 | -0.025726 | -0.00068151 | 0.066415 | 0.062551 | -0.0014633 | -0.0030483 | 0.0021511 | -0.009155 | 63.1714 | -59.2346 | 63.0432 | -58.9188 | -59.3308 | -59.2758 | 0.041199 | 63.1348 | -59.1248 | 63.1943 | -59.079 | -59.1705 | -58.8547 | -0.15106 | -0.013733 | -0.14191 | -0.12817 | 0.16479 | 0.31586 | -0.19226 | -0.24719 | -0.041199 | -0.096131 | -0.041199 | -0.050354 | 0.009155 | 0.059509 | 0.004578 | 0.045777 | 0.22888 | -0.086975 | 0.27008 | -0.045776
674.7812 | 122.4152 | -59.0881 | 63.1943 | -0.084763 | 5.9776e-05 | 0.070682 | 0.068919 | 0.0027493 | 0.0025166 | -0.00081179 | -0.028778 | 0.00017309 | 0.065134 | 0.063833 | 0.00024568 | -0.0019803 | 0.00044212 | -0.11444 | 63.1805 | -59.2621 | 62.9517 | -59.0286 | -59.4406 | -59.3719 | -0.15106 | 63.1393 | -59.1476 | 63.1897 | -58.9645 | -59.3308 | -59.0195 | -0.17395 | -0.013733 | -0.24262 | -0.22888 | 0.05951 | 0.23346 | -0.28381 | -0.35248 | -0.10986 | -0.17853 | -0.059509 | -0.054932 | -0.004578 | 0.050354 | 0.1236 | 0.1831 | 0.068665 | -0.24262 | 0.12817 | -0.18311
675.7812 | 122.4014 | -59.0836 | 63.1714 | -0.10002 | 0.0013416 | 0.068332 | 0.06956 | 0.00039951 | 0.0025166 | 0.0019654 | -0.028778 | 0.0033773 | 0.065561 | 0.062978 | 3.2078e-05 | 0.00079686 | 0.0030057 | -0.050354 | 63.1714 | -59.2255 | 63.0112 | -58.9737 | -59.3948 | -59.3308 | -0.086975 | 63.1348 | -59.1431 | 63.1393 | -59.1202 | -59.2667 | -58.9645 | -0.14191 | 0 | -0.16022 | -0.16022 | 0.10986 | 0.25177 | -0.24719 | -0.31128 | -0.10529 | -0.16937 | -0.059509 | -0.036621 | -0.032044 | 0.004577 | -0.036621 | 0.022888 | 0.11902 | -0.18311 | 0.17853 | -0.1236
676.7812 | 122.438 | -59.0515 | 63.1943 | -0.054245 | -0.0014356 | 0.068973 | 0.066355 | 0.0010403 | -0.00026063 | -0.00038459 | -0.053192 | 0.00060029 | 0.065561 | 0.066182 | -0.0010361 | -0.0013393 | -0.0025485 | 0.009155 | 63.1714 | -59.2026 | 63.0615 | -58.9096 | -59.2987 | -59.2804 | 0.041199 | 63.1714 | -59.0927 | 63.1851 | -59.0332 | -59.1934 | -58.9142 | -0.15106 | -0.022888 | -0.13275 | -0.10986 | 0.14191 | 0.29297 | -0.22888 | -0.24719 | -0.07782 | -0.09613 | -0.041198 | -0.022888 | -0.009155 | 0.013733 | 0.018311 | 0.059509 | 0.13733 | -0.14191 | 0.17853 | -0.10071
677.7812 | 122.4655 | -59.0744 | 63.1622 | -0.046616 | 0.00070078 | 0.067904 | 0.066355 | -0.00088229 | 0.00038017 | 0.002179 | -0.025726 | 0.0018821 | 0.067056 | 0.063833 | 0.00045928 | -0.00048494 | 0.0034329 | -0.09613 | 63.1531 | -59.2484 | 62.9883 | -58.9783 | -59.4223 | -59.3216 | -0.10529 | 63.1027 | -59.1476 | 63.1668 | -59.0927 | -59.2209 | -58.9142 | -0.17395 | -0.009155 | -0.17395 | -0.16479 | 0.096131 | 0.27008 | -0.24719 | -0.3479 | -0.073242 | -0.17395 | -0.073242 | -0.059509 | 0.004578 | 0.064087 | -0.01831 | 0.054932 | 0.16022 | -0.14648 | 0.23346 | -0.073242
678.7812 | 122.4701 | -59.0607 | 63.208 | -0.046616 | -0.00015382 | 0.066836 | 0.066996 | -0.0021641 | 0.0025166 | 0.0013244 | -0.016571 | 0.00060029 | 0.066629 | 0.060628 | -0.00018152 | -0.0039029 | -0.0014805 | -0.018311 | 63.1943 | -59.2026 | 63.0569 | -58.9462 | -59.3124 | -59.2667 | 0.018311 | 63.1714 | -59.0973 | 63.1989 | -59.0286 | -59.1614 | -58.8959 | -0.14191 | -0.013733 | -0.15106 | -0.13733 | 0.11444 | 0.25635 | -0.20599 | -0.25177 | -0.064087 | -0.10986 | -0.036621 | -0.036621 | -0.009155 | 0.027466 | 0.032044 | 0.068665 | 0.16479 | -0.10071 | 0.20142 | -0.064087
679.7812 | 122.4197 | -59.0652 | 63.1897 | -0.13054 | 0.00027338 | 0.06705 | 0.066569 | 0.0010403 | 0.0037982 | 0.00047001 | 0.0078429 | 0.0020957 | 0.067697 | 0.06426 | -0.00039512 | -0.0036893 | -0.00083948 | -0.032043 | 63.2126 | -59.1934 | 62.9608 | -58.905 | -59.3628 | -59.3216 | -0.022888 | 63.1805 | -59.111 | 63.1393 | -59.0607 | -59.2026 | -58.9417 | -0.12817 | 0.022888 | -0.22888 | -0.25177 | 0.16022 | 0.28839 | -0.25635 | -0.29755 | -0.12817 | -0.16937 | -0.045776 | -0.009155 | -0.050354 | -0.041199 | 0.004578 | 0.050354 | 0.1236 | -0.13733 | 0.16937 | -0.091553
680.7812 | 122.4838 | -59.0469 | 63.2355 | -0.077133 | -0.0029308 | 0.068759 | 0.06956 | -0.00045509 | 0.003371 | 0.00089721 | -0.013519 | -0.00025411 | 0.065134 | 0.06661 | -0.0018905 | -0.00069854 | -0.00041228 | -0.013733 | 63.2034 | -59.198 | 63.0478 | -58.96 | -59.3399 | -59.2758 | -0.009155 | 63.176 | -59.0927 | 63.176 | -59.0469 | -59.1751 | -58.8959 | -0.15106 | -0.032044 | -0.18768 | -0.15564 | 0.086975 | 0.23804 | -0.22888 | -0.29297 | -0.07782 | -0.14191 | -0.045776 | -0.05951 | -0.05951 | 0 | 0 | 0.045776 | 0.15106 | -0.12817 | 0.19684 | -0.082398
681.7812 | 122.4838 | -59.0698 | 63.2034 | -0.077133 | -0.00015382 | 0.066195 | 0.068278 | 0.0036039 | 0.0012348 | 0.002179 | -0.040985 | 0.0023093 | 0.062357 | 0.063833 | -0.0014633 | -0.0019803 | 0.0032193 | 0 | 63.1943 | -59.2163 | 63.0707 | -58.9142 | -59.3124 | -59.2712 | 0.045776 | 63.1714 | -59.0973 | 63.2034 | -59.0424 | -59.1705 | -58.8821 | -0.14649 | -0.009155 | -0.13275 | -0.1236 | 0.15564 | 0.30212 | -0.20142 | -0.24262 | -0.054931 | -0.09613 | -0.027466 | -0.032043 | 0 | 0.032043 | 0.027466 | 0.054932 | 0.18768 | -0.10071 | 0.21515 | -0.073242
682.7812 | 122.4472 | -59.0332 | 63.2172 | -0.077133 | 5.9776e-05 | 0.067691 | 0.068064 | -0.0023777 | 0.0027302 | -0.00081179 | -0.0074159 | 0.0018821 | 0.064066 | 0.071096 | -0.0025315 | -0.0030483 | 0.0015103 | 0.009155 | 63.1805 | -59.1934 | 63.0798 | -58.8867 | -59.2941 | -59.285 | 0.045776 | 63.1714 | -59.0836 | 63.1989 | -59.0286 | -59.2026 | -58.8867 | -0.16022 | -0.036621 | -0.13733 | -0.10071 | 0.14648 | 0.3067 | -0.25177 | -0.26093 | -0.091553 | -0.10071 | -0.050354 | -0.045776 | -0.01831 | 0.027466 | 0.004578 | 0.054932 | 0.14648 | -0.16937 | 0.19684 | -0.11902
683.7812 | 122.4838 | -59.0515 | 63.2492 | -0.038986 | 0.002196 | 0.068759 | 0.070201 | -0.0040867 | 0.00080757 | 0.00089721 | -0.028778 | 0.0012411 | 0.065561 | 0.069173 | 0.00024568 | -0.0034757 | -0.0019077 | 0.013733 | 63.1989 | -59.1888 | 63.0661 | -58.8913 | -59.2987 | -59.2758 | 0.045776 | 63.1851 | -59.1019 | 63.1851 | -59.024 | -59.1751 | -58.8867 | -0.13733 | -0.050354 | -0.18311 | -0.13275 | 0.16022 | 0.29755 | -0.2243 | -0.24719 | -0.086975 | -0.10986 | -0.050354 | -0.064087 | -0.064087 | 0 | 0.027466 | 0.07782 | 0.16479 | -0.1236 | 0.21515 | -0.073242
684.7812 | 122.4976 | -59.0378 | 63.176 | -0.054245 | -0.00079462 | 0.066195 | 0.066996 | 0.0014675 | 0.00038017 | 4.2808e-05 | 0.059723 | -4.0508e-05 | 0.063852 | 0.053579 | 3.2078e-05 | -0.0024075 | 0.0059963 | -0.086975 | 63.1531 | -59.1843 | 63.0753 | -58.9096 | -59.317 | -59.2484 | -0.041199 | 63.1302 | -59.0744 | 63.1851 | -59.0286 | -59.1431 | -58.8821 | -0.14648 | -0.022888 | -0.10071 | -0.07782 | 0.12817 | 0.27466 | -0.21057 | -0.27924 | -0.064087 | -0.13275 | -0.036621 | -0.045776 | 0.009156 | 0.054932 | 0.009156 | 0.045777 | 0.15564 | -0.10528 | 0.19226 | -0.068664
685.7812 | 122.4243 | -59.0378 | 63.2401 | -0.077133 | -0.00015382 | 0.067477 | 0.070414 | 0.0042447 | 0.00016657 | -0.0027344 | -0.0043641 | 0.0027365 | 0.065347 | 0.072378 | -0.0018905 | -0.0013393 | -0.00019868 | -0.009155 | 63.2263 | -59.198 | 63.0569 | -58.9554 | -59.3353 | -59.2896 | -0.018311 | 63.1989 | -59.0881 | 63.1714 | -59.0469 | -59.166 | -58.9096 | -0.16022 | -0.013733 | -0.1831 | -0.16937 | 0.082398 | 0.24262 | -0.25177 | -0.29755 | -0.091553 | -0.13733 | -0.050354 | -0.041198 | -0.068664 | -0.027466 | -0.009155 | 0.041199 | 0.12817 | -0.12817 | 0.17853 | -0.07782
686.7812 | 122.4426 | -59.0469 | 63.2355 | -0.038986 | 0.00048698 | 0.066409 | 0.06956 | -0.0010959 | -0.00068783 | 0.0019654 | 0.0017395 | 0.0014547 | 0.065561 | 0.063619 | -0.0010361 | -0.0036893 | -0.0014805 | -0.027466 | 63.2172 | -59.2117 | 63.0661 | -58.9462 | -59.3445 | -59.2804 | -0.022888 | 63.1989 | -59.0927 | 63.1805 | -59.0378 | -59.166 | -58.9096 | -0.16479 | -0.018311 | -0.16937 | -0.15106 | 0.10071 | 0.2655 | -0.23346 | -0.29755 | -0.068665 | -0.13275 | -0.045776 | -0.036621 | -0.054932 | -0.018311 | 0.009155 | 0.054931 | 0.13733 | -0.11902 | 0.1831 | -0.073243
687.7812 | 122.5433 | -59.0469 | 63.176 | -0.092392 | 0.00048698 | 0.0694 | 0.066783 | 0.0048855 | 0.0023028 | -0.001239 | 0.013947 | 0.00060029 | 0.065347 | 0.062764 | -0.0023177 | -0.0019803 | 0.0025785 | -0.045776 | 63.1256 | -59.2072 | 63.1027 | -58.8638 | -59.285 | -59.2392 | 0.004578 | 63.1073 | -59.079 | 63.2492 | -58.9828 | -59.1339 | -58.8272 | -0.16022 | -0.050354 | -0.073242 | -0.022888 | 0.18311 | 0.34332 | -0.19226 | -0.23804 | -0.032044 | -0.07782 | -0.032043 | -0.068664 | 0.073243 | 0.14191 | 0.064087 | 0.09613 | 0.21973 | -0.086975 | 0.25177 | -0.054932
688.7812 | 122.4564 | -59.0332 | 63.2309 | -0.084763 | 0.0015552 | 0.068973 | 0.069773 | -0.00088229 | 0.00038017 | 0.0011108 | -0.15085 | 0.00017309 | 0.065134 | 0.070241 | -0.0048813 | -0.0015529 | -0.0048983 | 0 | 63.1897 | -59.2163 | 63.0798 | -58.9005 | -59.3033 | -59.2804 | -0.068665 | 63.1851 | -59.1064 | 63.1989 | -59.024 | -59.1934 | -58.9188 | -0.18311 | -0.041199 | -0.15106 | -0.10986 | 0.13275 | 0.31586 | -0.24719 | -0.27008 | -0.064087 | -0.086975 | -0.073242 | -0.045776 | -0.032043 | 0.013733 | 0.009155 | 0.082397 | 0.11444 | -0.16022 | 0.18768 | -0.086975
689.2188 | 122.3785 | -59.0286 | 63.2446 | -0.054245 | -0.0018628 | 0.06705 | 0.068705 | -0.0021641 | 0.0010212 | 0.0017518 | -0.0043641 | 0.0016685 | 0.064066 | 0.061056 | 0.00045928 | -0.0043301 | 0.0012967 | -0.027466 | 63.2309 | -59.1751 | 63.0432 | -58.9188 | -59.317 | -59.2758 | 0 | 63.1943 | -59.0927 | 63.1668 | -58.8226 | -59.2026 | -59.0057 | -0.14648 | -0.013733 | -0.20142 | -0.18768 | 0.10986 | 0.25635 | -0.24719 | -0.28839 | -0.10071 | -0.14191 | -0.064087 | -0.050354 | -0.07782 | -0.027466 | 0.20599 | 0.27008 | 0.022888 | -0.17395 | 0.086975 | -0.10986
690.1719 | 122.4518 | -59.0424 | 63.2263 | -0.084763 | 0.00091438 | 0.068118 | 0.070414 | 0.0016813 | -0.0023968 | 0.0011108 | -0.077606 | 0.0025229 | 0.064706 | 0.065542 | -0.00082252 | -0.0032619 | -0.0012667 | -0.004578 | 63.2034 | -59.1934 | 63.089 | -58.8867 | -59.2896 | -59.2712 | 0.032043 | 63.1714 | -59.079 | 63.2263 | -59.0424 | -59.1431 | -58.8272 | -0.15106 | -0.022888 | -0.13733 | -0.11444 | 0.15564 | 0.3067 | -0.22888 | -0.24719 | -0.07782 | -0.096131 | -0.036621 | -0.054931 | 0 | 0.054931 | 0 | 0.036621 | 0.21515 | -0.10071 | 0.25177 | -0.064087
691.1719 | 122.4747 | -59.0378 | 63.2538 | -0.054245 | -0.00079462 | 0.067477 | 0.070201 | 0.0021085 | -0.00047423 | 0.001538 | -0.040985 | 0.0014547 | 0.065775 | 0.067251 | 0.00045928 | -0.0013393 | 0.0017239 | -0.013733 | 63.1943 | -59.1934 | 63.0981 | -58.8821 | -59.2941 | -59.2484 | 0.045776 | 63.1805 | -59.0881 | 63.1119 | -58.9371 | -59.1934 | -58.8867 | -0.15564 | -0.059509 | -0.15564 | -0.09613 | 0.15564 | 0.31128 | -0.21057 | -0.25635 | -0.054932 | -0.10071 | -0.050354 | -0.073242 | -0.14191 | -0.068665 | 0.10071 | 0.15106 | 0.15106 | -0.15564 | 0.20142 | -0.10529
692.1719 | 122.4792 | -59.0149 | 63.2217 | -0.016098 | -0.001222 | 0.064486 | 0.068705 | -0.0021641 | 0.0010212 | -0.00081179 | -0.013519 | 0.0010275 | 0.065988 | 0.065969 | -0.00082252 | -0.0026211 | -0.00083948 | -0.018311 | 63.2034 | -59.1705 | 63.089 | -58.9005 | -59.2896 | -59.2392 | 0.018311 | 63.2217 | -59.0927 | 63.2126 | -59.024 | -59.1385 | -58.8684 | -0.15564 | -0.018311 | -0.13275 | -0.11444 | 0.11444 | 0.27008 | -0.2243 | -0.27466 | -0.068665 | -0.11902 | -0.077819 | 0 | -0.009156 | -0.009156 | -0.009155 | 0.068664 | 0.14648 | -0.1236 | 0.2243 | -0.045777
693.1719 | 122.4609 | -59.0286 | 63.2401 | -0.10765 | 0.00048698 | 0.069613 | 0.068492 | -2.7692e-05 | 0.0012348 | -0.00017099 | 0.0017395 | 0.0020957 | 0.064493 | 0.068319 | -0.00018152 | -0.0017667 | 0.0066373 | -0.018311 | 63.2217 | -59.1751 | 63.0753 | -58.905 | -59.2804 | -59.2392 | 0.022888 | 63.2172 | -59.0561 | 63.2217 | -59.0057 | -59.166 | -58.873 | -0.14648 | -0.01831 | -0.16479 | -0.14648 | 0.1236 | 0.27008 | -0.21057 | -0.25177 | -0.064087 | -0.10529 | -0.027466 | -0.022888 | -0.01831 | 0.004578 | 0.022888 | 0.050354 | 0.15564 | -0.13733 | 0.18311 | -0.10986
694.1719 | 122.4747 | -59.0195 | 63.2263 | -0.17632 | -0.0020764 | 0.065555 | 0.067851 | 0.0014675 | -0.0021832 | -0.0031616 | -0.05014 | 0.0014547 | 0.065347 | 0.060415 | -0.0010361 | -0.0013393 | -0.0034031 | 0 | 63.1897 | -59.1934 | 63.0936 | -58.8913 | -59.3124 | -59.2667 | 0.013733 | 63.1943 | -59.0744 | 63.1989 | -59.0286 | -59.1568 | -58.8776 | -0.17395 | -0.036621 | -0.13275 | -0.09613 | 0.12817 | 0.30212 | -0.24719 | -0.29297 | -0.073243 | -0.11902 | -0.054932 | -0.032043 | -0.027465 | 0.004578 | -0.009155 | 0.045777 | 0.14191 | -0.13733 | 0.19684 | -0.082397
695.1719 | 122.4106 | -59.0378 | 63.2538 | -0.046616 | -0.0014356 | 0.068332 | 0.068705 | -0.0038729 | -0.00047423 | -0.0016662 | -0.028778 | -0.00089511 | 0.064493 | 0.062978 | -0.00039512 | -5.7742e-05 | 0.00022852 | -0.022888 | 63.208 | -59.1843 | 63.089 | -58.8913 | -59.2621 | -59.2484 | 0.022888 | 63.1989 | -59.0973 | 63.2034 | -59.0332 | -59.1797 | -58.8959 | -0.14648 | -0.045776 | -0.1648 | -0.11902 | 0.14649 | 0.29297 | -0.21057 | -0.2243 | -0.064087 | -0.07782 | -0.059509 | -0.054931 | -0.050354 | 0.004577 | 0.004578 | 0.064087 | 0.14191 | -0.14191 | 0.20142 | -0.082397
696.1719 | 122.4564 | -59.0012 | 63.2172 | -0.00083923 | -0.00079462 | 0.070254 | 0.066569 | 0.0053129 | 0.00080757 | -0.00059819 | -0.022675 | -0.00025411 | 0.065347 | 0.062764 | -0.0016769 | -0.0034757 | 0.0012967 | -0.022888 | 63.2126 | -59.1888 | 63.002 | -58.8089 | -59.3079 | -59.2987 | 0.004578 | 63.1851 | -59.0652 | 63.208 | -59.0286 | -59.1522 | -58.8684 | -0.18768 | -0.004578 | -0.21515 | -0.21057 | 0.19226 | 0.37994 | -0.29755 | -0.3067 | -0.10986 | -0.11902 | -0.064087 | -0.032043 | -0.009155 | 0.022888 | -0.027465 | 0.036622 | 0.13275 | -0.15106 | 0.19684 | -0.086975
697.1719 | 122.4426 | -59.0332 | 63.2217 | -0.084763 | -0.00079462 | 0.065982 | 0.068705 | 0.0048855 | 0.00059397 | 0.0011108 | -0.03183 | 0.0010275 | 0.064493 | 0.062124 | -0.00082252 | -0.0017667 | -0.0038303 | -0.009155 | 63.1897 | -59.1843 | 63.1027 | -58.8913 | -59.3857 | -59.2896 | 0.009155 | 63.1805 | -59.0515 | 63.2263 | -59.0057 | -59.1614 | -58.8684 | -0.15106 | -0.032044 | -0.11902 | -0.086975 | 0.14191 | 0.29297 | -0.25635 | -0.35248 | -0.10529 | -0.20142 | -0.018311 | -0.041199 | 0.004577 | 0.045776 | 0.027466 | 0.045777 | 0.16479 | -0.12817 | 0.18311 | -0.10986
698.1719 | 122.4472 | -59.0286 | 63.208 | -0.11528 | 0.00048698 | 0.068759 | 0.068919 | -2.7692e-05 | 0.0027302 | 4.2808e-05 | -0.062347 | 0.0016685 | 0.066202 | 0.063619 | -0.0018905 | -0.0032619 | -0.00062588 | -0.032043 | 63.1851 | -59.166 | 63.0569 | -58.9005 | -59.3262 | -59.285 | -0.050354 | 63.176 | -59.0881 | 63.1897 | -59.0561 | -59.2072 | -58.9096 | -0.13733 | -0.022888 | -0.15106 | -0.12817 | 0.12817 | 0.2655 | -0.25635 | -0.29755 | -0.11902 | -0.16022 | -0.05951 | -0.032044 | -0.018311 | 0.013733 | -0.027466 | 0.032044 | 0.11902 | -0.17853 | 0.17853 | -0.11902
699.1719 | 122.4655 | -59.0103 | 63.2401 | -0.054245 | -0.00058102 | 0.069186 | 0.06956 | 0.00039951 | 0.0031574 | 0.0028198 | -0.010467 | 0.00017309 | 0.066629 | 0.062764 | -0.00060892 | -0.0030483 | 0.0010831 | -0.009155 | 63.2126 | -59.1888 | 63.1027 | -58.873 | -59.2712 | -59.2392 | 0.018311 | 63.2126 | -59.079 | 63.2172 | -58.9874 | -59.1385 | -58.8593 | -0.17853 | -0.027466 | -0.13733 | -0.10986 | 0.13733 | 0.31586 | -0.22888 | -0.26093 | -0.050354 | -0.082397 | -0.068664 | -0.027466 | -0.022888 | 0.004578 | 0.022888 | 0.091552 | 0.15106 | -0.12817 | 0.21973 | -0.05951
700.1719 | 122.4701 | -59.024 | 63.2629 | -0.069504 | -0.00058102 | 0.067904 | 0.068492 | 0.0010403 | 0.0010212 | -0.00059819 | 0.0078429 | 0.0031637 | 0.06492 | 0.064687 | -0.0016769 | -0.0026211 | -0.00019868 | -0.018311 | 63.2217 | -59.1705 | 63.0981 | -58.8867 | -59.2575 | -59.2346 | 0.032043 | 63.2034 | -59.0469 | 63.2446 | -58.9874 | -59.1339 | -58.8593 | -0.14648 | -0.041198 | -0.16479 | -0.1236 | 0.13733 | 0.28381 | -0.21057 | -0.23346 | -0.064087 | -0.086975 | -0.022888 | -0.059509 | -0.01831 | 0.041199 | 0.036621 | 0.059509 | 0.16479 | -0.10986 | 0.18768 | -0.086975
701.1719 | 122.4289 | -59.0378 | 63.2355 | -0.069504 | 0.002196 | 0.067477 | 0.069346 | -0.0013095 | 0.003371 | 0.002179 | -0.022675 | 0.0018821 | 0.065347 | 0.062978 | -0.0014633 | -0.0030483 | 0.0034329 | -0.036621 | 63.2217 | -59.1797 | 63.1073 | -58.8776 | -59.2896 | -59.2392 | -0.004578 | 63.1989 | -59.0515 | 63.2126 | -59.0195 | -59.1843 | -58.8913 | -0.14191 | -0.013733 | -0.12817 | -0.11444 | 0.16022 | 0.30212 | -0.20142 | -0.25177 | -0.05951 | -0.10986 | -0.013733 | -0.036621 | -0.022889 | 0.013732 | 0.018311 | 0.032044 | 0.14649 | -0.14648 | 0.16022 | -0.13275
702.1719 | 122.4564 | -59.0012 | 63.2309 | -0.031357 | -0.001222 | 0.068759 | 0.070628 | 0.00039951 | 0.0010212 | 0.00025641 | -0.0013123 | 0.0023093 | 0.065347 | 0.062337 | -0.0021041 | -0.0013393 | -0.0019077 | -0.036621 | 63.2172 | -59.1705 | 63.1165 | -58.8684 | -59.2941 | -59.2575 | 0.009155 | 63.1897 | -59.0469 | 63.2126 | -59.0149 | -59.1522 | -58.873 | -0.16937 | -0.013733 | -0.11444 | -0.10071 | 0.13275 | 0.30212 | -0.25635 | -0.29297 | -0.086975 | -0.1236 | -0.045776 | -0.041199 | -0.018311 | 0.022888 | -0.013733 | 0.032043 | 0.12817 | -0.15106 | 0.17395 | -0.10529
703.1719 | 122.4655 | -59.0057 | 63.2309 | -0.023727 | 0.00027338 | 0.069613 | 0.068492 | -0.0013095 | 0.0010212 | -0.00038459 | -0.040985 | -4.0508e-05 | 0.063638 | 0.068746 | 0.00024568 | -0.00091214 | -0.00019868 | 0.07782 | 63.2217 | -59.1568 | 63.1027 | -58.8638 | -59.2804 | -59.2117 | 0.013733 | 63.2126 | -59.0332 | 63.2172 | -59.0195 | -59.1751 | -58.8867 | -0.15106 | -0.009155 | -0.12817 | -0.11902 | 0.14191 | 0.29297 | -0.20599 | -0.27466 | -0.054932 | -0.1236 | -0.027466 | -0.018311 | -0.013733 | 0.004578 | -0.013733 | 0.013733 | 0.11902 | -0.16937 | 0.14648 | -0.14191
704.1719 | 122.4655 | -59.0195 | 63.2263 | -0.10002 | -0.0014356 | 0.069186 | 0.066783 | -0.0025913 | -0.00090143 | -0.0033752 | -0.044037 | 0.0018821 | 0.064279 | 0.067251 | 0.00045928 | 0.00015586 | 0.00022852 | -0.022888 | 63.2126 | -59.166 | 62.9974 | -58.9737 | -59.1934 | -59.2621 | 0.041199 | 63.208 | -59.0332 | 63.2446 | -58.9874 | -59.1156 | -58.8455 | -0.14648 | -0.013733 | -0.22888 | -0.21515 | 0.045776 | 0.19226 | -0.24262 | -0.17395 | -0.09613 | -0.027465 | -0.013733 | -0.01831 | 0.018311 | 0.036621 | 0.032043 | 0.045776 | 0.17395 | -0.096131 | 0.18768 | -0.082398
705.1719 | 122.4472 | -59.0286 | 63.2492 | -0.092392 | 0.00091438 | 0.068759 | 0.069346 | 0.0023221 | 0.0040118 | 0.00047001 | -0.1539 | 0.00038669 | 0.064066 | 0.060842 | -0.0010361 | -0.0032619 | -0.0021213 | -0.009155 | 63.2172 | -59.1751 | 63.1165 | -58.8364 | -59.2758 | -59.2346 | 0.082397 | 63.1943 | -59.0698 | 63.2355 | -58.9783 | -59.1476 | -58.873 | -0.14648 | -0.032044 | -0.13275 | -0.10071 | 0.19226 | 0.33874 | -0.20599 | -0.24719 | -0.059509 | -0.10071 | -0.041199 | -0.054932 | -0.013733 | 0.041199 | 0.050354 | 0.091553 | 0.15564 | -0.11902 | 0.19684 | -0.07782
706.1719 | 122.4838 | -58.9874 | 63.2629 | -0.061874 | -0.00079462 | 0.065127 | 0.06956 | 0.0025357 | 0.001662 | -0.00017099 | 0.0017395 | 0.0014547 | 0.06492 | 0.056142 | 3.2078e-05 | -0.0024075 | 0.00065592 | -0.013733 | 63.2446 | -59.1431 | 63.1027 | -58.8776 | -59.2529 | -59.2163 | 0.018311 | 63.2126 | -59.0469 | 63.2446 | -58.9737 | -59.111 | -58.8272 | -0.15564 | -0.01831 | -0.16022 | -0.14191 | 0.10986 | 0.2655 | -0.22888 | -0.2655 | -0.073243 | -0.10986 | -0.059509 | -0.050354 | -0.01831 | 0.032044 | 0.013733 | 0.073242 | 0.16022 | -0.1236 | 0.21973 | -0.064087
707.1719 | 122.4289 | -59.0195 | 63.2401 | -0.054245 | 0.00027338 | 0.069613 | 0.067637 | -0.0023777 | 0.0023028 | -0.00017099 | -0.019623 | 0.0025229 | 0.065561 | 0.065114 | 0.00024568 | -0.0015529 | 0.0045011 | 0.009155 | 63.2309 | -59.166 | 63.0981 | -58.8547 | -59.2712 | -59.2438 | 0.032043 | 63.2172 | -59.0469 | 63.2446 | -58.9966 | -59.0286 | -58.9233 | -0.14648 | -0.009155 | -0.14191 | -0.13275 | 0.16479 | 0.31128 | -0.2243 | -0.25177 | -0.077819 | -0.10529 | -0.027466 | -0.022888 | 0.004578 | 0.027466 | 0.022888 | 0.050354 | 0.09613 | -0.009155 | 0.1236 | 0.018311
708.1719 | 122.4289 | -59.0103 | 63.2629 | -0.054245 | 5.9776e-05 | 0.067904 | 0.068278 | 0.0021085 | -4.7032e-05 | -0.0031616 | -0.019623 | 0.00017309 | 0.066202 | 0.062978 | -0.00082252 | -0.00048494 | 0.0017239 | 0.004578 | 63.208 | -59.1339 | 63.1165 | -58.8409 | -59.2896 | -59.2529 | 0.027466 | 63.2172 | -59.0469 | 63.2126 | -59.0103 | -59.1522 | -58.8501 | -0.1236 | -0.054931 | -0.14648 | -0.091553 | 0.16937 | 0.29297 | -0.24262 | -0.27924 | -0.11902 | -0.15564 | -0.036621 | -0.045776 | -0.050354 | -0.004578 | 0 | 0.036621 | 0.16022 | -0.14191 | 0.19684 | -0.10529
709.1719 | 122.4609 | -59.0012 | 63.2629 | -0.077133 | 0.00070078 | 0.069827 | 0.068064 | 0.0016813 | 0.00038017 | 0.00068361 | -0.022675 | 0.00017309 | 0.066629 | 0.061483 | -0.0010361 | -0.0017667 | 0.00086952 | -0.022888 | 63.2446 | -59.1568 | 63.1027 | -58.8181 | -59.2667 | -59.2117 | 0.027466 | 63.1897 | -59.0698 | 63.2446 | -59.0012 | -59.1522 | -58.8684 | -0.15564 | -0.01831 | -0.16022 | -0.14191 | 0.18311 | 0.33875 | -0.21057 | -0.2655 | -0.054932 | -0.10986 | -0.068664 | -0.073242 | -0.01831 | 0.054932 | 0 | 0.068664 | 0.13275 | -0.15106 | 0.20142 | -0.082398
710.1719 | 122.438 | -58.9874 | 63.2584 | -0.00083923 | 0.00048698 | 0.069827 | 0.066996 | 0.0014675 | -4.7032e-05 | 0.0011108 | 0.0047913 | -0.0011087 | 0.064706 | 0.059347 | -0.00018152 | -0.0021939 | 0.00022852 | -0.009155 | 63.208 | -59.1522 | 63.176 | -58.9645 | -59.3399 | -59.2209 | 0.022888 | 63.2172 | -59.0378 | 63.2355 | -58.9645 | -59.1339 | -58.8318 | -0.1648 | -0.050354 | -0.082398 | -0.032044 | 0.022888 | 0.18768 | -0.23346 | -0.35248 | -0.068664 | -0.18768 | -0.050354 | -0.041199 | -0.022888 | 0.018311 | 0.022888 | 0.073242 | 0.15564 | -0.14648 | 0.20599 | -0.09613
711.1719 | 122.5113 | -58.9783 | 63.2446 | -0.077133 | 0.0013416 | 0.070254 | 0.067424 | -2.7692e-05 | -0.00047423 | 0.0017518 | -0.025726 | 0.0020957 | 0.064066 | 0.067037 | -0.00082252 | -0.0034757 | 0.0038601 | -0.009155 | 63.2675 | -59.1156 | 63.1165 | -58.8547 | -59.2575 | -59.2117 | -0.050354 | 63.2126 | -59.0332 | 63.2446 | -58.9966 | -59.1064 | -58.8364 | -0.13733 | 0.022888 | -0.12817 | -0.15106 | 0.1236 | 0.26093 | -0.23346 | -0.27924 | -0.09613 | -0.14191 | -0.054932 | -0.032044 | 0 | 0.032044 | -0.018311 | 0.036621 | 0.14191 | -0.12817 | 0.19684 | -0.073242
712.1719 | 122.4197 | -59.0103 | 63.2492 | -0.054245 | 0.00048698 | 0.067477 | 0.066569 | -0.00045509 | 0.0029438 | 0.0013244 | 0.0047913 | 0.0020957 | 0.066629 | 0.067464 | 0.00088648 | -0.0017667 | 0.00065592 | -0.009155 | 63.2446 | -59.1385 | 63.1256 | -58.8364 | -59.2621 | -59.1934 | 0.050354 | 63.2126 | -59.0469 | 63.2172 | -58.9554 | -59.1522 | -58.8409 | -0.12817 | -0.004578 | -0.1236 | -0.11902 | 0.17395 | 0.30212 | -0.18311 | -0.25177 | -0.054931 | -0.1236 | -0.036621 | -0.036622 | -0.032044 | 0.004578 | 0.054932 | 0.091553 | 0.16937 | -0.14191 | 0.20599 | -0.10529
713.1719 | 122.5159 | -58.9645 | 63.2217 | -0.061874 | -0.00079462 | 0.067691 | 0.067424 | 0.0036039 | -0.00047423 | -0.00038459 | -0.0043641 | -0.00046791 | 0.063638 | 0.071309 | -0.00018152 | -0.0049709 | -0.0040439 | -0.009155 | 63.2538 | -59.1339 | 63.1302 | -58.8318 | -59.2438 | -59.2117 | 0.054932 | 63.208 | -59.0469 | 63.2492 | -58.9737 | -59.1156 | -58.8364 | -0.16937 | 0.032043 | -0.091553 | -0.1236 | 0.13275 | 0.30212 | -0.24719 | -0.27923 | -0.07782 | -0.10986 | -0.082397 | -0.013733 | 0.027466 | 0.041199 | -0.009155 | 0.073242 | 0.12817 | -0.15106 | 0.21057 | -0.068665
714.1719 | 122.4472 | -58.9737 | 63.2858 | -0.054245 | 0.0017688 | 0.067691 | 0.06721 | 0.0018949 | 0.0037982 | -0.001239 | -0.025726 | 0.0010275 | 0.063852 | 0.060628 | -0.00018152 | -0.0019803 | 0.0015103 | 0.018311 | 63.2584 | -59.1202 | 63.121 | -58.8226 | -59.2346 | -59.2072 | 0.045776 | 63.2263 | -59.0378 | 63.2492 | -58.9554 | -59.1293 | -58.8409 | -0.14648 | -0.027466 | -0.1648 | -0.13733 | 0.15106 | 0.29755 | -0.23346 | -0.26093 | -0.086975 | -0.11444 | -0.064087 | -0.05951 | -0.036621 | 0.022889 | 0.018311 | 0.082398 | 0.13275 | -0.15564 | 0.19684 | -0.091552
715.1719 | 122.4518 | -58.9966 | 63.2584 | -0.10002 | -0.001222 | 0.069186 | 0.070414 | -0.0036593 | -0.00026063 | 0.0030334 | -0.037933 | 0.00060029 | 0.064066 | 0.064901 | 0.00088648 | -0.0017667 | 0.00044212 | -0.004578 | 63.2538 | -59.1293 | 63.1302 | -58.8638 | -59.2255 | -59.2026 | 0.018311 | 63.2309 | -59.0332 | 63.2401 | -58.9874 | -59.1385 | -58.8547 | -0.13275 | -0.004578 | -0.12817 | -0.1236 | 0.13275 | 0.2655 | -0.20599 | -0.22888 | -0.073243 | -0.096131 | -0.036621 | -0.027466 | -0.018311 | 0.009155 | 0.009155 | 0.045776 | 0.14191 | -0.14191 | 0.17853 | -0.10529
716.1719 | 122.4243 | -58.992 | 63.2401 | -0.084763 | 0.002837 | 0.067904 | 0.067637 | -0.00045509 | -0.00090143 | 0.00025641 | -0.013519 | 0.00081389 | 0.063852 | 0.064473 | -0.00082252 | -0.0028347 | 0.00044212 | -0.004578 | 63.2401 | -59.1385 | 63.1073 | -58.9371 | -59.2758 | -59.2163 | 0.032043 | 63.2172 | -59.0469 | 63.2355 | -58.9691 | -59.1248 | -58.8272 | -0.14648 | 0 | -0.13275 | -0.13275 | 0.054931 | 0.20142 | -0.22431 | -0.28381 | -0.07782 | -0.13733 | -0.054932 | -0.022888 | -0.004577 | 0.018311 | 0.022888 | 0.07782 | 0.16479 | -0.13275 | 0.21973 | -0.07782
717.1719 | 122.4609 | -58.9966 | 63.2538 | -0.077133 | -0.00036742 | 0.0694 | 0.068705 | -0.0032321 | 0.00016657 | 4.2808e-05 | -0.095917 | 0.0044455 | 0.065561 | 0.065114 | -0.0016769 | -0.0032619 | 0.0012967 | -0.013733 | 63.2446 | -59.1431 | 63.1393 | -58.8272 | -59.2529 | -59.198 | -0.054932 | 63.2217 | -59.0378 | 63.2446 | -58.9554 | -59.1064 | -58.8501 | -0.14648 | -0.009155 | -0.11444 | -0.10529 | 0.16937 | 0.31586 | -0.20142 | -0.25635 | -0.054932 | -0.10986 | -0.041199 | -0.032043 | -0.009155 | 0.022888 | 0.041199 | 0.082398 | 0.14648 | -0.10986 | 0.18768 | -0.068664
718.1719 | 122.4335 | -59.0149 | 63.2629 | -0.11528 | 0.00027338 | 0.070254 | 0.068064 | -0.0032321 | 0.001662 | -0.00081179 | -0.013519 | 0.0016685 | 0.065134 | 0.061269 | 3.2078e-05 | -0.0041165 | -0.0014805 | -0.032043 | 63.2401 | -59.1339 | 63.1119 | -58.8455 | -59.2758 | -59.2255 | -0.018311 | 63.2401 | -59.0424 | 63.2309 | -58.992 | -59.1568 | -58.8638 | -0.11902 | -0.022888 | -0.15106 | -0.12817 | 0.16937 | 0.28839 | -0.21057 | -0.26093 | -0.091553 | -0.14191 | -0.027465 | -0.022888 | -0.032043 | -0.009155 | 0.022889 | 0.050354 | 0.15106 | -0.14191 | 0.17853 | -0.11444
719.1719 | 122.4792 | -58.9783 | 63.2675 | -0.038986 | 0.00048698 | 0.068973 | 0.067424 | -0.0040867 | 0.0023028 | -0.00038459 | -0.03183 | 0.0012411 | 0.066415 | 0.065114 | -0.0018905 | -0.0021939 | 0.0077053 | 0.004578 | 63.2721 | -59.1202 | 63.1439 | -58.8089 | -59.2484 | -59.1888 | 0.041199 | 63.2217 | -59.0103 | 63.2629 | -59.0561 | -59.1156 | -58.8638 | -0.14191 | 0.004578 | -0.1236 | -0.12817 | 0.16937 | 0.31128 | -0.21057 | -0.27008 | -0.068665 | -0.12817 | -0.032044 | -0.045776 | -0.004578 | 0.041198 | -0.07782 | -0.045776 | 0.11444 | -0.13733 | 0.14648 | -0.10529
720.1719 | 122.4472 | -58.9737 | 63.2904 | -0.016098 | -0.00058102 | 0.068545 | 0.066569 | 0.0016813 | 0.00038017 | -0.00081179 | 0.010895 | 0.00017309 | 0.066629 | 0.062337 | -0.00082252 | -0.0053983 | -0.00019868 | -0.018311 | 63.2767 | -59.111 | 63.1256 | -58.8364 | -59.2392 | -59.1888 | 0.004578 | 63.2629 | -59.0012 | 63.2629 | -58.9417 | -59.1019 | -58.8135 | -0.13733 | -0.013733 | -0.16479 | -0.15106 | 0.13733 | 0.27466 | -0.21515 | -0.2655 | -0.07782 | -0.12817 | -0.027466 | -0.027466 | -0.027466 | 0 | 0.032044 | 0.05951 | 0.16022 | -0.12817 | 0.18768 | -0.10071
721.1719 | 122.5021 | -58.992 | 63.2858 | -0.023727 | 0.002196 | 0.065555 | 0.068492 | 0.0021085 | 0.00038017 | 0.00047001 | 0.026154 | -0.00025411 | 0.063425 | 0.058919 | -0.0025315 | -0.0028347 | 0.0019375 | -0.013733 | 63.2767 | -59.1156 | 63.121 | -58.8547 | -59.23 | -59.1888 | 0.009155 | 63.2629 | -59.0057 | 63.2629 | -58.9279 | -59.0836 | -58.8135 | -0.1236 | -0.009156 | -0.1648 | -0.15564 | 0.13733 | 0.26093 | -0.19684 | -0.23804 | -0.073242 | -0.11444 | -0.013733 | -0.022889 | -0.022889 | 0 | 0.064087 | 0.07782 | 0.17853 | -0.091553 | 0.19226 | -0.07782
722.1719 | 122.4152 | -58.9783 | 63.2904 | -0.054245 | 5.9776e-05 | 0.068973 | 0.067637 | -0.00024129 | 0.0012348 | -0.0014526 | -0.037933 | 0.00081389 | 0.065988 | 0.063833 | 0.00045928 | -0.0026211 | 0.00065592 | -0.041199 | 63.2355 | -59.1248 | 63.1027 | -58.8684 | -59.2941 | -59.2438 | -0.054932 | 63.2355 | -59.0469 | 63.2217 | -59.0195 | -59.1751 | -58.873 | -0.14648 | -0.054931 | -0.18768 | -0.13275 | 0.10986 | 0.25635 | -0.2655 | -0.31586 | -0.11902 | -0.16937 | -0.068665 | -0.054931 | -0.068664 | -0.013733 | -0.041199 | 0.027466 | 0.10529 | -0.19684 | 0.17395 | -0.12817
723.1719 | 122.4792 | -58.9554 | 63.2675 | -0.016098 | -0.0010082 | 0.065982 | 0.069133 | -0.0010959 | 0.0010212 | 0.00089721 | -0.025726 | 0.00017309 | 0.063425 | 0.061696 | -0.0025315 | -0.0036893 | 0.0015103 | -0.013733 | 63.2446 | -59.1019 | 63.1622 | -58.8638 | -59.3445 | -59.2209 | 0.018311 | 63.2675 | -59.0012 | 63.2812 | -58.9508 | -59.1064 | -58.7997 | -0.14649 | -0.022888 | -0.10529 | -0.082398 | 0.091552 | 0.23804 | -0.2655 | -0.3891 | -0.11902 | -0.24261 | -0.045777 | 0 | 0.013733 | 0.013733 | 0.004577 | 0.050354 | 0.15564 | -0.15106 | 0.20142 | -0.10529
724.1719 | 122.4838 | -58.96 | 63.2904 | -0.077133 | 0.00048698 | 0.069186 | 0.069346 | 0.0014675 | 0.00016657 | 0.0011108 | 0.0078429 | 0.0016685 | 0.064493 | 0.064901 | -0.0029587 | -0.0026211 | 0.0019375 | -0.009155 | 63.2675 | -59.1202 | 63.1485 | -58.7997 | -59.2346 | -59.1843 | 0.10071 | 63.2538 | -59.0012 | 63.295 | -58.9508 | -59.079 | -58.7906 | -0.16022 | -0.022888 | -0.14191 | -0.11902 | 0.16022 | 0.32043 | -0.2243 | -0.27466 | -0.064087 | -0.11444 | -0.041199 | -0.036621 | 0.004578 | 0.041199 | 0.009155 | 0.050354 | 0.16937 | -0.11902 | 0.21057 | -0.077819
725.1719 | 122.4655 | -58.9462 | 63.2767 | -0.046616 | 5.9776e-05 | 0.067691 | 0.068705 | 0.0021085 | 0.0040118 | 0.0013244 | -0.016571 | 0.0016685 | 0.066202 | 0.061696 | 0.0011001 | -0.0032619 | 0.00086952 | 0 | 63.2721 | -59.1156 | 63.1531 | -58.8089 | -59.2255 | -59.1934 | 0.036621 | 63.2309 | -59.0057 | 63.414 | -59.0744 | -59.1614 | -58.8593 | -0.16937 | -0.004577 | -0.1236 | -0.11902 | 0.13733 | 0.3067 | -0.24719 | -0.27924 | -0.077819 | -0.10986 | -0.059509 | -0.045776 | 0.13733 | 0.1831 | -0.12817 | -0.068665 | 0.086975 | -0.21515 | 0.14648 | -0.15564
726.1719 | 122.4518 | -58.9645 | 63.2812 | -0.084763 | 0.001128 | 0.066836 | 0.06956 | -0.0010959 | 0.0012348 | -0.00038459 | -0.022675 | 0.0012411 | 0.064706 | 0.067037 | 0.00045928 | -0.0024075 | -0.0014805 | 0.009155 | 63.2675 | -59.1019 | 63.1668 | -58.7997 | -59.2163 | -59.166 | 0.041199 | 63.2446 | -59.0057 | 63.2721 | -58.9417 | -59.0881 | -58.7952 | -0.13733 | -0.013733 | -0.11444 | -0.10071 | 0.1648 | 0.30212 | -0.20142 | -0.25177 | -0.064087 | -0.11444 | -0.041198 | -0.036621 | -0.009155 | 0.027466 | 0.022889 | 0.064087 | 0.16937 | -0.1236 | 0.21057 | -0.082398
727.1719 | 122.4564 | -58.9508 | 63.3041 | -0.046616 | 0.00027338 | 0.065768 | 0.067851 | -0.0028049 | 0.00038017 | -0.00081179 | 0.0047913 | 0.0010275 | 0.066415 | 0.062978 | -0.0014633 | -0.0021939 | 0.0017239 | -0.027466 | 63.3041 | -59.1064 | 63.1256 | -58.8547 | -59.2346 | -59.1934 | -0.004578 | 63.2538 | -59.0103 | 63.2675 | -58.9188 | -59.0744 | -58.8135 | -0.15564 | 0 | -0.17853 | -0.17853 | 0.096131 | 0.25177 | -0.24261 | -0.28381 | -0.086975 | -0.12817 | -0.059509 | -0.050354 | -0.036621 | 0.013733 | 0.032044 | 0.091553 | 0.13733 | -0.1236 | 0.19684 | -0.064087
728.1719 | 122.438 | -58.9417 | 63.3133 | -0.069504 | -0.00079462 | 0.067691 | 0.066996 | 0.00018591 | -0.0013288 | -0.0018798 | 0.0017395 | -4.0508e-05 | 0.064706 | 0.059774 | -0.00060892 | -0.0039029 | -0.0010531 | -0.022888 | 63.2996 | -59.1202 | 63.1302 | -58.8593 | -59.2346 | -59.1888 | -0.009155 | 63.2629 | -58.9874 | 63.2584 | -58.9417 | -59.111 | -58.8135 | -0.17853 | -0.013732 | -0.18311 | -0.16937 | 0.082397 | 0.26093 | -0.24719 | -0.29297 | -0.068665 | -0.11444 | -0.045777 | -0.050354 | -0.054931 | -0.004577 | 0 | 0.045777 | 0.12817 | -0.16937 | 0.17395 | -0.1236
729.1719 | 122.4701 | -58.9554 | 63.2996 | -0.11528 | 0.00070078 | 0.069827 | 0.069133 | 0.00039951 | 0.00059397 | -0.0010254 | -0.028778 | 0.0010275 | 0.063852 | 0.062978 | -0.0012497 | -0.0021939 | -0.0023349 | 0 | 63.3087 | -59.0881 | 63.1668 | -58.786 | -59.2072 | -59.1797 | 0.032043 | 63.2675 | -59.0057 | 63.4094 | -59.0424 | -59.1797 | -58.8501 | -0.13275 | 0.009155 | -0.13275 | -0.14191 | 0.16937 | 0.30212 | -0.2243 | -0.25177 | -0.091552 | -0.11902 | -0.050354 | -0.032044 | 0.10986 | 0.14191 | -0.086975 | -0.036621 | 0.10528 | -0.2243 | 0.15564 | -0.17395
730.1719 | 122.4609 | -58.9691 | 63.3224 | -0.084763 | 5.9776e-05 | 0.070041 | 0.068919 | 0.0012539 | 0.0020892 | 0.00089721 | -0.028778 | 0.0025229 | 0.065988 | 0.065969 | -0.0010361 | -0.0026211 | 1.4916e-05 | 0.027466 | 63.2858 | -59.0927 | 63.1668 | -58.786 | -59.2117 | -59.1751 | 0.032043 | 63.2675 | -58.9966 | 63.2812 | -58.9188 | -59.0607 | -58.7906 | -0.1236 | -0.036621 | -0.15564 | -0.11902 | 0.1831 | 0.3067 | -0.20599 | -0.24262 | -0.082398 | -0.11902 | -0.027466 | -0.054932 | -0.041199 | 0.013733 | 0.050354 | 0.07782 | 0.17853 | -0.091553 | 0.20599 | -0.064087
731.1719 | 122.4243 | -58.9508 | 63.3133 | -0.023727 | 0.0015552 | 0.068118 | 0.067637 | 0.0031765 | 0.00059397 | -0.0027344 | -0.022675 | 0.0010275 | 0.064279 | 0.065969 | -0.0016769 | -0.0013393 | 0.00044212 | 0 | 63.2721 | -59.1156 | 63.1851 | -58.7997 | -59.2163 | -59.166 | 0.036621 | 63.2584 | -58.9737 | 63.2812 | -58.9325 | -59.0744 | -58.7952 | -0.16479 | -0.041198 | -0.12817 | -0.086975 | 0.15106 | 0.31586 | -0.21515 | -0.2655 | -0.050354 | -0.10071 | -0.022888 | -0.054931 | -0.032043 | 0.022888 | 0.018311 | 0.041199 | 0.15564 | -0.1236 | 0.17853 | -0.10071
732.1719 | 122.4518 | -58.9508 | 63.295 | -0.069504 | 0.0019824 | 0.067904 | 0.067637 | -2.7692e-05 | -4.7032e-05 | 0.0013244 | -0.022675 | 0.00081389 | 0.063852 | 0.067464 | -0.00060892 | -0.0043301 | -0.0010531 | 0 | 63.2721 | -59.0973 | 63.1805 | -58.7814 | -59.2163 | -59.2621 | 0.045776 | 63.2767 | -59.0149 | 63.3041 | -58.9325 | -59.0973 | -58.7906 | -0.14648 | -0.022888 | -0.11444 | -0.091553 | 0.16937 | 0.31586 | -0.31128 | -0.2655 | -0.16479 | -0.11902 | -0.064087 | -0.018311 | 0.009155 | 0.027466 | 0.018311 | 0.082398 | 0.16022 | -0.14648 | 0.22431 | -0.082397
733.1719 | 122.4426 | -58.9508 | 63.3133 | -0.061874 | -0.001222 | 0.066195 | 0.067424 | -0.00045509 | -0.0013288 | -0.0014526 | -0.019623 | 0.0012411 | 0.067056 | 0.065542 | -0.0018905 | -0.0026211 | 0.0015103 | -0.013733 | 63.3179 | -59.0698 | 63.1622 | -58.8135 | -59.2026 | -59.1568 | 0.032043 | 63.2721 | -58.992 | 63.2996 | -58.8959 | -59.0698 | -58.7723 | -0.11902 | 0.004578 | -0.15106 | -0.15564 | 0.13733 | 0.25635 | -0.20599 | -0.25177 | -0.086975 | -0.13275 | -0.041198 | -0.041198 | -0.013732 | 0.027466 | 0.054932 | 0.09613 | 0.17853 | -0.11902 | 0.21973 | -0.07782
734.1719 | 122.4426 | -58.9462 | 63.3316 | 0.13649 | -0.0029308 | 0.069186 | 0.067637 | 0.00039951 | 0.0025166 | -0.0018798 | 0.013947 | 0.00060029 | 0.065775 | 0.062124 | 0.0015275 | -0.0036893 | 0.0023647 | -0.027466 | 63.3408 | -59.0607 | 63.121 | -58.8364 | -59.2255 | -59.1751 | -0.036621 | 63.2996 | -58.9828 | 63.2446 | -58.9233 | -59.0973 | -58.7997 | -0.11444 | 0.009155 | -0.21057 | -0.21973 | 0.10986 | 0.2243 | -0.22888 | -0.27924 | -0.11444 | -0.1648 | -0.036621 | -0.032043 | -0.086975 | -0.054932 | 0.022888 | 0.059509 | 0.14648 | -0.15106 | 0.18311 | -0.11444
735.1719 | 122.4747 | -58.9462 | 63.3087 | -0.077133 | -0.0014356 | 0.067264 | 0.068064 | 0.0044583 | -0.00068783 | 4.2808e-05 | 0.016998 | 0.0023093 | 0.066843 | 0.063619 | 0.00045928 | -0.0028347 | 0.0015103 | -0.032043 | 63.327 | -59.0652 | 63.1714 | -58.8181 | -59.2163 | -59.1568 | 0.027466 | 63.295 | -58.9828 | 63.2996 | -58.9188 | -59.0652 | -58.7677 | -0.11902 | 0.01831 | -0.13733 | -0.15564 | 0.12817 | 0.24719 | -0.21057 | -0.27008 | -0.091552 | -0.15106 | -0.036621 | -0.013733 | -0.009155 | 0.004578 | 0.027466 | 0.064087 | 0.17853 | -0.11902 | 0.21515 | -0.082398
736.1719 | 122.4701 | -58.9462 | 63.3179 | -0.046616 | 0.00070078 | 0.068545 | 0.06721 | 0.00061311 | -0.00047423 | 0.0011108 | -0.037933 | 0.0020957 | 0.064706 | 0.068532 | 0.00067288 | -0.0032619 | -0.0042575 | 0.018311 | 63.3133 | -59.0698 | 63.176 | -58.7769 | -59.198 | -59.1568 | 0.041199 | 63.2812 | -58.9783 | 63.295 | -58.905 | -59.0836 | -58.7997 | -0.1236 | -0.004578 | -0.14191 | -0.13733 | 0.16937 | 0.29297 | -0.21057 | -0.25177 | -0.086975 | -0.12817 | -0.032043 | -0.036621 | -0.022888 | 0.013733 | 0.041199 | 0.073242 | 0.14648 | -0.13733 | 0.17853 | -0.10529
737.1719 | 122.4838 | -58.9325 | 63.3041 | -0.077133 | 0.00048698 | 0.068759 | 0.066996 | 0.0018949 | 0.00059397 | 0.001538 | -0.028778 | 0.0010275 | 0.064279 | 0.063833 | -0.0016769 | -0.0013393 | -0.00062588 | 0.018311 | 63.3133 | -59.0836 | 63.1851 | -58.7723 | -59.1888 | -59.1568 | 0.041199 | 63.2812 | -58.9645 | 63.3133 | -58.9005 | -59.0744 | -58.7814 | -0.15106 | 0.009155 | -0.11902 | -0.12817 | 0.16022 | 0.31128 | -0.2243 | -0.25635 | -0.073242 | -0.10529 | -0.032044 | -0.022888 | 0.009155 | 0.032043 | 0.032043 | 0.064087 | 0.15106 | -0.14191 | 0.18311 | -0.10986
738.1719 | 122.4701 | -58.9188 | 63.3179 | -0.092392 | -0.00015382 | 0.065768 | 0.069133 | 0.0023221 | 0.0031574 | 0.00047001 | -0.016571 | 0.0010275 | 0.065775 | 0.065328 | -0.0016769 | -0.0028347 | -0.0031895 | 0.009155 | 63.3133 | -59.0561 | 63.1897 | -58.7769 | -59.1888 | -59.1431 | 0.041199 | 63.295 | -58.9874 | 63.3041 | -58.8959 | -59.0744 | -58.7677 | -0.13733 | -0.004578 | -0.12817 | -0.1236 | 0.14191 | 0.27924 | -0.2243 | -0.27008 | -0.086975 | -0.13275 | -0.068665 | -0.022888 | -0.013733 | 0.009155 | 0.022888 | 0.091553 | 0.15106 | -0.15564 | 0.21973 | -0.086975
739.1719 | 122.4655 | -58.9142 | 63.3179 | -0.054245 | 0.0019824 | 0.068759 | 0.067851 | 0.00018591 | -0.001115 | -0.0020936 | -0.016571 | -4.0508e-05 | 0.064493 | 0.0696 | -0.0012497 | -0.0036893 | -0.0019077 | -0.004578 | 63.3133 | -59.0836 | 63.1714 | -58.786 | -59.1797 | -59.1476 | 0.054932 | 63.2812 | -58.96 | 63.295 | -58.9005 | -59.0424 | -58.7585 | -0.16937 | -0.004578 | -0.14648 | -0.14191 | 0.12817 | 0.29755 | -0.23346 | -0.2655 | -0.064087 | -0.09613 | -0.045776 | -0.036621 | -0.022888 | 0.013733 | 0.013733 | 0.059509 | 0.15564 | -0.12817 | 0.20142 | -0.082397
740.1719 | 122.4838 | -58.9279 | 63.3316 | -0.12291 | -0.00058102 | 0.068545 | 0.068705 | 0.0036039 | -4.7032e-05 | -0.0027344 | -0.03183 | 0.0016685 | 0.06727 | 0.065969 | -0.0025315 | -0.0036893 | 0.0034329 | -0.009155 | 63.2812 | -59.0881 | 63.1805 | -58.786 | -59.198 | -59.1476 | 0.050354 | 63.2812 | -58.9828 | 63.3041 | -58.8593 | -59.1614 | -58.905 | -0.16022 | -0.050354 | -0.15106 | -0.10071 | 0.14191 | 0.30212 | -0.21973 | -0.27008 | -0.059509 | -0.10986 | -0.054932 | -0.050354 | -0.027466 | 0.022888 | 0.068664 | 0.1236 | 0.022888 | -0.23346 | 0.07782 | -0.17853
741.1719 | 122.4426 | -58.9325 | 63.3408 | -0.016098 | -0.00015382 | 0.068332 | 0.065715 | -0.0036593 | 0.0037982 | 0.0023926 | -0.0074159 | 0.0012411 | 0.067911 | 0.060201 | -0.0012497 | -0.0024075 | 0.0019375 | -0.013733 | 63.3179 | -59.0698 | 63.0295 | -58.7906 | -59.1797 | -59.2072 | -0.032043 | 63.3041 | -58.9645 | 63.2721 | -58.9142 | -59.0698 | -58.8181 | -0.13733 | -0.022888 | -0.31128 | -0.28839 | 0.14191 | 0.27924 | -0.27466 | -0.24719 | -0.13733 | -0.10986 | -0.032044 | -0.036621 | -0.068664 | -0.032043 | 0.01831 | 0.050354 | 0.11444 | -0.13733 | 0.14649 | -0.10528
742.1719 | 122.4426 | -58.905 | 63.3453 | -0.016098 | 5.9776e-05 | 0.068332 | 0.067424 | 0.0021085 | 0.00059397 | 4.2808e-05 | -0.0013123 | 0.00038669 | 0.065134 | 0.055288 | -0.0027451 | -0.0028347 | -0.0012667 | -0.013733 | 63.327 | -59.0836 | 63.1714 | -58.8135 | -59.2163 | -59.1614 | -0.004578 | 63.3224 | -58.9691 | 63.2996 | -58.8867 | -59.0378 | -58.7585 | -0.17853 | -0.018311 | -0.17395 | -0.15564 | 0.091552 | 0.27008 | -0.25635 | -0.31128 | -0.07782 | -0.13275 | -0.064087 | -0.022888 | -0.045776 | -0.022888 | 0.01831 | 0.082397 | 0.14648 | -0.13275 | 0.21057 | -0.068665
743.1719 | 122.4701 | -58.9142 | 63.3362 | -0.061874 | -0.00229 | 0.068118 | 0.067637 | -0.0010959 | -0.001756 | -0.0020936 | -0.010467 | 0.0020957 | 0.067484 | 0.061696 | -0.00082252 | -0.0017667 | 0.00022852 | 0.027466 | 63.3133 | -59.079 | 63.1943 | -58.7769 | -59.1751 | -59.1248 | 0.045776 | 63.2812 | -58.9737 | 63.3179 | -58.8821 | -59.0469 | -58.7723 | -0.16479 | -0.022889 | -0.14191 | -0.11902 | 0.13733 | 0.30212 | -0.21057 | -0.26092 | -0.045777 | -0.096131 | -0.059509 | -0.054932 | -0.018311 | 0.036621 | 0.032044 | 0.091553 | 0.14191 | -0.13275 | 0.20142 | -0.073242
744.1719 | 122.4426 | -58.9279 | 63.3362 | -0.016098 | -0.001222 | 0.067477 | 0.066355 | 0.00061311 | 0.00016657 | 0.0013244 | -0.034882 | 0.0023093 | 0.067056 | 0.071096 | -0.00018152 | -0.0021939 | -0.00062588 | 0.009155 | 63.3224 | -59.0698 | 63.1805 | -58.7494 | -59.1705 | -59.1202 | 0.045776 | 63.3224 | -58.9554 | 63.3179 | -58.8684 | -59.0378 | -58.7585 | -0.14191 | -0.013733 | -0.15564 | -0.14191 | 0.17853 | 0.32043 | -0.19226 | -0.24262 | -0.050354 | -0.10071 | -0.027466 | -0.013733 | -0.018311 | -0.004578 | 0.059509 | 0.086975 | 0.16937 | -0.10986 | 0.19684 | -0.082398
745.1719 | 122.5113 | -58.9188 | 63.3545 | -0.077133 | -0.00058102 | 0.069827 | 0.068705 | 0.0010403 | -0.00047423 | -0.00081179 | -0.016571 | 0.00017309 | 0.06492 | 0.064046 | 0.00024568 | -0.0028347 | 0.0015103 | -0.018311 | 63.327 | -59.0515 | 63.1943 | -58.7677 | -59.1934 | -59.1339 | 0.018311 | 63.295 | -58.9554 | 63.3133 | -58.8684 | -59.024 | -58.7585 | -0.13275 | -0.027466 | -0.16022 | -0.13275 | 0.15106 | 0.28381 | -0.21515 | -0.27466 | -0.082397 | -0.14191 | -0.036621 | -0.059509 | -0.041199 | 0.01831 | 0.050354 | 0.086975 | 0.16022 | -0.10529 | 0.19684 | -0.068665
746.1719 | 122.4335 | -58.8776 | 63.3545 | -0.023727 | -0.0010082 | 0.070041 | 0.067637 | -0.0045139 | 0.0029438 | 0.00068361 | -0.0013123 | 0.0027365 | 0.067484 | 0.06426 | -0.0018905 | -0.0026211 | 0.00086952 | -0.013733 | 63.3499 | -59.0607 | 63.1943 | -58.7814 | -59.1843 | -59.1339 | 0.032043 | 63.3362 | -58.9462 | 63.3179 | -58.8638 | -59.0286 | -58.7494 | -0.18311 | -0.004577 | -0.16022 | -0.15564 | 0.09613 | 0.27924 | -0.25635 | -0.3067 | -0.073242 | -0.1236 | -0.068665 | -0.01831 | -0.036621 | -0.018311 | 0.013732 | 0.082397 | 0.12817 | -0.15106 | 0.19684 | -0.082397
747.1719 | 122.493 | -58.9142 | 63.3453 | -0.038986 | -0.00015382 | 0.066623 | 0.069133 | 0.0010403 | 0.001662 | -0.0014526 | -0.010467 | 0.0020957 | 0.068338 | 0.065542 | -0.0016769 | -0.0034757 | 0.0019375 | -0.018311 | 63.3499 | -59.0744 | 63.1714 | -58.7997 | -59.1751 | -59.1385 | -0.004578 | 63.3179 | -58.9508 | 63.3041 | -58.873 | -59.0515 | -58.7631 | -0.16022 | 0.004578 | -0.17395 | -0.17853 | 0.11444 | 0.27466 | -0.2243 | -0.26092 | -0.064087 | -0.10071 | -0.036621 | -0.027466 | -0.041199 | -0.013733 | 0.041199 | 0.07782 | 0.15106 | -0.13733 | 0.18768 | -0.10071
748.1719 | 122.3648 | -58.9325 | 63.3591 | 0.022049 | -0.001222 | 0.066623 | 0.068064 | 0.0010403 | 0.00016657 | 0.00068361 | 0.0047913 | 0.0018821 | 0.065347 | 0.05956 | -0.00060892 | -0.0021939 | 0.0040737 | -0.032043 | 63.3591 | -59.0424 | 63.1485 | -58.7997 | -59.198 | -59.1339 | -0.009155 | 63.3362 | -58.9462 | 63.1989 | -58.9188 | -59.0286 | -58.7585 | -0.10986 | 0 | -0.21057 | -0.21057 | 0.13275 | 0.24261 | -0.20142 | -0.2655 | -0.091553 | -0.15564 | -0.013733 | -0.022888 | -0.16022 | -0.13733 | 0.013733 | 0.027466 | 0.17395 | -0.09613 | 0.18768 | -0.082397
749.1719 | 122.4701 | -58.8959 | 63.382 | -0.054245 | -0.0014356 | 0.067264 | 0.070842 | -0.0015231 | -0.00026063 | -0.0027344 | -0.010467 | 0.00060029 | 0.064706 | 0.060628 | -0.00082252 | -0.0030483 | -0.0029757 | -0.045776 | 63.3087 | -59.0744 | 63.176 | -58.7997 | -59.198 | -59.1614 | -0.009155 | 63.3316 | -58.9783 | 63.3087 | -58.9142 | -59.0515 | -58.7631 | -0.17853 | -0.073242 | -0.20599 | -0.13275 | 0.09613 | 0.27466 | -0.2655 | -0.30212 | -0.086975 | -0.1236 | -0.082397 | -0.050354 | -0.073242 | -0.022888 | -0.018311 | 0.064086 | 0.13275 | -0.15564 | 0.21515 | -0.073243
750.1719 | 122.4655 | -58.8913 | 63.3545 | -0.077133 | -0.00015382 | 0.069827 | 0.068705 | 0.0029629 | -0.00090143 | 0.001538 | -0.010467 | 0.00060029 | 0.065988 | 0.068319 | -0.0033859 | -0.0030483 | 0.0042875 | 0.004578 | 63.3408 | -59.0652 | 63.208 | -58.7769 | -59.1568 | -59.1248 | 0.036621 | 63.3362 | -58.9417 | 63.3133 | -58.873 | -59.0195 | -58.7357 | -0.17395 | -0.013733 | -0.14648 | -0.13275 | 0.11444 | 0.28839 | -0.23346 | -0.2655 | -0.059509 | -0.091552 | -0.050354 | -0.01831 | -0.041199 | -0.022889 | 0.01831 | 0.068664 | 0.15564 | -0.12817 | 0.20599 | -0.07782
751.1719 | 122.4289 | -58.8959 | 63.3408 | -0.11528 | 0.00048698 | 0.067691 | 0.068705 | 0.00018591 | -0.00026063 | -0.002948 | -0.0043641 | 0.0027365 | 0.065561 | 0.062337 | -0.0018905 | -0.0030483 | -0.0023349 | -0.004578 | 63.3545 | -59.0103 | 63.1943 | -58.7814 | -59.1843 | -59.1385 | 0.013733 | 63.3179 | -58.9645 | 63.3133 | -58.8776 | -59.024 | -58.7448 | -0.11444 | 0.013733 | -0.14648 | -0.16022 | 0.11444 | 0.22888 | -0.24262 | -0.28839 | -0.12817 | -0.17395 | -0.068665 | -0.022888 | -0.027466 | -0.004578 | 0.018311 | 0.086976 | 0.15106 | -0.12817 | 0.21973 | -0.059509
752.1719 | 122.438 | -58.8867 | 63.3591 | -0.069504 | -0.0016492 | 0.069186 | 0.068705 | -0.0010959 | 0.0014484 | -0.00038459 | -0.016571 | 0.00038669 | 0.065561 | 0.066182 | 0.00088648 | -0.0013393 | 0.0015103 | -0.013733 | 63.3408 | -59.0698 | 63.1851 | -58.7906 | -59.1797 | -59.1339 | -0.009155 | 63.3499 | -58.9371 | 63.3133 | -58.8821 | -59.0286 | -58.754 | -0.1831 | -0.018311 | -0.17395 | -0.15564 | 0.096131 | 0.27924 | -0.24719 | -0.29297 | -0.064087 | -0.10986 | -0.050354 | -0.009155 | -0.045777 | -0.036622 | 0.004578 | 0.054932 | 0.13275 | -0.14191 | 0.18311 | -0.091552
753.1719 | 122.4289 | -58.9005 | 63.3728 | -0.084763 | -0.00015382 | 0.067691 | 0.069987 | 0.0027493 | 0.001662 | 0.0028198 | -0.0013123 | -4.0508e-05 | 0.064066 | 0.06661 | 0.0015275 | -0.0041165 | 0.00065592 | -0.009155 | 63.3408 | -59.0561 | 63.1714 | -58.786 | -59.1751 | -59.1385 | -0.013733 | 63.3499 | -58.9508 | 63.2904 | -58.8776 | -59.0607 | -58.7677 | -0.15564 | -0.032044 | -0.20142 | -0.16937 | 0.11444 | 0.27008 | -0.23804 | -0.27466 | -0.082398 | -0.11902 | -0.050354 | -0.022888 | -0.082398 | -0.05951 | 0.022889 | 0.073243 | 0.13275 | -0.16022 | 0.18311 | -0.10986
754.1719 | 122.4243 | -58.9005 | 63.3636 | -0.016098 | -0.0029308 | 0.06705 | 0.068278 | -0.0040867 | 0.001662 | -0.0010254 | 0.0017395 | -0.00068151 | 0.06492 | 0.062764 | -0.0014633 | -0.0045437 | 0.0010831 | 0.15564 | 63.4186 | -59.0195 | 63.1439 | -58.7997 | -59.1888 | -59.1293 | -0.004578 | 63.3362 | -58.9233 | 63.3133 | -58.8913 | -59.0424 | -58.7631 | -0.11902 | 0.054932 | -0.21973 | -0.27466 | 0.10071 | 0.21973 | -0.22888 | -0.28839 | -0.10986 | -0.16937 | -0.022888 | -0.027465 | -0.050354 | -0.022889 | 0.009156 | 0.032044 | 0.13733 | -0.14191 | 0.16022 | -0.11902
755.1719 | 122.4335 | -58.9096 | 63.3682 | -0.10002 | -0.00079462 | 0.067477 | 0.065287 | -0.0015231 | 0.0031574 | -0.002948 | 0.035309 | 0.00038669 | 0.064706 | 0.062124 | -0.00018152 | -0.0024075 | 0.00065592 | 0 | 63.3453 | -59.0378 | 63.2172 | -58.7585 | -59.1705 | -59.0927 | 0.009155 | 63.3408 | -58.9417 | 63.2721 | -58.8913 | -59.0149 | -58.7723 | -0.12817 | -0.022888 | -0.15106 | -0.12817 | 0.15106 | 0.27924 | -0.1831 | -0.26093 | -0.054931 | -0.13275 | -0.032043 | -0.027466 | -0.09613 | -0.068664 | 0.018311 | 0.050354 | 0.13733 | -0.10529 | 0.16937 | -0.073243
756.1719 | 122.4152 | -58.9096 | 63.3957 | 0.029678 | -0.00015382 | 0.066409 | 0.068919 | -0.00066869 | 0.0014484 | -0.0038024 | 0.0047913 | -4.0508e-05 | 0.064706 | 0.066396 | 0.0028091 | -0.0024075 | -0.0021213 | -0.027466 | 63.3636 | -59.0561 | 63.1805 | -58.7723 | -59.1614 | -59.1385 | 0.009155 | 63.3408 | -58.9142 | 63.2675 | -58.8409 | -59.0652 | -58.7723 | -0.14648 | -0.032044 | -0.21515 | -0.1831 | 0.13733 | 0.28381 | -0.22888 | -0.25177 | -0.082398 | -0.10529 | -0.004578 | -0.054932 | -0.12817 | -0.073242 | 0.068665 | 0.073243 | 0.13733 | -0.15564 | 0.14191 | -0.15106
757.1719 | 122.4152 | -58.9005 | 63.3636 | -0.031357 | 0.0019824 | 0.06705 | 0.068278 | -2.7692e-05 | 0.00059397 | -0.00059819 | -0.013519 | -0.0017495 | 0.065561 | 0.062764 | 3.2078e-05 | -0.0019803 | -0.0023349 | -0.041199 | 63.3453 | -59.0607 | 63.1989 | -58.7997 | -59.1751 | -59.1248 | 0.013733 | 63.3408 | -58.9325 | 63.3179 | -58.8776 | -59.024 | -58.7677 | -0.16022 | -0.01831 | -0.16479 | -0.14648 | 0.10071 | 0.26093 | -0.2243 | -0.27466 | -0.064087 | -0.11444 | -0.032043 | -0.022888 | -0.045776 | -0.022888 | 0.022889 | 0.054932 | 0.13275 | -0.1236 | 0.1648 | -0.091553
758.1719 | 122.4472 | -58.8776 | 63.4048 | -0.016098 | -0.00036742 | 0.067264 | 0.068919 | -0.0028049 | -0.0013288 | -0.0031616 | -0.025726 | 0.0031637 | 0.064493 | 0.061269 | -0.0012497 | -0.0026211 | -0.00083948 | -0.059509 | 63.4003 | -58.992 | 63.2034 | -58.8272 | -59.1843 | -59.111 | -0.013733 | 63.3316 | -58.9325 | 63.3087 | -58.8867 | -59.0378 | -58.7723 | -0.11444 | -0.004577 | -0.20142 | -0.19684 | 0.050354 | 0.16479 | -0.23346 | -0.3067 | -0.11902 | -0.19226 | -0.054932 | -0.073242 | -0.09613 | -0.022888 | -0.009156 | 0.045776 | 0.10529 | -0.16022 | 0.16022 | -0.10529
759.1719 | 122.4014 | -58.873 | 63.3728 | -0.016098 | -0.0018628 | 0.065982 | 0.067851 | 0.00061311 | -4.7032e-05 | 0.001538 | -0.08371 | 0.0016685 | 0.065561 | 0.06426 | -0.0018905 | -0.0019803 | -0.00083948 | -0.013733 | 63.382 | -59.0332 | 63.2034 | -58.7494 | -58.8959 | -59.0195 | 0 | 63.3408 | -58.9325 | 63.3133 | -58.8593 | -59.0424 | -58.7723 | -0.16022 | 0.009155 | -0.16937 | -0.17853 | 0.1236 | 0.28381 | -0.14648 | -0.022888 | 0.013733 | 0.13733 | -0.059509 | -0.032044 | -0.05951 | -0.027466 | 0.013733 | 0.073242 | 0.10071 | -0.16937 | 0.16022 | -0.10986
760.1719 | 122.4335 | -58.8684 | 63.414 | -0.084763 | -0.00079462 | 0.067691 | 0.067851 | 0.0021085 | -0.001756 | 0.0019654 | -0.0074159 | 0.0014547 | 0.063211 | 0.064901 | 3.2078e-05 | -0.00091214 | 1.4916e-05 | -0.004578 | 63.3545 | -59.0378 | 63.1989 | -58.7814 | -59.166 | -59.1248 | -0.013733 | 63.3499 | -58.9096 | 63.3362 | -58.8364 | -59.0149 | -58.7448 | -0.16937 | -0.059509 | -0.21515 | -0.15564 | 0.086975 | 0.25635 | -0.25635 | -0.29755 | -0.086975 | -0.12817 | -0.041199 | -0.064086 | -0.077819 | -0.013733 | 0.032043 | 0.073242 | 0.1236 | -0.14648 | 0.16479 | -0.10529
761.1719 | 122.3236 | -58.8547 | 63.3591 | -0.061874 | -0.00015382 | 0.066409 | 0.068278 | -0.0017367 | 0.0010212 | 0.00047001 | -0.0043641 | -0.00046791 | 0.064279 | 0.062764 | 0.00045928 | -0.0045437 | -0.00019868 | -0.018311 | 63.3636 | -59.0149 | 63.1851 | -58.7585 | -59.1522 | -59.1156 | 0.004578 | 63.3316 | -58.9233 | 63.2721 | -58.7402 | -58.8501 | -58.7631 | -0.16022 | 0.004577 | -0.17395 | -0.17853 | 0.09613 | 0.25635 | -0.26093 | -0.29755 | -0.10071 | -0.13733 | -0.068665 | -0.027466 | -0.086975 | -0.059509 | 0.11444 | 0.18311 | 0.091552 | 0.004577 | 0.16022 | 0.073242
762.1719 | 122.4792 | -58.8776 | 63.3636 | -0.046616 | 0.0013416 | 0.066409 | 0.065715 | 0.0023221 | -0.0013288 | -0.0014526 | 0.038361 | -4.0508e-05 | 0.065561 | 0.063619 | 3.2078e-05 | -0.0043301 | 1.4916e-05 | 0.032043 | 63.3316 | -59.0286 | 63.2858 | -58.6578 | -59.0698 | -59.0286 | 0.10071 | 63.327 | -58.9005 | 63.3362 | -58.8318 | -58.96 | -58.7357 | -0.15106 | -0.032043 | -0.077819 | -0.045776 | 0.21973 | 0.37079 | -0.15106 | -0.19226 | 0 | -0.041199 | -0.022889 | -0.036621 | -0.027465 | 0.009156 | 0.045776 | 0.068665 | 0.14191 | -0.082398 | 0.16479 | -0.059509
763.1719 | 122.3694 | -58.905 | 63.4003 | -0.061874 | -0.00036742 | 0.067691 | 0.067851 | -0.00088229 | -0.00090143 | -0.001239 | 0.0017395 | 0.00081389 | 0.066415 | 0.061056 | -0.00060892 | -0.0024075 | 0.0012967 | 0.004578 | 63.3682 | -58.9966 | 63.2172 | -58.7494 | -59.1522 | -59.1156 | 0.013733 | 63.3545 | -58.9188 | 63.3316 | -58.873 | -59.024 | -58.7631 | -0.091553 | -0.032044 | -0.18311 | -0.15106 | 0.15564 | 0.24719 | -0.21057 | -0.24719 | -0.11902 | -0.15564 | -0.013733 | -0.045777 | -0.068665 | -0.022888 | 0.032043 | 0.045776 | 0.14191 | -0.11902 | 0.15564 | -0.10529
764.1719 | 122.4106 | -58.8959 | 63.4003 | -0.069504 | -0.0014356 | 0.067264 | 0.066355 | 0.00018591 | 0.00080757 | -0.0027344 | -0.022675 | 0.0023093 | 0.06492 | 0.064046 | -0.00039512 | -0.0032619 | 0.00022852 | -0.027466 | 63.3545 | -59.0057 | 63.1943 | -58.7036 | -59.2026 | -59.1339 | -0.009155 | 63.3545 | -58.9279 | 63.3179 | -58.8684 | -59.0286 | -58.7494 | -0.10986 | -0.045777 | -0.20599 | -0.16022 | 0.19226 | 0.30212 | -0.23804 | -0.3067 | -0.12817 | -0.19684 | -0.032043 | -0.045777 | -0.082398 | -0.036621 | 0.027466 | 0.059509 | 0.14648 | -0.13275 | 0.17853 | -0.10071
765.1719 | 122.3923 | -58.8593 | 63.3911 | 0.0067903 | -0.0010082 | 0.069186 | 0.068705 | 0.0014675 | -0.0015424 | -0.002948 | -0.013519 | -0.00068151 | 0.063211 | 0.06191 | 3.2078e-05 | -0.0041165 | -0.0010531 | -0.018311 | 63.382 | -59.0424 | 63.3087 | -58.9005 | -59.3582 | -59.1568 | -0.022888 | 63.382 | -58.9325 | 63.3316 | -58.8547 | -59.0103 | -58.7219 | -0.1831 | -0.009155 | -0.082397 | -0.073242 | -0.041199 | 0.14191 | -0.29755 | -0.49896 | -0.11444 | -0.31586 | -0.073242 | -0.009155 | -0.059509 | -0.050354 | 0.004578 | 0.07782 | 0.13733 | -0.15106 | 0.21057 | -0.07782
766.1719 | 122.4243 | -58.8638 | 63.3911 | -0.061874 | -0.00036742 | 0.068759 | 0.068919 | -0.00088229 | -0.00090143 | 0.002179 | -0.077606 | -4.0508e-05 | 0.065775 | 0.063619 | 3.2078e-05 | -0.0015529 | 0.0027921 | -0.018311 | 63.4003 | -59.0195 | 63.1668 | -58.6075 | -58.9508 | -59.2346 | 0.059509 | 63.3545 | -58.9417 | 63.3224 | -58.8455 | -59.0286 | -58.7585 | -0.15564 | 0.009156 | -0.2243 | -0.23346 | 0.25635 | 0.41199 | -0.37079 | -0.086975 | -0.21515 | 0.068664 | -0.077819 | -0.036621 | -0.068664 | -0.032043 | 0.018311 | 0.09613 | 0.10529 | -0.16479 | 0.18311 | -0.086975
767.1719 | 122.4243 | -58.8638 | 63.3957 | -0.12291 | -0.0020764 | 0.066409 | 0.067424 | -0.0015231 | -0.001115 | 4.2808e-05 | 0.035309 | 0.0016685 | 0.066415 | 0.065755 | 3.2078e-05 | -0.0041165 | 0.0055691 | -0.032043 | 63.3728 | -59.0057 | 63.208 | -58.7585 | -59.1522 | -58.9828 | -0.15564 | 63.382 | -58.9188 | 63.327 | -58.8684 | -59.0332 | -58.7723 | -0.14191 | -0.022888 | -0.18768 | -0.16479 | 0.10529 | 0.24719 | -0.11902 | -0.28839 | 0.022888 | -0.14648 | -0.054931 | -0.013733 | -0.068665 | -0.054932 | -0.004577 | 0.050354 | 0.091553 | -0.16937 | 0.14648 | -0.11444
768.1719 | 122.4243 | -58.873 | 63.4048 | -0.038986 | 0.001128 | 0.066623 | 0.067851 | -0.0013095 | -0.0013288 | -0.00081179 | -0.016571 | -0.00089511 | 0.066415 | 0.064687 | -0.00039512 | -0.0043301 | -0.0027621 | -0.032043 | 63.4048 | -59.0195 | 63.2217 | -58.7448 | -59.1476 | -59.0881 | 0.10529 | 63.4048 | -58.8501 | 63.3224 | -58.905 | -58.9828 | -58.7311 | -0.14648 | 0 | -0.1831 | -0.1831 | 0.12817 | 0.27466 | -0.21515 | -0.27466 | -0.068665 | -0.12817 | 0.022888 | 0 | -0.082397 | -0.082397 | -0.032043 | -0.054931 | 0.14191 | -0.10986 | 0.11902 | -0.13275
769.1719 | 122.3877 | -58.8684 | 63.382 | -0.046616 | 0.0026234 | 0.067477 | 0.064006 | -0.0010959 | -0.00068783 | -0.00081179 | 0.035309 | 0.00038669 | 0.063852 | 0.060842 | -0.0012497 | -0.0024075 | 0.0019375 | 0.036621 | 63.3316 | -59.0012 | 63.2629 | -58.6807 | -59.0973 | -59.0515 | 0.045776 | 63.3774 | -58.9142 | 63.327 | -58.8364 | -58.9737 | -58.7311 | -0.13275 | -0.050354 | -0.11902 | -0.068665 | 0.18768 | 0.32043 | -0.18311 | -0.22888 | -0.050354 | -0.09613 | -0.045777 | -0.004578 | -0.054932 | -0.050354 | 0.032043 | 0.07782 | 0.13733 | -0.10529 | 0.18311 | -0.059509
770.1719 | 122.406 | -58.8684 | 63.382 | -0.023727 | -0.0020764 | 0.068118 | 0.068064 | -0.00045509 | -0.00047423 | 0.00068361 | 0.0047913 | 0.0018821 | 0.064066 | 0.064046 | -0.0027451 | -0.0024075 | -0.0025485 | -0.022888 | 63.382 | -59.0378 | 63.2263 | -58.7402 | -59.1431 | -59.1064 | 0.009155 | 63.4003 | -58.8959 | 63.2629 | -58.9096 | -58.9828 | -58.699 | -0.16937 | 0 | -0.15564 | -0.15564 | 0.12817 | 0.29755 | -0.23804 | -0.27466 | -0.068664 | -0.10528 | -0.027466 | 0.018311 | -0.11902 | -0.13733 | -0.041199 | -0.013733 | 0.16937 | -0.11444 | 0.19684 | -0.086975
771.1719 | 122.4655 | -58.8501 | 63.4003 | -0.0084687 | -0.00015382 | 0.068545 | 0.067851 | 0.0025357 | 0.0020892 | -0.0025208 | -0.03183 | 0.0027365 | 0.065775 | 0.062337 | -0.0016769 | -0.0015529 | -0.0010531 | 0.004578 | 63.3957 | -59.024 | 63.2263 | -58.754 | -59.1522 | -59.0927 | 0 | 63.3728 | -58.905 | 63.3499 | -58.8318 | -58.992 | -58.7082 | -0.17395 | -0.004578 | -0.17395 | -0.16937 | 0.096131 | 0.27008 | -0.24261 | -0.30212 | -0.068664 | -0.12817 | -0.054931 | -0.027466 | -0.050354 | -0.022888 | 0.018311 | 0.073242 | 0.14191 | -0.14191 | 0.19684 | -0.086975
772.1719 | 122.4426 | -58.8684 | 63.382 | -0.10002 | 5.9776e-05 | 0.069613 | 0.070842 | 0.0023221 | 0.0031574 | -0.0014526 | -0.056244 | 0.0010275 | 0.066843 | 0.064473 | -0.00018152 | -0.0015529 | -0.00041228 | 0.013733 | 63.3957 | -59.0195 | 63.2675 | -58.6853 | -59.1019 | -59.0607 | 0.064087 | 63.3636 | -58.9096 | 63.3911 | -58.8181 | -58.9737 | -58.6899 | -0.15106 | 0.013733 | -0.11444 | -0.12817 | 0.18311 | 0.33417 | -0.19226 | -0.23346 | -0.041199 | -0.082398 | -0.041199 | -0.018311 | 0.009155 | 0.027466 | 0.050354 | 0.091553 | 0.17853 | -0.10529 | 0.21973 | -0.064087
773.1719 | 122.4609 | -58.8638 | 63.4094 | -0.10002 | -0.00229 | 0.067691 | 0.069346 | 0.0012539 | 0.0010212 | -0.002948 | -0.022675 | 0.0010275 | 0.066843 | 0.067464 | -0.00039512 | -0.0028347 | 0.0012967 | -0.027466 | 63.3865 | -59.0332 | 63.2309 | -58.7448 | -59.1248 | -59.0927 | 0.004578 | 63.382 | -58.9005 | 63.3499 | -58.8318 | -58.992 | -58.7082 | -0.16937 | -0.022888 | -0.17853 | -0.15564 | 0.11902 | 0.28839 | -0.22888 | -0.26093 | -0.059509 | -0.091553 | -0.036621 | -0.027466 | -0.059509 | -0.032043 | 0.032044 | 0.068665 | 0.15564 | -0.12817 | 0.19226 | -0.091552
774.1719 | 122.4197 | -58.8593 | 63.446 | -0.069504 | 0.0017688 | 0.067691 | 0.068278 | 0.0018949 | 0.00016657 | -0.0023072 | -0.028778 | 0.0023093 | 0.066202 | 0.06191 | -0.00039512 | -0.00091214 | 0.0012967 | -0.050354 | 63.3911 | -59.0103 | 63.2263 | -58.7585 | -59.1293 | -59.0836 | 0 | 63.3774 | -58.9005 | 63.3636 | -58.8409 | -58.9874 | -58.7357 | -0.15106 | -0.054932 | -0.21973 | -0.16479 | 0.10071 | 0.25177 | -0.2243 | -0.27008 | -0.073242 | -0.11902 | -0.041199 | -0.068665 | -0.082398 | -0.013733 | 0.018311 | 0.05951 | 0.1236 | -0.12817 | 0.16479 | -0.086975
775.1719 | 122.4701 | -58.8638 | 63.4506 | -0.061874 | -0.0010082 | 0.068332 | 0.069987 | -0.0015231 | -0.00026063 | -0.00038459 | -0.0074159 | 0.00017309 | 0.063425 | 0.070028 | -0.0025315 | -0.0041165 | 0.00065592 | -0.018311 | 63.4186 | -59.0103 | 63.2492 | -58.7311 | -59.1293 | -59.0973 | -0.027466 | 63.3865 | -58.8959 | 63.3545 | -58.8272 | -58.992 | -58.7311 | -0.14648 | -0.032044 | -0.20142 | -0.16937 | 0.13275 | 0.27924 | -0.23346 | -0.2655 | -0.086975 | -0.11902 | -0.032043 | -0.064087 | -0.096131 | -0.032044 | 0.036622 | 0.068665 | 0.13275 | -0.12817 | 0.16479 | -0.09613
776.1719 | 122.4014 | -58.8409 | 63.4369 | -0.084763 | 5.9776e-05 | 0.066409 | 0.06721 | -0.0019503 | 0.00016657 | -0.002948 | -0.0074159 | 0.0042319 | 0.068124 | 0.061696 | 3.2078e-05 | -0.0041165 | -0.0012667 | -0.018311 | 63.382 | -58.992 | 63.2217 | -58.7128 | -59.1385 | -59.0927 | 0.013733 | 63.3682 | -58.8867 | 63.3453 | -58.8226 | -58.9828 | -58.7265 | -0.15106 | -0.054932 | -0.21515 | -0.16022 | 0.12817 | 0.27923 | -0.25177 | -0.29755 | -0.10071 | -0.14648 | -0.045777 | -0.068665 | -0.091553 | -0.022888 | 0.01831 | 0.064087 | 0.11444 | -0.14191 | 0.16022 | -0.09613
777.1719 | 122.4701 | -58.8135 | 63.414 | -0.061874 | -0.00015382 | 0.067904 | 0.069987 | 0.0023221 | -0.001115 | -0.0010254 | -0.028778 | 0.0020957 | 0.06492 | 0.064473 | -0.00082252 | -0.0019803 | 0.0017239 | 0 | 63.4003 | -58.992 | 63.2629 | -58.7036 | -59.0973 | -59.0607 | 0.032043 | 63.3957 | -58.8959 | 63.4003 | -58.786 | -58.9783 | -58.6853 | -0.17853 | -0.013732 | -0.15106 | -0.13733 | 0.10986 | 0.28839 | -0.24719 | -0.28381 | -0.068665 | -0.10529 | -0.082397 | -0.01831 | -0.013732 | 0.004578 | 0.027466 | 0.10986 | 0.12817 | -0.16479 | 0.21057 | -0.082397
778.1719 | 122.4426 | -58.8409 | 63.4094 | -0.046616 | 5.9776e-05 | 0.067477 | 0.068064 | 0.0025357 | 0.0012348 | -0.00059819 | -0.062347 | 0.00017309 | 0.066202 | 0.064473 | -0.00018152 | -0.0030483 | -0.0023349 | -0.009155 | 63.4003 | -58.9783 | 63.2538 | -58.7173 | -59.0973 | -59.0561 | 0.013733 | 63.3957 | -58.8959 | 63.3865 | -58.8043 | -58.96 | -58.6761 | -0.13733 | -0.009155 | -0.15564 | -0.14648 | 0.1236 | 0.26093 | -0.21515 | -0.25635 | -0.07782 | -0.11902 | -0.054932 | -0.013733 | -0.022888 | -0.009155 | 0.036621 | 0.091553 | 0.16479 | -0.11902 | 0.21973 | -0.064087
779.1719 | 122.4747 | -58.8455 | 63.3911 | -0.031357 | 0.00070078 | 0.068759 | 0.067851 | -0.0015231 | 0.00080757 | -0.00038459 | -0.010467 | 0.0012411 | 0.064706 | 0.066182 | -0.0012497 | -0.0030483 | -0.0014805 | -0.036621 | 63.414 | -58.9828 | 63.2584 | -58.7082 | -59.1156 | -59.0652 | 0.009155 | 63.3911 | -58.8684 | 63.3865 | -58.8089 | -58.9737 | -58.6853 | -0.13733 | 0.022888 | -0.13275 | -0.15564 | 0.13733 | 0.27466 | -0.21973 | -0.27008 | -0.082398 | -0.13275 | -0.022888 | 0 | -0.004577 | -0.004577 | 0.036621 | 0.059509 | 0.16022 | -0.12817 | 0.18311 | -0.10529
780.1719 | 122.4243 | -58.8272 | 63.4369 | -0.077133 | -0.00058102 | 0.06705 | 0.070414 | -0.0028049 | 0.00038017 | 0.00025641 | -0.025726 | 0.00017309 | 0.064706 | 0.061696 | -0.0012497 | -0.0039029 | -0.0023349 | -0.032043 | 63.4094 | -58.9966 | 63.2538 | -58.7173 | -59.1156 | -59.0652 | -0.009155 | 63.4003 | -58.8913 | 63.4003 | -58.8089 | -58.9737 | -58.6716 | -0.16937 | -0.027466 | -0.18311 | -0.15564 | 0.10986 | 0.27924 | -0.23804 | -0.28839 | -0.068665 | -0.11902 | -0.064087 | -0.036621 | -0.036621 | 0 | 0.01831 | 0.082397 | 0.15564 | -0.14648 | 0.21973 | -0.082398
781.1719 | 122.4472 | -58.8318 | 63.446 | -0.069504 | -0.00036742 | 0.065768 | 0.068064 | -0.0019503 | 0.0029438 | -0.0025208 | -0.034882 | 0.00017309 | 0.065561 | 0.067037 | -0.0012497 | -0.0021939 | 0.0017239 | -0.004578 | 63.4003 | -59.0012 | 63.2492 | -58.7128 | -59.1248 | -59.0744 | 0.004578 | 63.4003 | -58.8821 | 63.3682 | -58.7952 | -58.9645 | -58.699 | -0.16937 | -0.045776 | -0.19684 | -0.15106 | 0.11902 | 0.28839 | -0.24262 | -0.29297 | -0.073242 | -0.1236 | -0.050354 | -0.045776 | -0.07782 | -0.032044 | 0.036621 | 0.086975 | 0.13275 | -0.13275 | 0.1831 | -0.082398
782.1719 | 122.4838 | -58.8043 | 63.3911 | -0.038986 | 0.001128 | 0.068545 | 0.067637 | -0.0015231 | 0.00080757 | 0.003247 | -0.022675 | 0.00081389 | 0.064279 | 0.063833 | 3.2078e-05 | -0.0026211 | 0.00022852 | 0 | 63.4094 | -58.9691 | 63.2858 | -58.667 | -59.0881 | -59.0561 | 0.059509 | 63.3774 | -58.8913 | 63.4003 | -58.7723 | -58.96 | -58.6899 | -0.16479 | 0.018311 | -0.10528 | -0.1236 | 0.13733 | 0.30212 | -0.25177 | -0.28381 | -0.086975 | -0.11902 | -0.086975 | -0.013733 | 0.009156 | 0.022889 | 0.032043 | 0.11902 | 0.11444 | -0.15564 | 0.20142 | -0.068665
783.1719 | 122.4335 | -58.8135 | 63.4369 | -0.031357 | 5.9776e-05 | 0.067904 | 0.068492 | -0.0023777 | -0.00068783 | -0.00038459 | 0.0017395 | 0.0033773 | 0.064279 | 0.058278 | 0.0017411 | -0.0036893 | -0.0031895 | -0.027466 | 63.4186 | -58.96 | 63.2355 | -58.7128 | -59.1019 | -59.0469 | 0.013733 | 63.414 | -58.8593 | 63.3636 | -58.8089 | -58.9691 | -58.7036 | -0.14648 | -0.018311 | -0.20142 | -0.1831 | 0.10071 | 0.24719 | -0.23346 | -0.28839 | -0.086975 | -0.14191 | -0.045776 | -0.022889 | -0.073243 | -0.050354 | 0.004578 | 0.050354 | 0.10986 | -0.15564 | 0.15564 | -0.10986
784.1719 | 122.4747 | -58.8089 | 63.4277 | -0.061874 | 0.001128 | 0.068332 | 0.06956 | 0.0025357 | -0.0015424 | -0.002948 | -0.016571 | 0.00038669 | 0.064493 | 0.066182 | -0.0010361 | -0.0030483 | 0.00086952 | -0.022888 | 63.4232 | -58.9828 | 63.2767 | -58.6807 | -59.0927 | -59.0469 | 0.036621 | 63.4094 | -58.8501 | 63.3957 | -58.7677 | -58.9508 | -58.6533 | -0.17395 | -0.004577 | -0.15106 | -0.14649 | 0.12817 | 0.30212 | -0.23804 | -0.28381 | -0.064087 | -0.10986 | -0.041199 | -0.01831 | -0.032043 | -0.013733 | 0.041199 | 0.082398 | 0.15564 | -0.14191 | 0.19684 | -0.10071
785.1719 | 122.4289 | -58.8181 | 63.4277 | -0.10765 | 0.0015552 | 0.066623 | 0.067424 | 0.0029629 | 0.00059397 | 0.0019654 | -0.03183 | 0.0010275 | 0.065347 | 0.064901 | 3.2078e-05 | -0.0024075 | 0.00086952 | 0 | 63.4323 | -58.9783 | 63.2629 | -58.6899 | -59.0927 | -59.0378 | 0.004578 | 63.4232 | -58.8547 | 63.414 | -58.7952 | -58.9417 | -58.6761 | -0.16022 | 0.004578 | -0.16479 | -0.16937 | 0.12817 | 0.28839 | -0.21973 | -0.27466 | -0.05951 | -0.11444 | -0.036621 | -0.004577 | -0.013733 | -0.009156 | 0.022888 | 0.059509 | 0.14191 | -0.1236 | 0.17853 | -0.086975
786.1719 | 122.4609 | -58.8181 | 63.4415 | -0.046616 | 0.00048698 | 0.066836 | 0.067637 | 0.0014675 | 0.0012348 | 0.0011108 | -0.0074159 | -0.00068151 | 0.064066 | 0.061483 | -0.00039512 | -0.0011257 | 0.00086952 | -0.013733 | 63.4186 | -58.992 | 63.2629 | -58.6853 | -59.1019 | -59.0515 | 0.009155 | 63.4048 | -58.8867 | 63.4003 | -58.7814 | -58.9417 | -58.6533 | -0.17395 | -0.022888 | -0.17853 | -0.15564 | 0.13275 | 0.3067 | -0.23346 | -0.28381 | -0.05951 | -0.10986 | -0.068665 | -0.036621 | -0.041198 | -0.004577 | 0.036621 | 0.10529 | 0.16479 | -0.1236 | 0.23346 | -0.054931
787.1719 | 122.4289 | -58.8089 | 63.4369 | -0.046616 | -0.0025036 | 0.068545 | 0.068492 | 0.00039951 | -0.0013288 | 0.00025641 | 0.0047913 | 0.0014547 | 0.066843 | 0.061056 | -0.00018152 | -0.0028347 | 0.0021511 | -0.018311 | 63.4232 | -58.9737 | 63.2629 | -58.6945 | -59.1064 | -59.0561 | 0 | 63.414 | -58.8501 | 63.382 | -58.786 | -58.9508 | -58.6716 | -0.1648 | -0.013733 | -0.17395 | -0.16022 | 0.11444 | 0.27924 | -0.24719 | -0.29755 | -0.082397 | -0.13275 | -0.041199 | -0.022889 | -0.054932 | -0.032043 | 0.022888 | 0.064087 | 0.13733 | -0.14191 | 0.17853 | -0.10071
788.1719 | 122.4518 | -58.7997 | 63.4369 | -0.0084687 | 0.0013416 | 0.068118 | 0.06721 | -0.00024129 | -0.00026063 | -0.00059819 | -0.03183 | 0.0012411 | 0.066415 | 0.064687 | -0.0010361 | -0.0034757 | 0.0021511 | 0.009155 | 63.4186 | -58.96 | 63.2858 | -58.6533 | -59.0607 | -59.0286 | 0.059509 | 63.4048 | -58.8547 | 63.414 | -58.7677 | -58.9417 | -58.6441 | -0.16022 | -0.018311 | -0.15106 | -0.13275 | 0.14648 | 0.3067 | -0.22888 | -0.26093 | -0.068664 | -0.10071 | -0.054931 | -0.032044 | -0.022889 | 0.009155 | 0.032044 | 0.086975 | 0.15564 | -0.14191 | 0.21057 | -0.086975
789.4062 | 122.438 | -58.8318 | 63.4048 | -0.046616 | -0.0014356 | 0.067264 | 0.069773 | 0.0033901 | 0.0010212 | -0.00059819 | -0.059296 | 0.0018821 | 0.064279 | 0.064901 | -0.00082252 | -0.0036893 | 0.0012967 | -0.1236 | 63.3865 | -58.9966 | 63.2217 | -58.7311 | -59.1614 | -59.0973 | -0.11444 | 63.3911 | -58.9096 | 63.3957 | -58.8272 | -59.0332 | -58.7082 | -0.16479 | -0.01831 | -0.1831 | -0.16479 | 0.10071 | 0.2655 | -0.2655 | -0.32959 | -0.10071 | -0.1648 | -0.07782 | -0.013733 | -0.009155 | 0.004578 | 0.004578 | 0.082398 | 0.1236 | -0.20142 | 0.20142 | -0.1236
790.2969 | 122.4335 | -58.8364 | 63.4186 | -0.077133 | -0.00079462 | 0.068759 | 0.069987 | -0.0013095 | -0.00068783 | -0.00017099 | -0.089813 | -0.0021769 | 0.063638 | 0.067251 | 0.0011001 | -0.0013393 | 0.0012967 | -0.09613 | 63.3865 | -59.0103 | 63.2126 | -58.7402 | -59.1843 | -59.3308 | 0.10529 | 63.3911 | -58.8776 | 63.3774 | -58.8226 | -59.0378 | -58.7036 | -0.17395 | -0.032043 | -0.20599 | -0.17395 | 0.096131 | 0.27008 | -0.49439 | -0.3479 | -0.32044 | -0.17395 | -0.041198 | -0.027466 | -0.041199 | -0.013733 | 0.013733 | 0.054931 | 0.13275 | -0.20142 | 0.17395 | -0.16022
791.1875 | 122.4472 | -58.8089 | 63.4003 | -0.077133 | 5.9776e-05 | 0.070468 | 0.066996 | 0.00061311 | 0.00080757 | 4.2808e-05 | -0.0013123 | -0.00089511 | 0.064706 | 0.062764 | -0.0014633 | 0.00036966 | -0.0016941 | -0.041199 | 63.3408 | -58.9691 | 63.3224 | -58.6029 | -59.0515 | -58.9691 | 0.009155 | 63.3179 | -58.8593 | 63.3591 | -58.58 | -58.7036 | -58.5114 | -0.16022 | -0.05951 | -0.07782 | -0.01831 | 0.20599 | 0.36621 | -0.16022 | -0.24262 | 0 | -0.082398 | -0.050354 | -0.082398 | -0.041199 | 0.041199 | 0.22888 | 0.27924 | 0.29755 | 0.10529 | 0.3479 | 0.15564
792.3125 | 122.5479 | -58.7997 | 63.3636 | -0.1458 | -0.00058102 | 0.068545 | 0.068919 | 0.0057401 | -0.00026063 | 4.2808e-05 | -0.0074159 | -0.00089511 | 0.063638 | 0.055929 | -0.0029587 | -0.0049709 | -0.00019868 | -0.068665 | 63.3636 | -58.9417 | 63.3362 | -58.5938 | -59.079 | -59.0057 | -0.013733 | 63.3682 | -58.8547 | 63.4781 | -58.7265 | -58.8913 | -58.5846 | -0.14191 | 0 | -0.027465 | -0.027465 | 0.20599 | 0.3479 | -0.20599 | -0.27923 | -0.064087 | -0.13733 | -0.054931 | 0.004578 | 0.11444 | 0.10986 | 0.073243 | 0.12817 | 0.21515 | -0.091552 | 0.27008 | -0.036621
793.1875 | 122.5067 | -58.7814 | 63.382 | -0.13054 | 0.00070078 | 0.070254 | 0.068705 | 0.0010403 | 0.0012348 | -0.00038459 | -0.019623 | -0.0011087 | 0.06492 | 0.060628 | -0.0023177 | -0.0032619 | -0.0016941 | -0.059509 | 63.3499 | -58.9828 | 63.3453 | -58.6029 | -59.0561 | -59.0012 | -0.013733 | 63.3545 | -58.873 | 63.4735 | -58.7173 | -58.8867 | -58.5892 | -0.20142 | -0.032043 | -0.036621 | -0.004578 | 0.17853 | 0.37994 | -0.21973 | -0.27466 | -0.018311 | -0.073242 | -0.091553 | -0.027466 | 0.091553 | 0.11902 | 0.064087 | 0.15564 | 0.19226 | -0.10529 | 0.28381 | -0.013733
794.1875 | 122.4197 | -58.7997 | 63.4369 | -0.046616 | -0.0016492 | 0.068759 | 0.069133 | -0.00066869 | 0.0042254 | -0.00038459 | -0.065399 | -4.0508e-05 | 0.065134 | 0.069387 | -0.0018905 | -0.0011257 | 0.0038601 | -0.009155 | 63.4232 | -58.9508 | 63.2172 | -58.5892 | -59.0698 | -59.0652 | 0.013733 | 63.4323 | -58.8501 | 63.4415 | -58.7769 | -58.9279 | -58.6212 | -0.15106 | -0.013733 | -0.21973 | -0.20599 | 0.21057 | 0.36163 | -0.2655 | -0.27008 | -0.11444 | -0.11902 | -0.050354 | -0.004578 | 0.004577 | 0.009155 | 0.022889 | 0.073243 | 0.17853 | -0.12817 | 0.22888 | -0.077819
795.2344 | 122.4884 | -58.7997 | 63.4369 | -0.016098 | 0.0019824 | 0.066836 | 0.069133 | 0.0010403 | -0.00090143 | -0.00081179 | -0.037933 | 0.00038669 | 0.061929 | 0.066823 | -0.0016769 | -0.0028347 | -0.0010531 | 0.009155 | 63.4689 | -58.9554 | 63.3133 | -58.6395 | -59.0515 | -59.0149 | 0.064087 | 63.4415 | -58.8638 | 63.4552 | -58.7494 | -58.9279 | -58.6487 | -0.15564 | 0.032043 | -0.1236 | -0.15564 | 0.16022 | 0.31586 | -0.21515 | -0.25177 | -0.05951 | -0.096131 | -0.064087 | 0.004577 | 0.01831 | 0.013733 | 0.050354 | 0.11444 | 0.15106 | -0.12817 | 0.21515 | -0.064086
796.1875 | 122.438 | -58.7952 | 63.4644 | -0.077133 | 0.00091438 | 0.066836 | 0.067424 | -0.00088229 | 0.0018756 | 0.00068361 | -0.016571 | -0.0011087 | 0.065134 | 0.063192 | -0.00039512 | -0.0041165 | 0.0036465 | -0.027466 | 63.4369 | -58.96 | 63.2858 | -58.6807 | -59.0469 | -59.0195 | -0.004578 | 63.4186 | -58.8364 | 63.4323 | -58.7631 | -58.9417 | -58.6624 | -0.16479 | -0.027465 | -0.17853 | -0.15106 | 0.11444 | 0.27924 | -0.2243 | -0.25177 | -0.059509 | -0.086975 | -0.041199 | -0.045776 | -0.032043 | 0.013733 | 0.032043 | 0.073242 | 0.13275 | -0.14648 | 0.17395 | -0.10529
797.2969 | 122.4426 | -58.7906 | 63.4552 | -0.046616 | -0.0018628 | 0.070895 | 0.068919 | -0.00066869 | -0.00026063 | 0.00068361 | -0.040985 | 0.0012411 | 0.066415 | 0.060628 | -0.0012497 | -0.0032619 | 0.0023647 | -0.013733 | 63.4415 | -58.9417 | 63.3224 | -58.6121 | -59.0607 | -59.0057 | 0.045776 | 63.4277 | -58.8226 | 63.4964 | -58.8821 | -58.9279 | -58.6166 | -0.15106 | -0.013733 | -0.13275 | -0.11902 | 0.17853 | 0.32959 | -0.21515 | -0.27008 | -0.064087 | -0.11902 | -0.032044 | -0.027466 | 0.041199 | 0.068665 | -0.091553 | -0.059509 | 0.17395 | -0.13733 | 0.20599 | -0.10529
798.1875 | 122.4152 | -58.7906 | 63.4323 | -0.061874 | -0.00015382 | 0.068759 | 0.06956 | 0.00039951 | -0.00047423 | -0.0025208 | -0.044037 | 0.0029501 | 0.063638 | 0.062978 | -0.0012497 | -0.0032619 | 0.0010831 | -0.009155 | 63.4369 | -58.9508 | 63.3224 | -58.6166 | -59.0607 | -59.0195 | 0.013733 | 63.4415 | -58.8318 | 63.4186 | -58.7723 | -58.96 | -58.667 | -0.16022 | 0.004578 | -0.10986 | -0.11444 | 0.17395 | 0.33417 | -0.22888 | -0.27008 | -0.068664 | -0.10986 | -0.041199 | 0.009155 | -0.013733 | -0.022888 | 0.01831 | 0.059509 | 0.1236 | -0.16937 | 0.16479 | -0.12817
799.2969 | 122.4243 | -58.786 | 63.4598 | -0.016098 | 0.00091438 | 0.068973 | 0.069346 | 0.00018591 | -0.00026063 | -0.0016662 | -0.044037 | 0.0027365 | 0.064493 | 0.062551 | -0.00018152 | -0.0026211 | 0.0042875 | 0.013733 | 63.4689 | -58.9096 | 63.3362 | -58.6166 | -59.0424 | -58.9828 | 0.050354 | 63.414 | -58.8409 | 63.4644 | -58.7265 | -58.9096 | -58.6121 | -0.1236 | 0.009155 | -0.1236 | -0.13275 | 0.16937 | 0.29297 | -0.19684 | -0.25635 | -0.073242 | -0.13275 | -0.054931 | -0.045777 | 0.004577 | 0.050354 | 0.05951 | 0.11444 | 0.17395 | -0.1236 | 0.22888 | -0.068665
800.1719 | 122.4289 | -58.8318 | 63.4277 | -0.11528 | -0.00229 | 0.068973 | 0.067851 | 0.0033901 | 0.0023028 | -0.0016662 | 0.016998 | 0.0012411 | 0.067484 | 0.065969 | 0.00067288 | -0.0019803 | 0.0012967 | -0.10529 | 63.4232 | -58.9691 | 63.4277 | -58.4473 | -59.1797 | -59.0927 | -0.11902 | 63.4186 | -58.8409 | 63.4048 | -58.8089 | -58.9966 | -58.6807 | -0.13733 | -0.004577 | 0 | 0.004577 | 0.38452 | 0.52185 | -0.26093 | -0.3479 | -0.1236 | -0.21057 | -0.009155 | -0.009155 | -0.022888 | -0.013733 | 0.022888 | 0.032043 | 0.15106 | -0.16479 | 0.16022 | -0.15564
801.1719 | 122.4609 | -58.7952 | 63.5056 | -0.084763 | -0.00015382 | 0.0694 | 0.06721 | -0.0023777 | 0.00059397 | -0.0020936 | -0.0013123 | 0.0020957 | 0.066202 | 0.061483 | 0.0013139 | -0.0049709 | -0.0053257 | -0.032043 | 63.4735 | -58.9325 | 63.2904 | -58.6487 | -59.0515 | -59.0057 | -0.027466 | 63.4598 | -58.8089 | 63.4232 | -58.7311 | -58.9233 | -58.6395 | -0.13733 | -0.032043 | -0.21515 | -0.18311 | 0.14648 | 0.28381 | -0.21057 | -0.25635 | -0.073242 | -0.11902 | -0.013733 | -0.045776 | -0.082397 | -0.036621 | 0.064087 | 0.07782 | 0.15564 | -0.12817 | 0.16937 | -0.11444
802.1719 | 122.4106 | -58.7906 | 63.4552 | -0.12291 | -0.00079462 | 0.068118 | 0.070201 | 0.0018949 | -0.00068783 | 0.00089721 | -0.040985 | 0.0016685 | 0.064066 | 0.065328 | -0.0012497 | -0.0030483 | -0.00019868 | -0.027466 | 63.4277 | -58.9188 | 63.3224 | -58.6166 | -59.0744 | -59.0149 | 0.004578 | 63.446 | -58.8455 | 63.4369 | -58.7585 | -58.9233 | -58.6349 | -0.12817 | -0.027466 | -0.13275 | -0.10529 | 0.17395 | 0.30212 | -0.22431 | -0.28381 | -0.096131 | -0.15564 | -0.054932 | -0.009155 | -0.01831 | -0.009155 | 0.032043 | 0.086975 | 0.15564 | -0.13275 | 0.21057 | -0.07782
803.1719 | 122.4014 | -58.786 | 63.501 | -0.054245 | -0.0020764 | 0.067477 | 0.067637 | -0.0032321 | -0.00068783 | 0.00068361 | 0.038361 | 0.0027365 | 0.064706 | 0.065114 | -0.00039512 | -0.0053983 | 0.0036465 | -0.013733 | 63.4552 | -58.905 | 63.295 | -58.6349 | -59.0652 | -58.9874 | 0.013733 | 63.4781 | -58.7906 | 63.4323 | -58.754 | -58.905 | -58.6533 | -0.11902 | -0.045777 | -0.20599 | -0.16022 | 0.15106 | 0.27008 | -0.20142 | -0.27924 | -0.082398 | -0.16022 | -0.004577 | -0.022889 | -0.068665 | -0.045776 | 0.032044 | 0.036621 | 0.13275 | -0.11902 | 0.13733 | -0.11444
804.1719 | 122.374 | -58.7723 | 63.533 | -0.046616 | -0.00015382 | 0.069613 | 0.066569 | -0.0043003 | -0.0015424 | -0.0025208 | 0.023102 | 0.0016685 | 0.06492 | 0.059987 | -0.00039512 | -0.0045437 | 0.00044212 | -0.018311 | 63.4598 | -58.9142 | 63.3179 | -58.6029 | -59.0286 | -58.9966 | 0.013733 | 63.4827 | -58.8226 | 63.4323 | -58.7357 | -58.9096 | -58.6533 | -0.14191 | -0.073242 | -0.21515 | -0.14191 | 0.16937 | 0.31128 | -0.2243 | -0.25635 | -0.082397 | -0.11444 | -0.050354 | -0.050354 | -0.10071 | -0.050354 | 0.036621 | 0.086975 | 0.11902 | -0.13733 | 0.16937 | -0.086975
805.1719 | 122.5067 | -58.7585 | 63.4964 | -0.061874 | -0.0010082 | 0.068332 | 0.067851 | 0.00061311 | 0.0014484 | 0.00025641 | -0.0013123 | 0.0010275 | 0.065561 | 0.065755 | -0.0021041 | -0.0017667 | 0.0038601 | -0.022888 | 63.4872 | -58.9096 | 63.3362 | -58.6349 | -59.0378 | -58.9783 | 0.013733 | 63.4735 | -58.8043 | 63.4644 | -58.7128 | -58.8821 | -58.6075 | -0.15106 | -0.009155 | -0.16022 | -0.15106 | 0.1236 | 0.27466 | -0.21973 | -0.27924 | -0.068664 | -0.12817 | -0.045776 | -0.022888 | -0.032044 | -0.009156 | 0.045776 | 0.091552 | 0.15106 | -0.1236 | 0.19684 | -0.07782
806.1719 | 122.4243 | -58.7723 | 63.501 | -0.084763 | 0.00048698 | 0.067904 | 0.066783 | 0.00018591 | -0.0023968 | 0.001538 | -0.0043641 | -0.0013223 | 0.065775 | 0.064046 | 0.00045928 | -0.0039029 | -0.0014805 | 0 | 63.5147 | -58.9005 | 63.3362 | -58.6258 | -59.0378 | -58.992 | 0.009155 | 63.4689 | -58.7997 | 63.4598 | -58.7219 | -58.8821 | -58.6029 | -0.12817 | 0.013732 | -0.16479 | -0.17853 | 0.14648 | 0.27466 | -0.21973 | -0.2655 | -0.091552 | -0.13733 | -0.027466 | -0.032044 | -0.041199 | -0.009155 | 0.050354 | 0.07782 | 0.16937 | -0.10986 | 0.19684 | -0.082397
807.1719 | 122.4426 | -58.754 | 63.5101 | -0.038986 | -0.00015382 | 0.067264 | 0.069773 | 0.0010403 | -0.002824 | 4.2808e-05 | -0.025726 | 0.0025229 | 0.066202 | 0.063192 | -0.0016769 | -0.0041165 | 0.0010831 | -0.018311 | 63.4918 | -58.905 | 63.3224 | -58.6304 | -59.0332 | -58.9828 | 0.004578 | 63.4872 | -58.8089 | 63.4598 | -58.7082 | -58.8913 | -58.6212 | -0.15106 | -0.018311 | -0.18768 | -0.16937 | 0.1236 | 0.27466 | -0.22888 | -0.27924 | -0.07782 | -0.12817 | -0.054932 | -0.022888 | -0.050354 | -0.027466 | 0.045776 | 0.10071 | 0.13275 | -0.13733 | 0.18768 | -0.082397
808.1719 | 122.4152 | -58.7494 | 63.5239 | -0.10765 | -0.0025036 | 0.067691 | 0.069773 | 0.0044583 | 0.00080757 | -0.0035888 | -0.019623 | 0.0018821 | 0.065561 | 0.06426 | -0.0016769 | -0.0045437 | 0.00044212 | -0.036621 | 63.4918 | -58.8913 | 63.3453 | -58.6395 | -59.0286 | -58.9737 | -0.009155 | 63.4827 | -58.8135 | 63.446 | -58.7082 | -58.8867 | -58.6121 | -0.14191 | -0.032044 | -0.17853 | -0.14648 | 0.10986 | 0.25177 | -0.2243 | -0.27923 | -0.082398 | -0.13733 | -0.064087 | -0.041199 | -0.07782 | -0.036621 | 0.041199 | 0.10529 | 0.13733 | -0.13733 | 0.20142 | -0.073242
809.1719 | 122.4609 | -58.7585 | 63.4918 | -0.11528 | 5.9776e-05 | 0.068545 | 0.068492 | -0.00066869 | 0.0023028 | -0.001239 | 0.0047913 | -4.0508e-05 | 0.065561 | 0.067251 | 0.00067288 | -0.0019803 | -0.0029757 | -0.091553 | 63.501 | -58.9096 | 63.3133 | -58.6166 | -59.0286 | -58.9828 | 0.009155 | 63.4552 | -58.7906 | 63.4644 | -58.7173 | -58.8821 | -58.6029 | -0.15106 | 0.009156 | -0.17853 | -0.18768 | 0.14191 | 0.29297 | -0.2243 | -0.27008 | -0.073242 | -0.11902 | -0.032043 | -0.036621 | -0.027466 | 0.009155 | 0.041199 | 0.073242 | 0.15564 | -0.1236 | 0.18768 | -0.091553
810.1719 | 122.3877 | -58.7357 | 63.5147 | -0.023727 | -0.001222 | 0.067904 | 0.065501 | 0.00018591 | -4.7032e-05 | -0.0014526 | 0.053619 | 0.0016685 | 0.06727 | 0.061696 | 0.00088648 | -0.0053983 | 0.0015103 | 0.004578 | 63.5101 | -58.8867 | 63.3545 | -58.5617 | -59.024 | -58.9554 | 0.009155 | 63.4735 | -58.7769 | 63.4552 | -58.7173 | -58.9005 | -58.6258 | -0.15106 | -0.004577 | -0.16022 | -0.15564 | 0.17395 | 0.32501 | -0.21973 | -0.28839 | -0.068664 | -0.13733 | -0.041198 | -0.041198 | -0.059509 | -0.018311 | 0.018311 | 0.059509 | 0.10986 | -0.16479 | 0.15106 | -0.1236
811.1719 | 122.4472 | -58.7265 | 63.5284 | -0.084763 | 0.0019824 | 0.068759 | 0.066996 | -0.00066869 | -0.00068783 | 0.0017518 | 0.016998 | 0.00017309 | 0.06492 | 0.064046 | -0.00039512 | -0.0028347 | 0.0021511 | 0.036621 | 63.4872 | -58.8593 | 63.3728 | -58.5571 | -58.9828 | -58.9188 | 0.045776 | 63.5056 | -58.7677 | 63.4689 | -58.6853 | -58.8821 | -58.6212 | -0.13275 | -0.041198 | -0.15564 | -0.11444 | 0.16937 | 0.30212 | -0.19226 | -0.25635 | -0.059509 | -0.1236 | -0.041199 | -0.022888 | -0.059509 | -0.036621 | 0.041198 | 0.082397 | 0.10529 | -0.15564 | 0.14648 | -0.11444
812.1719 | 122.4106 | -58.7494 | 63.5513 | -0.12291 | -0.001222 | 0.067904 | 0.066996 | 0.00061311 | 0.00016657 | -0.0016662 | -0.028778 | 0.00081389 | 0.068124 | 0.067678 | -0.0018905 | -0.0030483 | -0.0044711 | 0.013733 | 63.5101 | -58.8913 | 63.3453 | -58.5938 | -59.0103 | -58.96 | 0 | 63.4918 | -58.7769 | 63.4598 | -58.7128 | -58.8776 | -58.5892 | -0.14191 | -0.041199 | -0.20599 | -0.16479 | 0.15564 | 0.29755 | -0.21057 | -0.26093 | -0.068665 | -0.11902 | -0.027465 | -0.05951 | -0.091553 | -0.032043 | 0.036621 | 0.064086 | 0.16022 | -0.12817 | 0.18768 | -0.10071
813.1719 | 122.438 | -58.7265 | 63.5422 | -0.18394 | 0.00070078 | 0.06705 | 0.067851 | 0.0016813 | 0.00016657 | 0.0034606 | -0.016571 | 0.0014547 | 0.064706 | 0.069814 | -0.0021041 | -0.0030483 | 0.00065592 | -0.022888 | 63.501 | -58.873 | 63.327 | -58.58 | -59.0149 | -58.9737 | 0 | 63.5147 | -58.7631 | 63.4644 | -58.7036 | -58.8593 | -58.6029 | -0.14648 | -0.041198 | -0.21515 | -0.17395 | 0.14648 | 0.29297 | -0.24719 | -0.28839 | -0.10071 | -0.14191 | -0.036622 | -0.027466 | -0.07782 | -0.050354 | 0.022888 | 0.05951 | 0.1236 | -0.13275 | 0.16022 | -0.09613
814.1719 | 122.4243 | -58.7036 | 63.5193 | -0.11528 | -0.0016492 | 0.066836 | 0.068705 | -0.0019503 | -0.001756 | -0.00038459 | -0.03183 | 0.0018821 | 0.065134 | 0.058278 | -0.0016769 | -0.0028347 | 0.0015103 | -0.018311 | 63.5239 | -58.8593 | 63.2629 | -58.754 | -59.0012 | -58.9691 | 0.013733 | 63.4872 | -58.786 | 63.4735 | -58.699 | -58.8593 | -58.5892 | -0.15564 | 0.004578 | -0.25635 | -0.26093 | -0.050354 | 0.10529 | -0.2655 | -0.29755 | -0.10986 | -0.14191 | -0.082398 | -0.032043 | -0.045776 | -0.013733 | 0.004577 | 0.086975 | 0.11444 | -0.15564 | 0.19684 | -0.073242
815.1719 | 122.4243 | -58.754 | 63.5376 | -0.031357 | -0.00058102 | 0.068545 | 0.067424 | -0.0040867 | 0.0010212 | -0.0018798 | -0.0013123 | 0.0010275 | 0.065347 | 0.060842 | 3.2078e-05 | -0.0039029 | 0.0045011 | -0.036621 | 63.533 | -58.8593 | 63.3682 | -58.5754 | -58.9874 | -58.96 | -0.054932 | 63.533 | -58.754 | 63.4918 | -58.7402 | -58.8684 | -58.58 | -0.10529 | -0.004578 | -0.16937 | -0.16479 | 0.17853 | 0.28381 | -0.20599 | -0.23346 | -0.10071 | -0.12817 | 0 | -0.004578 | -0.045777 | -0.041199 | 0.013733 | 0.013733 | 0.17395 | -0.11444 | 0.17395 | -0.11444
816.1719 | 122.4289 | -58.7357 | 63.501 | -0.11528 | 5.9776e-05 | 0.068759 | 0.070842 | -0.0015231 | 0.00059397 | 0.002179 | -0.013519 | -0.00089511 | 0.064706 | 0.061483 | -0.00082252 | -0.0021939 | 0.0021511 | -0.09613 | 63.4872 | -58.9279 | 63.3087 | -58.6578 | -59.2346 | -59.2392 | 0.027466 | 63.5056 | -58.8135 | 63.4506 | -58.7311 | -58.9325 | -58.6121 | -0.19226 | -0.013733 | -0.19226 | -0.17853 | 0.07782 | 0.27008 | -0.50354 | -0.49896 | -0.31128 | -0.3067 | -0.07782 | 0.004577 | -0.050354 | -0.054931 | 0.004578 | 0.082398 | 0.1236 | -0.19684 | 0.20142 | -0.11902
817.1719 | 122.4472 | -58.699 | 63.5239 | -0.031357 | 5.9776e-05 | 0.068545 | 0.068705 | -0.0021641 | -0.0023968 | 0.001538 | -0.040985 | 0.00017309 | 0.062997 | 0.067037 | -0.0010361 | -0.0032619 | 0.0023647 | -0.013733 | 63.5284 | -58.873 | 63.3865 | -58.5617 | -58.9691 | -58.9417 | 0.018311 | 63.5239 | -58.7585 | 63.5284 | -58.6395 | -58.8318 | -58.5709 | -0.17395 | 0.004577 | -0.13733 | -0.14191 | 0.13733 | 0.31128 | -0.24261 | -0.27008 | -0.068664 | -0.09613 | -0.059509 | 0 | 0.004577 | 0.004577 | 0.05951 | 0.11902 | 0.12817 | -0.13275 | 0.18768 | -0.073242
818.1719 | 122.3877 | -58.7402 | 63.5696 | -0.016098 | 0.00027338 | 0.068973 | 0.066783 | -0.0045139 | -0.00068783 | 0.0011108 | 0.047516 | -0.0013223 | 0.064493 | 0.061056 | -0.0014633 | -0.0021939 | 0.0030057 | 0.022888 | 63.5147 | -58.8593 | 63.4003 | -58.5205 | -58.96 | -58.9096 | 0.064087 | 63.501 | -58.7494 | 63.4689 | -58.6716 | -58.8364 | -58.6075 | -0.11902 | -0.054932 | -0.16937 | -0.11444 | 0.21973 | 0.33875 | -0.16937 | -0.21973 | -0.050354 | -0.10071 | -0.009156 | -0.068664 | -0.10071 | -0.032044 | 0.068664 | 0.07782 | 0.13275 | -0.096131 | 0.14191 | -0.086975
819.1719 | 122.3831 | -58.7036 | 63.533 | -0.046616 | 0.00091438 | 0.066836 | 0.066142 | -0.0010959 | 0.0010212 | -0.0038024 | 0.013947 | 0.00060029 | 0.064493 | 0.062124 | 0.0011001 | -0.0058255 | 0.0021511 | -0.036621 | 63.5422 | -58.8547 | 63.3682 | -58.5617 | -58.992 | -58.9188 | 0.013733 | 63.5376 | -58.7357 | 63.4506 | -58.4976 | -58.7082 | -58.5846 | -0.15106 | 0.009155 | -0.16479 | -0.17395 | 0.14191 | 0.29297 | -0.21515 | -0.28839 | -0.064087 | -0.13733 | -0.032044 | 0.004578 | -0.082397 | -0.086975 | 0.20599 | 0.23804 | 0.11902 | -0.004578 | 0.15106 | 0.027466
820.1719 | 122.4014 | -58.6945 | 63.5513 | -0.038986 | -0.00079462 | 0.068545 | 0.068278 | 0.00039951 | 0.00038017 | -0.0035888 | 0.0078429 | 0.0018821 | 0.066415 | 0.063192 | 3.2078e-05 | -0.0036893 | -0.0066073 | -0.004578 | 63.5468 | -58.8364 | 63.3865 | -58.5526 | -58.9737 | -58.9233 | 0.027466 | 63.533 | -58.7402 | 63.4964 | -58.6716 | -58.8409 | -58.58 | -0.14191 | -0.004578 | -0.16479 | -0.16022 | 0.14191 | 0.28381 | -0.22888 | -0.27924 | -0.086975 | -0.13733 | -0.045776 | -0.018311 | -0.054932 | -0.036621 | 0.022888 | 0.068664 | 0.11444 | -0.14648 | 0.16022 | -0.10071
821.1875 | 122.438 | -58.7402 | 63.5239 | -0.18394 | 0.00027338 | 0.065982 | 0.068278 | -0.00088229 | 0.00059397 | -0.00081179 | -0.05014 | 0.0014547 | 0.06492 | 0.068746 | 0.00067288 | -0.0017667 | -0.0025485 | 0.086975 | 63.5284 | -58.8867 | 63.3499 | -58.5892 | -59.079 | -59.024 | -0.10986 | 63.4872 | -58.7631 | 63.4918 | -58.7128 | -58.8959 | -58.5846 | -0.14648 | 0.004577 | -0.17395 | -0.17853 | 0.15106 | 0.29755 | -0.28381 | -0.33874 | -0.13733 | -0.19226 | -0.022889 | -0.036621 | -0.032044 | 0.004577 | 0.027465 | 0.050354 | 0.15564 | -0.15564 | 0.17853 | -0.13275
822.1719 | 122.4426 | -58.699 | 63.5468 | -0.023727 | 0.0013416 | 0.068545 | 0.067851 | -0.0025913 | -0.0015424 | 0.0011108 | -0.065399 | 0.00081389 | 0.064066 | 0.067251 | -0.0025315 | -0.0024075 | 0.0023647 | -0.018311 | 63.533 | -58.8638 | 63.4369 | -58.5068 | -58.9737 | -58.9142 | 0.036621 | 63.5101 | -58.7082 | 63.5513 | -58.6487 | -58.7997 | -58.5068 | -0.16479 | -0.013733 | -0.10986 | -0.09613 | 0.19226 | 0.35706 | -0.21515 | -0.27466 | -0.050354 | -0.10986 | -0.009155 | -0.036621 | 0.004578 | 0.041199 | 0.050354 | 0.059509 | 0.19226 | -0.10071 | 0.20142 | -0.091553
823.1719 | 122.4838 | -58.6761 | 63.5422 | -0.061874 | -0.00015382 | 0.067904 | 0.067424 | 0.00039951 | 0.001662 | -0.00017099 | 0.035309 | -0.00089511 | 0.06257 | 0.064901 | -0.00018152 | -0.0034757 | 0.0017239 | 0.027466 | 63.5422 | -58.8089 | 63.4598 | -58.4473 | -58.8821 | -58.8272 | 0.10529 | 63.5056 | -58.7082 | 63.5605 | -58.5846 | -58.7631 | -58.4976 | -0.13275 | 0 | -0.082397 | -0.082397 | 0.22888 | 0.36163 | -0.15106 | -0.20599 | -0.01831 | -0.073242 | -0.032044 | -0.036621 | 0.018311 | 0.054932 | 0.091552 | 0.1236 | 0.17853 | -0.086976 | 0.21057 | -0.054932
824.1719 | 122.4792 | -58.6441 | 63.5056 | -0.12291 | 0.00070078 | 0.066195 | 0.066783 | -0.0036593 | 0.00059397 | -0.002948 | 0.053619 | -0.00068151 | 0.063425 | 0.062551 | 0.00088648 | -0.0056119 | 0.0032193 | 0.045776 | 63.4964 | -58.7952 | 63.5239 | -58.374 | -58.8501 | -58.8089 | 0.14191 | 63.4644 | -58.7036 | 63.588 | -58.2504 | -58.7631 | -58.5846 | -0.15106 | -0.009155 | 0.018311 | 0.027466 | 0.27008 | 0.42114 | -0.16479 | -0.20599 | -0.013733 | -0.054932 | -0.059509 | -0.041199 | 0.082398 | 0.1236 | 0.39368 | 0.45319 | 0.059509 | -0.11902 | 0.11902 | -0.05951
825.1719 | 122.525 | -58.6533 | 63.5513 | -0.10002 | 0.0015552 | 0.068759 | 0.065501 | 0.0044583 | 0.0012348 | -0.00038459 | 0.035309 | 0.00017309 | 0.066202 | 0.058492 | -0.0010361 | -0.0049709 | 0.0021511 | 0.2243 | 63.5788 | -58.7631 | 63.4827 | -58.461 | -58.8043 | -58.7952 | 0.15106 | 63.4964 | -58.6853 | 63.6063 | -58.5434 | -58.7082 | -58.4564 | -0.10986 | 0.027465 | -0.068665 | -0.09613 | 0.19226 | 0.30212 | -0.14191 | -0.15106 | -0.032043 | -0.041198 | -0.032044 | -0.054932 | 0.054931 | 0.10986 | 0.10986 | 0.14191 | 0.19684 | -0.054932 | 0.22888 | -0.022888
826.1875 | 122.3923 | -58.699 | 63.5651 | -0.046616 | 0.00027338 | 0.068973 | 0.070201 | 0.00061311 | -0.00026063 | 0.0019654 | -0.092865 | 0.0014547 | 0.065134 | 0.067891 | -0.0025315 | -0.0015529 | -0.0010531 | -0.054932 | 63.5239 | -58.8593 | 63.4186 | -58.5342 | -59.0103 | -58.9508 | -0.068665 | 63.5468 | -58.7448 | 63.5101 | -58.6945 | -58.8776 | -58.5754 | -0.16022 | -0.041198 | -0.14648 | -0.10529 | 0.16479 | 0.32501 | -0.25177 | -0.31128 | -0.091553 | -0.15106 | -0.045776 | -0.01831 | -0.054931 | -0.036621 | 0.004578 | 0.050354 | 0.1236 | -0.17853 | 0.16937 | -0.13275
827.1719 | 122.438 | -58.667 | 63.5513 | -0.054245 | 0.001128 | 0.070895 | 0.070628 | 0.0023221 | 0.00038017 | -0.0033752 | -0.068451 | 0.00038669 | 0.065988 | 0.066396 | -0.00082252 | -0.0026211 | -0.00083948 | -0.004578 | 63.5605 | -58.8272 | 63.4506 | -58.5022 | -58.9279 | -58.8684 | 0.045776 | 63.5513 | -58.7082 | 63.5696 | -58.6121 | -58.7997 | -58.5251 | -0.16022 | 0.009155 | -0.10071 | -0.10986 | 0.16479 | 0.32501 | -0.20142 | -0.26093 | -0.041199 | -0.10071 | -0.041199 | 0 | 0.01831 | 0.01831 | 0.054931 | 0.09613 | 0.14191 | -0.13275 | 0.18311 | -0.091553
828.1719 | 122.4838 | -58.6395 | 63.5284 | -0.061874 | 0.001128 | 0.067904 | 0.067851 | -0.0010959 | -0.0013288 | 0.00068361 | 0.0047913 | 0.00038669 | 0.064493 | 0.064046 | -0.0010361 | -0.0041165 | -0.0012667 | 0.018311 | 63.5605 | -58.786 | 63.3408 | -58.5663 | -58.7631 | -58.786 | 0.15564 | 63.5193 | -58.6899 | 63.62 | -58.5068 | -58.699 | -58.429 | -0.14649 | 0.032044 | -0.18768 | -0.21973 | 0.073242 | 0.21973 | -0.14649 | -0.1236 | 0 | 0.022888 | -0.050354 | -0.009155 | 0.091553 | 0.10071 | 0.13275 | 0.18311 | 0.21057 | -0.05951 | 0.26093 | -0.009156
829.1719 | 122.5479 | -58.6395 | 63.5376 | -0.061874 | -0.00058102 | 0.067477 | 0.064647 | 0.00061311 | 0.0018756 | -0.00038459 | 0.062775 | 0.0014547 | 0.065561 | 0.067251 | 0.00088648 | -0.0036893 | -0.00019868 | 0.041199 | 63.5422 | -58.7952 | 63.533 | -58.3694 | -58.8089 | -58.7723 | 0.16937 | 63.5193 | -58.6624 | 63.6566 | -58.5068 | -58.754 | -58.461 | -0.15564 | 0.004577 | -0.004578 | -0.009155 | 0.27008 | 0.42572 | -0.13275 | -0.16937 | 0.022888 | -0.013733 | -0.022889 | -0.018311 | 0.11902 | 0.13733 | 0.13275 | 0.15564 | 0.17853 | -0.11444 | 0.20142 | -0.091552
830.1719 | 122.4976 | -58.6487 | 63.5422 | -0.10765 | 0.00070078 | 0.06705 | 0.065287 | -0.0010959 | 0.0018756 | -0.001239 | 0.059723 | -0.0017495 | 0.066843 | 0.064687 | -0.0025315 | -0.0045437 | -0.00019868 | 0.032043 | 63.5468 | -58.7814 | 63.5422 | -58.3603 | -58.7997 | -58.7448 | 0.14648 | 63.5056 | -58.667 | 63.6383 | -58.4885 | -58.6624 | -58.4198 | -0.13275 | 0.004578 | 0 | -0.004578 | 0.28839 | 0.42114 | -0.09613 | -0.15106 | 0.036621 | -0.018311 | -0.01831 | -0.036621 | 0.096131 | 0.13275 | 0.16022 | 0.17853 | 0.22888 | -0.013733 | 0.24719 | 0.004577
831.1719 | 122.5479 | -58.6166 | 63.5651 | -0.084763 | -0.00015382 | 0.069186 | 0.066569 | -0.0023777 | 0.00059397 | -0.0027344 | 0.02005 | -0.00068151 | 0.065775 | 0.062764 | -0.0014633 | -0.0028347 | 0.0010831 | -0.036621 | 63.5376 | -58.7952 | 63.4827 | -58.4473 | -58.8821 | -58.8135 | 0.050354 | 63.5284 | -58.6853 | 63.62 | -58.5342 | -58.6945 | -58.4427 | -0.17853 | -0.027465 | -0.082397 | -0.054932 | 0.16937 | 0.3479 | -0.19684 | -0.2655 | -0.018311 | -0.086975 | -0.068665 | -0.036621 | 0.054932 | 0.091553 | 0.082397 | 0.15106 | 0.17395 | -0.07782 | 0.24262 | -0.009155
832.1719 | 122.4655 | -58.6624 | 63.6063 | -0.084763 | -0.0010082 | 0.065341 | 0.072123 | -0.0025913 | 0.0010212 | 0.002179 | -0.071503 | 0.0014547 | 0.065561 | 0.071523 | -0.0010361 | -0.0041165 | 0.0027921 | 0 | 63.6017 | -58.7997 | 63.4506 | -58.5251 | -58.9279 | -58.8776 | -0.004578 | 63.5788 | -58.6899 | 63.5925 | -58.5938 | -58.786 | -58.4885 | -0.13733 | -0.004577 | -0.15564 | -0.15106 | 0.13733 | 0.27466 | -0.21515 | -0.2655 | -0.077819 | -0.12817 | -0.027465 | -0.027466 | -0.013733 | 0.013733 | 0.068665 | 0.09613 | 0.17395 | -0.1236 | 0.20142 | -0.096131
833.1719 | 122.5571 | -58.6075 | 63.5605 | -0.084763 | 0.00027338 | 0.067264 | 0.066783 | -2.7692e-05 | 0.00016657 | 0.00025641 | 0.038361 | 0.00060029 | 0.063638 | 0.061269 | -0.00060892 | -0.0034757 | -0.00083948 | 0.032043 | 63.5605 | -58.7585 | 63.5696 | -58.3374 | -58.7997 | -58.7357 | 0.14648 | 63.5468 | -58.6533 | 63.6658 | -58.4747 | -58.6487 | -58.4152 | -0.15106 | 0 | 0.009155 | 0.009155 | 0.27008 | 0.42114 | -0.12817 | -0.19226 | 0.022888 | -0.041199 | -0.045776 | -0.013733 | 0.10529 | 0.11902 | 0.13275 | 0.17853 | 0.19226 | -0.041199 | 0.23804 | 0.004577
834.1719 | 122.5571 | -58.6029 | 63.5834 | -0.038986 | -0.001222 | 0.066195 | 0.066142 | 0.0010403 | -0.00047423 | -0.00038459 | 0.023102 | 0.0010275 | 0.062357 | 0.068319 | -0.0018905 | -0.0041165 | 0.0021511 | 0.009155 | 63.5468 | -58.7723 | 63.5605 | -58.3328 | -58.8135 | -58.7357 | 0.15564 | 63.5468 | -58.6624 | 63.6749 | -58.4518 | -58.6258 | -58.3969 | -0.16937 | -0.036621 | -0.022888 | 0.013733 | 0.27008 | 0.43945 | -0.13275 | -0.21057 | 0.036621 | -0.041199 | -0.05951 | -0.036621 | 0.091553 | 0.12817 | 0.15106 | 0.21057 | 0.20599 | -0.022888 | 0.2655 | 0.036622
835.1719 | 122.5342 | -58.5983 | 63.5742 | -0.084763 | 0.002196 | 0.06705 | 0.068064 | -2.7692e-05 | -0.00090143 | 4.2808e-05 | 0.0078429 | -0.00046791 | 0.066415 | 0.061483 | 3.2078e-05 | -0.0041165 | -0.00041228 | -0.022888 | 63.5559 | -58.754 | 63.5468 | -58.374 | -58.8455 | -58.7723 | 0.11902 | 63.5376 | -58.6624 | 63.6795 | -58.4747 | -58.6441 | -58.3923 | -0.15564 | -0.018311 | -0.027466 | -0.009155 | 0.22431 | 0.37994 | -0.17395 | -0.24719 | -0.018311 | -0.091553 | -0.064087 | -0.036621 | 0.10529 | 0.14191 | 0.1236 | 0.18768 | 0.20599 | -0.045776 | 0.27008 | 0.018311
836.1719 | 122.5296 | -58.6212 | 63.5834 | -0.069504 | 0.0013416 | 0.067477 | 0.06721 | -0.0021641 | 0.00038017 | -0.0010254 | 0.0078429 | -0.0011087 | 0.065134 | 0.068746 | -0.0010361 | -0.0060391 | -0.00041228 | -0.032043 | 63.5513 | -58.7906 | 63.5284 | -58.4015 | -58.8364 | -58.7814 | 0.073242 | 63.5513 | -58.667 | 63.6612 | -58.4885 | -58.6578 | -58.3969 | -0.16937 | -0.032043 | -0.054932 | -0.022889 | 0.21973 | 0.3891 | -0.16022 | -0.21515 | 0.009155 | -0.045777 | -0.045776 | -0.032043 | 0.07782 | 0.10986 | 0.13275 | 0.17853 | 0.2243 | -0.036621 | 0.27008 | 0.009155
837.1719 | 122.5021 | -58.6075 | 63.6063 | -0.092392 | 0.00048698 | 0.067477 | 0.06721 | 0.00018591 | 0.001662 | -0.00081179 | 0.038361 | 0.0014547 | 0.063638 | 0.065114 | 0.00045928 | -0.0039029 | -0.0014805 | 0.032043 | 63.5834 | -58.754 | 63.7848 | -58.4885 | -59.1156 | -58.7769 | 0.064087 | 63.5376 | -58.6349 | 63.6978 | -58.4839 | -58.6349 | -58.3878 | -0.14648 | -0.022888 | 0.17853 | 0.20142 | 0.11902 | 0.2655 | -0.16937 | -0.50812 | -0.022888 | -0.36163 | -0.027466 | -0.068664 | 0.091553 | 0.16022 | 0.1236 | 0.15106 | 0.21973 | -0.027466 | 0.24719 | 0
838.1719 | 122.4792 | -58.6075 | 63.5834 | -0.12291 | -0.00015382 | 0.065768 | 0.067637 | -2.7692e-05 | 0.00059397 | -0.0010254 | -0.010467 | -0.0017495 | 0.065134 | 0.055501 | 0.00024568 | -0.0041165 | -0.0021213 | -0.054932 | 63.5559 | -58.7677 | 63.5056 | -58.4244 | -58.8776 | -58.8089 | -0.022888 | 63.5376 | -58.6533 | 63.6108 | -58.5983 | -58.58 | -58.3237 | -0.16022 | -0.027466 | -0.07782 | -0.050354 | 0.18311 | 0.34332 | -0.20142 | -0.27008 | -0.041199 | -0.10986 | -0.045776 | -0.045776 | 0.027466 | 0.073242 | 0.009155 | 0.054931 | 0.28381 | 0.027466 | 0.32959 | 0.073242
839.1719 | 122.4884 | -58.6166 | 63.6063 | -0.11528 | -0.0014356 | 0.067477 | 0.066355 | -0.0025913 | -0.0013288 | 0.00047001 | -0.010467 | -0.0023905 | 0.063852 | 0.069173 | -0.0033859 | -0.0030483 | 0.00065592 | 0.045776 | 63.5605 | -58.7585 | 63.501 | -58.4061 | -58.8913 | -58.8135 | -0.009155 | 63.5651 | -58.6945 | 63.6292 | -58.5114 | -58.6899 | -58.4335 | -0.14191 | -0.045776 | -0.10529 | -0.059509 | 0.21057 | 0.35248 | -0.19684 | -0.27466 | -0.054932 | -0.13275 | -0.07782 | -0.041199 | 0.022888 | 0.064087 | 0.10529 | 0.1831 | 0.18311 | -0.073242 | 0.26093 | 0.004578
840.1719 | 122.5021 | -58.5983 | 63.5971 | -0.10765 | -0.0018628 | 0.068332 | 0.068064 | -0.00088229 | 0.0025166 | -0.00038459 | 0.0047913 | -0.0023905 | 0.065347 | 0.064901 | -0.00018152 | -0.0045437 | -0.0023349 | -0.022888 | 63.5742 | -58.7723 | 63.5513 | -58.3878 | -58.8318 | -58.7631 | 0.041199 | 63.5788 | -58.6533 | 63.6795 | -58.4702 | -58.6578 | -58.4061 | -0.17395 | -0.022888 | -0.045776 | -0.022888 | 0.21057 | 0.38452 | -0.16479 | -0.23346 | 0.009155 | -0.059509 | -0.054931 | -0.018311 | 0.082397 | 0.10071 | 0.12817 | 0.1831 | 0.19226 | -0.059509 | 0.24719 | -0.004578
841.1719 | 122.4976 | -58.5754 | 63.6017 | -0.10765 | 0.00027338 | 0.066409 | 0.067637 | 0.0018949 | 0.0014484 | -0.00017099 | -0.013519 | 0.00060029 | 0.065775 | 0.060842 | -0.00082252 | -0.0043301 | 0.0015103 | -0.050354 | 63.5742 | -58.754 | 63.5422 | -58.3878 | -58.8364 | -58.7631 | 0.009155 | 63.5742 | -58.6487 | 63.6703 | -58.4839 | -58.6761 | -58.4061 | -0.17853 | -0.027466 | -0.05951 | -0.032044 | 0.18768 | 0.36621 | -0.18768 | -0.26093 | -0.009156 | -0.082398 | -0.073243 | -0.027466 | 0.068664 | 0.09613 | 0.091552 | 0.16479 | 0.16937 | -0.10071 | 0.24262 | -0.027465
842.1719 | 122.4014 | -58.5938 | 63.6658 | -0.061874 | 0.0013416 | 0.067904 | 0.068705 | -0.00088229 | 0.00059397 | -0.0014526 | -0.071503 | 0.0010275 | 0.061929 | 0.064687 | -0.0010361 | -0.0028347 | -0.0014805 | 0.032043 | 63.6063 | -58.7357 | 63.5376 | -58.4015 | -58.8501 | -58.7677 | 0.036621 | 63.6475 | -58.6304 | 63.6063 | -58.5297 | -58.7357 | -58.493 | -0.14191 | -0.059509 | -0.12817 | -0.068664 | 0.19226 | 0.33417 | -0.17395 | -0.25635 | -0.032043 | -0.11444 | -0.036621 | -0.01831 | -0.059509 | -0.041199 | 0.064087 | 0.10071 | 0.10071 | -0.14191 | 0.13733 | -0.10529
843.1719 | 122.4976 | -58.5846 | 63.6154 | -0.077133 | 0.0017688 | 0.067477 | 0.06721 | -0.0025913 | 0.00080757 | 0.0026062 | -0.010467 | -0.0015359 | 0.066202 | 0.059987 | -0.0010361 | -0.0056119 | -0.0016941 | -0.059509 | 63.5788 | -58.7311 | 63.4964 | -58.4381 | -58.8364 | -58.7723 | -0.022888 | 63.5605 | -58.6624 | 63.6658 | -58.5022 | -58.6945 | -58.4244 | -0.14648 | -0.036621 | -0.11902 | -0.082397 | 0.14648 | 0.29297 | -0.18768 | -0.25177 | -0.041199 | -0.10529 | -0.07782 | -0.054931 | 0.050354 | 0.10529 | 0.082398 | 0.16022 | 0.16022 | -0.10986 | 0.23804 | -0.032043
844.1719 | 122.4838 | -58.5892 | 63.6017 | -0.092392 | 0.00091438 | 0.069613 | 0.068492 | 0.0012539 | 0.0012348 | 0.00089721 | -0.019623 | 0.0033773 | 0.065134 | 0.06661 | -0.00039512 | -0.0024075 | 0.0010831 | -0.07782 | 63.5788 | -58.7357 | 63.5101 | -58.4198 | -58.8684 | -58.7952 | -0.045776 | 63.5559 | -58.6395 | 63.6612 | -58.5159 | -58.7036 | -58.4152 | -0.14649 | -0.022889 | -0.091553 | -0.068664 | 0.16937 | 0.31586 | -0.20599 | -0.27924 | -0.059509 | -0.13275 | -0.050354 | -0.045777 | 0.059509 | 0.10529 | 0.073242 | 0.1236 | 0.17395 | -0.11444 | 0.2243 | -0.064087
845.1719 | 122.4655 | -58.5938 | 63.5971 | -0.092392 | -0.00015382 | 0.067904 | 0.06956 | -0.0013095 | 0.00038017 | 0.0013244 | -0.074554 | 0.0020957 | 0.065347 | 0.064901 | -0.00082252 | -0.0034757 | -0.00083948 | -0.050354 | 63.5559 | -58.7585 | 63.5605 | -58.3694 | -58.8593 | -58.786 | -0.013733 | 63.5651 | -58.6304 | 63.6612 | -58.5068 | -58.699 | -58.4198 | -0.16479 | -0.041199 | -0.036621 | 0.004578 | 0.2243 | 0.3891 | -0.19226 | -0.2655 | -0.027466 | -0.10071 | -0.036621 | -0.032044 | 0.064087 | 0.096131 | 0.086975 | 0.1236 | 0.17395 | -0.10529 | 0.21057 | -0.068665
846.1719 | 122.4838 | -58.5892 | 63.6108 | -0.11528 | 0.00070078 | 0.068332 | 0.068278 | -0.00024129 | 0.00059397 | 0.00047001 | -0.040985 | 0.00017309 | 0.064066 | 0.063833 | -0.0018905 | -0.0021939 | 0.00044212 | -0.073242 | 63.5696 | -58.7311 | 63.533 | -58.4106 | -58.8638 | -58.7769 | -0.036621 | 63.5788 | -58.6487 | 63.6703 | -58.5068 | -58.667 | -58.4106 | -0.14191 | -0.041199 | -0.07782 | -0.036621 | 0.17853 | 0.32043 | -0.18768 | -0.27466 | -0.045776 | -0.13275 | -0.05951 | -0.032044 | 0.059509 | 0.091553 | 0.082397 | 0.14191 | 0.17853 | -0.07782 | 0.23804 | -0.01831
847.1719 | 122.4655 | -58.5571 | 63.6887 | -0.077133 | -0.0018628 | 0.068332 | 0.066142 | -0.0019503 | 0.00016657 | 0.0017518 | -0.022675 | 0.0012411 | 0.06257 | 0.061269 | 3.2078e-05 | -0.0043301 | -0.0021213 | 0.032043 | 63.6292 | -58.7128 | 63.6246 | -58.3054 | -58.7585 | -58.7219 | 0.10071 | 63.652 | -58.6029 | 63.707 | -58.4518 | -58.6349 | -58.3878 | -0.15564 | -0.05951 | -0.064087 | -0.004577 | 0.25177 | 0.40741 | -0.16479 | -0.20142 | -0.009155 | -0.045776 | -0.045776 | -0.036621 | 0.01831 | 0.054931 | 0.10529 | 0.15106 | 0.16937 | -0.07782 | 0.21515 | -0.032044
848.1719 | 122.4884 | -58.5754 | 63.6292 | -0.11528 | 0.0015552 | 0.066836 | 0.066996 | 0.00082671 | -0.00068783 | 0.0017518 | -0.047089 | -0.0017495 | 0.064493 | 0.06661 | -0.0010361 | -0.0034757 | -0.0036167 | -0.064087 | 63.5925 | -58.7311 | 63.5605 | -58.3832 | -58.8455 | -58.7677 | -0.027466 | 63.6063 | -58.6349 | 63.6703 | -58.4839 | -58.6761 | -58.3923 | -0.15564 | -0.036621 | -0.068664 | -0.032043 | 0.19226 | 0.3479 | -0.19226 | -0.27008 | -0.036621 | -0.11444 | -0.05951 | -0.022888 | 0.041199 | 0.064087 | 0.091552 | 0.15106 | 0.18311 | -0.10071 | 0.24262 | -0.041198
849.1719 | 122.4472 | -58.6075 | 63.6154 | -0.11528 | 5.9776e-05 | 0.067904 | 0.06721 | -0.00066869 | 0.00080757 | 4.2808e-05 | -0.044037 | 0.00081389 | 0.063638 | 0.061056 | -0.0023177 | -0.0032619 | -0.0025485 | -0.073242 | 63.588 | -58.7082 | 63.5468 | -58.3649 | -58.8638 | -58.7631 | -0.027466 | 63.62 | -58.6166 | 63.6795 | -58.5022 | -58.6807 | -58.4015 | -0.10071 | -0.027465 | -0.068664 | -0.041199 | 0.24262 | 0.34332 | -0.15564 | -0.25635 | -0.054932 | -0.15564 | -0.009155 | 0.004578 | 0.064087 | 0.059509 | 0.10529 | 0.11444 | 0.20599 | -0.073242 | 0.21515 | -0.064087
850.1719 | 122.4747 | -58.5754 | 63.62 | -0.10002 | 0.001128 | 0.070254 | 0.068492 | -0.0023777 | -0.0013288 | 0.0026062 | -0.040985 | 0.00081389 | 0.064279 | 0.061056 | -0.0012497 | -0.0015529 | 0.00022852 | -0.054932 | 63.5742 | -58.7128 | 63.5559 | -58.3282 | -58.8272 | -58.7677 | 0 | 63.5925 | -58.5983 | 63.6932 | -58.4702 | -58.6761 | -58.3923 | -0.13733 | -0.045776 | -0.064087 | -0.018311 | 0.24719 | 0.38452 | -0.19226 | -0.25177 | -0.054931 | -0.11444 | -0.022889 | -0.027466 | 0.073242 | 0.10071 | 0.10529 | 0.12817 | 0.18311 | -0.10071 | 0.20599 | -0.077819
851.1719 | 122.4701 | -58.5526 | 63.5788 | -0.10002 | -0.00058102 | 0.070468 | 0.070414 | 0.0025357 | 0.0014484 | -0.00059819 | -0.056244 | -4.0508e-05 | 0.063638 | 0.065542 | -0.0018905 | -0.0036893 | -0.0029757 | -0.045776 | 63.6154 | -58.6945 | 63.6017 | -58.3191 | -58.8043 | -58.7357 | 0.004578 | 63.588 | -58.6075 | 63.7436 | -58.4518 | -58.6395 | -58.3786 | -0.14191 | 0.036621 | 0.022889 | -0.013732 | 0.23346 | 0.37537 | -0.18311 | -0.25177 | -0.041199 | -0.10986 | -0.054932 | 0.009156 | 0.1648 | 0.15564 | 0.10071 | 0.15564 | 0.17395 | -0.086975 | 0.22888 | -0.032043
852.1719 | 122.493 | -58.5526 | 63.6429 | -0.15343 | 0.0013416 | 0.0694 | 0.06721 | -0.00024129 | 0.0031574 | 0.00089721 | -0.05014 | 0.0014547 | 0.065134 | 0.06426 | -0.0010361 | -0.00069854 | -0.00062588 | -0.054932 | 63.6108 | -58.7036 | 63.5788 | -58.3099 | -58.8181 | -58.7402 | -0.004578 | 63.5834 | -58.6075 | 63.7161 | -58.461 | -58.6761 | -58.3878 | -0.15106 | -0.032043 | -0.064087 | -0.032044 | 0.24261 | 0.39368 | -0.18768 | -0.2655 | -0.036621 | -0.11444 | -0.054932 | -0.059509 | 0.073242 | 0.13275 | 0.091552 | 0.14648 | 0.16479 | -0.1236 | 0.21973 | -0.068664
853.1719 | 122.4655 | -58.5754 | 63.5834 | -0.16106 | 0.0026234 | 0.068759 | 0.06956 | 0.00039951 | 0.0044392 | 0.0013244 | -0.10812 | 0.00060029 | 0.063425 | 0.060842 | -0.0016769 | -0.0030483 | -0.00041228 | -0.041199 | 63.588 | -58.7036 | 63.6063 | -58.3191 | -58.7997 | -58.7311 | -0.004578 | 63.6337 | -58.5892 | 63.7482 | -58.4244 | -58.6899 | -58.4244 | -0.12817 | 0.004578 | 0.022888 | 0.01831 | 0.25635 | 0.38452 | -0.15564 | -0.2243 | -0.027466 | -0.096131 | -0.013733 | 0.050354 | 0.16479 | 0.11444 | 0.15106 | 0.16479 | 0.15106 | -0.11444 | 0.16479 | -0.10071
854.1719 | 122.5204 | -58.5342 | 63.6108 | -0.16869 | 5.9776e-05 | 0.06705 | 0.068492 | 0.0036039 | 0.0037982 | 0.0030334 | -0.08371 | 0.00060029 | 0.063425 | 0.064901 | -0.0029587 | -0.0015529 | 0.00065592 | -0.045776 | 63.6108 | -58.7036 | 63.4827 | -58.1955 | -58.5571 | -58.6487 | 0.041199 | 63.6063 | -58.58 | 63.739 | -58.4244 | -58.6258 | -58.3557 | -0.16937 | 0 | -0.12817 | -0.12817 | 0.33875 | 0.50812 | -0.11444 | -0.022888 | 0.054931 | 0.14648 | -0.045776 | -0.004578 | 0.12817 | 0.13275 | 0.10986 | 0.15564 | 0.17853 | -0.091552 | 0.2243 | -0.045776
855.1719 | 122.5616 | -58.5571 | 63.62 | -0.13817 | -0.0014356 | 0.068973 | 0.069773 | 0.0023221 | -4.7032e-05 | -0.00081179 | -0.095917 | 0.00060029 | 0.065561 | 0.064901 | -0.0014633 | -0.00091214 | -0.0044711 | -0.041199 | 63.6017 | -58.7173 | 63.6108 | -58.3191 | -58.7906 | -58.7265 | 0 | 63.6292 | -58.5892 | 63.7482 | -58.429 | -58.8684 | -58.4702 | -0.16022 | -0.01831 | -0.009155 | 0.009155 | 0.23804 | 0.39825 | -0.16937 | -0.23346 | -0.009155 | -0.073242 | -0.032043 | 0.009155 | 0.12817 | 0.11902 | 0.12817 | 0.16022 | 0.086975 | -0.31128 | 0.11902 | -0.27924
856.1719 | 122.5342 | -58.5342 | 63.6154 | -0.15343 | -0.0010082 | 0.068545 | 0.069133 | -0.0038729 | 0.0010212 | -0.0010254 | -0.068451 | 0.00017309 | 0.063211 | 0.066823 | -0.0018905 | -0.0028347 | 1.4916e-05 | -0.082397 | 63.6246 | -58.7128 | 63.588 | -58.3511 | -58.8043 | -58.7128 | -0.027466 | 63.6246 | -58.5938 | 63.7573 | -58.4152 | -58.6258 | -58.3054 | -0.17853 | 0.009156 | -0.027465 | -0.036621 | 0.18311 | 0.36163 | -0.17853 | -0.27008 | 0 | -0.091552 | -0.059509 | 0.009156 | 0.14191 | 0.13275 | 0.11902 | 0.17853 | 0.22888 | -0.091552 | 0.28839 | -0.032043
857.1719 | 122.4792 | -58.5434 | 63.6566 | -0.13817 | -0.00058102 | 0.066836 | 0.069346 | -0.0034457 | 0.0020892 | -0.0016662 | -0.0043641 | 0.0018821 | 0.062357 | 0.065969 | 0.0021683 | -0.0045437 | -0.0016941 | -0.091553 | 63.6337 | -58.6853 | 63.5696 | -58.3191 | -58.7952 | -58.7311 | -0.041199 | 63.6292 | -58.6029 | 63.7024 | -58.4106 | -58.6212 | -58.374 | -0.14191 | -0.022888 | -0.086975 | -0.064087 | 0.2243 | 0.36621 | -0.18768 | -0.25177 | -0.045776 | -0.10986 | -0.059509 | -0.027466 | 0.045777 | 0.073243 | 0.13275 | 0.19226 | 0.16937 | -0.07782 | 0.22888 | -0.018311
858.1719 | 122.4701 | -58.5388 | 63.652 | -0.10765 | -0.00036742 | 0.068759 | 0.069773 | 0.0010403 | -4.7032e-05 | 0.00047001 | -0.10507 | 0.00017309 | 0.064706 | 0.067891 | 3.2078e-05 | -0.0013393 | 0.00065592 | -0.10986 | 63.6337 | -58.7173 | 63.533 | -58.3923 | -58.8593 | -58.8043 | -0.09613 | 63.652 | -58.5892 | 63.7161 | -58.4747 | -58.6899 | -58.3786 | -0.17853 | -0.018311 | -0.11902 | -0.10071 | 0.14648 | 0.32501 | -0.2655 | -0.32044 | -0.086975 | -0.14191 | -0.050354 | 0 | 0.064086 | 0.064086 | 0.064087 | 0.11444 | 0.16022 | -0.15106 | 0.21057 | -0.10071
859.1719 | 122.4472 | -58.5846 | 63.7024 | -0.1458 | 0.00091438 | 0.068973 | 0.069133 | -0.0036593 | 0.001662 | 0.0026062 | -0.11728 | 0.0023093 | 0.063211 | 0.071309 | 3.2078e-05 | -5.7742e-05 | 0.00086952 | -0.11444 | 63.6887 | -58.7128 | 63.5056 | -58.461 | -58.9142 | -58.8272 | -0.15106 | 63.6841 | -58.6121 | 63.7665 | -58.786 | -58.8318 | -58.3832 | -0.12817 | -0.013733 | -0.19684 | -0.18311 | 0.1236 | 0.25177 | -0.24261 | -0.32959 | -0.11444 | -0.20142 | -0.027466 | -0.018311 | 0.064086 | 0.082397 | -0.20142 | -0.17395 | 0.20142 | -0.24719 | 0.22888 | -0.21973
860.1719 | 122.4289 | -58.5571 | 63.6795 | -0.10765 | -0.00079462 | 0.067691 | 0.073191 | 0.00018591 | 0.00059397 | 0.00047001 | -0.13559 | 0.0010275 | 0.064279 | 0.065755 | -0.00082252 | -0.0019803 | 0.0012967 | -0.12817 | 63.6887 | -58.699 | 63.5468 | -58.4976 | -58.7814 | -58.7311 | -0.1236 | 63.6749 | -58.6029 | 63.6658 | -58.4976 | -58.7082 | -58.4152 | -0.14191 | 0.009156 | -0.13275 | -0.14191 | 0.059509 | 0.20142 | -0.17395 | -0.2243 | -0.032043 | -0.082397 | -0.045776 | -0.004577 | -0.013733 | -0.009156 | 0.059509 | 0.10529 | 0.14191 | -0.15106 | 0.18768 | -0.10529
861.1719 | 122.4472 | -58.5434 | 63.652 | -0.10002 | -0.00058102 | 0.066195 | 0.067851 | -0.00066869 | 0.00080757 | -0.0025208 | -0.077606 | 0.0014547 | 0.065988 | 0.057851 | -0.00018152 | -0.0013393 | -0.0010531 | -0.11444 | 63.6749 | -58.6807 | 63.5284 | -58.3786 | -58.8455 | -58.7952 | -0.10529 | 63.652 | -58.6029 | 63.6978 | -58.4702 | -58.6853 | -58.3878 | -0.13733 | 0.022888 | -0.1236 | -0.14648 | 0.16479 | 0.30212 | -0.25177 | -0.30212 | -0.11444 | -0.16479 | -0.059509 | 0 | 0.045776 | 0.045776 | 0.073242 | 0.13275 | 0.15564 | -0.14191 | 0.21515 | -0.082398
862.1719 | 122.4701 | -58.5205 | 63.6566 | -0.10765 | -0.00229 | 0.068759 | 0.064647 | -0.0013095 | 0.00080757 | -0.0020936 | -0.0043641 | 0.00081389 | 0.065775 | 0.057851 | 0.00088648 | -0.0060391 | 0.0034329 | -0.086975 | 63.6612 | -58.6533 | 63.5971 | -58.3008 | -58.7814 | -58.7036 | -0.068665 | 63.6246 | -58.58 | 63.739 | -58.4106 | -58.5938 | -58.342 | -0.13275 | 0.004578 | -0.059509 | -0.064087 | 0.21973 | 0.35248 | -0.1831 | -0.26093 | -0.050354 | -0.12817 | -0.059509 | -0.032043 | 0.082398 | 0.11444 | 0.10986 | 0.16937 | 0.17853 | -0.073242 | 0.23804 | -0.013733
863.1719 | 122.406 | -58.548 | 63.7161 | -0.077133 | 0.0017688 | 0.0694 | 0.069987 | 0.0021085 | 0.00038017 | 0.0028198 | -0.10507 | -4.0508e-05 | 0.06492 | 0.069814 | -0.0010361 | -0.00091214 | 0.00086952 | -0.10071 | 63.6978 | -58.7128 | 63.4323 | -58.6761 | -58.7311 | -58.5754 | -0.14648 | 63.6978 | -58.5983 | 63.6841 | -58.5022 | -58.699 | -58.3832 | -0.16479 | -0.01831 | -0.28381 | -0.2655 | -0.12817 | 0.036622 | -0.027465 | -0.1831 | 0.13733 | -0.01831 | -0.050354 | -0.01831 | -0.032043 | -0.013733 | 0.045777 | 0.096131 | 0.1648 | -0.15106 | 0.21515 | -0.10071
864.1719 | 122.4976 | -58.5297 | 63.7161 | -0.092392 | 0.00091438 | 0.069827 | 0.069773 | -0.0013095 | -0.0015424 | -0.0010254 | -0.071503 | 0.0020957 | 0.064066 | 0.06191 | -0.00039512 | -0.0011257 | -0.0029757 | -0.12817 | 63.707 | -58.6899 | 63.5101 | -58.4335 | -58.8776 | -58.7906 | -0.14191 | 63.6932 | -58.5709 | 63.7985 | -58.786 | -58.8776 | -58.4061 | -0.16022 | -0.009155 | -0.20599 | -0.19684 | 0.09613 | 0.25635 | -0.26093 | -0.3479 | -0.10071 | -0.18768 | -0.041199 | -0.022888 | 0.082398 | 0.10529 | -0.25635 | -0.21515 | 0.1236 | -0.3479 | 0.16479 | -0.3067
865.1719 | 122.3648 | -58.5297 | 63.7024 | -0.092392 | 0.0015552 | 0.070468 | 0.068705 | -0.0013095 | 0.00016657 | 0.002179 | -0.095917 | 0.0027365 | 0.064279 | 0.065328 | 0.00024568 | -0.0013393 | 0.0030057 | -0.091553 | 63.7161 | -58.6624 | 63.5513 | -58.4061 | -58.8501 | -58.7952 | -0.15564 | 63.7161 | -58.5663 | 63.6612 | -58.5251 | -58.7173 | -58.4061 | -0.13275 | 0.013732 | -0.15106 | -0.16479 | 0.1236 | 0.25635 | -0.2655 | -0.32044 | -0.13275 | -0.18768 | -0.036621 | 0.013732 | -0.041199 | -0.054931 | 0.004578 | 0.041199 | 0.1236 | -0.18768 | 0.16022 | -0.15106
866.1719 | 122.438 | -58.5068 | 63.7024 | -0.10765 | -0.0018628 | 0.068759 | 0.068705 | 0.0014675 | 0.0010212 | -0.0023072 | -0.077606 | 0.00017309 | 0.065134 | 0.066823 | -0.00018152 | -0.0024075 | -0.00041228 | -0.004578 | 63.7207 | -58.6533 | 63.6063 | -58.2916 | -58.7677 | -58.7128 | 0 | 63.7299 | -58.5526 | 63.7344 | -58.4335 | -58.6624 | -58.3603 | -0.14648 | 0.01831 | -0.096131 | -0.11444 | 0.21515 | 0.36163 | -0.20599 | -0.26093 | -0.05951 | -0.11444 | -0.045776 | 0.027465 | 0.032043 | 0.004578 | 0.073242 | 0.11902 | 0.14648 | -0.15564 | 0.19226 | -0.10986
867.1719 | 122.4243 | -58.5159 | 63.7115 | -0.10002 | 0.0015552 | 0.069613 | 0.069133 | 0.00082671 | 0.0014484 | -0.0020936 | -0.065399 | 0.0029501 | 0.06727 | 0.059987 | 0.0011001 | -0.0034757 | -0.0012667 | -0.33417 | 63.6566 | -58.6807 | 63.5788 | -58.3878 | -58.8409 | -58.7723 | -0.11902 | 63.6978 | -58.5617 | 63.7161 | -58.4702 | -58.6807 | -58.3603 | -0.1648 | -0.054932 | -0.13275 | -0.07782 | 0.12817 | 0.29297 | -0.25635 | -0.32501 | -0.091553 | -0.16022 | -0.045777 | -0.013733 | 0.004577 | 0.01831 | 0.045776 | 0.091553 | 0.15564 | -0.1648 | 0.20142 | -0.11902
868.1719 | 122.3511 | -58.5022 | 63.7436 | -0.069504 | -0.00036742 | 0.069186 | 0.06721 | 0.0014675 | 0.001662 | -0.00081179 | -0.089813 | 0.0016685 | 0.065561 | 0.068746 | 0.0011001 | -0.0045437 | 1.4916e-05 | -0.07782 | 63.7436 | -58.6395 | 63.5559 | -58.3511 | -58.8272 | -58.7723 | -0.13275 | 63.7253 | -58.5663 | 63.6841 | -58.4885 | -58.7082 | -58.4015 | -0.13733 | 0 | -0.18768 | -0.18768 | 0.15106 | 0.28839 | -0.27008 | -0.32501 | -0.13275 | -0.18768 | -0.064087 | -0.01831 | -0.059509 | -0.041199 | 0.013733 | 0.07782 | 0.10071 | -0.20599 | 0.1648 | -0.14191
869.1719 | 122.4426 | -58.5205 | 63.7665 | -0.11528 | 0.0019824 | 0.068759 | 0.069346 | 0.00039951 | 0.00016657 | 0.0019654 | -0.077606 | -0.00089511 | 0.06492 | 0.065542 | -0.0018905 | -0.00091214 | 0.00086952 | -0.059509 | 63.7344 | -58.6258 | 63.6017 | -58.3099 | -58.7997 | -58.7265 | -0.07782 | 63.7299 | -58.5617 | 63.7253 | -58.461 | -58.6716 | -58.3557 | -0.10529 | -0.032043 | -0.16479 | -0.13275 | 0.21057 | 0.31586 | -0.20599 | -0.27924 | -0.10071 | -0.17395 | -0.041199 | -0.036621 | -0.041198 | -0.004577 | 0.059509 | 0.10071 | 0.16479 | -0.15106 | 0.20599 | -0.10986
870.1719 | 122.4014 | -58.4793 | 63.7573 | -0.054245 | 0.00027338 | 0.068973 | 0.06721 | 0.0033901 | -0.0015424 | 0.00047001 | -0.062347 | 0.00017309 | 0.063425 | 0.063192 | 3.2078e-05 | -0.0026211 | 0.0025785 | -0.027466 | 63.7527 | -58.6166 | 63.6108 | -58.2825 | -58.7585 | -58.7036 | -0.018311 | 63.739 | -58.5297 | 63.739 | -58.429 | -58.6349 | -58.3557 | -0.13733 | -0.004577 | -0.14648 | -0.14191 | 0.19684 | 0.33417 | -0.2243 | -0.27924 | -0.086975 | -0.14191 | -0.050354 | -0.01831 | -0.01831 | 0 | 0.050354 | 0.10071 | 0.1236 | -0.15564 | 0.17395 | -0.10529
871.1719 | 122.4426 | -58.493 | 63.7756 | -0.10765 | 0.00027338 | 0.066409 | 0.068492 | 0.0050991 | 0.0029438 | -0.0010254 | -0.05014 | 0.0014547 | 0.065988 | 0.062337 | 3.2078e-05 | -0.0015529 | 0.0015103 | -0.11902 | 63.7573 | -58.6029 | 63.6429 | -58.2825 | -58.7219 | -58.6807 | 0 | 63.7436 | -58.5159 | 63.7619 | -58.3923 | -58.6121 | -58.3145 | -0.10986 | -0.018311 | -0.13275 | -0.11444 | 0.21057 | 0.32043 | -0.18768 | -0.22888 | -0.07782 | -0.11902 | -0.022888 | -0.032044 | -0.013733 | 0.018311 | 0.10071 | 0.1236 | 0.17853 | -0.11902 | 0.20142 | -0.096131
872.1719 | 122.525 | -58.4839 | 63.7161 | -0.13817 | 5.9776e-05 | 0.068118 | 0.068705 | 0.0012539 | 0.00059397 | 0.0023926 | 0.0017395 | 0.0014547 | 0.06492 | 0.06661 | 0.00045928 | -0.0024075 | 0.0023647 | -0.10529 | 63.7482 | -58.6258 | 63.5925 | -58.3237 | -58.7997 | -58.7265 | -0.10529 | 63.7207 | -58.5342 | 63.7482 | -58.3969 | -58.6075 | -58.2962 | -0.14191 | 0.032044 | -0.1236 | -0.15564 | 0.16022 | 0.30212 | -0.24261 | -0.31586 | -0.10071 | -0.17395 | -0.050354 | 0.004578 | 0.032044 | 0.027466 | 0.086975 | 0.13733 | 0.18768 | -0.1236 | 0.23804 | -0.073242
873.1719 | 122.4838 | -58.4702 | 63.7665 | -0.077133 | 0.0013416 | 0.068759 | 0.06956 | 0.0021085 | -4.7032e-05 | -0.0018798 | -0.053192 | 0.0023093 | 0.066415 | 0.065328 | 0.00067288 | -0.0028347 | 0.0025785 | 0.018311 | 63.7939 | -58.6029 | 63.6612 | -58.2504 | -58.7219 | -58.6761 | 0.041199 | 63.7573 | -58.5114 | 63.7848 | -58.3603 | -58.5846 | -58.3008 | -0.13275 | 0.027466 | -0.10528 | -0.13275 | 0.21973 | 0.35248 | -0.20599 | -0.25177 | -0.073242 | -0.11902 | -0.041199 | -0.009155 | 0.018311 | 0.027466 | 0.10986 | 0.15106 | 0.16937 | -0.11444 | 0.21057 | -0.073242
874.1719 | 122.4335 | -58.4518 | 63.7939 | -0.061874 | 0.00091438 | 0.068545 | 0.065501 | 0.00061311 | 0.001662 | -0.004657 | -0.0074159 | 0.00038669 | 0.067697 | 0.060842 | -0.00060892 | -0.0024075 | 0.0010831 | 0.004578 | 63.7802 | -58.5846 | 63.6566 | -58.2596 | -58.7036 | -58.6578 | 0.009155 | 63.739 | -58.493 | 63.7985 | -58.3466 | -58.5709 | -58.2962 | -0.13275 | -0.013733 | -0.13733 | -0.1236 | 0.19226 | 0.32501 | -0.20599 | -0.25177 | -0.073242 | -0.11902 | -0.041199 | -0.054931 | 0.004578 | 0.059509 | 0.10528 | 0.14648 | 0.15564 | -0.11902 | 0.19684 | -0.07782
875.1719 | 122.4747 | -58.4473 | 63.8077 | -0.046616 | -0.00015382 | 0.068759 | 0.069346 | 0.0021085 | -4.7032e-05 | 0.00047001 | -0.03183 | 0.0012411 | 0.067697 | 0.065969 | 0.00045928 | -0.0024075 | 0.0051419 | 0.013733 | 63.7802 | -58.5892 | 63.6658 | -58.2184 | -58.7128 | -58.667 | 0.045776 | 63.7619 | -58.4839 | 63.7939 | -58.3603 | -58.5754 | -58.2825 | -0.14191 | -0.027466 | -0.14191 | -0.11444 | 0.22888 | 0.37079 | -0.21973 | -0.2655 | -0.07782 | -0.1236 | -0.036621 | -0.045776 | -0.013733 | 0.032043 | 0.086975 | 0.1236 | 0.16479 | -0.12817 | 0.20142 | -0.091552
876.1719 | 122.4289 | -58.4152 | 63.8214 | -0.023727 | -0.00058102 | 0.070468 | 0.068492 | -2.7692e-05 | 0.0020892 | -0.0014526 | 0.0047913 | 0.0010275 | 0.065561 | 0.069387 | 0.00088648 | -0.0026211 | 0.0021511 | 0 | 63.7894 | -58.5983 | 63.652 | -58.255 | -58.699 | -58.6533 | 0.032043 | 63.7756 | -58.4747 | 63.8031 | -58.3466 | -58.5617 | -58.255 | -0.18311 | -0.032043 | -0.16937 | -0.13733 | 0.16022 | 0.34332 | -0.23804 | -0.28381 | -0.054931 | -0.10071 | -0.059509 | -0.045776 | -0.01831 | 0.027466 | 0.068664 | 0.12817 | 0.16022 | -0.14648 | 0.21973 | -0.086976
877.1719 | 122.4472 | -58.493 | 63.7711 | -0.12291 | -0.0014356 | 0.069186 | 0.068064 | 0.0021085 | -0.0013288 | -0.0010254 | -0.037933 | 0.00081389 | 0.065775 | 0.068532 | 0.0015275 | -0.0041165 | 0.0012967 | -0.091553 | 63.7619 | -58.6121 | 63.7253 | -58.3878 | -58.8181 | -58.6487 | -0.1236 | 63.7527 | -58.5159 | 63.7299 | -58.429 | -58.6304 | -58.3374 | -0.11902 | -0.009155 | -0.045776 | -0.036621 | 0.10529 | 0.2243 | -0.15564 | -0.32501 | -0.036621 | -0.20599 | -0.022888 | -0.01831 | -0.041199 | -0.022889 | 0.064087 | 0.086975 | 0.15564 | -0.13733 | 0.17853 | -0.11444
878.1719 | 122.3923 | -58.4473 | 63.826 | -0.092392 | 0.00070078 | 0.067264 | 0.067851 | -0.0013095 | -0.0023968 | 0.0011108 | -0.0013123 | 0.00081389 | 0.062784 | 0.062551 | 3.2078e-05 | -0.0034757 | 0.0062101 | -0.036621 | 63.8306 | -58.5846 | 63.5468 | -58.4518 | -58.7265 | -58.6716 | -0.013733 | 63.7848 | -58.493 | 63.7573 | -58.3603 | -58.5617 | -58.3008 | -0.13733 | 0.004577 | -0.27924 | -0.28381 | -0.004577 | 0.13275 | -0.2243 | -0.27923 | -0.086975 | -0.14191 | -0.045776 | -0.041199 | -0.068665 | -0.027466 | 0.086975 | 0.13275 | 0.14648 | -0.11444 | 0.19226 | -0.068665
879.1719 | 122.4014 | -58.4335 | 63.8672 | -0.077133 | -0.00015382 | 0.068759 | 0.066783 | -0.0010959 | -0.001756 | -0.0016662 | 0.035309 | 0.0023093 | 0.065134 | 0.058706 | -0.00060892 | -0.0039029 | -0.00019868 | -0.045776 | 63.8168 | -58.5159 | 63.6429 | -58.2367 | -58.6945 | -58.6487 | 0.013733 | 63.8031 | -58.4564 | 63.7619 | -58.3832 | -58.5526 | -58.3054 | -0.082397 | -0.050354 | -0.2243 | -0.17395 | 0.19684 | 0.27924 | -0.21515 | -0.26093 | -0.13275 | -0.17853 | -0.022888 | -0.064086 | -0.10529 | -0.041199 | 0.050354 | 0.073242 | 0.12817 | -0.11902 | 0.15106 | -0.09613
880.1719 | 122.4655 | -58.4061 | 63.8168 | -0.069504 | 0.00048698 | 0.069613 | 0.065501 | -0.0021641 | 0.00080757 | -0.00038459 | 0.047516 | -0.0011087 | 0.066202 | 0.064473 | 3.2078e-05 | -0.0026211 | 0.0027921 | 0.018311 | 63.7802 | -58.5571 | 63.7619 | -58.1223 | -58.6029 | -58.548 | 0.11902 | 63.7619 | -58.4335 | 63.8489 | -58.2596 | -58.461 | -58.2047 | -0.15106 | -0.036621 | -0.054931 | -0.01831 | 0.28381 | 0.43488 | -0.14191 | -0.19684 | 0.009155 | -0.045776 | -0.027466 | -0.054931 | 0.032044 | 0.086975 | 0.14648 | 0.17395 | 0.20142 | -0.054932 | 0.22888 | -0.027466
881.1719 | 122.4747 | -58.4061 | 63.8397 | -0.092392 | 0.00070078 | 0.068118 | 0.065928 | -0.0023777 | -0.001756 | -0.0023072 | 0.056671 | 0.00017309 | 0.063852 | 0.061483 | -0.00018152 | -0.0043301 | -0.00019868 | 0.022888 | 63.7894 | -58.5388 | 63.7344 | -58.136 | -58.5938 | -58.548 | 0.10071 | 63.7848 | -58.4473 | 63.8214 | -58.287 | -58.4656 | -58.2138 | -0.13275 | -0.050354 | -0.10529 | -0.054932 | 0.27008 | 0.40283 | -0.14191 | -0.18768 | -0.009156 | -0.054932 | -0.041199 | -0.054932 | -0.018311 | 0.036621 | 0.11902 | 0.16022 | 0.19226 | -0.059509 | 0.23346 | -0.01831
882.1875 | 122.4838 | -58.4198 | 63.7985 | -0.13817 | -0.00015382 | 0.068545 | 0.070842 | 0.0059537 | 0.001662 | -0.00081179 | -0.044037 | 0.00060029 | 0.065988 | 0.068319 | 0.00067288 | -0.0045437 | -0.0019077 | -0.091553 | 63.7894 | -58.6166 | 63.6292 | -58.2825 | -58.7677 | -58.6945 | -0.07782 | 63.739 | -58.4656 | 63.8031 | -58.374 | -58.5571 | -58.2825 | -0.19684 | -0.009155 | -0.16937 | -0.16022 | 0.13733 | 0.33417 | -0.27466 | -0.3479 | -0.07782 | -0.15106 | -0.045776 | -0.059509 | 0.004578 | 0.064087 | 0.045777 | 0.091553 | 0.13733 | -0.13733 | 0.1831 | -0.091553
883.1719 | 122.4838 | -58.4106 | 63.8214 | -0.069504 | 0.00070078 | 0.068759 | 0.067424 | -0.0013095 | 0.00038017 | 4.2808e-05 | 0.032257 | -0.00089511 | 0.065347 | 0.058278 | -0.00082252 | -0.0039029 | -0.0016941 | 0.032043 | 63.8123 | -58.5114 | 63.7665 | -58.1177 | -58.5846 | -58.5251 | 0.10529 | 63.7939 | -58.4244 | 63.8351 | -58.2642 | -58.4564 | -58.2001 | -0.10071 | -0.009155 | -0.054932 | -0.045777 | 0.29297 | 0.39368 | -0.11444 | -0.17395 | -0.013732 | -0.073242 | -0.013732 | -0.027466 | 0.013733 | 0.041199 | 0.14649 | 0.16022 | 0.21057 | -0.045776 | 0.2243 | -0.032044
884.1719 | 122.5021 | -58.3694 | 63.826 | -0.10002 | -0.00058102 | 0.067691 | 0.06486 | -0.0017367 | 0.00038017 | 0.00025641 | 0.038361 | 0.00038669 | 0.065561 | 0.062551 | -0.00082252 | -0.0047573 | -0.00083948 | 0.24261 | 63.8306 | -58.4702 | 63.7436 | -58.1497 | -58.5342 | -58.5114 | 0.13275 | 63.7756 | -58.4335 | 63.8855 | -58.2458 | -58.4244 | -58.1772 | -0.10071 | 0.004577 | -0.082398 | -0.086975 | 0.21973 | 0.32044 | -0.14191 | -0.16479 | -0.041199 | -0.064087 | -0.064087 | -0.050354 | 0.059509 | 0.10986 | 0.1236 | 0.18768 | 0.19226 | -0.054931 | 0.25635 | 0.009156
885.1719 | 122.5021 | -58.3694 | 63.8077 | -0.084763 | -0.00015382 | 0.069186 | 0.068919 | 0.0021085 | 0.00080757 | -0.0025208 | 0.047516 | -0.00089511 | 0.065561 | 0.055715 | -0.00018152 | -0.00069854 | 0.00022852 | 0.050354 | 63.8031 | -58.5342 | 63.7985 | -58.0856 | -58.5663 | -58.5022 | 0.14191 | 63.7802 | -58.3969 | 63.7756 | -58.1039 | -58.4473 | -58.2001 | -0.16479 | -0.004577 | -0.009155 | -0.004578 | 0.28381 | 0.44861 | -0.13275 | -0.19684 | 0.032044 | -0.032043 | -0.027466 | -0.027466 | -0.032043 | -0.004577 | 0.2655 | 0.29297 | 0.16937 | -0.07782 | 0.19684 | -0.050354
886.1719 | 122.5479 | -58.3786 | 63.8168 | -0.061874 | 5.9776e-05 | 0.067691 | 0.066783 | -2.7692e-05 | 0.001662 | -0.0014526 | 0.062775 | -0.00068151 | 0.064706 | 0.061269 | -0.0029587 | -0.0043301 | -0.0038303 | 0.013733 | 63.7894 | -58.4976 | 63.7939 | -58.0719 | -58.5526 | -58.4885 | 0.16022 | 63.7573 | -58.4198 | 63.9267 | -58.1909 | -58.3832 | -58.1451 | -0.11902 | -0.027465 | -0.022888 | 0.004577 | 0.3067 | 0.42572 | -0.10986 | -0.17395 | 0.009156 | -0.054931 | -0.041199 | -0.059509 | 0.10986 | 0.16937 | 0.18768 | 0.22888 | 0.23346 | -0.004578 | 0.27466 | 0.036621
887.1719 | 122.4518 | -58.3923 | 63.8443 | -0.061874 | -0.0014356 | 0.068545 | 0.068278 | 0.0029629 | 0.0014484 | -0.001239 | -0.059296 | 0.0018821 | 0.065561 | 0.068746 | -0.00060892 | -0.0015529 | 0.00086952 | 0.013733 | 63.826 | -58.5571 | 63.7253 | -58.1726 | -58.6716 | -58.6121 | 0.041199 | 63.8168 | -58.4381 | 63.8489 | -58.3008 | -58.5159 | -58.2458 | -0.1648 | -0.01831 | -0.11902 | -0.10071 | 0.21973 | 0.38452 | -0.21973 | -0.27924 | -0.054932 | -0.11444 | -0.045776 | -0.027466 | 0.004578 | 0.032044 | 0.091553 | 0.13733 | 0.14648 | -0.1236 | 0.19226 | -0.07782
888.1719 | 122.5159 | -58.3786 | 63.826 | -0.10765 | 0.0019824 | 0.068545 | 0.066355 | 0.0029629 | -0.00090143 | -0.00081179 | 0.068878 | -0.00089511 | 0.063638 | 0.062124 | 0.0017411 | -0.0034757 | 0.00086952 | -0.032043 | 63.8031 | -58.5159 | 63.7619 | -58.1451 | -58.6029 | -58.5251 | 0.068665 | 63.7619 | -58.4152 | 63.8947 | -58.2184 | -58.4106 | -58.1589 | -0.13733 | -0.022888 | -0.064087 | -0.041199 | 0.23346 | 0.37079 | -0.14648 | -0.2243 | -0.009155 | -0.086975 | -0.036621 | -0.064087 | 0.068664 | 0.13275 | 0.16022 | 0.19684 | 0.21973 | -0.032044 | 0.25635 | 0.004577
889.1719 | 122.5113 | -58.3649 | 63.8214 | -0.038986 | -0.0016492 | 0.065555 | 0.065501 | -0.0017367 | 0.0018756 | -0.00038459 | 0.032257 | -0.00089511 | 0.064706 | 0.061483 | -0.0010361 | -0.0049709 | -0.0016941 | 0.004578 | 63.7894 | -58.5159 | 63.7939 | -58.0994 | -58.5663 | -58.493 | 0.14191 | 63.7985 | -58.4061 | 63.9175 | -58.1863 | -58.3878 | -58.136 | -0.15106 | -0.032043 | -0.027466 | 0.004577 | 0.2655 | 0.41656 | -0.12817 | -0.20142 | 0.022888 | -0.050354 | -0.041199 | -0.022888 | 0.096131 | 0.11902 | 0.17853 | 0.21973 | 0.22888 | -0.022888 | 0.27008 | 0.018311
890.1719 | 122.4609 | -58.3374 | 63.8168 | -0.084763 | 0.00048698 | 0.070468 | 0.067637 | -0.0047275 | 0.00080757 | 4.2808e-05 | 0.035309 | -0.00068151 | 0.063211 | 0.059987 | -0.0014633 | -0.0049709 | 0.00022852 | 0.036621 | 63.7939 | -58.4793 | 63.8123 | -58.0444 | -58.5297 | -58.4747 | 0.14191 | 63.7939 | -58.4015 | 63.8077 | -57.9117 | -58.3466 | -58.2321 | -0.14191 | -0.022888 | -0.004577 | 0.018311 | 0.29297 | 0.43487 | -0.13733 | -0.19226 | 0.004578 | -0.050354 | -0.064087 | -0.022888 | -0.009155 | 0.013733 | 0.42572 | 0.48981 | 0.10528 | -0.009156 | 0.16937 | 0.054931
891.1719 | 122.5388 | -58.3374 | 63.858 | -0.077133 | 0.00091438 | 0.067904 | 0.066569 | 0.00018591 | 0.001662 | -0.00081179 | 0.029205 | 0.00060029 | 0.065988 | 0.062551 | 3.2078e-05 | -0.0041165 | -0.0016941 | -0.027466 | 63.8031 | -58.5388 | 63.7802 | -58.1314 | -58.5938 | -58.5297 | 0.068665 | 63.8031 | -58.4061 | 63.8947 | -58.2138 | -58.4198 | -58.1589 | -0.20142 | -0.054931 | -0.07782 | -0.022889 | 0.20599 | 0.40741 | -0.19226 | -0.25635 | 0.009155 | -0.054932 | -0.068665 | -0.054931 | 0.036621 | 0.091552 | 0.1236 | 0.19226 | 0.17853 | -0.082398 | 0.24719 | -0.013733
892.1875 | 122.4426 | -58.3878 | 63.8489 | -0.046616 | 0.00091438 | 0.06705 | 0.070628 | 0.0023221 | 0.0037982 | 0.00089721 | -0.098968 | 0.00038669 | 0.064706 | 0.068105 | 0.00024568 | -0.0028347 | 0.00086952 | -0.032043 | 63.8123 | -58.5297 | 63.7527 | -58.1497 | -58.667 | -58.6121 | 0.022888 | 63.8443 | -58.4152 | 63.858 | -58.3008 | -58.5205 | -58.2367 | -0.14191 | -0.036621 | -0.09613 | -0.059509 | 0.23804 | 0.37994 | -0.2243 | -0.27924 | -0.082398 | -0.13733 | -0.027466 | -0.004578 | 0.009155 | 0.013733 | 0.086975 | 0.11444 | 0.15106 | -0.13275 | 0.17853 | -0.10529
893.1719 | 122.4747 | -58.3694 | 63.8077 | -0.069504 | 0.0013416 | 0.068332 | 0.067637 | 0.0025357 | -0.00090143 | -0.00038459 | 0.02005 | -0.0015359 | 0.065775 | 0.058492 | 0.00067288 | -0.0045437 | -0.00019868 | 0.064087 | 63.8077 | -58.4793 | 63.7527 | -58.1497 | -58.6395 | -58.5526 | -0.027466 | 63.8123 | -58.4061 | 63.8718 | -58.2367 | -58.4518 | -58.1863 | -0.10986 | 0 | -0.054931 | -0.054931 | 0.21973 | 0.32959 | -0.1831 | -0.27008 | -0.073242 | -0.16022 | -0.036621 | 0.004578 | 0.064087 | 0.059509 | 0.13275 | 0.16937 | 0.18311 | -0.082397 | 0.21973 | -0.045776
894.1719 | 122.4838 | -58.3603 | 63.8626 | -0.10002 | 0.00027338 | 0.067477 | 0.06721 | 0.00061311 | 0.0010212 | -0.001239 | 0.0047913 | 0.00081389 | 0.066415 | 0.062124 | -0.0023177 | -0.0032619 | -0.00041228 | -0.064087 | 63.8123 | -58.5114 | 63.7665 | -58.1497 | -58.6212 | -58.4839 | -0.027466 | 63.8077 | -58.4106 | 63.8443 | -58.2504 | -58.4381 | -58.1772 | -0.15106 | -0.050354 | -0.096131 | -0.045777 | 0.21057 | 0.36163 | -0.1236 | -0.26093 | 0.027466 | -0.10986 | -0.050354 | -0.054932 | -0.018311 | 0.036621 | 0.10986 | 0.16022 | 0.18311 | -0.077819 | 0.23346 | -0.027465
895.1719 | 122.5067 | -58.3466 | 63.8443 | -0.13054 | -0.001222 | 0.067904 | 0.06956 | -0.0047275 | 0.0014484 | 0.00068361 | 0.0078429 | -0.00089511 | 0.065347 | 0.062551 | -0.0031723 | -0.0032619 | -0.00041228 | -0.073242 | 63.8123 | -58.5114 | 63.7756 | -58.1314 | -58.6075 | -58.5342 | 0 | 63.8306 | -58.4106 | 63.9084 | -58.1955 | -58.4198 | -58.1772 | -0.16479 | -0.032043 | -0.068664 | -0.036621 | 0.21515 | 0.37994 | -0.18768 | -0.26093 | -0.022888 | -0.09613 | -0.064087 | -0.013733 | 0.064087 | 0.07782 | 0.15106 | 0.21515 | 0.16937 | -0.073242 | 0.23346 | -0.009155
896.1719 | 122.4747 | -58.342 | 63.8397 | 0.13649 | -0.0020764 | 0.068973 | 0.068919 | -0.00088229 | 0.0037982 | -0.0014526 | 0.0017395 | -0.00089511 | 0.063425 | 0.059133 | -0.0021041 | -0.0047573 | -0.0027621 | 0.009155 | 63.8214 | -58.5022 | 63.7665 | -58.1543 | -58.6121 | -58.5205 | -0.013733 | 63.8077 | -58.3878 | 63.8901 | -58.2047 | -58.4427 | -58.1772 | -0.16022 | -0.018311 | -0.073243 | -0.054932 | 0.18768 | 0.3479 | -0.17853 | -0.27008 | -0.018311 | -0.10986 | -0.045776 | -0.032044 | 0.050354 | 0.082398 | 0.13733 | 0.18311 | 0.16479 | -0.10071 | 0.21057 | -0.054932
897.1719 | 122.5479 | -58.3191 | 63.8672 | -0.10002 | -0.00058102 | 0.066195 | 0.066783 | -2.7692e-05 | 0.00038017 | 0.00068361 | 0.010895 | 0.0010275 | 0.064493 | 0.065542 | -0.0023177 | -0.0045437 | -0.0010531 | 0.013733 | 63.8397 | -58.4747 | 63.8535 | -58.0032 | -58.5022 | -58.4564 | 0.16479 | 63.8077 | -58.3603 | 63.9862 | -58.136 | -58.3511 | -58.1039 | -0.15564 | -0.027465 | -0.013732 | 0.013733 | 0.31586 | 0.4715 | -0.13733 | -0.18311 | 0.01831 | -0.027466 | -0.041199 | -0.059509 | 0.11902 | 0.17853 | 0.18311 | 0.2243 | 0.21515 | -0.032043 | 0.25635 | 0.009156
898.1719 | 122.4426 | -58.3282 | 63.8443 | -0.10765 | -0.00079462 | 0.068759 | 0.068278 | -0.0028049 | 0.00080757 | 0.0017518 | -0.034882 | 0.0010275 | 0.061716 | 0.058065 | -0.00060892 | -0.00069854 | 0.0012967 | -0.064087 | 63.8443 | -58.5114 | 63.7665 | -58.1451 | -58.6029 | -58.5434 | -0.041199 | 63.8123 | -58.3786 | 63.8947 | -58.2047 | -58.4473 | -58.2092 | -0.18311 | 0 | -0.07782 | -0.07782 | 0.1831 | 0.36621 | -0.21515 | -0.27466 | -0.032043 | -0.091552 | -0.050354 | -0.032043 | 0.050354 | 0.082397 | 0.1236 | 0.17395 | 0.11902 | -0.11902 | 0.16937 | -0.068665
899.1719 | 122.5342 | -58.3374 | 63.8626 | -0.10765 | -0.00058102 | 0.068759 | 0.069133 | 0.0014675 | 0.00038017 | -0.0033752 | -0.03183 | -0.0013223 | 0.06492 | 0.059347 | -0.00082252 | -0.0043301 | -0.0012667 | -0.07782 | 63.8443 | -58.4839 | 63.7894 | -58.1177 | -58.5846 | -58.5388 | -0.027466 | 63.8443 | -58.3694 | 63.8947 | -58.2184 | -58.429 | -58.2458 | -0.14649 | -0.018311 | -0.073242 | -0.054931 | 0.21973 | 0.36621 | -0.20142 | -0.24719 | -0.054931 | -0.10071 | -0.032044 | -0.018311 | 0.032043 | 0.050354 | 0.11902 | 0.15106 | 0.091552 | -0.091553 | 0.1236 | -0.059509
900.1719 | 122.4747 | -58.3282 | 63.858 | -0.061874 | 0.0019824 | 0.068118 | 0.069773 | 0.00018591 | 0.00016657 | 0.0030334 | -0.037933 | 0.00017309 | 0.065988 | 0.062124 | -0.0025315 | -0.0049709 | 0.0010831 | -0.086975 | 63.8397 | -58.4885 | 63.7802 | -58.1223 | -58.5983 | -58.5251 | -0.032043 | 63.8214 | -58.3786 | 63.913 | -58.2184 | -58.4473 | -58.1726 | -0.16022 | -0.01831 | -0.07782 | -0.05951 | 0.20599 | 0.36621 | -0.19684 | -0.27008 | -0.036621 | -0.10986 | -0.050354 | -0.036621 | 0.054932 | 0.091553 | 0.10986 | 0.16022 | 0.15564 | -0.11902 | 0.20599 | -0.068665
901.1719 | 122.4838 | -58.3145 | 63.8718 | -0.15343 | 0.00027338 | 0.069186 | 0.068492 | 0.0038175 | 0.0020892 | 0.0019654 | -0.05014 | 0.00060029 | 0.062784 | 0.062764 | -0.0012497 | -0.0039029 | -0.0016941 | -0.091553 | 63.8397 | -58.4747 | 63.7848 | -58.1085 | -58.6121 | -58.548 | -0.045776 | 63.826 | -58.4015 | 63.913 | -58.2184 | -58.4152 | -58.1543 | -0.16022 | -0.032043 | -0.086975 | -0.054932 | 0.20599 | 0.36621 | -0.23346 | -0.29755 | -0.073243 | -0.13733 | -0.086975 | -0.045776 | 0.041199 | 0.086975 | 0.09613 | 0.1831 | 0.16022 | -0.10071 | 0.24719 | -0.013733
902.1719 | 122.5388 | -58.2687 | 63.8855 | -0.10002 | -0.00036742 | 0.070468 | 0.068064 | -0.0030185 | 0.0020892 | -0.00038459 | -0.16916 | 0.0012411 | 0.060648 | 0.063833 | 3.2078e-05 | -0.0026211 | -0.00083948 | 0.032043 | 63.858 | -58.4518 | 63.8763 | -58.0124 | -58.0856 | -58.3649 | 0.18768 | 63.8397 | -58.3466 | 63.9679 | -58.1085 | -58.3328 | -58.0948 | -0.1831 | -0.027466 | -0.009155 | 0.018311 | 0.25635 | 0.43945 | -0.09613 | 0.18311 | 0.086975 | 0.36621 | -0.07782 | -0.045776 | 0.082398 | 0.12817 | 0.16022 | 0.23804 | 0.17395 | -0.064087 | 0.25177 | 0.013733
903.1719 | 122.4838 | -58.2825 | 63.8672 | -0.092392 | 0.002196 | 0.069186 | 0.068919 | -0.0013095 | 0.00080757 | -0.00081179 | -0.037933 | -0.0015359 | 0.062784 | 0.060628 | -0.00039512 | -0.0043301 | -0.0053257 | -0.068665 | 63.8809 | -58.4793 | 63.7802 | -58.1085 | -58.5892 | -58.5114 | -0.027466 | 63.8626 | -58.3878 | 63.8306 | -58.049 | -58.3328 | -58.1223 | -0.19684 | 0.013733 | -0.086975 | -0.10071 | 0.17395 | 0.37079 | -0.22888 | -0.3067 | -0.032044 | -0.10986 | -0.10529 | -0.004577 | -0.036621 | -0.032044 | 0.23346 | 0.33875 | 0.16022 | -0.050354 | 0.2655 | 0.054931
904.1719 | 122.4792 | -58.3145 | 63.858 | -0.10765 | 0.00070078 | 0.067477 | 0.06956 | -0.00045509 | 0.0010212 | -0.00059819 | -0.047089 | 0.0012411 | 0.064493 | 0.062764 | -0.00039512 | -0.0011257 | 0.0017239 | -0.091553 | 63.8535 | -58.4702 | 63.7894 | -58.0994 | -58.5938 | -58.5068 | -0.045776 | 63.8489 | -58.3923 | 63.9221 | -58.2138 | -58.4198 | -58.1451 | -0.15564 | -0.004577 | -0.068664 | -0.064087 | 0.21515 | 0.37079 | -0.19226 | -0.27924 | -0.036621 | -0.1236 | -0.07782 | -0.009155 | 0.064087 | 0.073242 | 0.10071 | 0.17853 | 0.16937 | -0.10529 | 0.24719 | -0.027466
905.1719 | 122.4518 | -58.3145 | 63.8901 | -0.10002 | 0.00048698 | 0.066836 | 0.069133 | -0.00045509 | 0.0014484 | -0.0010254 | -0.03183 | 0.00017309 | 0.062357 | 0.068319 | -0.0025315 | -0.0028347 | 0.0012967 | -0.07782 | 63.8306 | -58.4656 | 63.7985 | -58.1131 | -58.5663 | -58.5022 | -0.041199 | 63.826 | -58.374 | 63.9359 | -58.1955 | -58.4106 | -58.1635 | -0.15106 | -0.05951 | -0.091553 | -0.032043 | 0.20142 | 0.35248 | -0.18768 | -0.25177 | -0.036621 | -0.10071 | -0.059509 | -0.064087 | 0.045776 | 0.10986 | 0.11902 | 0.17853 | 0.15106 | -0.096131 | 0.21057 | -0.036622
906.1719 | 122.4838 | -58.2962 | 63.8855 | -0.11528 | 0.0019824 | 0.067264 | 0.069133 | -0.0019503 | -0.0015424 | 0.00068361 | -0.019623 | -0.0011087 | 0.065988 | 0.070241 | -0.00039512 | -0.0028347 | -0.0021213 | -0.082397 | 63.8214 | -58.4564 | 63.8214 | -58.0948 | -58.5754 | -58.5205 | -0.036621 | 63.8489 | -58.3511 | 63.945 | -58.1726 | -58.4015 | -58.1451 | -0.16022 | -0.064087 | -0.064087 | 0 | 0.20142 | 0.36163 | -0.2243 | -0.27923 | -0.064087 | -0.11902 | -0.054931 | -0.036621 | 0.059509 | 0.09613 | 0.1236 | 0.17853 | 0.15106 | -0.10528 | 0.20599 | -0.050354
907.1719 | 122.4289 | -58.2916 | 63.9313 | -0.046616 | 0.001128 | 0.068545 | 0.067424 | -0.0015231 | 0.00080757 | 4.2808e-05 | -0.053192 | 0.00081389 | 0.063638 | 0.06426 | -0.0012497 | -0.0039029 | -0.00083948 | 0.036621 | 63.8947 | -58.429 | 63.8718 | -58.017 | -58.4839 | -58.4564 | 0.09613 | 63.9084 | -58.3237 | 63.9404 | -58.1635 | -58.3832 | -58.1268 | -0.13733 | -0.036621 | -0.059509 | -0.022888 | 0.27466 | 0.41199 | -0.16479 | -0.19226 | -0.027466 | -0.054932 | -0.032043 | -0.022888 | 0.009156 | 0.032044 | 0.12817 | 0.16022 | 0.16479 | -0.091553 | 0.19684 | -0.05951
908.1719 | 122.4747 | -58.3008 | 63.8809 | -0.12291 | 0.00027338 | 0.070468 | 0.070628 | -0.00088229 | -4.7032e-05 | 0.001538 | -0.059296 | 0.0018821 | 0.065775 | 0.062337 | -0.0027451 | -0.0026211 | -0.00019868 | -0.054932 | 63.8535 | -58.4518 | 63.8168 | -58.0719 | -58.5846 | -58.4793 | -0.022888 | 63.8351 | -58.3603 | 63.9313 | -58.2001 | -58.3878 | -58.1131 | -0.15106 | -0.027465 | -0.064087 | -0.036622 | 0.22888 | 0.37994 | -0.17853 | -0.28381 | -0.027466 | -0.13275 | -0.05951 | -0.045776 | 0.050354 | 0.09613 | 0.10071 | 0.16022 | 0.18768 | -0.086975 | 0.24719 | -0.027465
909.1719 | 122.4289 | -58.2962 | 63.9084 | -0.10765 | 0.00091438 | 0.068973 | 0.069133 | -0.00088229 | 0.0010212 | -0.00038459 | -0.044037 | -4.0508e-05 | 0.064493 | 0.06426 | -0.00082252 | -0.0039029 | 0.0010831 | -0.073242 | 63.8901 | -58.4427 | 63.8306 | -58.0765 | -58.548 | -58.4885 | -0.054932 | 63.8763 | -58.3328 | 63.9267 | -58.0582 | -58.3694 | -58.1039 | -0.14648 | -0.01831 | -0.07782 | -0.05951 | 0.21973 | 0.36621 | -0.19226 | -0.25177 | -0.045776 | -0.10529 | -0.036621 | -0.032043 | 0.018311 | 0.050354 | 0.23804 | 0.27466 | 0.19226 | -0.073242 | 0.22888 | -0.036621
910.1719 | 122.4838 | -58.2779 | 63.9084 | -0.10002 | -0.0014356 | 0.068332 | 0.066355 | -0.0013095 | 0.0010212 | 0.00089721 | -0.028778 | -4.0508e-05 | 0.063211 | 0.06426 | -0.0012497 | -0.0045437 | 0.0023647 | -0.086975 | 63.8901 | -58.4335 | 63.8351 | -58.0765 | -58.5663 | -58.4702 | -0.013733 | 63.8718 | -58.3328 | 63.9542 | -58.1772 | -58.3832 | -58.1131 | -0.15564 | -0.01831 | -0.073242 | -0.054932 | 0.20142 | 0.35706 | -0.19226 | -0.28839 | -0.036621 | -0.13275 | -0.054932 | -0.036621 | 0.045777 | 0.082398 | 0.10071 | 0.15564 | 0.16479 | -0.10529 | 0.21973 | -0.050354
911.1719 | 122.4426 | -58.2825 | 63.913 | -0.12291 | 0.00027338 | 0.068973 | 0.068278 | -0.0034457 | 0.0012348 | -0.00081179 | -0.053192 | 0.0018821 | 0.064706 | 0.06661 | 0.00088648 | -0.0043301 | 0.0027921 | -0.064087 | 63.8947 | -58.4381 | 63.8306 | -58.0673 | -58.5526 | -58.4656 | -0.022888 | 63.8672 | -58.342 | 63.9587 | -58.1589 | -58.374 | -58.1177 | -0.15564 | -0.018311 | -0.082398 | -0.064087 | 0.21515 | 0.37079 | -0.1831 | -0.27008 | -0.027466 | -0.11444 | -0.059509 | -0.045777 | 0.045776 | 0.091553 | 0.1236 | 0.1831 | 0.16479 | -0.091552 | 0.2243 | -0.032043
912.1719 | 122.4838 | -58.255 | 63.9175 | -0.016098 | 0.00048698 | 0.068118 | 0.068919 | 0.00082671 | 0.00038017 | 0.00089721 | -0.037933 | 0.0016685 | 0.063852 | 0.055074 | -0.00039512 | -0.0039029 | 0.00086952 | -0.022888 | 63.8901 | -58.4152 | 63.8992 | -57.9803 | -58.4518 | -58.4061 | 0.10986 | 63.9175 | -58.3099 | 64.0137 | -58.0948 | -58.3282 | -58.0582 | -0.16022 | -0.027466 | -0.018311 | 0.009155 | 0.27466 | 0.43487 | -0.15106 | -0.19684 | 0.009155 | -0.036621 | -0.054932 | 0 | 0.09613 | 0.09613 | 0.16022 | 0.21515 | 0.19684 | -0.073242 | 0.25177 | -0.01831
913.1719 | 122.4518 | -58.2733 | 63.9084 | -0.12291 | 5.9776e-05 | 0.067904 | 0.068064 | -0.0060093 | 0.0023028 | 0.002179 | -0.040985 | 0.0010275 | 0.063852 | 0.06661 | -0.00039512 | -0.0028347 | 0.0017239 | -0.082397 | 63.8626 | -58.429 | 63.8397 | -58.0399 | -58.5526 | -58.461 | -0.032043 | 63.8672 | -58.3511 | 63.9679 | -58.1497 | -58.3694 | -58.1039 | -0.15564 | -0.045776 | -0.068664 | -0.022888 | 0.23346 | 0.3891 | -0.18768 | -0.27924 | -0.032044 | -0.1236 | -0.07782 | -0.041199 | 0.05951 | 0.10071 | 0.1236 | 0.20142 | 0.16937 | -0.096131 | 0.24719 | -0.018311
914.1719 | 122.4747 | -58.2687 | 63.9175 | -0.12291 | 0.001128 | 0.070041 | 0.068064 | -0.0047275 | 0.0014484 | -0.00038459 | -0.077606 | -0.0013223 | 0.062784 | 0.065755 | -0.0012497 | -0.0017667 | -0.0010531 | -0.054932 | 63.8901 | -58.429 | 63.8672 | -57.9941 | -58.5159 | -58.4427 | 0.004578 | 63.8809 | -58.287 | 63.9954 | -58.1497 | -58.3511 | -58.0994 | -0.16022 | -0.027466 | -0.050355 | -0.022889 | 0.27466 | 0.43488 | -0.17395 | -0.24719 | -0.013733 | -0.086975 | -0.01831 | -0.036622 | 0.077819 | 0.11444 | 0.11902 | 0.13733 | 0.16937 | -0.082397 | 0.18768 | -0.064087
915.1719 | 122.4701 | -58.2458 | 63.9221 | -0.061874 | 0.0019824 | 0.068332 | 0.068919 | -0.0028049 | -0.00068783 | -0.002948 | -0.028778 | -0.00046791 | 0.064493 | 0.059347 | -0.0014633 | -0.0021939 | 0.00065592 | -0.082397 | 63.8901 | -58.4381 | 63.8351 | -58.0673 | -58.5526 | -58.461 | -0.036621 | 63.8855 | -58.3145 | 63.9679 | -58.1451 | -58.3603 | -58.0948 | -0.19226 | -0.032043 | -0.086975 | -0.054932 | 0.17853 | 0.37079 | -0.21515 | -0.3067 | -0.022889 | -0.11444 | -0.068664 | -0.036621 | 0.045777 | 0.082398 | 0.10071 | 0.16937 | 0.15106 | -0.11444 | 0.21973 | -0.045777
916.1719 | 122.4838 | -58.2596 | 63.913 | -0.11528 | -0.00058102 | 0.069613 | 0.070628 | 0.00061311 | 0.00080757 | 0.0017518 | 0.13297 | 0.0010275 | 0.064706 | 0.059774 | 0.0011001 | -0.0013393 | 0.0055691 | -0.082397 | 63.9221 | -58.4198 | 63.8443 | -58.0399 | -58.5114 | -58.4564 | -0.018311 | 63.9313 | -58.3191 | 63.9908 | -58.1406 | -58.342 | -58.0536 | -0.16022 | 0.009155 | -0.068665 | -0.07782 | 0.21973 | 0.37994 | -0.19684 | -0.25177 | -0.036621 | -0.091553 | -0.059509 | 0.01831 | 0.07782 | 0.05951 | 0.11902 | 0.17853 | 0.20599 | -0.082397 | 0.2655 | -0.022888
917.1719 | 122.5067 | -58.2321 | 63.9496 | -0.10002 | -0.00015382 | 0.068332 | 0.06721 | 0.0025357 | 0.00016657 | -0.00017099 | -0.14474 | 0.0023093 | 0.065134 | 0.06191 | -0.0029587 | -0.0028347 | -0.00083948 | 0.036621 | 63.945 | -58.3649 | 63.9496 | -57.9117 | -58.1726 | -58.1818 | 0.013733 | 63.9038 | -58.2687 | 64.0457 | -58.0627 | -58.2642 | -58.0124 | -0.13275 | -0.004578 | 0 | 0.004578 | 0.32044 | 0.45319 | 0.050354 | 0.05951 | 0.18311 | 0.19226 | -0.036621 | -0.045776 | 0.09613 | 0.14191 | 0.16937 | 0.20599 | 0.21973 | -0.032043 | 0.25635 | 0.004578
918.1719 | 122.438 | -58.2413 | 63.9267 | -0.10002 | 0.00048698 | 0.068332 | 0.068278 | -0.0019503 | 0.00059397 | -0.0031616 | -0.022675 | 0.00017309 | 0.061716 | 0.069173 | -0.00082252 | -0.0015529 | -0.0021213 | -0.068665 | 63.945 | -58.4061 | 63.8535 | -58.0261 | -58.5251 | -58.4427 | -0.018311 | 63.9313 | -58.3099 | 63.9908 | -58.1039 | -58.2825 | -57.93 | -0.16479 | 0.01831 | -0.073242 | -0.091552 | 0.21515 | 0.37994 | -0.20142 | -0.28381 | -0.036621 | -0.11902 | -0.068665 | 0.004577 | 0.064087 | 0.05951 | 0.13733 | 0.20599 | 0.31128 | -0.041199 | 0.37994 | 0.027466
919.1719 | 122.4747 | -58.2596 | 63.9542 | -0.10765 | -0.00058102 | 0.066836 | 0.068064 | 0.0029629 | 0.0027302 | -0.0033752 | -0.040985 | -0.00025411 | 0.066202 | 0.065114 | -0.0027451 | -0.0045437 | -0.0010531 | -0.082397 | 63.9313 | -58.4106 | 63.8626 | -58.017 | -58.4976 | -58.4381 | -0.027466 | 63.9175 | -58.2962 | 63.9908 | -58.1223 | -58.3466 | -58.0673 | -0.15106 | -0.022889 | -0.091553 | -0.068664 | 0.24262 | 0.39368 | -0.17853 | -0.23804 | -0.027465 | -0.086975 | -0.036621 | -0.036621 | 0.036621 | 0.073242 | 0.13733 | 0.17395 | 0.19226 | -0.086975 | 0.22888 | -0.050354
920.1719 | 122.4838 | -58.2321 | 63.945 | -0.10765 | -0.0016492 | 0.068545 | 0.068278 | 0.0018949 | 0.0048664 | -0.00081179 | -0.019623 | -0.00046791 | 0.06492 | 0.061696 | -0.0010361 | -0.0062527 | 0.0030057 | -0.073242 | 63.9313 | -58.3832 | 63.8763 | -58.0032 | -58.5068 | -58.4564 | 0 | 63.9175 | -58.3008 | 64.0182 | -58.0856 | -58.3008 | -58.0536 | -0.15106 | -0.013733 | -0.068664 | -0.054931 | 0.22888 | 0.37994 | -0.2243 | -0.27466 | -0.073242 | -0.1236 | -0.068664 | -0.027465 | 0.073243 | 0.10071 | 0.14649 | 0.21515 | 0.17853 | -0.068664 | 0.24719 | 0
921.1719 | 122.4792 | -58.2458 | 63.9404 | -0.10002 | 0.0024096 | 0.068545 | 0.070201 | 0.0031765 | 0.0023028 | 0.0013244 | -0.077606 | 0.00081389 | 0.06492 | 0.065542 | -0.0016769 | -0.0019803 | 0.0025785 | -0.059509 | 63.9038 | -58.4015 | 63.8992 | -57.9666 | -58.493 | -58.4061 | 0.018311 | 63.8901 | -58.287 | 64.0366 | -58.1223 | -58.3282 | -58.0536 | -0.15564 | -0.036621 | -0.041199 | -0.004578 | 0.27924 | 0.43487 | -0.16022 | -0.24719 | -0.004578 | -0.091553 | -0.041198 | -0.050354 | 0.09613 | 0.14648 | 0.1236 | 0.16479 | 0.19226 | -0.082397 | 0.23346 | -0.041199
922.1719 | 122.493 | -58.223 | 63.9771 | -0.092392 | 0.00048698 | 0.069613 | 0.067637 | -0.0043003 | 0.00059397 | -0.0035888 | -0.010467 | -4.0508e-05 | 0.064066 | 0.064687 | 0.0013139 | -0.0049709 | 0.00022852 | -0.045776 | 63.9542 | -58.3832 | 63.9038 | -57.9712 | -58.461 | -58.3923 | 0.027466 | 63.9404 | -58.287 | 64.0503 | -58.049 | -58.2779 | -58.0124 | -0.16022 | -0.022888 | -0.073242 | -0.050354 | 0.25177 | 0.41199 | -0.16937 | -0.23804 | -0.009155 | -0.07782 | -0.064087 | -0.036621 | 0.073242 | 0.10986 | 0.17395 | 0.23804 | 0.21057 | -0.054932 | 0.27466 | 0.009155
923.1719 | 122.525 | -58.2092 | 63.9725 | -0.16106 | 5.9776e-05 | 0.068118 | 0.068064 | 0.0012539 | 0.0014484 | 0.00068361 | -0.047089 | -4.0508e-05 | 0.064706 | 0.060842 | -0.0014633 | -0.0011257 | 0.0012967 | -0.073242 | 63.9587 | -58.3969 | 63.8809 | -58.017 | -58.5022 | -58.4244 | -0.027466 | 63.945 | -58.2962 | 64.0137 | -58.0811 | -58.287 | -57.9712 | -0.18768 | -0.013733 | -0.091553 | -0.07782 | 0.19226 | 0.37994 | -0.21515 | -0.29297 | -0.027465 | -0.10529 | -0.086975 | -0.027466 | 0.041199 | 0.068665 | 0.12817 | 0.21515 | 0.23804 | -0.077819 | 0.32501 | 0.009156
924.1719 | 122.4747 | -58.2138 | 63.9496 | -0.077133 | -0.00058102 | 0.065982 | 0.066569 | -0.00066869 | 0.00016657 | -0.0018798 | -0.028778 | 0.0020957 | 0.063852 | 0.06661 | -0.0016769 | -0.0034757 | -0.0044711 | -0.091553 | 63.9587 | -58.3832 | 63.9038 | -57.9941 | -58.493 | -58.4015 | -0.027466 | 63.9404 | -58.2733 | 64.0182 | -58.0994 | -58.3145 | -58.0673 | -0.16937 | 0.009155 | -0.045776 | -0.054931 | 0.21973 | 0.3891 | -0.18768 | -0.27924 | -0.01831 | -0.10986 | -0.059509 | -0.009155 | 0.068665 | 0.07782 | 0.11444 | 0.17395 | 0.14648 | -0.10071 | 0.20599 | -0.041199
925.1719 | 122.493 | -58.223 | 63.9633 | -0.13054 | 0.00027338 | 0.070895 | 0.068705 | -0.0017367 | 0.0020892 | 0.0034606 | -0.056244 | -0.00046791 | 0.063638 | 0.062337 | 3.2078e-05 | -0.0028347 | -0.0036167 | -0.091553 | 63.9359 | -58.374 | 63.8718 | -58.0078 | -58.4793 | -58.4152 | -0.032043 | 63.9175 | -58.2596 | 64.0228 | -58.0948 | -58.3282 | -58.049 | -0.15106 | -0.027466 | -0.091553 | -0.064087 | 0.21515 | 0.36621 | -0.19226 | -0.25635 | -0.041199 | -0.10529 | -0.036622 | -0.045776 | 0.059509 | 0.10529 | 0.12817 | 0.16479 | 0.17395 | -0.10529 | 0.21057 | -0.068664
926.1719 | 122.4564 | -58.223 | 63.9954 | -0.077133 | 0.00091438 | 0.068973 | 0.066783 | -0.0049411 | 0.00080757 | 0.0013244 | -0.040985 | 0.00017309 | 0.062997 | 0.065114 | -0.00039512 | -0.0039029 | -0.0012667 | -0.07782 | 63.9404 | -58.3694 | 63.8947 | -57.9849 | -58.4793 | -58.4106 | -0.050354 | 63.9496 | -58.2687 | 64.0274 | -58.0902 | -58.3237 | -58.0582 | -0.14649 | -0.054931 | -0.10071 | -0.045777 | 0.23804 | 0.38452 | -0.18768 | -0.25635 | -0.041199 | -0.10986 | -0.045777 | -0.045776 | 0.032044 | 0.07782 | 0.13275 | 0.17853 | 0.16479 | -0.10071 | 0.21057 | -0.054931
927.1719 | 122.4701 | -58.223 | 64.0228 | -0.11528 | -0.00015382 | 0.068332 | 0.06721 | -2.7692e-05 | 0.0014484 | 0.00089721 | -0.040985 | 0.00017309 | 0.063852 | 0.067464 | -0.00060892 | -0.0028347 | -0.0051121 | 0.036621 | 63.9999 | -58.3649 | 63.9725 | -57.8979 | -58.3923 | -58.3511 | 0.091553 | 63.9725 | -58.2413 | 64.064 | -58.0307 | -58.2642 | -58.017 | -0.14191 | -0.022888 | -0.050354 | -0.027466 | 0.32501 | 0.46692 | -0.12817 | -0.16937 | 0.013733 | -0.027466 | -0.018311 | -0.050354 | 0.041199 | 0.091553 | 0.19226 | 0.21057 | 0.20599 | -0.041199 | 0.2243 | -0.022888
928.1719 | 122.5021 | -58.1726 | 63.9679 | -0.12291 | 0.002196 | 0.067264 | 0.069133 | -0.0015231 | 0.0046528 | 0.0017518 | -0.11423 | 0.00017309 | 0.062357 | 0.067891 | -0.0014633 | -0.0026211 | -0.0023349 | -0.032043 | 63.9313 | -58.3832 | 63.8031 | -57.9575 | -58.1863 | -58.2687 | 0.045776 | 63.9313 | -58.2413 | 64.0686 | -58.0444 | -58.2916 | -58.017 | -0.21057 | -0.036622 | -0.1648 | -0.12817 | 0.21515 | 0.42572 | -0.096131 | -0.013733 | 0.11444 | 0.19684 | -0.068665 | -0.036622 | 0.10071 | 0.13733 | 0.12817 | 0.19684 | 0.15564 | -0.11902 | 0.2243 | -0.050354
929.1719 | 122.4701 | -58.2275 | 63.9542 | -0.18394 | 0.0013416 | 0.068973 | 0.065715 | -0.0023777 | 0.0012348 | -0.00059819 | -0.053192 | 0.0027365 | 0.064279 | 0.064687 | -0.0012497 | -0.0024075 | -0.0019077 | -0.068665 | 63.945 | -58.3237 | 63.9496 | -57.9208 | -58.461 | -58.3832 | 0.009155 | 63.9542 | -58.2458 | 64.0732 | -58.0627 | -58.2916 | -58.0078 | -0.09613 | -0.009156 | -0.004578 | 0.004578 | 0.3067 | 0.40283 | -0.15564 | -0.23346 | -0.05951 | -0.13733 | -0.018311 | 0 | 0.11902 | 0.11902 | 0.16479 | 0.18311 | 0.21973 | -0.064087 | 0.23804 | -0.045776
930.1719 | 122.4884 | -58.1955 | 63.9954 | -0.092392 | -0.0027172 | 0.067264 | 0.068278 | -0.00066869 | -0.00068783 | 0.0017518 | -0.034882 | -0.00068151 | 0.065988 | 0.066823 | 0.00067288 | -0.0028347 | 0.0025785 | -0.068665 | 63.9496 | -58.342 | 63.9038 | -57.9712 | -58.4702 | -58.4015 | -0.054932 | 63.9542 | -58.2458 | 64.0366 | -58.0536 | -58.3054 | -58.0399 | -0.14648 | -0.045776 | -0.091552 | -0.045776 | 0.22431 | 0.37079 | -0.20599 | -0.27466 | -0.059509 | -0.12817 | -0.050354 | -0.041198 | 0.041199 | 0.082397 | 0.14191 | 0.19226 | 0.15564 | -0.10986 | 0.20599 | -0.059509
931.1719 | 122.4564 | -58.2001 | 63.9954 | -0.069504 | 0.00091438 | 0.067264 | 0.069346 | 0.0036039 | -0.0013288 | -0.00059819 | -0.037933 | 0.00017309 | 0.061929 | 0.058919 | -0.00082252 | -0.0011257 | 0.00044212 | -0.068665 | 63.9725 | -58.3649 | 63.9221 | -57.9391 | -58.4702 | -58.3832 | -0.032043 | 63.9587 | -58.255 | 64.0457 | -58.0765 | -58.3282 | -58.0261 | -0.16479 | -0.022888 | -0.073242 | -0.050354 | 0.26093 | 0.42572 | -0.18311 | -0.27008 | -0.018311 | -0.10529 | -0.054932 | -0.036621 | 0.050354 | 0.086975 | 0.1236 | 0.17853 | 0.17395 | -0.12817 | 0.22888 | -0.073242
932.1719 | 122.4564 | -58.168 | 64.0274 | -0.10002 | 0.0026234 | 0.068759 | 0.066569 | -0.0023777 | 0.0025166 | 0.0030334 | -0.013519 | -0.00089511 | 0.063852 | 0.058278 | -0.0027451 | -0.0013393 | 0.0012967 | -0.004578 | 63.9725 | -58.3145 | 63.9816 | -57.8796 | -58.3557 | -58.3099 | 0.10071 | 64.0091 | -58.2047 | 64.0915 | -57.9987 | -58.2321 | -57.9712 | -0.14648 | -0.054932 | -0.045777 | 0.009155 | 0.28839 | 0.43488 | -0.14191 | -0.18768 | 0.004577 | -0.041199 | -0.036621 | -0.018311 | 0.064087 | 0.082398 | 0.16937 | 0.20599 | 0.19684 | -0.064087 | 0.23346 | -0.027466
933.1719 | 122.4884 | -58.2092 | 63.9816 | -0.10002 | -0.0014356 | 0.070254 | 0.068064 | 0.0042447 | 0.0014484 | -0.0018798 | -0.05014 | -0.0011087 | 0.063852 | 0.065542 | -0.00018152 | -0.0021939 | -0.0034031 | -0.045776 | 63.9542 | -58.3237 | 63.9587 | -57.9163 | -58.429 | -58.3603 | 0.009155 | 63.9587 | -58.2367 | 64.0732 | -58.0627 | -58.287 | -58.0032 | -0.11444 | -0.027465 | -0.022888 | 0.004577 | 0.29297 | 0.40741 | -0.15106 | -0.21973 | -0.036622 | -0.10529 | -0.027465 | -0.022888 | 0.091553 | 0.11444 | 0.14648 | 0.17395 | 0.20599 | -0.077819 | 0.23346 | -0.050354
934.1719 | 122.5021 | -58.1635 | 63.9633 | -0.054245 | 0.00048698 | 0.068973 | 0.068278 | 0.0016813 | 0.001662 | 0.0028198 | -0.062347 | 0.0027365 | 0.063852 | 0.060201 | -0.0010361 | -0.0034757 | 0.00086952 | -0.050354 | 63.9587 | -58.3054 | 63.9679 | -57.8888 | -58.4198 | -58.3466 | 0.009155 | 63.9679 | -58.2367 | 64.0961 | -58.0307 | -58.2596 | -57.9941 | -0.14191 | -0.004578 | 0.004578 | 0.009156 | 0.27466 | 0.41657 | -0.18311 | -0.25635 | -0.041199 | -0.11444 | -0.073242 | 0.004578 | 0.13275 | 0.12817 | 0.13275 | 0.20599 | 0.16937 | -0.096131 | 0.24261 | -0.022889
935.1719 | 122.5113 | -58.1543 | 63.9816 | -0.077133 | 0.0032642 | 0.070895 | 0.068064 | -0.0015231 | 0.00080757 | -0.0014526 | -0.10507 | 0.0025229 | 0.064493 | 0.072591 | -0.0014633 | 0.0014377 | -0.0019077 | -0.059509 | 63.9725 | -58.3282 | 63.9587 | -57.8979 | -58.429 | -58.3557 | 0 | 63.9542 | -58.2275 | 64.1235 | -58.0215 | -58.2367 | -57.9483 | -0.17395 | -0.009155 | -0.022888 | -0.013733 | 0.25635 | 0.4303 | -0.20142 | -0.27466 | -0.027466 | -0.10071 | -0.073242 | -0.027465 | 0.14191 | 0.16937 | 0.13275 | 0.20599 | 0.20599 | -0.082397 | 0.27924 | -0.009155
936.1719 | 122.5204 | -58.1726 | 63.9954 | -0.13054 | -0.001222 | 0.070041 | 0.06956 | 0.0012539 | 0.0020892 | 0.0019654 | -0.092865 | 0.0010275 | 0.064066 | 0.063192 | -0.00018152 | -0.0032619 | 0.0023647 | -0.036621 | 63.9679 | -58.2962 | 63.8626 | -58.0719 | -58.2001 | -58.1955 | 0.018311 | 63.9862 | -58.223 | 64.119 | -58.017 | -58.2367 | -57.9758 | -0.1236 | -0.027465 | -0.13275 | -0.10529 | 0.10071 | 0.22431 | -0.022889 | -0.027466 | 0.10071 | 0.096131 | -0.050354 | -0.009155 | 0.1236 | 0.13275 | 0.15564 | 0.20599 | 0.19684 | -0.064087 | 0.24719 | -0.013733
937.1719 | 122.5342 | -58.1268 | 64.0457 | -0.054245 | -0.001222 | 0.066409 | 0.068492 | 0.0027493 | 0.00016657 | 0.0011108 | 0.029205 | -0.00025411 | 0.064493 | 0.056783 | -0.0025315 | -0.0017667 | -0.0021213 | 0.004578 | 64.032 | -58.2779 | 64.0182 | -57.811 | -58.3099 | -58.2779 | 0.14648 | 63.9771 | -58.1543 | 64.1785 | -57.9254 | -58.136 | -57.9208 | -0.15106 | -0.013733 | -0.027465 | -0.013732 | 0.31586 | 0.46692 | -0.15106 | -0.18311 | 0 | -0.032044 | -0.027466 | -0.068664 | 0.13275 | 0.20142 | 0.20142 | 0.22888 | 0.20599 | -0.009155 | 0.23346 | 0.018311
938.1719 | 122.4701 | -58.1406 | 64.0778 | -0.10765 | -0.003999 | 0.069827 | 0.068278 | -0.0015231 | 0.0020892 | -0.0020936 | -0.013519 | -0.0015359 | 0.064066 | 0.064687 | -0.0014633 | -0.0030483 | -0.0044711 | 0.027466 | 64.0594 | -58.2962 | 63.9954 | -57.8567 | -58.3603 | -58.3145 | 0.10529 | 64.0274 | -58.1909 | 64.0961 | -57.9895 | -58.2184 | -57.9666 | -0.15564 | -0.018311 | -0.082398 | -0.064087 | 0.28381 | 0.43945 | -0.17395 | -0.21973 | -0.01831 | -0.064087 | -0.050354 | -0.050354 | 0.01831 | 0.068664 | 0.15106 | 0.20142 | 0.17395 | -0.07782 | 0.2243 | -0.027466
939.1719 | 122.4976 | -58.136 | 64.0366 | -0.092392 | 0.0015552 | 0.070254 | 0.071055 | -0.0047275 | -0.0015424 | -0.00038459 | -0.065399 | 0.00060029 | 0.063852 | 0.062764 | -0.0031723 | -0.0028347 | 0.0023647 | -0.050354 | 63.9908 | -58.3191 | 64.0457 | -57.9163 | -58.3649 | -58.3054 | 0.027466 | 64.0137 | -58.1955 | 64.1464 | -58.017 | -58.2321 | -57.9483 | -0.18311 | -0.045776 | 0.009155 | 0.054931 | 0.21973 | 0.40283 | -0.16937 | -0.22888 | 0.013733 | -0.045776 | -0.05951 | -0.022888 | 0.10986 | 0.13275 | 0.11902 | 0.17853 | 0.18768 | -0.096131 | 0.24719 | -0.036621
940.1719 | 122.5342 | -58.1314 | 64.0228 | -0.11528 | -0.00036742 | 0.068973 | 0.06956 | 0.0031765 | 0.00038017 | -0.0014526 | -0.089813 | 0.0010275 | 0.067056 | 0.0696 | -0.00082252 | -0.0030483 | -0.0038303 | -0.041199 | 63.9725 | -58.3054 | 63.9771 | -57.8705 | -58.3969 | -58.3282 | -0.086975 | 63.9587 | -58.2321 | 64.1418 | -57.9987 | -58.2138 | -57.9117 | -0.17395 | -0.050354 | -0.045776 | 0.004578 | 0.26093 | 0.43488 | -0.19684 | -0.2655 | -0.022888 | -0.091553 | -0.10071 | -0.064087 | 0.11902 | 0.18311 | 0.13275 | 0.23346 | 0.21973 | -0.082397 | 0.32044 | 0.018311
941.1719 | 122.58 | -58.1406 | 64.0366 | -0.12291 | 0.0013416 | 0.069186 | 0.068278 | 0.0010403 | -4.7032e-05 | 0.001538 | -0.068451 | -0.0013223 | 0.06492 | 0.060201 | -0.00060892 | -0.0041165 | 0.0019375 | -0.050354 | 63.9954 | -58.3191 | 63.9954 | -57.8796 | -58.4015 | -58.3237 | -0.009155 | 64.0182 | -58.2047 | 64.2288 | -58.2001 | -58.3191 | -57.9025 | -0.17853 | -0.041199 | -0.041199 | 0 | 0.26093 | 0.43945 | -0.18311 | -0.26093 | -0.004577 | -0.082397 | -0.064087 | -0.01831 | 0.19226 | 0.21057 | -0.059509 | 0.004578 | 0.23804 | -0.17853 | 0.30212 | -0.11444
942.1719 | 122.5525 | -58.1131 | 64.032 | -0.084763 | 0.00027338 | 0.067904 | 0.066996 | -2.7692e-05 | -0.00026063 | 0.00025641 | 0.016998 | 0.0014547 | 0.063852 | 0.064046 | -0.00039512 | -0.0049709 | -0.00083948 | -0.009155 | 64.0503 | -58.2687 | 64.0091 | -57.8339 | -58.342 | -58.2642 | 0.09613 | 64.0411 | -58.168 | 64.1647 | -57.9163 | -58.1497 | -57.9071 | -0.15564 | 0.018311 | -0.022888 | -0.041199 | 0.27924 | 0.43488 | -0.15106 | -0.22888 | 0.004578 | -0.073242 | -0.054932 | 0.009156 | 0.13275 | 0.1236 | 0.19684 | 0.25177 | 0.20599 | -0.036621 | 0.26093 | 0.018311
943.1719 | 122.5983 | -58.0627 | 64.1144 | -0.10002 | -0.00015382 | 0.066623 | 0.065074 | 0.00018591 | -0.00026063 | -0.0025208 | -0.0013123 | -0.00025411 | 0.065134 | 0.061056 | 0.00067288 | -0.0017667 | 0.0019375 | 0.018311 | 64.032 | -58.2504 | 64.0594 | -57.7698 | -58.2642 | -58.2275 | 0.13275 | 64.032 | -58.1543 | 64.1693 | -57.93 | -58.2733 | -57.8751 | -0.18768 | -0.082398 | -0.054932 | 0.027466 | 0.29297 | 0.48065 | -0.16479 | -0.20142 | 0.022888 | -0.013733 | -0.091553 | -0.082398 | 0.054932 | 0.13733 | 0.13275 | 0.2243 | 0.18768 | -0.21057 | 0.27924 | -0.11902
944.1719 | 122.5159 | -58.1223 | 64.0366 | -0.13817 | -0.0010082 | 0.0694 | 0.072764 | 0.0038175 | 0.00059397 | 0.0017518 | -0.03183 | 0.00038669 | 0.065134 | 0.06426 | -0.00039512 | -0.0028347 | -0.0038303 | -0.050354 | 63.9999 | -58.3054 | 64.0137 | -57.8567 | -58.3786 | -58.3191 | 0.082397 | 63.9725 | -58.1909 | 64.1144 | -57.9529 | -58.1726 | -57.9346 | -0.18311 | -0.036621 | -0.022888 | 0.013733 | 0.2655 | 0.44861 | -0.19684 | -0.25635 | -0.013733 | -0.073242 | -0.068665 | -0.064087 | 0.07782 | 0.14191 | 0.16937 | 0.23804 | 0.18768 | -0.050354 | 0.25635 | 0.018311
945.1719 | 122.4884 | -58.1497 | 63.9771 | -0.16106 | -0.00015382 | 0.0694 | 0.068705 | -0.00066869 | 0.0014484 | -0.0025208 | -0.092865 | 0.00038669 | 0.063425 | 0.071096 | -0.0016769 | -0.00048494 | -0.0036167 | -0.054932 | 64.0045 | -58.3237 | 64.0182 | -57.843 | -58.3832 | -58.3008 | 0 | 64.0091 | -58.1863 | 64.1556 | -57.9895 | -58.2047 | -57.962 | -0.17395 | 0.027466 | 0.041199 | 0.013733 | 0.3067 | 0.48065 | -0.15106 | -0.23346 | 0.022888 | -0.05951 | -0.036621 | 0.032043 | 0.17853 | 0.14648 | 0.16022 | 0.19684 | 0.18768 | -0.054932 | 0.2243 | -0.018311
946.1719 | 122.5021 | -58.1223 | 64.0274 | -0.1458 | -0.00036742 | 0.067264 | 0.069346 | 0.0021085 | 0.0012348 | 0.0030334 | -0.077606 | 0.00038669 | 0.065988 | 0.0696 | 0.00045928 | -0.0021939 | 0.0010831 | -0.064087 | 63.9999 | -58.2779 | 64.0045 | -57.8705 | -58.3969 | -58.3099 | -0.013733 | 64.0228 | -58.1909 | 64.1602 | -57.9758 | -58.2138 | -57.9071 | -0.15564 | -0.027466 | -0.022888 | 0.004578 | 0.25177 | 0.40741 | -0.18768 | -0.27466 | -0.032044 | -0.11902 | -0.068665 | -0.004578 | 0.13275 | 0.13733 | 0.14648 | 0.21515 | 0.21515 | -0.091553 | 0.28381 | -0.022888
947.1719 | 122.4564 | -58.1314 | 64.0594 | -0.092392 | -0.00036742 | 0.068759 | 0.066996 | 0.00082671 | -0.0013288 | -0.00017099 | -0.010467 | 0.00017309 | 0.065134 | 0.061696 | -0.00082252 | -0.0026211 | 0.0064237 | -0.091553 | 64.0411 | -58.2733 | 63.9862 | -57.8751 | -58.3649 | -58.3099 | -0.064087 | 64.0594 | -58.1497 | 64.1235 | -57.9895 | -58.2321 | -57.9575 | -0.14191 | -0.01831 | -0.073242 | -0.054932 | 0.25635 | 0.39825 | -0.17853 | -0.23346 | -0.036622 | -0.091553 | -0.01831 | 0 | 0.064087 | 0.064087 | 0.14191 | 0.16022 | 0.17395 | -0.10071 | 0.19226 | -0.082398
948.1719 | 122.3648 | -58.1085 | 64.1235 | -0.077133 | 0.00048698 | 0.064914 | 0.06956 | 0.00039951 | -0.00068783 | -0.0010254 | -0.065399 | 0.0016685 | 0.065134 | 0.065542 | -0.00060892 | -0.0026211 | -0.0010531 | 0 | 64.0869 | -58.255 | 63.9908 | -57.9163 | -58.3923 | -58.3282 | 0.009155 | 64.1098 | -58.1543 | 64.0823 | -58.0215 | -58.2596 | -57.9941 | -0.14648 | -0.036621 | -0.13275 | -0.09613 | 0.19226 | 0.33874 | -0.21973 | -0.28381 | -0.073242 | -0.13733 | -0.045776 | -0.013733 | -0.041199 | -0.027466 | 0.086976 | 0.13275 | 0.11444 | -0.15106 | 0.16022 | -0.10529
949.1719 | 122.406 | -58.1451 | 64.1144 | -0.092392 | -0.00058102 | 0.06705 | 0.06956 | -0.0017367 | 0.00080757 | 0.0017518 | -0.089813 | 0.00060029 | 0.064279 | 0.067678 | -0.0018905 | -0.0021939 | -0.00083948 | -0.09613 | 64.0869 | -58.3008 | 63.9175 | -57.9758 | -58.4839 | -58.4061 | -0.10529 | 64.0823 | -58.1909 | 64.0732 | -58.049 | -58.3008 | -57.9941 | -0.15564 | -0.027466 | -0.19684 | -0.16937 | 0.16937 | 0.32501 | -0.26093 | -0.33875 | -0.10529 | -0.18311 | -0.045776 | -0.032044 | -0.041199 | -0.009155 | 0.096131 | 0.14191 | 0.15106 | -0.15564 | 0.19684 | -0.10986
950.1719 | 122.4655 | -58.1085 | 64.0778 | -0.13817 | 0.00091438 | 0.067904 | 0.070842 | 0.00039951 | 0.0010212 | -0.00059819 | -0.056244 | 0.0020957 | 0.063852 | 0.064687 | 0.00067288 | -0.0028347 | -0.0025485 | -0.11902 | 64.0778 | -58.2687 | 63.9496 | -57.9025 | -58.4061 | -58.3511 | -0.073242 | 64.032 | -58.1863 | 64.1556 | -57.9666 | -58.2321 | -57.9254 | -0.16022 | 0 | -0.12817 | -0.12817 | 0.20599 | 0.36621 | -0.24261 | -0.29755 | -0.082397 | -0.13733 | -0.077819 | -0.045777 | 0.07782 | 0.1236 | 0.14191 | 0.21973 | 0.18311 | -0.1236 | 0.26093 | -0.045777
951.1719 | 122.5204 | -58.1314 | 63.9908 | -0.077133 | 0.002196 | 0.068759 | 0.070201 | 0.0033901 | 0.0014484 | -0.0016662 | -0.047089 | 0.00038669 | 0.066843 | 0.062551 | -0.00039512 | -0.0034757 | -0.00083948 | -0.059509 | 64.0457 | -58.2687 | 64.0274 | -57.8247 | -58.3328 | -58.2733 | -0.022888 | 64.0274 | -58.1543 | 64.1876 | -57.9391 | -58.1726 | -57.811 | -0.13733 | 0.054931 | 0.036621 | -0.01831 | 0.3067 | 0.44403 | -0.14191 | -0.20142 | -0.004577 | -0.064087 | -0.022888 | 0.036621 | 0.19684 | 0.16022 | 0.19226 | 0.21515 | 0.32043 | -0.041198 | 0.34332 | -0.01831
952.1719 | 122.5296 | -58.0673 | 64.0778 | -0.061874 | -0.0027172 | 0.068118 | 0.065715 | -0.0021641 | -0.00068783 | -0.0025208 | 0.029205 | -0.00068151 | 0.065134 | 0.060628 | -0.0023177 | -0.0045437 | -0.0021213 | -0.032043 | 64.0823 | -58.2458 | 64.0732 | -57.756 | -58.2825 | -58.2138 | 0.10071 | 64.0686 | -58.1268 | 64.2014 | -57.8613 | -58.0856 | -57.8522 | -0.17853 | 0.004577 | -0.004578 | -0.009155 | 0.31128 | 0.48981 | -0.14648 | -0.21515 | 0.032044 | -0.036621 | -0.059509 | -0.009155 | 0.1236 | 0.13275 | 0.20599 | 0.2655 | 0.21515 | -0.01831 | 0.27466 | 0.041199
953.1719 | 122.5342 | -58.0994 | 64.0732 | -0.15343 | 0.0015552 | 0.071536 | 0.070201 | 0.0010403 | 0.00059397 | 0.002179 | -0.059296 | 0.0033773 | 0.065988 | 0.070028 | -0.0025315 | -0.0039029 | -0.0016941 | -0.054932 | 64.064 | -58.2687 | 64.0274 | -57.7927 | -58.3282 | -58.2687 | -0.004578 | 64.0366 | -58.1406 | 64.183 | -57.9391 | -58.1772 | -57.8751 | -0.16937 | -0.009155 | -0.045776 | -0.036621 | 0.3067 | 0.47607 | -0.16937 | -0.22888 | 0 | -0.059509 | -0.041199 | -0.036621 | 0.10986 | 0.14648 | 0.16022 | 0.20142 | 0.2243 | -0.07782 | 0.2655 | -0.036621
954.1719 | 122.5021 | -58.1039 | 64.0778 | -0.1992 | -0.00079462 | 0.067904 | 0.069133 | 0.00061311 | 0.0020892 | 0.00089721 | -0.065399 | -0.00025411 | 0.064279 | 0.064046 | -0.00018152 | -0.0030483 | -0.0025485 | -0.064087 | 64.0366 | -58.255 | 64.0411 | -57.811 | -58.3466 | -58.2642 | -0.004578 | 64.0594 | -58.1406 | 64.1876 | -57.9208 | -58.1726 | -57.8705 | -0.15106 | -0.041199 | -0.036621 | 0.004578 | 0.29297 | 0.44403 | -0.16022 | -0.24262 | -0.009155 | -0.091553 | -0.036621 | -0.018311 | 0.10986 | 0.12817 | 0.18311 | 0.21973 | 0.23346 | -0.068664 | 0.27008 | -0.032043
955.1719 | 122.525 | -58.0765 | 64.0869 | -0.084763 | 0.0019824 | 0.067904 | 0.068492 | -0.0017367 | -0.0013288 | 0.00047001 | -0.037933 | 0.0012411 | 0.063638 | 0.074727 | -0.00039512 | -0.0019803 | -0.0014805 | -0.064087 | 64.064 | -58.2367 | 64.0411 | -57.8018 | -58.3328 | -58.2687 | -0.032043 | 64.0549 | -58.1451 | 64.2105 | -57.9117 | -58.1589 | -57.8659 | -0.16022 | -0.022888 | -0.045776 | -0.022888 | 0.27466 | 0.43487 | -0.19226 | -0.25635 | -0.032044 | -0.096131 | -0.068665 | -0.032043 | 0.1236 | 0.15564 | 0.16479 | 0.23346 | 0.21057 | -0.082398 | 0.27924 | -0.013733
956.1719 | 122.4106 | -58.1223 | 64.1006 | -0.13054 | 0.001128 | 0.068118 | 0.067637 | 0.00082671 | 0.00016657 | 0.001538 | -0.086761 | 0.0010275 | 0.065561 | 0.070028 | -0.00039512 | -0.0036893 | 0.0021511 | -0.10071 | 64.1464 | -58.255 | 63.9542 | -57.8888 | -58.4198 | -58.3466 | -0.09613 | 64.1006 | -58.168 | 64.151 | -57.9941 | -58.2367 | -57.9391 | -0.13275 | 0.045776 | -0.14648 | -0.19226 | 0.23346 | 0.36621 | -0.22431 | -0.29755 | -0.091553 | -0.1648 | -0.045777 | 0 | 0.050354 | 0.050354 | 0.12817 | 0.17395 | 0.1831 | -0.11444 | 0.22888 | -0.068664
957.1719 | 122.4426 | -58.0719 | 64.1052 | -0.12291 | 0.00070078 | 0.066409 | 0.065928 | -0.0028049 | 0.0027302 | -0.0025208 | 0.013947 | -0.00025411 | 0.065775 | 0.063619 | 0.00045928 | -0.0039029 | -0.0016941 | -0.086975 | 64.1098 | -58.2504 | 64.0137 | -57.811 | -58.3282 | -58.2413 | -0.050354 | 64.1006 | -58.1131 | 64.1464 | -57.93 | -58.1726 | -57.9025 | -0.17853 | 0.004577 | -0.091553 | -0.09613 | 0.26093 | 0.43945 | -0.16937 | -0.25635 | 0.009155 | -0.07782 | -0.041199 | -0.004578 | 0.041198 | 0.045776 | 0.14191 | 0.1831 | 0.16937 | -0.10071 | 0.21057 | -0.059509
958.1719 | 122.4792 | -58.0673 | 64.0961 | -0.13054 | 0.00027338 | 0.068759 | 0.068919 | 0.0014675 | 0.0037982 | -0.00059819 | -0.013519 | -0.00025411 | 0.063425 | 0.071737 | -0.00018152 | -0.0024075 | -0.00062588 | -0.059509 | 64.0594 | -58.2413 | 64.1785 | -57.637 | -58.3511 | -58.2413 | 0.004578 | 64.064 | -58.1177 | 64.2197 | -57.9025 | -58.1451 | -57.8751 | -0.17395 | -0.036621 | 0.082398 | 0.11902 | 0.4303 | 0.60425 | -0.17395 | -0.28381 | 0 | -0.10986 | -0.050354 | -0.032043 | 0.1236 | 0.15564 | 0.16479 | 0.21515 | 0.19226 | -0.07782 | 0.24262 | -0.027466
959.1719 | 122.5204 | -58.049 | 64.1006 | -0.16106 | 0.00070078 | 0.068545 | 0.068278 | 0.0050991 | 0.0023028 | 0.0036744 | -0.037933 | 0.00017309 | 0.06492 | 0.065328 | -0.0023177 | -0.0024075 | -0.0019077 | -0.054932 | 64.0778 | -58.223 | 64.0503 | -57.5912 | -58.1726 | -58.2184 | 0.009155 | 64.0823 | -58.1039 | 64.2059 | -57.8888 | -58.1543 | -57.8659 | -0.17395 | -0.022888 | -0.050354 | -0.027466 | 0.45776 | 0.63171 | -0.16937 | -0.1236 | 0.004577 | 0.050354 | -0.054932 | -0.018311 | 0.10529 | 0.1236 | 0.16022 | 0.21515 | 0.1831 | -0.10529 | 0.23804 | -0.050354
960.1719 | 122.5159 | -58.0536 | 64.1052 | -0.077133 | 0.00070078 | 0.068332 | 0.068492 | -0.0019503 | 0.00016657 | -0.00081179 | -0.062347 | 0.0023093 | 0.064493 | 0.06896 | -0.00018152 | -0.0024075 | -0.00041228 | -0.045776 | 64.064 | -58.2092 | 63.9038 | -57.8888 | -58.0399 | -58.0902 | 0.032043 | 64.0823 | -58.0994 | 64.1968 | -57.8888 | -58.1085 | -57.8659 | -0.15564 | -0.041199 | -0.20142 | -0.16022 | 0.1648 | 0.32044 | -0.036621 | 0.013733 | 0.11902 | 0.16937 | -0.045776 | -0.022889 | 0.091552 | 0.11444 | 0.1648 | 0.21057 | 0.18768 | -0.054932 | 0.23346 | -0.009156
961.1719 | 122.5067 | -58.0582 | 64.1052 | -0.10765 | 0.002196 | 0.070041 | 0.068919 | 0.0010403 | 0.0025166 | -0.00059819 | 0.044464 | -0.00046791 | 0.065775 | 0.066823 | -0.00060892 | -0.0026211 | 0.0021511 | -0.054932 | 64.0915 | -58.2047 | 64.0823 | -57.7332 | -58.5709 | -58.3511 | 0.022888 | 64.0961 | -58.0811 | 64.2471 | -57.8979 | -58.1268 | -57.8567 | -0.14648 | -0.013733 | -0.022889 | -0.009156 | 0.32501 | 0.4715 | -0.29297 | -0.5127 | -0.14648 | -0.36621 | -0.022888 | -0.009156 | 0.14191 | 0.15106 | 0.16022 | 0.18311 | 0.20142 | -0.068664 | 0.22431 | -0.045776
962.1719 | 122.4976 | -58.0124 | 64.1418 | -0.061874 | 0.002837 | 0.066195 | 0.067851 | 0.00082671 | 0.0031574 | -0.00081179 | 0.035309 | 0.0012411 | 0.066202 | 0.065542 | 0.0011001 | -0.0047573 | 0.00086952 | 0.004578 | 64.1418 | -58.1543 | 64.1281 | -57.6965 | -58.1909 | -58.1451 | 0.15564 | 64.119 | -58.0582 | 64.2517 | -57.811 | -58.0399 | -57.811 | -0.14191 | 0 | -0.013733 | -0.013733 | 0.31586 | 0.45776 | -0.13275 | -0.17853 | 0.009155 | -0.036621 | -0.045777 | -0.022888 | 0.10986 | 0.13275 | 0.20142 | 0.24719 | 0.20142 | -0.027466 | 0.24719 | 0.018311
963.1719 | 122.493 | -58.0536 | 64.1006 | -0.11528 | -0.00015382 | 0.069827 | 0.06956 | 0.0018949 | 0.003371 | 0.0019654 | -0.025726 | 0.00060029 | 0.064066 | 0.062124 | -0.00082252 | -0.0021939 | -0.00062588 | -0.041199 | 64.0915 | -58.2001 | 64.0961 | -57.756 | -58.2916 | -58.2184 | -0.009155 | 64.0778 | -58.1085 | 64.2563 | -57.9117 | -58.0902 | -57.8384 | -0.14648 | -0.009155 | -0.004578 | 0.004577 | 0.29755 | 0.44403 | -0.16479 | -0.23804 | -0.018311 | -0.091553 | -0.054932 | -0.022888 | 0.15564 | 0.17853 | 0.14191 | 0.19684 | 0.21515 | -0.036621 | 0.27008 | 0.018311
964.1719 | 122.5159 | -58.0399 | 64.1281 | -0.13054 | 5.9776e-05 | 0.072177 | 0.06956 | -0.0025913 | 0.00080757 | 0.0017518 | -0.044037 | 0.0010275 | 0.064706 | 0.069387 | -0.00082252 | -0.0011257 | -0.0019077 | -0.050354 | 64.0869 | -58.223 | 64.1006 | -57.7423 | -58.2825 | -58.2184 | 0.004578 | 64.0961 | -58.1131 | 64.1876 | -57.7057 | -58.0078 | -57.7972 | -0.1831 | -0.041199 | -0.027466 | 0.013733 | 0.29755 | 0.48065 | -0.17853 | -0.24262 | 0.004577 | -0.05951 | -0.073242 | -0.032044 | 0.059509 | 0.091553 | 0.33417 | 0.40741 | 0.24262 | 0.032044 | 0.31586 | 0.10529
965.1719 | 122.4655 | -58.0444 | 64.1373 | -0.16869 | -0.001222 | 0.0694 | 0.06721 | -0.0038729 | 0.0020892 | -0.00081179 | -0.03183 | 0.00081389 | 0.066202 | 0.067464 | -0.0025315 | -0.0034757 | 0.0051419 | -0.041199 | 64.119 | -58.1909 | 64.0869 | -57.7332 | -58.2962 | -58.1863 | 0 | 64.1235 | -58.0948 | 64.2242 | -57.8796 | -58.1268 | -57.8613 | -0.14648 | -0.01831 | -0.050354 | -0.032044 | 0.31128 | 0.45776 | -0.14191 | -0.25177 | 0.004578 | -0.10529 | -0.050354 | -0.013733 | 0.086975 | 0.10071 | 0.1648 | 0.21515 | 0.18311 | -0.082397 | 0.23346 | -0.032043
966.1719 | 122.5067 | -58.0307 | 64.1235 | -0.16869 | 0.00070078 | 0.0694 | 0.069346 | 0.0014675 | 0.0018756 | -0.001239 | -0.034882 | -0.0015359 | 0.065561 | 0.069387 | -0.0012497 | -0.0039029 | 0.00022852 | -0.027466 | 64.119 | -58.1726 | 64.1052 | -57.7194 | -58.2733 | -58.1955 | 0.009155 | 64.1144 | -58.0765 | 64.2471 | -57.8751 | -58.1039 | -57.8293 | -0.14191 | -0.004577 | -0.01831 | -0.013733 | 0.31128 | 0.45319 | -0.16479 | -0.24261 | -0.022889 | -0.10071 | -0.045776 | -0.009155 | 0.1236 | 0.13275 | 0.15564 | 0.20142 | 0.20142 | -0.073242 | 0.24719 | -0.027466
967.1719 | 122.493 | -58.0261 | 64.151 | -0.10765 | -0.00058102 | 0.068332 | 0.06721 | 0.00039951 | 0.001662 | 0.0028198 | -0.010467 | -0.0013223 | 0.063638 | 0.05956 | -0.0014633 | -0.0017667 | 0.00044212 | -0.073242 | 64.1098 | -58.1818 | 64.0823 | -57.7469 | -58.2916 | -58.2184 | -0.036621 | 64.119 | -58.0948 | 64.2105 | -57.9163 | -58.136 | -57.8705 | -0.15564 | -0.041199 | -0.068665 | -0.027466 | 0.27924 | 0.43488 | -0.19226 | -0.2655 | -0.036621 | -0.10986 | -0.068665 | -0.032043 | 0.059509 | 0.091552 | 0.10986 | 0.17853 | 0.15564 | -0.10986 | 0.22431 | -0.041198
968.1719 | 122.5159 | -58.017 | 64.1235 | -0.1458 | -0.0014356 | 0.071322 | 0.068278 | 0.0012539 | 0.0010212 | 0.0019654 | -0.028778 | 0.00017309 | 0.065775 | 0.057851 | -0.00018152 | -0.0034757 | -0.0025485 | -0.12817 | 64.0961 | -58.2275 | 64.1235 | -57.6828 | -58.2642 | -58.2001 | -0.004578 | 64.1098 | -58.0673 | 64.2517 | -57.8476 | -58.0948 | -57.8064 | -0.21057 | -0.027466 | 0 | 0.027466 | 0.33417 | 0.54474 | -0.18311 | -0.24719 | 0.027466 | -0.036621 | -0.050354 | -0.013733 | 0.12817 | 0.14191 | 0.16937 | 0.21973 | 0.21057 | -0.07782 | 0.26093 | -0.027466
969.1719 | 122.5067 | -57.9758 | 64.151 | -0.13817 | 0.00091438 | 0.068545 | 0.070842 | 0.00082671 | -0.0013288 | 0.00025641 | -0.025726 | 0.00038669 | 0.064706 | 0.067464 | 0.00045928 | -0.0039029 | 0.00086952 | -0.041199 | 64.1144 | -58.168 | 64.2471 | -57.7148 | -58.3054 | -58.1451 | 0.013733 | 64.1098 | -58.0719 | 64.2563 | -57.843 | -58.0765 | -57.8018 | -0.19226 | -0.036621 | 0.09613 | 0.13275 | 0.26093 | 0.45319 | -0.16937 | -0.32959 | 0.022888 | -0.13733 | -0.09613 | -0.041199 | 0.10529 | 0.14648 | 0.13275 | 0.22888 | 0.17395 | -0.10071 | 0.27008 | -0.004578
970.1719 | 122.5159 | -57.9895 | 64.1647 | -0.10765 | 0.0030506 | 0.068332 | 0.06721 | -0.0034457 | 0.0010212 | 0.0028198 | -0.037933 | 0.00081389 | 0.064279 | 0.067037 | 3.2078e-05 | -0.0030483 | 0.0019375 | -0.064087 | 64.1144 | -58.1589 | 64.064 | -57.9208 | -58.1543 | -58.0215 | 0.009155 | 64.1327 | -58.049 | 64.2517 | -57.8247 | -58.0765 | -57.811 | -0.16937 | -0.050354 | -0.10071 | -0.050354 | 0.068665 | 0.23804 | -0.032043 | -0.16479 | 0.13733 | 0.004578 | -0.059509 | -0.032044 | 0.086975 | 0.11902 | 0.1648 | 0.2243 | 0.17853 | -0.086975 | 0.23804 | -0.027466
971.1719 | 122.4792 | -58.0078 | 64.151 | -0.16869 | 5.9776e-05 | 0.068973 | 0.070628 | 0.0014675 | 0.00080757 | 0.00047001 | -0.028778 | -4.0508e-05 | 0.064279 | 0.062551 | -0.0023177 | -0.0036893 | 0.0049283 | -0.064087 | 64.1693 | -58.1726 | 64.1281 | -57.7011 | -58.255 | -58.1818 | -0.004578 | 64.1006 | -58.0582 | 64.238 | -57.8567 | -58.0994 | -57.8476 | -0.16479 | 0.018311 | -0.022888 | -0.041199 | 0.3067 | 0.4715 | -0.17395 | -0.24719 | -0.009156 | -0.082398 | -0.050355 | -0.050354 | 0.086975 | 0.13733 | 0.15106 | 0.20142 | 0.16022 | -0.091553 | 0.21057 | -0.041198
972.1719 | 122.525 | -57.9666 | 64.2014 | -0.10765 | -0.0014356 | 0.068545 | 0.067424 | -0.0015231 | 0.0012348 | 0.00089721 | 0.050568 | 0.00081389 | 0.064066 | 0.066823 | -0.00082252 | -0.0041165 | -0.0019077 | 0 | 64.1647 | -58.1177 | 64.1739 | -57.6416 | -58.1635 | -58.1085 | 0.16479 | 64.1785 | -58.0261 | 64.3158 | -57.756 | -58.0032 | -57.7606 | -0.15106 | -0.036621 | -0.027466 | 0.009155 | 0.32501 | 0.47607 | -0.14191 | -0.19684 | 0.009155 | -0.045776 | -0.059509 | -0.022888 | 0.11444 | 0.13733 | 0.21057 | 0.27008 | 0.20599 | -0.036621 | 0.2655 | 0.022888
973.1719 | 122.493 | -58.0124 | 64.1739 | -0.13054 | 0.0013416 | 0.067904 | 0.068278 | 0.0012539 | 0.003371 | -0.00038459 | -0.062347 | 0.00060029 | 0.062143 | 0.062124 | 0.00088648 | -0.0024075 | -0.0010531 | -0.068665 | 64.1464 | -58.1497 | 64.1281 | -57.692 | -58.2413 | -58.1726 | 0 | 64.1464 | -58.0627 | 64.2792 | -57.8476 | -58.0719 | -57.8156 | -0.13733 | -0.027466 | -0.045776 | -0.01831 | 0.32043 | 0.45776 | -0.16022 | -0.22888 | -0.022888 | -0.091553 | -0.050354 | -0.027466 | 0.10529 | 0.13275 | 0.1648 | 0.21515 | 0.19684 | -0.059509 | 0.24719 | -0.009155
974.1719 | 122.4838 | -57.9849 | 64.1739 | -0.077133 | 0.001128 | 0.069186 | 0.068492 | -0.00045509 | 0.0023028 | 4.2808e-05 | -0.040985 | -0.0013223 | 0.06492 | 0.05721 | -0.0014633 | -0.0019803 | -0.0038303 | -0.068665 | 64.1373 | -58.1177 | 64.1693 | -57.6828 | -58.2367 | -58.1589 | 0.009155 | 64.151 | -58.0444 | 64.2929 | -57.8247 | -58.0627 | -57.7744 | -0.13275 | -0.036621 | -0.004577 | 0.032044 | 0.30212 | 0.43488 | -0.17395 | -0.25177 | -0.041199 | -0.11902 | -0.05951 | -0.022888 | 0.11902 | 0.14191 | 0.16022 | 0.21973 | 0.21057 | -0.07782 | 0.27008 | -0.01831
975.1719 | 122.5067 | -57.9666 | 64.1922 | -0.16869 | 0.0026234 | 0.06705 | 0.067851 | 0.00061311 | -0.00026063 | 0.00047001 | -0.028778 | -0.00025411 | 0.06257 | 0.069387 | -0.0023177 | -0.0021939 | -0.00019868 | 0.091553 | 64.1556 | -58.1268 | 64.1647 | -57.6599 | -58.2321 | -58.1589 | 0.022888 | 64.1602 | -58.049 | 64.2792 | -57.7927 | -58.0627 | -57.8018 | -0.16022 | -0.036621 | -0.027466 | 0.009155 | 0.3067 | 0.46692 | -0.19226 | -0.2655 | -0.032044 | -0.10529 | -0.082397 | -0.032044 | 0.086975 | 0.11902 | 0.17395 | 0.25635 | 0.16479 | -0.09613 | 0.24719 | -0.013733
976.1719 | 122.5113 | -57.9666 | 64.1785 | -0.13054 | -0.00079462 | 0.0694 | 0.071055 | 0.0025357 | -0.0013288 | -0.00081179 | -0.047089 | -4.0508e-05 | 0.063638 | 0.065542 | 0.0011001 | -0.0015529 | 0.00065592 | -0.054932 | 64.151 | -58.136 | 64.2151 | -57.7194 | -58.2001 | -58.1085 | 0.032043 | 64.1602 | -58.017 | 64.325 | -57.7927 | -58.0399 | -57.756 | -0.16937 | -0.027466 | 0.036621 | 0.064087 | 0.24719 | 0.41657 | -0.14191 | -0.23346 | 0.027465 | -0.064087 | -0.050354 | -0.018311 | 0.14648 | 0.16479 | 0.17395 | 0.2243 | 0.21057 | -0.073242 | 0.26093 | -0.022888
977.1719 | 122.4976 | -57.962 | 64.2059 | -0.069504 | 0.0024096 | 0.067904 | 0.066783 | -0.00024129 | 0.0023028 | 0.00025641 | 0.026154 | -4.0508e-05 | 0.065134 | 0.062978 | -0.0016769 | -0.0026211 | -0.0042575 | 0.032043 | 64.2151 | -58.0948 | 64.2151 | -57.5821 | -58.0994 | -58.0582 | 0.1236 | 64.2105 | -57.9895 | 64.3295 | -57.7194 | -57.962 | -57.7332 | -0.13275 | 0.009155 | 0.009155 | 0 | 0.37994 | 0.5127 | -0.096131 | -0.13733 | 0.036621 | -0.004577 | -0.027466 | 0.004577 | 0.1236 | 0.11902 | 0.24262 | 0.27008 | 0.22888 | 0 | 0.25635 | 0.027466
978.1719 | 122.4884 | -57.9391 | 64.2059 | -0.10765 | -0.0014356 | 0.069613 | 0.068492 | 0.0025357 | 0.0035846 | 0.00047001 | -0.092865 | -0.00025411 | 0.06257 | 0.06661 | -0.0038131 | -0.0026211 | 0.0017239 | -0.045776 | 64.1785 | -58.1039 | 64.183 | -57.6553 | -58.2092 | -58.136 | 0.068665 | 64.1739 | -58.0627 | 64.3066 | -57.756 | -58.0215 | -57.7698 | -0.16479 | -0.027466 | -0.022889 | 0.004577 | 0.28381 | 0.44861 | -0.19684 | -0.27008 | -0.032043 | -0.10529 | -0.1236 | -0.032044 | 0.10071 | 0.13275 | 0.18311 | 0.3067 | 0.16937 | -0.082397 | 0.29297 | 0.041199
979.1719 | 122.4884 | -57.9437 | 64.1968 | -0.10002 | 0.00027338 | 0.070041 | 0.068064 | -0.0017367 | 0.0029438 | 0.00068361 | -0.05014 | 0.00038669 | 0.062784 | 0.062124 | -0.00039512 | -0.0026211 | 1.4916e-05 | -0.059509 | 64.1968 | -58.1131 | 64.2014 | -57.6508 | -58.2138 | -58.136 | 0.013733 | 64.1693 | -58.0078 | 64.325 | -57.7835 | -58.0353 | -57.756 | -0.16937 | 0 | 0.004578 | 0.004578 | 0.29297 | 0.46234 | -0.19226 | -0.27008 | -0.022888 | -0.10071 | -0.064086 | -0.027465 | 0.12817 | 0.15564 | 0.16022 | 0.2243 | 0.18768 | -0.091552 | 0.25177 | -0.027466
980.1719 | 122.4747 | -57.9437 | 64.2059 | -0.17632 | 0.0013416 | 0.068759 | 0.068705 | 0.0033901 | -4.7032e-05 | 0.0019654 | -0.047089 | 0.00017309 | 0.065775 | 0.062551 | -0.0021041 | -0.0011257 | -0.0029757 | 0.086975 | 64.1693 | -58.1039 | 64.1922 | -57.6324 | -58.1818 | -58.1177 | 0.004578 | 64.1876 | -58.0353 | 64.3204 | -57.7744 | -58.0261 | -57.7469 | -0.16022 | -0.036621 | -0.013733 | 0.022888 | 0.31128 | 0.4715 | -0.17395 | -0.23804 | -0.013733 | -0.07782 | -0.091552 | -0.018311 | 0.11444 | 0.13275 | 0.16937 | 0.26093 | 0.19684 | -0.082397 | 0.28839 | 0.009155
981.1719 | 122.5296 | -57.962 | 64.1968 | -0.16106 | 0.001128 | 0.069613 | 0.067851 | -0.00024129 | -4.7032e-05 | 0.00068361 | -0.034882 | -0.00025411 | 0.065134 | 0.067891 | -0.0025315 | -0.0011257 | -0.0025485 | -0.054932 | 64.2014 | -58.0948 | 64.3066 | -57.5684 | -58.255 | -58.0902 | 0.009155 | 64.2059 | -57.9666 | 64.3387 | -57.7606 | -58.017 | -57.7286 | -0.13275 | 0.004578 | 0.10986 | 0.10529 | 0.39368 | 0.52643 | -0.12817 | -0.29297 | 0.004578 | -0.16022 | -0.004578 | 0.009156 | 0.14191 | 0.13275 | 0.20142 | 0.20599 | 0.23346 | -0.054932 | 0.23804 | -0.050354
982.1719 | 122.5296 | -57.9254 | 64.2471 | -0.084763 | 0.00070078 | 0.067477 | 0.06721 | -0.0051547 | 0.0018756 | 0.00089721 | 0.0047913 | 0.0010275 | 0.065134 | 0.064046 | 3.2078e-05 | -0.0036893 | 0.0012967 | -0.032043 | 64.2563 | -58.0811 | 64.183 | -57.6233 | -58.1635 | -58.0948 | 0.045776 | 64.2197 | -57.9803 | 64.3433 | -57.7103 | -57.9575 | -57.724 | -0.15564 | 0.009156 | -0.064087 | -0.073243 | 0.30212 | 0.45776 | -0.16937 | -0.23804 | -0.013733 | -0.082397 | -0.054932 | -0.027465 | 0.096131 | 0.1236 | 0.21515 | 0.27008 | 0.20142 | -0.032043 | 0.25635 | 0.022889
983.1719 | 122.5662 | -57.9208 | 64.2242 | -0.13817 | 0.00070078 | 0.068545 | 0.068492 | -2.7692e-05 | 0.0014484 | 0.001538 | -0.074554 | -4.0508e-05 | 0.066202 | 0.069173 | -0.0021041 | -0.0041165 | -0.0031895 | -0.050354 | 64.2014 | -58.1039 | 64.2105 | -57.637 | -58.1772 | -58.1223 | -0.009155 | 64.2014 | -57.9941 | 64.4531 | -57.7835 | -58.017 | -57.7606 | -0.18311 | -0.022888 | -0.013733 | 0.009155 | 0.28381 | 0.46692 | -0.20142 | -0.25635 | -0.01831 | -0.073242 | -0.073243 | -0.022888 | 0.22888 | 0.25177 | 0.13733 | 0.21057 | 0.16022 | -0.096131 | 0.23346 | -0.022888
984.1719 | 122.5296 | -57.9437 | 64.151 | -0.10765 | 0.00070078 | 0.070468 | 0.068064 | 0.0014675 | 0.00080757 | 0.00047001 | -0.062347 | 0.00081389 | 0.064066 | 0.068532 | -0.0018905 | -0.0039029 | 0.0030057 | -0.059509 | 64.2151 | -58.1131 | 64.2014 | -57.6553 | -58.1863 | -58.1177 | -0.009155 | 64.2197 | -57.9895 | 64.3387 | -57.7469 | -57.9895 | -57.7103 | -0.16937 | 0.064087 | 0.050354 | -0.013733 | 0.28839 | 0.45776 | -0.17395 | -0.24261 | -0.004578 | -0.073242 | -0.045776 | 0.068665 | 0.18768 | 0.11902 | 0.19684 | 0.24262 | 0.23346 | -0.045776 | 0.27924 | 0
985.1719 | 122.5525 | -57.9346 | 64.2105 | -0.084763 | 0.0013416 | 0.068118 | 0.069987 | -2.7692e-05 | 0.001662 | -0.00059819 | -0.080658 | 0.00017309 | 0.063211 | 0.068746 | 0.00088648 | -0.0019803 | -0.00083948 | -0.050354 | 64.2426 | -58.0948 | 64.2014 | -57.6416 | -58.1772 | -58.1039 | -0.013733 | 64.2105 | -58.0078 | 64.3616 | -57.7515 | -58.0078 | -57.7011 | -0.16022 | 0.032044 | -0.009155 | -0.041199 | 0.29297 | 0.45319 | -0.16937 | -0.24262 | -0.009155 | -0.082397 | -0.073242 | 0 | 0.15106 | 0.15106 | 0.1831 | 0.25635 | 0.23346 | -0.073242 | 0.3067 | 0
986.1719 | 122.5067 | -57.9163 | 64.238 | -0.15343 | 0.00070078 | 0.068759 | 0.068278 | -0.0023777 | 0.0010212 | -0.00081179 | -0.062347 | -0.0019631 | 0.065347 | 0.06191 | 0.00067288 | -0.0030483 | 0.0032193 | -0.068665 | 64.2242 | -58.0994 | 64.0411 | -57.7377 | -57.9208 | -57.93 | 0.022888 | 64.2242 | -57.9666 | 64.3524 | -57.7423 | -57.9758 | -57.7057 | -0.1831 | -0.013733 | -0.19684 | -0.1831 | 0.17853 | 0.36163 | -0.013733 | -0.004577 | 0.16937 | 0.17853 | -0.050354 | -0.013733 | 0.11444 | 0.12817 | 0.17395 | 0.2243 | 0.21057 | -0.059509 | 0.26093 | -0.009155
987.1719 | 122.5525 | -57.8842 | 64.2471 | -0.084763 | 5.9776e-05 | 0.068545 | 0.067851 | -0.00066869 | 0.0010212 | 0.00089721 | 0.016998 | 0.00017309 | 0.06492 | 0.063833 | -0.00082252 | -0.0058255 | 0.0023647 | -0.004578 | 64.2609 | -58.0307 | 64.2609 | -57.5363 | -58.0765 | -58.0124 | 0.14191 | 64.238 | -57.9437 | 64.4028 | -57.6599 | -57.9117 | -57.6645 | -0.14648 | 0.013733 | 0.013733 | 0 | 0.3479 | 0.49439 | -0.12817 | -0.19226 | 0.018311 | -0.045776 | -0.05951 | -0.009155 | 0.15564 | 0.16479 | 0.2243 | 0.28381 | 0.21973 | -0.027466 | 0.27924 | 0.032044
988.1719 | 122.5983 | -57.9163 | 64.2609 | -0.12291 | 0.00048698 | 0.068332 | 0.068705 | 0.0029629 | 0.00080757 | 0.001538 | -0.062347 | -0.00025411 | 0.064493 | 0.070455 | -0.0023177 | -0.0015529 | 0.0015103 | -0.054932 | 64.2334 | -58.0627 | 64.2426 | -57.6233 | -58.1543 | -58.0719 | -0.004578 | 64.1968 | -57.9437 | 64.4073 | -57.8888 | -58.1451 | -57.6828 | -0.14648 | -0.027466 | -0.01831 | 0.009156 | 0.29297 | 0.43945 | -0.15564 | -0.23804 | -0.009155 | -0.091553 | -0.027466 | -0.064087 | 0.14648 | 0.21057 | 0.027466 | 0.054932 | 0.23346 | -0.22888 | 0.26093 | -0.20142
989.1719 | 122.5159 | -57.9117 | 64.2426 | -0.11528 | 5.9776e-05 | 0.070468 | 0.068705 | 0.00018591 | 0.0010212 | -0.00081179 | -0.065399 | -0.00089511 | 0.066202 | 0.060415 | -0.0029587 | -0.0015529 | -0.0023349 | -0.073242 | 64.2471 | -58.0353 | 64.2151 | -57.6096 | -58.1635 | -58.0765 | -0.018311 | 64.2242 | -57.9483 | 64.3707 | -57.7286 | -57.9758 | -57.6828 | -0.1236 | 0.004577 | -0.027466 | -0.032043 | 0.30212 | 0.42572 | -0.16479 | -0.25177 | -0.041199 | -0.12817 | -0.036621 | -0.018311 | 0.12817 | 0.14648 | 0.1831 | 0.21973 | 0.22888 | -0.064087 | 0.2655 | -0.027466
990.1719 | 122.5296 | -57.9025 | 64.2654 | -0.084763 | 5.9776e-05 | 0.069186 | 0.068705 | -0.0010959 | 0.0037982 | -0.0016662 | -0.065399 | 0.0016685 | 0.064493 | 0.068532 | 3.2078e-05 | -0.0019803 | 0.00044212 | -0.073242 | 64.2792 | -58.0673 | 64.2059 | -57.6462 | -58.1543 | -58.0673 | -0.032043 | 64.2471 | -57.962 | 64.389 | -57.7194 | -57.9758 | -57.692 | -0.16479 | 0.013733 | -0.059509 | -0.073242 | 0.25635 | 0.42114 | -0.16479 | -0.25177 | 0 | -0.086975 | -0.059509 | -0.018311 | 0.1236 | 0.14191 | 0.18311 | 0.24262 | 0.21057 | -0.073242 | 0.27008 | -0.013733
991.1719 | 122.493 | -57.9025 | 64.2883 | -0.10765 | -0.0018628 | 0.068332 | 0.069987 | 0.0036039 | 0.00059397 | 0.001538 | -0.080658 | -0.0011087 | 0.064279 | 0.061269 | -0.0012497 | -0.0011257 | -0.0053257 | -0.086975 | 64.2746 | -58.0902 | 64.1922 | -57.6279 | -58.1726 | -58.0994 | -0.032043 | 64.2563 | -57.9483 | 64.3707 | -57.7103 | -57.9803 | -57.6691 | -0.18768 | -0.013733 | -0.09613 | -0.082397 | 0.27466 | 0.46234 | -0.19684 | -0.27008 | -0.009155 | -0.082397 | -0.045776 | -0.032043 | 0.082398 | 0.11444 | 0.19226 | 0.23804 | 0.23346 | -0.07782 | 0.27924 | -0.032044
992.1719 | 122.493 | -57.8567 | 64.3021 | -0.12291 | -0.00058102 | 0.068973 | 0.067851 | -0.0045139 | -0.00047423 | -0.001239 | 0.026154 | 0.00060029 | 0.066202 | 0.061483 | 0.00067288 | -0.0041165 | 0.0055691 | -0.050354 | 64.2975 | -57.9987 | 64.2288 | -57.5867 | -58.1131 | -58.0536 | 0.027466 | 64.2838 | -57.93 | 64.389 | -57.6645 | -57.9163 | -57.6645 | -0.14191 | -0.004578 | -0.073242 | -0.068664 | 0.27008 | 0.41199 | -0.19684 | -0.25635 | -0.054932 | -0.11444 | -0.073243 | -0.018311 | 0.086975 | 0.10529 | 0.19226 | 0.2655 | 0.19226 | -0.05951 | 0.2655 | 0.013733
993.1719 | 122.493 | -57.8705 | 64.2746 | -0.054245 | 0.00091438 | 0.0694 | 0.068919 | -2.7692e-05 | 0.0012348 | -0.0031616 | -0.053192 | 0.0012411 | 0.063638 | 0.065755 | -0.0027451 | -0.0036893 | -0.00083948 | -0.041199 | 64.27 | -58.0307 | 64.2563 | -57.5592 | -58.1223 | -58.0627 | -0.004578 | 64.27 | -57.9391 | 64.3799 | -57.6874 | -57.9437 | -57.6782 | -0.16022 | -0.004577 | -0.01831 | -0.013733 | 0.31128 | 0.4715 | -0.19226 | -0.25177 | -0.032043 | -0.091552 | -0.068665 | -0.004577 | 0.10529 | 0.10986 | 0.1831 | 0.25177 | 0.19226 | -0.073243 | 0.26093 | -0.004578
994.1719 | 122.5388 | -57.8567 | 64.27 | -0.092392 | 0.00070078 | 0.069186 | 0.069773 | -0.00066869 | 0.0012348 | 0.0011108 | -0.074554 | 0.0012411 | 0.06492 | 0.06191 | -0.00060892 | -0.0019803 | -0.00041228 | -0.082397 | 64.2838 | -58.0399 | 64.238 | -57.5821 | -58.1085 | -58.0673 | -0.018311 | 64.2792 | -57.93 | 64.4073 | -57.692 | -57.9391 | -57.637 | -0.18311 | 0.013732 | -0.032044 | -0.045776 | 0.27466 | 0.45776 | -0.21057 | -0.25177 | -0.027466 | -0.068665 | -0.073243 | 0.009155 | 0.13733 | 0.12817 | 0.16479 | 0.23804 | 0.21973 | -0.082398 | 0.29297 | -0.009155
995.1719 | 122.4701 | -57.8705 | 64.2517 | -0.084763 | -0.00079462 | 0.0694 | 0.068492 | 0.00061311 | 0.0014484 | 0.0017518 | -0.065399 | 0.00060029 | 0.065561 | 0.069814 | 0.00024568 | -0.0017667 | 0.0015103 | -0.11444 | 64.3021 | -58.0444 | 64.2471 | -57.6233 | -58.1543 | -58.0627 | -0.041199 | 64.2746 | -57.9208 | 64.3982 | -57.7011 | -57.9483 | -57.6508 | -0.17395 | 0.050354 | -0.004578 | -0.054932 | 0.24719 | 0.42114 | -0.19226 | -0.28381 | -0.01831 | -0.10986 | -0.050354 | 0.022888 | 0.14648 | 0.1236 | 0.16937 | 0.21973 | 0.21973 | -0.07782 | 0.27008 | -0.027466
996.1719 | 122.4792 | -57.8842 | 64.2883 | -0.10002 | 0.00091438 | 0.068332 | 0.06956 | -0.00066869 | 0.001662 | -0.0018798 | -0.068451 | 0.0020957 | 0.065134 | 0.064473 | 3.2078e-05 | -0.00069854 | -0.0025485 | -0.10071 | 64.3204 | -58.0444 | 64.1922 | -57.6324 | -58.1909 | -58.0948 | -0.07782 | 64.3021 | -57.93 | 64.3936 | -57.692 | -57.9712 | -57.692 | -0.16022 | 0.032044 | -0.09613 | -0.12817 | 0.25177 | 0.41199 | -0.21057 | -0.3067 | -0.050354 | -0.14648 | -0.045777 | 0.013733 | 0.10529 | 0.091553 | 0.19226 | 0.23804 | 0.19226 | -0.086975 | 0.23804 | -0.041198
997.1719 | 122.5296 | -57.8613 | 64.2746 | -0.13054 | -0.00058102 | 0.070468 | 0.065074 | -0.0032321 | 0.0027302 | 0.00025641 | -0.019623 | 0.00017309 | 0.063638 | 0.063833 | -0.00060892 | -0.0021939 | 0.0010831 | -0.068665 | 64.2838 | -58.0307 | 64.238 | -57.5455 | -58.0444 | -58.0307 | -0.004578 | 64.27 | -57.9208 | 64.4073 | -57.6599 | -57.9208 | -57.6508 | -0.16937 | 0.009155 | -0.036621 | -0.045776 | 0.31586 | 0.48523 | -0.16937 | -0.18311 | 0 | -0.013733 | -0.059509 | -0.004577 | 0.13275 | 0.13733 | 0.20142 | 0.26093 | 0.21057 | -0.059509 | 0.27008 | 0
998.1719 | 122.4289 | -57.8751 | 64.3158 | -0.092392 | 0.0013416 | 0.068759 | 0.070201 | -0.0023777 | 0.00059397 | 0.002179 | -0.098968 | -4.0508e-05 | 0.064066 | 0.067251 | 0.00045928 | -0.0026211 | 0.00086952 | -0.09613 | 64.3478 | -58.0215 | 64.1647 | -57.6691 | -58.1909 | -58.1223 | -0.10986 | 64.3204 | -57.9346 | 64.3295 | -57.7469 | -58.0399 | -57.7423 | -0.14648 | 0.032043 | -0.15106 | -0.1831 | 0.20599 | 0.35248 | -0.24719 | -0.31586 | -0.10071 | -0.16937 | -0.059509 | 0.004578 | 0.013733 | 0.009155 | 0.12817 | 0.18768 | 0.13275 | -0.16479 | 0.19226 | -0.10529
999.1719 | 122.4884 | -57.8705 | 64.2975 | -0.077133 | 0.00091438 | 0.069613 | 0.066569 | -0.00024129 | 0.00016657 | -0.0025208 | 0.0047913 | -0.00068151 | 0.066629 | 0.062337 | 0.00088648 | -0.0058255 | -0.0019077 | -0.082397 | 64.3295 | -57.9849 | 64.2288 | -57.5729 | -58.1177 | -58.0536 | -0.045776 | 64.2975 | -57.8888 | 64.3799 | -57.6965 | -57.9437 | -57.6645 | -0.11444 | 0.032044 | -0.068664 | -0.10071 | 0.29755 | 0.41199 | -0.18311 | -0.24719 | -0.068665 | -0.13275 | -0.018311 | 0 | 0.082398 | 0.082398 | 0.17395 | 0.19226 | 0.20599 | -0.073243 | 0.2243 | -0.054932
1000.1719 | 122.4426 | -57.8751 | 64.3524 | -0.084763 | -0.00058102 | 0.070041 | 0.069773 | 0.00061311 | -0.00068783 | -0.00081179 | -0.080658 | -0.00046791 | 0.064066 | 0.064901 | -0.00082252 | -0.0026211 | -0.00019868 | -0.11444 | 64.3295 | -58.0444 | 64.1922 | -57.6462 | -58.1909 | -58.1039 | -0.10071 | 64.3295 | -57.93 | 64.3707 | -57.7377 | -57.9987 | -57.7057 | -0.16937 | -0.022888 | -0.16022 | -0.13733 | 0.22888 | 0.39826 | -0.22888 | -0.31586 | -0.059509 | -0.14648 | -0.054932 | -0.022888 | 0.018311 | 0.041199 | 0.13733 | 0.19226 | 0.16937 | -0.1236 | 0.22431 | -0.068664
1001.1719 | 122.4243 | -57.8613 | 64.325 | -0.0084687 | -0.0014356 | 0.068332 | 0.068919 | -0.0017367 | -0.00068783 | 0.0023926 | -0.086761 | -0.00068151 | 0.065134 | 0.068746 | 0.00045928 | -0.0019803 | 0.0019375 | -0.082397 | 64.3524 | -58.0215 | 64.1876 | -57.6462 | -58.1955 | -58.1268 | -0.1236 | 64.3066 | -57.9071 | 64.3524 | -57.7423 | -58.0215 | -57.7332 | -0.16022 | 0.027466 | -0.13733 | -0.16479 | 0.21515 | 0.37537 | -0.2655 | -0.33417 | -0.10529 | -0.17395 | -0.045776 | -0.01831 | 0.027466 | 0.045776 | 0.11902 | 0.16479 | 0.12817 | -0.16022 | 0.17395 | -0.11444
1002.1719 | 122.4976 | -57.8476 | 64.3158 | -0.084763 | 0.00091438 | 0.070895 | 0.068064 | -0.0043003 | 0.0020892 | -0.0027344 | -0.028778 | 0.0023093 | 0.064066 | 0.062337 | -0.00082252 | -0.0039029 | 0.00022852 | -0.073242 | 64.2929 | -57.9987 | 64.2975 | -57.5455 | -58.0673 | -58.017 | -0.018311 | 64.2654 | -57.9025 | 64.4302 | -57.6462 | -57.9254 | -57.6187 | -0.15106 | -0.022888 | -0.018311 | 0.004577 | 0.30212 | 0.45319 | -0.16937 | -0.21973 | -0.018311 | -0.068665 | -0.054932 | -0.050354 | 0.11444 | 0.1648 | 0.20142 | 0.25635 | 0.22888 | -0.07782 | 0.28381 | -0.022888
1003.1719 | 122.374 | -57.8613 | 64.3616 | -0.046616 | 0.00027338 | 0.067691 | 0.069133 | -0.0010959 | -0.0023968 | -0.00017099 | -0.074554 | 0.0031637 | 0.064706 | 0.063833 | -0.0016769 | -5.7742e-05 | 0.0017239 | -0.10071 | 64.3616 | -58.0353 | 64.1876 | -57.6691 | -58.2047 | -58.1268 | -0.027466 | 64.325 | -57.9391 | 64.2883 | -57.692 | -58.0124 | -57.7332 | -0.17395 | 0 | -0.17395 | -0.17395 | 0.19226 | 0.36621 | -0.2655 | -0.34332 | -0.091553 | -0.16937 | -0.07782 | -0.036621 | -0.073242 | -0.036621 | 0.16937 | 0.24719 | 0.12817 | -0.15106 | 0.20599 | -0.073242
1004.1719 | 122.5159 | -57.8201 | 64.3662 | -0.077133 | -0.00015382 | 0.06705 | 0.065287 | -0.0019503 | 0.0020892 | 0.0030334 | 0.047516 | 0.00060029 | 0.066629 | 0.062337 | 0.0019547 | -0.0068935 | 0.0012967 | 0.013733 | 64.3616 | -57.93 | 64.3387 | -57.4493 | -57.9758 | -57.9391 | 0.09613 | 64.3341 | -57.8384 | 64.5172 | -57.5958 | -57.9025 | -57.6233 | -0.10986 | -0.004578 | -0.027466 | -0.022888 | 0.37079 | 0.48065 | -0.11902 | -0.15564 | -0.009155 | -0.045776 | -0.018311 | -0.032044 | 0.15106 | 0.18311 | 0.2243 | 0.24262 | 0.19684 | -0.082398 | 0.21515 | -0.064087
1005.1719 | 122.3877 | -57.8659 | 64.3982 | -0.10002 | -0.00015382 | 0.068973 | 0.070201 | -0.00024129 | -0.00090143 | 0.0030334 | -0.10507 | 0.0014547 | 0.063638 | 0.06661 | -0.0010361 | -0.0026211 | -0.0019077 | -0.086975 | 64.3845 | -58.0078 | 64.2014 | -57.6324 | -58.1772 | -58.1314 | -0.13733 | 64.3662 | -57.8934 | 64.3478 | -57.7835 | -58.0627 | -57.7377 | -0.14191 | -0.013733 | -0.19684 | -0.1831 | 0.23346 | 0.37537 | -0.2655 | -0.31128 | -0.1236 | -0.16937 | -0.027466 | -0.032043 | -0.050354 | -0.018311 | 0.082398 | 0.10986 | 0.12817 | -0.19684 | 0.15564 | -0.16937
1006.1719 | 122.4197 | -57.8476 | 64.389 | -0.077133 | 0.00091438 | 0.0694 | 0.070842 | 0.0014675 | 0.0012348 | 0.00068361 | -0.071503 | 0.0031637 | 0.065988 | 0.064687 | -0.00060892 | -0.0041165 | -0.00019868 | -0.11902 | 64.3707 | -58.0078 | 64.2014 | -57.6462 | -58.1818 | -58.1085 | -0.16022 | 64.3845 | -57.9208 | 64.4989 | -57.8247 | -57.9987 | -57.6736 | -0.16022 | -0.01831 | -0.18768 | -0.16937 | 0.20142 | 0.36163 | -0.26093 | -0.33417 | -0.10071 | -0.17395 | -0.073242 | -0.004578 | 0.10986 | 0.11444 | 0.022888 | 0.09613 | 0.17395 | -0.15106 | 0.24719 | -0.07782
1007.1719 | 122.4564 | -57.8339 | 64.3616 | -0.15343 | 5.9776e-05 | 0.067904 | 0.067637 | 0.0012539 | 0.0040118 | 4.2808e-05 | -0.019623 | -0.00025411 | 0.066629 | 0.066182 | 0.00024568 | -0.0041165 | 0.0036465 | -0.045776 | 64.3433 | -57.9529 | 64.3158 | -57.4631 | -58.0444 | -57.9849 | 0 | 64.3158 | -57.8796 | 64.4531 | -57.6324 | -57.8979 | -57.6004 | -0.11902 | -0.01831 | -0.045776 | -0.027466 | 0.37079 | 0.48981 | -0.15106 | -0.21057 | -0.032043 | -0.091553 | -0.045777 | -0.045776 | 0.091553 | 0.13733 | 0.20142 | 0.24719 | 0.23346 | -0.064087 | 0.27924 | -0.01831
1008.1719 | 122.3694 | -57.843 | 64.4165 | -0.084763 | 0.00070078 | 0.070682 | 0.069133 | -0.0023777 | -0.001115 | 4.2808e-05 | -0.068451 | 0.00038669 | 0.064493 | 0.062337 | 0.00024568 | -0.0024075 | 0.0047147 | -0.068665 | 64.4073 | -57.9941 | 64.2059 | -57.6279 | -58.1543 | -58.1039 | -0.14191 | 64.389 | -57.9163 | 64.3341 | -57.7606 | -58.0353 | -57.724 | -0.15106 | -0.009155 | -0.21057 | -0.20142 | 0.21515 | 0.36621 | -0.26093 | -0.31128 | -0.10986 | -0.16022 | -0.073242 | -0.027466 | -0.082398 | -0.054932 | 0.082398 | 0.15564 | 0.11902 | -0.19226 | 0.19226 | -0.11902
1009.1719 | 122.4701 | -57.8018 | 64.3753 | -0.13817 | -0.00058102 | 0.068332 | 0.067637 | -0.0040867 | 0.0018756 | -0.00081179 | 0.038361 | 0.00060029 | 0.066629 | 0.060415 | 3.2078e-05 | -0.0056119 | -0.0027621 | -0.082397 | 64.3982 | -57.9391 | 64.2792 | -57.5226 | -58.0811 | -57.9849 | -0.041199 | 64.3524 | -57.8613 | 64.4302 | -57.6324 | -57.8888 | -57.6233 | -0.13733 | 0.022888 | -0.09613 | -0.11902 | 0.27924 | 0.41657 | -0.1831 | -0.27924 | -0.045776 | -0.14191 | -0.059509 | -0.022888 | 0.054932 | 0.07782 | 0.16937 | 0.22888 | 0.17853 | -0.086975 | 0.23804 | -0.027466
1010.1719 | 122.3694 | -57.811 | 64.4119 | -0.084763 | 0.00048698 | 0.068973 | 0.070628 | 0.00082671 | 0.00080757 | -0.00059819 | -0.065399 | 0.0023093 | 0.066629 | 0.058919 | -0.00082252 | -0.0024075 | -0.00062588 | -0.045776 | 64.389 | -57.9666 | 64.2471 | -57.5684 | -58.1085 | -58.0673 | -0.073242 | 64.3799 | -57.843 | 64.3707 | -57.7103 | -57.9895 | -57.7057 | -0.15564 | -0.022888 | -0.16479 | -0.14191 | 0.24262 | 0.39825 | -0.25635 | -0.29755 | -0.10071 | -0.14191 | -0.032044 | -0.032043 | -0.041198 | -0.009155 | 0.10071 | 0.13275 | 0.10529 | -0.17853 | 0.13733 | -0.14648
1011.1719 | 122.4884 | -57.7927 | 64.444 | -0.046616 | 0.00027338 | 0.069827 | 0.068492 | 0.0014675 | 0.0014484 | -0.0040162 | -0.047089 | 0.0042319 | 0.065775 | 0.068319 | 0.00024568 | -0.0026211 | 0.00022852 | -0.022888 | 64.4119 | -57.9437 | 64.3066 | -57.4905 | -58.049 | -58.0032 | -0.11902 | 64.3936 | -57.8384 | 64.4348 | -57.637 | -57.8796 | -57.6324 | -0.15106 | -0.032044 | -0.13733 | -0.10528 | 0.30212 | 0.45319 | -0.21057 | -0.25635 | -0.059509 | -0.10529 | -0.045776 | -0.050354 | -0.009156 | 0.041198 | 0.15564 | 0.20142 | 0.16022 | -0.086975 | 0.20599 | -0.041199
1012.1719 | 122.5159 | -57.7698 | 64.3753 | -0.092392 | 0.001128 | 0.0694 | 0.069346 | 0.0016813 | 0.00016657 | 0.00089721 | -0.025726 | 0.00081389 | 0.066843 | 0.057424 | -0.0023177 | -0.0024075 | -0.0019077 | -0.09613 | 64.3982 | -57.9483 | 64.2929 | -57.4951 | -58.0444 | -57.9758 | -0.022888 | 64.3662 | -57.8522 | 64.4897 | -57.6004 | -57.8705 | -57.5684 | -0.17853 | 0.022888 | -0.082397 | -0.10529 | 0.27466 | 0.45319 | -0.20599 | -0.27466 | -0.027466 | -0.096131 | -0.082398 | -0.009155 | 0.11444 | 0.1236 | 0.16937 | 0.25177 | 0.20142 | -0.10071 | 0.28381 | -0.01831
1013.1719 | 122.3831 | -57.7744 | 64.4348 | -0.016098 | 0.00070078 | 0.068545 | 0.069773 | 0.0010403 | 0.00038017 | 0.0017518 | -0.077606 | 0.00081389 | 0.065988 | 0.058919 | -0.0014633 | -0.0045437 | 0.0025785 | -0.054932 | 64.4394 | -57.9712 | 64.2654 | -57.5317 | -58.0856 | -58.0353 | -0.068665 | 64.4348 | -57.8659 | 64.4073 | -57.6782 | -57.9575 | -57.6645 | -0.19684 | 0.004578 | -0.16937 | -0.17395 | 0.24262 | 0.43945 | -0.26093 | -0.31128 | -0.064087 | -0.11444 | -0.091553 | 0 | -0.027465 | -0.027465 | 0.09613 | 0.18768 | 0.10986 | -0.18311 | 0.20142 | -0.091552
1014.1719 | 122.4289 | -57.8018 | 64.4394 | -0.054245 | -0.0016492 | 0.069186 | 0.071482 | 0.0016813 | 0.00016657 | 0.0028198 | -0.056244 | -0.00068151 | 0.064066 | 0.072164 | -0.00018152 | -0.0028347 | 1.4916e-05 | -0.036621 | 64.4302 | -57.9208 | 64.2883 | -57.5134 | -58.0765 | -58.0215 | -0.032043 | 64.4073 | -57.8339 | 64.4211 | -57.6508 | -57.9437 | -57.6462 | -0.11902 | -0.009155 | -0.15106 | -0.14191 | 0.28839 | 0.40741 | -0.21973 | -0.27466 | -0.10071 | -0.15564 | -0.032043 | -0.032043 | -0.01831 | 0.013733 | 0.15106 | 0.18311 | 0.15564 | -0.14191 | 0.18768 | -0.10986
1015.1719 | 122.4106 | -57.7881 | 64.4257 | -0.084763 | 0.0019824 | 0.0694 | 0.068492 | -0.0038729 | -0.00047423 | 0.0013244 | -0.040985 | 0.0012411 | 0.063638 | 0.064473 | 0.0017411 | -0.0021939 | 0.0038601 | -0.027466 | 64.4577 | -57.9208 | 64.3204 | -57.5089 | -58.0353 | -57.9987 | -0.009155 | 64.4211 | -57.8201 | 64.4394 | -57.6416 | -57.9117 | -57.6187 | -0.13275 | 0.032044 | -0.10528 | -0.13733 | 0.27924 | 0.41199 | -0.21057 | -0.24719 | -0.07782 | -0.11444 | -0.032043 | -0.004577 | 0.013733 | 0.01831 | 0.14648 | 0.17853 | 0.16937 | -0.1236 | 0.20142 | -0.091553
1016.1719 | 122.4335 | -57.7789 | 64.4623 | -0.046616 | -0.0025036 | 0.067904 | 0.068705 | 0.00082671 | 0.0012348 | 0.00089721 | -0.05014 | 0.00081389 | 0.066202 | 0.060415 | -0.00018152 | -0.0024075 | 0.0051419 | -0.082397 | 64.444 | -57.9163 | 64.3341 | -57.4768 | -57.9987 | -57.962 | 0 | 64.4257 | -57.8018 | 64.4623 | -57.6096 | -57.8934 | -57.6096 | -0.13733 | -0.01831 | -0.12817 | -0.10986 | 0.30212 | 0.43945 | -0.1831 | -0.21973 | -0.045776 | -0.082397 | -0.022888 | -0.036621 | 0 | 0.036621 | 0.16937 | 0.19226 | 0.16937 | -0.11444 | 0.19226 | -0.091553
1017.1719 | 122.4152 | -57.7835 | 64.4623 | -0.046616 | 0.0013416 | 0.068118 | 0.068705 | 0.0029629 | 0.00059397 | 0.00025641 | -0.0043641 | -0.00089511 | 0.062357 | 0.062764 | -0.00018152 | -0.0043301 | 0.0015103 | -0.10986 | 64.4531 | -57.9483 | 64.4348 | -57.431 | -58.2001 | -58.0032 | -0.15564 | 64.4257 | -57.8201 | 64.4302 | -57.6782 | -57.9575 | -57.6508 | -0.1648 | -0.009155 | -0.027466 | -0.018311 | 0.35248 | 0.51727 | -0.21973 | -0.41657 | -0.054932 | -0.25177 | -0.036621 | -0.036621 | -0.032043 | 0.004578 | 0.10528 | 0.14191 | 0.13275 | -0.17395 | 0.16937 | -0.13733
1018.1719 | 122.4518 | -57.7469 | 64.4806 | -0.10002 | 0.00048698 | 0.069613 | 0.067851 | 0.0027493 | 0.0012348 | 0.0013244 | -0.0074159 | 0.00060029 | 0.065347 | 0.061483 | -0.00060892 | -0.0019803 | -0.00019868 | -0.013733 | 64.5035 | -57.8979 | 64.3295 | -57.4768 | -57.9758 | -57.9575 | 0.036621 | 64.4806 | -57.811 | 64.4852 | -57.5867 | -57.8567 | -57.5775 | -0.15106 | 0.022888 | -0.15106 | -0.17395 | 0.27008 | 0.42114 | -0.21057 | -0.22888 | -0.059509 | -0.07782 | -0.064087 | 0 | 0.004577 | 0.004577 | 0.16022 | 0.2243 | 0.16937 | -0.10986 | 0.23346 | -0.045776
1019.1719 | 122.4289 | -57.7423 | 64.4806 | -0.069504 | -0.001222 | 0.067904 | 0.068492 | 0.0036039 | 0.00059397 | -0.0018798 | -0.047089 | 0.0014547 | 0.064279 | 0.057424 | -0.00060892 | -0.0030483 | 0.0012967 | 0.004578 | 64.4531 | -57.8979 | 64.3616 | -57.4539 | -57.9941 | -57.9391 | 0.041199 | 64.444 | -57.7927 | 64.4943 | -57.5638 | -57.8705 | -57.5775 | -0.15564 | -0.027466 | -0.11902 | -0.091553 | 0.28839 | 0.44403 | -0.19684 | -0.25177 | -0.041199 | -0.096131 | -0.050354 | -0.036621 | 0.013733 | 0.050354 | 0.17853 | 0.22888 | 0.1648 | -0.12817 | 0.21515 | -0.077819
1020.1719 | 122.4289 | -57.724 | 64.5035 | -0.084763 | 0.00070078 | 0.068973 | 0.06956 | 0.0016813 | 0.0014484 | 0.002179 | -0.14474 | 0.00017309 | 0.065988 | 0.062124 | 0.00088648 | -0.00069854 | -0.00062588 | 0 | 64.4943 | -57.8842 | 64.3616 | -57.4356 | -57.9895 | -57.9437 | 0.068665 | 64.476 | -57.7835 | 64.4989 | -57.5775 | -57.8613 | -57.5912 | -0.16022 | -0.009155 | -0.14191 | -0.13275 | 0.28839 | 0.44861 | -0.21973 | -0.2655 | -0.05951 | -0.10529 | -0.059509 | -0.027466 | -0.004578 | 0.022888 | 0.14648 | 0.20599 | 0.13275 | -0.13733 | 0.19226 | -0.07782
1021.1719 | 122.3969 | -57.7194 | 64.5035 | -0.069504 | -0.00036742 | 0.070041 | 0.069346 | 0.00082671 | -0.00026063 | 4.2808e-05 | -0.019623 | 0.0020957 | 0.064493 | 0.062764 | 3.2078e-05 | -0.0021939 | 0.00044212 | -0.013733 | 64.4852 | -57.8888 | 64.3616 | -57.4219 | -57.9758 | -57.9346 | 0.041199 | 64.4669 | -57.7744 | 64.5447 | -57.6828 | -57.8934 | -57.5684 | -0.16937 | -0.018311 | -0.14191 | -0.1236 | 0.29755 | 0.46692 | -0.21515 | -0.25635 | -0.045776 | -0.086975 | -0.054932 | -0.036621 | 0.041199 | 0.07782 | 0.036621 | 0.091553 | 0.15106 | -0.17395 | 0.20599 | -0.11902
1022.1719 | 122.5342 | -57.692 | 64.4531 | -0.10765 | 0.00027338 | 0.067904 | 0.069133 | 0.0016813 | 0.0027302 | 0.00047001 | -0.0043641 | 0.00060029 | 0.066202 | 0.071523 | -0.00082252 | -0.0030483 | 0.0021511 | -0.09613 | 64.4257 | -57.8796 | 64.3662 | -57.4448 | -57.9987 | -57.9254 | -0.022888 | 64.4073 | -57.7286 | 64.5309 | -57.5363 | -57.7927 | -57.5134 | -0.18768 | -0.027466 | -0.086975 | -0.059509 | 0.24719 | 0.43488 | -0.23346 | -0.3067 | -0.045776 | -0.11902 | -0.036621 | -0.045776 | 0.07782 | 0.1236 | 0.15564 | 0.19226 | 0.17853 | -0.10071 | 0.21515 | -0.064087
1023.1719 | 122.438 | -57.7103 | 64.5309 | -0.10002 | 0.00027338 | 0.069827 | 0.069773 | 0.0027493 | 0.00038017 | 0.0017518 | -0.0043641 | 0.00081389 | 0.064493 | 0.067678 | 0.00088648 | -0.0043301 | 0.0030057 | -0.013733 | 64.5035 | -57.8613 | 64.3021 | -57.431 | -57.9483 | -57.9712 | 0.022888 | 64.4806 | -57.7423 | 64.5081 | -57.5592 | -57.8201 | -57.5409 | -0.15106 | -0.027466 | -0.22888 | -0.20142 | 0.27924 | 0.4303 | -0.26093 | -0.23804 | -0.10986 | -0.086975 | -0.032044 | -0.050354 | -0.022888 | 0.027466 | 0.15106 | 0.18311 | 0.16937 | -0.10986 | 0.20142 | -0.077819
1024.1719 | 122.4792 | -57.6782 | 64.5355 | -0.069504 | 0.0015552 | 0.067904 | 0.067637 | -0.0019503 | 0.0012348 | 4.2808e-05 | -0.0013123 | 0.0025229 | 0.065988 | 0.06426 | -0.00018152 | -0.0013393 | -0.00083948 | -0.004578 | 64.5172 | -57.8247 | 64.3707 | -57.4081 | -57.9391 | -57.9117 | 0.045776 | 64.4852 | -57.7423 | 64.4943 | -57.5363 | -57.811 | -57.5272 | -0.14648 | -0.01831 | -0.16479 | -0.14648 | 0.27008 | 0.41656 | -0.23346 | -0.26093 | -0.086975 | -0.11444 | -0.064087 | -0.050354 | -0.041198 | 0.009156 | 0.14191 | 0.20599 | 0.15106 | -0.13275 | 0.21515 | -0.068664
1025.1719 | 122.4701 | -57.7148 | 64.5218 | -0.077133 | 0.00091438 | 0.069613 | 0.070201 | -2.7692e-05 | 0.0014484 | -0.00059819 | -0.0074159 | 0.0014547 | 0.064066 | 0.068532 | -0.0018905 | -0.0049709 | -0.0042575 | -0.027466 | 64.5172 | -57.8384 | 64.3845 | -57.4219 | -57.93 | -57.9071 | 0.013733 | 64.5081 | -57.7194 | 64.5172 | -57.518 | -57.7972 | -57.5272 | -0.1236 | -0.004578 | -0.13733 | -0.13275 | 0.29297 | 0.41656 | -0.19226 | -0.21515 | -0.068664 | -0.091553 | -0.004577 | -0.013733 | -0.004578 | 0.009155 | 0.19684 | 0.20142 | 0.18768 | -0.082397 | 0.19226 | -0.07782
1026.1719 | 122.4197 | -57.692 | 64.5172 | -0.054245 | 0.00070078 | 0.0694 | 0.068705 | -0.00024129 | -0.001115 | 0.001538 | -0.013519 | -0.00046791 | 0.066202 | 0.062764 | 0.00024568 | -0.0024075 | 0.00044212 | 0.004578 | 64.5309 | -57.8522 | 64.4073 | -57.399 | -57.9346 | -57.9025 | 0.041199 | 64.4943 | -57.724 | 64.5401 | -57.5317 | -57.7972 | -57.5317 | -0.16022 | 0.013733 | -0.10986 | -0.1236 | 0.29297 | 0.45319 | -0.21057 | -0.24261 | -0.050354 | -0.082397 | -0.032043 | -0.022888 | 0.022888 | 0.045776 | 0.16022 | 0.19226 | 0.16022 | -0.10529 | 0.19226 | -0.073242
1027.1719 | 122.3877 | -57.7103 | 64.5493 | -0.077133 | 0.00091438 | 0.070682 | 0.069346 | -0.00088229 | -0.00090143 | 0.00068361 | -0.034882 | 0.0014547 | 0.067697 | 0.064046 | -0.00082252 | -0.0032619 | 0.0034329 | -0.09613 | 64.5172 | -57.8659 | 64.3295 | -57.5134 | -58.0627 | -58.0032 | -0.13275 | 64.5126 | -57.7652 | 64.4852 | -57.6096 | -57.8888 | -57.5775 | -0.15564 | -0.032043 | -0.21973 | -0.18768 | 0.19684 | 0.35248 | -0.29297 | -0.35248 | -0.13733 | -0.19684 | -0.054932 | -0.036621 | -0.064087 | -0.027466 | 0.10071 | 0.15564 | 0.13275 | -0.17853 | 0.18768 | -0.1236
1028.1719 | 122.4335 | -57.6736 | 64.5493 | -0.061874 | 0.0030506 | 0.069186 | 0.068492 | 0.00061311 | 0.0020892 | -0.0040162 | 0.0047913 | 0.00081389 | 0.065775 | 0.061269 | -0.0012497 | -0.0036893 | -0.00019868 | 0 | 64.5264 | -57.8156 | 64.4028 | -57.3807 | -57.9163 | -57.8751 | 0.041199 | 64.5035 | -57.7057 | 64.5401 | -57.5043 | -57.8018 | -57.5134 | -0.14191 | -0.022888 | -0.14648 | -0.1236 | 0.29297 | 0.43488 | -0.20142 | -0.24262 | -0.059509 | -0.10071 | -0.032043 | -0.045776 | -0.009155 | 0.036621 | 0.16937 | 0.20142 | 0.16022 | -0.12817 | 0.19226 | -0.096131
1029.1719 | 122.4518 | -57.637 | 64.5447 | -0.00083923 | 0.0013416 | 0.067904 | 0.06721 | -2.7692e-05 | -0.001115 | 0.00089721 | 0.0047913 | 0.0023093 | 0.065347 | 0.065755 | -0.0029587 | -0.0030483 | 0.0025785 | -0.004578 | 64.563 | -57.8156 | 64.4028 | -57.3761 | -57.9208 | -57.8705 | 0.050354 | 64.5218 | -57.7148 | 64.5584 | -57.4768 | -57.8018 | -57.5043 | -0.17853 | 0.01831 | -0.14191 | -0.16022 | 0.26092 | 0.43945 | -0.23346 | -0.28381 | -0.054931 | -0.10529 | -0.07782 | -0.022888 | 0.013733 | 0.036621 | 0.16022 | 0.23804 | 0.13275 | -0.1648 | 0.21057 | -0.086975
1030.1719 | 122.3877 | -57.6508 | 64.5813 | -0.077133 | 0.00091438 | 0.070468 | 0.068919 | -2.7692e-05 | 0.0010212 | -0.00017099 | 0.016998 | 0.0020957 | 0.067484 | 0.066823 | -0.0029587 | -0.0060391 | 0.0012967 | -0.004578 | 64.5767 | -57.8018 | 64.4028 | -57.3853 | -57.9071 | -57.8842 | 0.027466 | 64.5309 | -57.7011 | 64.563 | -57.4768 | -57.7698 | -57.486 | -0.15106 | -0.004578 | -0.17853 | -0.17395 | 0.2655 | 0.41656 | -0.23346 | -0.25635 | -0.082397 | -0.10528 | -0.050354 | -0.050354 | -0.018311 | 0.032043 | 0.17395 | 0.2243 | 0.16479 | -0.11902 | 0.21515 | -0.068664
1031.1719 | 122.4518 | -57.6645 | 64.595 | -0.046616 | -0.00036742 | 0.068545 | 0.069346 | -0.00088229 | 0.00038017 | -0.0025208 | 0.0047913 | 0.0023093 | 0.066843 | 0.062978 | -0.0023177 | -0.0024075 | 0.00065592 | -0.018311 | 64.5813 | -57.8247 | 64.4165 | -57.3944 | -57.8934 | -57.8705 | 0.027466 | 64.5538 | -57.6874 | 64.5584 | -57.4722 | -57.7606 | -57.486 | -0.16022 | -0.013733 | -0.17853 | -0.16479 | 0.27008 | 0.4303 | -0.20599 | -0.22888 | -0.045776 | -0.068665 | -0.022888 | -0.041199 | -0.036621 | 0.004578 | 0.19226 | 0.21515 | 0.17853 | -0.09613 | 0.20142 | -0.073242
1032.1719 | 122.3923 | -57.6599 | 64.6042 | -0.054245 | -0.00058102 | 0.0694 | 0.065715 | 0.0023221 | -0.00090143 | 0.0013244 | -0.0043641 | 0.0016685 | 0.066202 | 0.062764 | -0.00039512 | -0.0024075 | 0.0034329 | -0.059509 | 64.563 | -57.7972 | 64.4028 | -57.2388 | -57.7469 | -58.0078 | 0.050354 | 64.5767 | -57.7011 | 64.5172 | -57.4905 | -57.7744 | -57.5089 | -0.13733 | -0.041199 | -0.20142 | -0.16022 | 0.42114 | 0.55847 | -0.3479 | -0.086975 | -0.21057 | 0.050354 | -0.041199 | -0.027466 | -0.086975 | -0.059509 | 0.16937 | 0.21057 | 0.15106 | -0.11444 | 0.19226 | -0.073242
1033.1719 | 122.4609 | -57.6416 | 64.6042 | -0.069504 | 0.00070078 | 0.066623 | 0.06956 | -0.0025913 | -0.0023968 | -0.0014526 | -0.016571 | -0.00089511 | 0.066415 | 0.065114 | -0.0018905 | -0.0030483 | 0.0021511 | -0.022888 | 64.5721 | -57.7881 | 64.4211 | -57.3807 | -57.9117 | -57.8567 | -0.013733 | 64.5584 | -57.7011 | 64.5538 | -57.4768 | -57.7515 | -57.4905 | -0.14648 | -0.032043 | -0.1831 | -0.15106 | 0.26093 | 0.40741 | -0.21515 | -0.27008 | -0.068664 | -0.1236 | -0.059509 | -0.045776 | -0.050354 | -0.004578 | 0.16479 | 0.2243 | 0.15106 | -0.10986 | 0.21057 | -0.050354
1034.1719 | 122.438 | -57.6324 | 64.6042 | -0.054245 | -0.0010082 | 0.068332 | 0.067851 | -0.00066869 | 0.00080757 | -0.0038024 | -0.010467 | -0.00025411 | 0.063638 | 0.067678 | -0.00039512 | -0.0015529 | 1.4916e-05 | 0.004578 | 64.6088 | -57.7927 | 64.4577 | -57.3257 | -57.8796 | -57.843 | 0.064087 | 64.5813 | -57.6553 | 64.5767 | -57.4631 | -57.7423 | -57.4631 | -0.16022 | 0.004578 | -0.14648 | -0.15106 | 0.3067 | 0.46692 | -0.21057 | -0.24719 | -0.050354 | -0.086975 | -0.022888 | -0.022888 | -0.027466 | -0.004578 | 0.16937 | 0.19226 | 0.16937 | -0.10986 | 0.19226 | -0.086976
1035.1719 | 122.4564 | -57.6004 | 64.6179 | -0.069504 | -0.00058102 | 0.06705 | 0.068064 | -0.00045509 | 0.0020892 | -0.00017099 | -0.040985 | 0.00081389 | 0.065134 | 0.061056 | 0.00045928 | -0.0030483 | 0.00065592 | -0.009155 | 64.6042 | -57.7881 | 64.4623 | -57.3349 | -57.8705 | -57.843 | 0.045776 | 64.5584 | -57.6645 | 64.5905 | -57.4493 | -57.7286 | -57.4677 | -0.18768 | -0.013733 | -0.15564 | -0.14191 | 0.2655 | 0.45319 | -0.24262 | -0.27008 | -0.054932 | -0.082397 | -0.064087 | -0.059509 | -0.027466 | 0.032043 | 0.15106 | 0.21515 | 0.13275 | -0.12817 | 0.19684 | -0.064087
1036.1719 | 122.4564 | -57.6279 | 64.6317 | -0.061874 | -0.00079462 | 0.068545 | 0.068492 | 0.0010403 | -0.00026063 | -0.0040162 | -0.028778 | 0.0010275 | 0.064493 | 0.064687 | 0.00024568 | -0.0017667 | 0.0023647 | -0.018311 | 64.595 | -57.7698 | 64.4485 | -57.3303 | -57.8659 | -57.843 | 0.018311 | 64.5996 | -57.6691 | 64.5813 | -57.4493 | -57.7423 | -57.4677 | -0.14191 | -0.036621 | -0.18311 | -0.14648 | 0.29755 | 0.43945 | -0.21515 | -0.23804 | -0.073243 | -0.096131 | -0.041198 | -0.032044 | -0.050354 | -0.01831 | 0.17853 | 0.21973 | 0.16022 | -0.11444 | 0.20142 | -0.073243
1037.1719 | 122.4426 | -57.6553 | 64.5538 | -0.061874 | -0.00036742 | 0.068118 | 0.068278 | 0.0025357 | 0.0010212 | 0.00025641 | -0.019623 | 0.00060029 | 0.066202 | 0.068746 | -0.00039512 | -0.0024075 | 0.00086952 | -0.091553 | 64.5767 | -57.8293 | 64.3799 | -57.4539 | -58.0032 | -57.93 | -0.11902 | 64.5767 | -57.692 | 64.5447 | -57.5455 | -57.8293 | -57.5317 | -0.17395 | 0.022888 | -0.17395 | -0.19684 | 0.20142 | 0.37537 | -0.27466 | -0.3479 | -0.10071 | -0.17395 | -0.036622 | 0.022888 | -0.009155 | -0.032043 | 0.10986 | 0.14648 | 0.1236 | -0.17395 | 0.16022 | -0.13733
1038.1719 | 122.4152 | -57.605 | 64.6362 | -0.016098 | 0.00027338 | 0.068118 | 0.068064 | 0.0029629 | 0.00038017 | -0.0010254 | 0.016998 | -0.00046791 | 0.065134 | 0.062551 | -0.0012497 | -0.0039029 | 1.4916e-05 | -0.013733 | 64.6271 | -57.7744 | 64.4302 | -57.3532 | -57.8796 | -57.843 | -0.018311 | 64.595 | -57.6599 | 64.5813 | -57.4493 | -57.7423 | -57.4585 | -0.16937 | -0.009155 | -0.20599 | -0.19684 | 0.25177 | 0.42114 | -0.23804 | -0.27466 | -0.068665 | -0.10529 | -0.054932 | -0.041198 | -0.054931 | -0.013733 | 0.15564 | 0.21057 | 0.14648 | -0.13733 | 0.20142 | -0.082398
1039.1719 | 122.4197 | -57.5958 | 64.6683 | -0.054245 | -0.0033582 | 0.068118 | 0.066996 | -0.0057955 | 0.00038017 | -0.0042298 | 0.044464 | 0.0010275 | 0.065347 | 0.063192 | 0.0013139 | -0.0032619 | 0.0021511 | -0.041199 | 64.6133 | -57.7652 | 64.444 | -57.3532 | -57.8934 | -57.8384 | 0.018311 | 64.6133 | -57.6645 | 64.5493 | -57.4585 | -57.5958 | -57.3303 | -0.16937 | -0.054932 | -0.2243 | -0.16937 | 0.24262 | 0.41199 | -0.24262 | -0.29755 | -0.073242 | -0.12817 | -0.068665 | -0.054932 | -0.11902 | -0.064087 | 0.13733 | 0.20599 | 0.2655 | 0 | 0.33417 | 0.068665
1040.1719 | 122.4197 | -57.605 | 64.65 | -0.016098 | 0.00048698 | 0.064914 | 0.069346 | -0.0025913 | 0.0020892 | -0.001239 | 0.0047913 | -0.0011087 | 0.065134 | 0.062337 | 3.2078e-05 | -0.0041165 | -0.00062588 | 0.15564 | 64.6774 | -57.7057 | 64.4073 | -57.3395 | -57.8705 | -57.8247 | 0.004578 | 64.6133 | -57.6599 | 64.5676 | -57.4585 | -57.7377 | -57.4722 | -0.10071 | 0.027466 | -0.24261 | -0.27008 | 0.2655 | 0.36621 | -0.21973 | -0.2655 | -0.11902 | -0.16479 | -0.054932 | -0.036621 | -0.082397 | -0.045776 | 0.14648 | 0.20142 | 0.13275 | -0.13275 | 0.18768 | -0.07782
1041.1719 | 122.4197 | -57.6004 | 64.6545 | -0.046616 | -0.00079462 | 0.068545 | 0.067851 | -0.0021641 | 0.0031574 | 0.001538 | -0.019623 | 0.0023093 | 0.063852 | 0.067037 | -0.00039512 | -0.0045437 | -0.0025485 | -0.013733 | 64.6271 | -57.7652 | 64.4577 | -57.3441 | -57.8522 | -57.8247 | 0.009155 | 64.6271 | -57.6462 | 64.6133 | -57.4219 | -57.7103 | -57.4356 | -0.16479 | -0.027466 | -0.19684 | -0.16937 | 0.25635 | 0.42114 | -0.2243 | -0.25177 | -0.059509 | -0.086975 | -0.045776 | -0.027466 | -0.041199 | -0.013733 | 0.17853 | 0.2243 | 0.16479 | -0.10986 | 0.21057 | -0.064087
1042.1719 | 122.3373 | -57.6187 | 64.6408 | -0.069504 | 0.0019824 | 0.066623 | 0.070842 | 0.0025357 | -0.00026063 | -0.0025208 | -0.053192 | 0.0010275 | 0.067056 | 0.071523 | -0.00018152 | -0.0036893 | 0.00044212 | -0.07782 | 64.6133 | -57.7789 | 64.4165 | -57.3715 | -57.9575 | -57.8751 | -0.15106 | 64.6042 | -57.6691 | 64.5493 | -57.5363 | -57.8156 | -57.5043 | -0.16022 | -0.027466 | -0.2243 | -0.19684 | 0.24719 | 0.40741 | -0.25635 | -0.33875 | -0.09613 | -0.17853 | -0.050354 | -0.036621 | -0.091553 | -0.054932 | 0.082397 | 0.13275 | 0.11444 | -0.19684 | 0.16479 | -0.14648
1043.1719 | 122.406 | -57.5821 | 64.6545 | -0.1458 | 0.002837 | 0.067691 | 0.068705 | -0.00024129 | -0.00090143 | -0.0020936 | 0.010895 | 0.00060029 | 0.064706 | 0.060415 | 0.00067288 | -0.0051845 | -0.0019077 | -0.091553 | 64.6545 | -57.7377 | 64.444 | -57.3212 | -57.8567 | -57.8018 | -0.013733 | 64.6454 | -57.6141 | 64.5996 | -57.4173 | -57.6965 | -57.4356 | -0.15564 | 0 | -0.21057 | -0.21057 | 0.26093 | 0.41656 | -0.21973 | -0.27466 | -0.064087 | -0.11902 | -0.032044 | -0.009155 | -0.054932 | -0.045777 | 0.1648 | 0.19684 | 0.14648 | -0.11444 | 0.17853 | -0.082397
1044.1719 | 122.4197 | -57.5684 | 64.6729 | -0.061874 | 5.9776e-05 | 0.067264 | 0.06721 | -0.0028049 | -0.0013288 | 0.00068361 | 0.0017395 | 0.0023093 | 0.065347 | 0.059774 | 0.00024568 | -0.0021939 | -0.0019077 | -0.032043 | 64.6637 | -57.7332 | 64.476 | -57.4722 | -57.9803 | -57.811 | -0.013733 | 64.6317 | -57.6416 | 64.5905 | -57.4356 | -57.6965 | -57.4356 | -0.16479 | -0.009156 | -0.19684 | -0.18768 | 0.09613 | 0.26093 | -0.24262 | -0.41199 | -0.07782 | -0.24719 | -0.073243 | -0.041199 | -0.082398 | -0.041199 | 0.13275 | 0.20599 | 0.13275 | -0.12817 | 0.20599 | -0.054931
1045.1719 | 122.4106 | -57.6004 | 64.6774 | -0.077133 | -0.00058102 | 0.067904 | 0.069773 | -0.00066869 | 0.0012348 | 0.00068361 | 0.032257 | 0.0031637 | 0.065561 | 0.061483 | -0.00060892 | -0.0021939 | 0.0017239 | -0.036621 | 64.6729 | -57.7194 | 64.4714 | -57.2983 | -57.8293 | -57.7881 | 0.004578 | 64.6729 | -57.6004 | 64.5813 | -57.399 | -57.7515 | -57.4356 | -0.11902 | -0.004577 | -0.20599 | -0.20142 | 0.30212 | 0.42114 | -0.18768 | -0.22888 | -0.068665 | -0.10986 | 0 | -0.004577 | -0.09613 | -0.091553 | 0.20142 | 0.20142 | 0.16479 | -0.15106 | 0.16479 | -0.15106
1046.1719 | 122.4472 | -57.5546 | 64.6729 | -0.077133 | -0.00058102 | 0.066836 | 0.06721 | 0.00039951 | 0.0014484 | -0.0016662 | 0.035309 | 0.0010275 | 0.064066 | 0.063619 | -0.0012497 | -0.0047573 | 0.0017239 | 0.032043 | 64.6179 | -57.692 | 64.5813 | -57.1976 | -57.756 | -57.6965 | 0.09613 | 64.6179 | -57.5958 | 64.6317 | -57.3715 | -57.6508 | -57.399 | -0.13733 | -0.054932 | -0.091553 | -0.036621 | 0.35705 | 0.49438 | -0.14191 | -0.20142 | -0.004577 | -0.064086 | -0.041199 | -0.054932 | -0.041199 | 0.013733 | 0.1831 | 0.2243 | 0.15564 | -0.096131 | 0.19684 | -0.054932
1047.1719 | 122.4106 | -57.5821 | 64.6637 | -0.13054 | 5.9776e-05 | 0.070682 | 0.072123 | 0.0036039 | 0.001662 | -0.0016662 | -0.077606 | 0.00038669 | 0.063852 | 0.065328 | -0.0014633 | -0.0034757 | -0.00083948 | -0.054932 | 64.6133 | -57.724 | 64.5035 | -57.312 | -57.8888 | -57.8156 | -0.068665 | 64.6088 | -57.6279 | 64.6042 | -57.4631 | -57.7606 | -57.4493 | -0.14191 | -0.050354 | -0.16022 | -0.10986 | 0.27008 | 0.41199 | -0.23346 | -0.3067 | -0.091553 | -0.16479 | -0.045777 | -0.054931 | -0.059509 | -0.004578 | 0.11902 | 0.16479 | 0.13275 | -0.17853 | 0.17853 | -0.13275
1048.1719 | 122.3877 | -57.5729 | 64.7003 | 0.022049 | 0.0017688 | 0.068332 | 0.066996 | -0.0015231 | 0.00016657 | -0.00017099 | 0.02005 | 0.00081389 | 0.064493 | 0.062764 | 0.00024568 | -0.0021939 | 0.0027921 | -0.027466 | 64.682 | -57.7011 | 64.5035 | -57.2754 | -57.8247 | -57.7744 | -0.004578 | 64.6683 | -57.5821 | 64.5218 | -57.3853 | -57.6645 | -57.3944 | -0.12817 | -0.01831 | -0.19684 | -0.17853 | 0.29755 | 0.42572 | -0.20142 | -0.25177 | -0.073242 | -0.1236 | -0.009155 | -0.032043 | -0.17853 | -0.14648 | 0.18768 | 0.19684 | 0.17853 | -0.091553 | 0.18768 | -0.082398
1049.1719 | 122.3923 | -57.5455 | 64.7095 | -0.061874 | -0.0020764 | 0.068118 | 0.067424 | -0.00045509 | 0.00016657 | 0.0011108 | 0.016998 | 0.0010275 | 0.06727 | 0.059774 | -0.0031723 | -0.0034757 | -0.0016941 | -0.032043 | 64.682 | -57.6645 | 64.5218 | -57.2571 | -57.8201 | -57.756 | 0.013733 | 64.6729 | -57.5867 | 64.6042 | -57.3761 | -57.6782 | -57.4173 | -0.11902 | -0.027466 | -0.18768 | -0.16022 | 0.28839 | 0.40741 | -0.21057 | -0.27466 | -0.091552 | -0.15564 | -0.041199 | -0.036621 | -0.10529 | -0.068665 | 0.16937 | 0.21057 | 0.12817 | -0.13275 | 0.16937 | -0.091553
1050.1719 | 122.4197 | -57.5546 | 64.7278 | -0.092392 | -0.0014356 | 0.068759 | 0.068064 | -0.00024129 | 0.00080757 | -0.0018798 | 0.026154 | 0.0020957 | 0.065775 | 0.061269 | -0.00082252 | -0.0049709 | 0.00044212 | 0.013733 | 64.6637 | -57.6782 | 64.5264 | -57.1976 | -57.7789 | -57.7423 | 0.036621 | 64.682 | -57.5729 | 64.6225 | -57.3761 | -57.6736 | -57.4219 | -0.1236 | -0.064087 | -0.20142 | -0.13733 | 0.35705 | 0.48065 | -0.18768 | -0.22431 | -0.064087 | -0.10071 | -0.018311 | -0.045776 | -0.10529 | -0.059509 | 0.17853 | 0.19684 | 0.13275 | -0.11902 | 0.15106 | -0.10071
1051.1719 | 122.3831 | -57.5226 | 64.7232 | -0.046616 | 0.0017688 | 0.067264 | 0.067851 | -2.7692e-05 | 0.00059397 | -0.00038459 | 0.02005 | -0.00089511 | 0.065561 | 0.067464 | -0.0016769 | -0.0043301 | 0.0012967 | 0.027466 | 64.682 | -57.6645 | 64.5813 | -57.193 | -57.7377 | -57.7057 | 0.045776 | 64.6774 | -57.55 | 64.6591 | -57.3669 | -57.6416 | -57.4081 | -0.14191 | -0.041199 | -0.14191 | -0.10071 | 0.32959 | 0.4715 | -0.18311 | -0.21515 | -0.041198 | -0.073242 | -0.027466 | -0.045777 | -0.064087 | -0.01831 | 0.15564 | 0.18311 | 0.11444 | -0.11902 | 0.14191 | -0.091553
1052.1719 | 122.438 | -57.5317 | 64.7003 | -0.031357 | 0.00070078 | 0.0694 | 0.069346 | -0.0015231 | 0.00038017 | 4.2808e-05 | -0.059296 | 0.0010275 | 0.064279 | 0.059774 | -0.00039512 | -0.0021939 | 0.00086952 | -0.018311 | 64.6637 | -57.7057 | 64.5767 | -57.2342 | -57.7881 | -57.7515 | -0.004578 | 64.6729 | -57.5912 | 64.6729 | -57.3853 | -57.6874 | -57.3715 | -0.17395 | -0.036621 | -0.1236 | -0.086975 | 0.29755 | 0.4715 | -0.21973 | -0.25635 | -0.045777 | -0.082398 | -0.05951 | -0.027465 | -0.027465 | 0 | 0.14648 | 0.20599 | 0.16022 | -0.15564 | 0.21973 | -0.09613
1053.1719 | 122.4701 | -57.4951 | 64.7141 | -0.023727 | 0.00027338 | 0.066836 | 0.065074 | -0.0021641 | 0.0012348 | -0.001239 | 0.032257 | -0.00025411 | 0.063638 | 0.062764 | -0.00039512 | -0.0041165 | 0.0027921 | 0.032043 | 64.682 | -57.637 | 64.6362 | -57.1289 | -57.6782 | -57.6279 | 0.10071 | 64.6774 | -57.5272 | 64.7095 | -57.2708 | -57.5729 | -57.3166 | -0.14191 | -0.032043 | -0.07782 | -0.045777 | 0.36621 | 0.50812 | -0.13275 | -0.18311 | 0.009155 | -0.041199 | -0.032044 | -0.036621 | -0.004577 | 0.032044 | 0.2243 | 0.25635 | 0.17853 | -0.07782 | 0.21057 | -0.045776
1054.1719 | 122.4884 | -57.4951 | 64.7095 | 0.0067903 | -0.001222 | 0.070682 | 0.066783 | -0.0021641 | -0.00026063 | -0.0048706 | 0.026154 | -0.00046791 | 0.065775 | 0.060415 | 0.00024568 | -0.0030483 | 0.0034329 | -0.027466 | 64.6683 | -57.6324 | 64.6683 | -57.106 | -57.6828 | -57.6141 | 0.12817 | 64.6591 | -57.5409 | 64.7644 | -57.2525 | -57.5363 | -57.2845 | -0.13733 | -0.041199 | -0.041199 | 0 | 0.3891 | 0.52643 | -0.11902 | -0.18768 | 0.01831 | -0.050354 | -0.045777 | -0.050354 | 0.054931 | 0.10528 | 0.24262 | 0.28839 | 0.21057 | -0.041199 | 0.25635 | 0.004578
1055.1719 | 122.4564 | -57.4677 | 64.7507 | -0.054245 | 0.00070078 | 0.066836 | 0.06721 | 0.00082671 | -4.7032e-05 | 4.2808e-05 | 0.016998 | -0.00089511 | 0.063852 | 0.06426 | -0.0021041 | -0.0026211 | 0.0023647 | 0.041199 | 64.682 | -57.6187 | 64.6408 | -57.1106 | -57.6965 | -57.6508 | 0.10071 | 64.682 | -57.5272 | 64.7003 | -57.2754 | -57.5729 | -57.3257 | -0.15106 | -0.068664 | -0.10986 | -0.041199 | 0.35705 | 0.50812 | -0.18311 | -0.22888 | -0.032044 | -0.07782 | -0.05951 | -0.068664 | -0.050354 | 0.01831 | 0.19226 | 0.25177 | 0.14191 | -0.10529 | 0.20142 | -0.045776
1056.1719 | 122.438 | -57.4585 | 64.7324 | -0.077133 | 5.9776e-05 | 0.068118 | 0.068064 | 0.00082671 | 0.00080757 | 4.2808e-05 | 0.026154 | -0.00089511 | 0.066843 | 0.062337 | 3.2078e-05 | -0.0036893 | 0.00065592 | 0.041199 | 64.6866 | -57.6279 | 64.6454 | -57.1014 | -57.6599 | -57.6279 | 0.10986 | 64.6866 | -57.518 | 64.7369 | -57.2617 | -57.5455 | -57.2983 | -0.16937 | -0.045777 | -0.086975 | -0.041198 | 0.35706 | 0.52643 | -0.16937 | -0.20142 | 0 | -0.032043 | -0.059509 | -0.045777 | 0.004577 | 0.050354 | 0.19684 | 0.25635 | 0.16022 | -0.086975 | 0.21973 | -0.027466
1057.1719 | 122.4655 | -57.4493 | 64.7324 | -0.10765 | 0.0019824 | 0.0694 | 0.067424 | 0.0023221 | 0.0025166 | -0.0027344 | 0.016998 | 0.00017309 | 0.064279 | 0.065969 | -0.00060892 | -0.0045437 | -0.0016941 | 0.013733 | 64.6912 | -57.637 | 64.6729 | -57.0969 | -57.6599 | -57.6096 | 0.1236 | 64.6866 | -57.5043 | 64.769 | -57.2342 | -57.5089 | -57.2754 | -0.18768 | -0.041199 | -0.059509 | -0.01831 | 0.35248 | 0.54016 | -0.16022 | -0.21057 | 0.027466 | -0.022888 | -0.054931 | -0.045777 | 0.036621 | 0.082398 | 0.21515 | 0.27008 | 0.17395 | -0.059509 | 0.22888 | -0.004578
1058.1719 | 122.4884 | -57.4539 | 64.7415 | -0.061874 | -0.00058102 | 0.066623 | 0.066996 | 0.00082671 | 0.0014484 | 0.0026062 | 0.041412 | -0.0015359 | 0.066202 | 0.062551 | -0.00060892 | -0.0053983 | -0.0044711 | 0.032043 | 64.682 | -57.605 | 64.6088 | -56.9824 | -57.692 | -57.6782 | 0.1236 | 64.6683 | -57.518 | 64.7919 | -57.2067 | -57.5089 | -57.2708 | -0.15106 | -0.059509 | -0.13275 | -0.073242 | 0.4715 | 0.62256 | -0.22431 | -0.23804 | -0.073243 | -0.086976 | -0.064087 | -0.073242 | 0.050354 | 0.1236 | 0.24719 | 0.31128 | 0.18311 | -0.054932 | 0.24719 | 0.009155
1059.1719 | 122.5616 | -57.4997 | 64.7095 | -0.077133 | 5.9776e-05 | 0.068332 | 0.06721 | 0.00061311 | -0.0013288 | -0.0025208 | -0.010467 | -0.0013223 | 0.065561 | 0.067891 | -0.0016769 | -0.0053983 | 0.00044212 | 0.027466 | 64.7049 | -57.6096 | 64.7003 | -57.0419 | -57.6324 | -57.5775 | 0.11902 | 64.7186 | -57.4631 | 64.8468 | -57.2754 | -57.4905 | -57.2159 | -0.10986 | -0.004578 | -0.009156 | -0.004578 | 0.45776 | 0.56763 | -0.07782 | -0.13275 | 0.032043 | -0.022888 | 0.036621 | 0.009155 | 0.13733 | 0.12817 | 0.2243 | 0.18768 | 0.28381 | 0.009155 | 0.24719 | -0.027466
1060.1719 | 122.5296 | -57.4631 | 64.7232 | -0.038986 | -0.00015382 | 0.067477 | 0.067424 | -0.0010959 | -0.00090143 | 0.001538 | 0.029205 | 0.00060029 | 0.064066 | 0.061056 | -0.0021041 | -0.0019803 | 0.0017239 | -0.018311 | 64.7324 | -57.5958 | 64.6683 | -57.0923 | -57.6691 | -57.6004 | 0.07782 | 64.682 | -57.4951 | 64.801 | -57.2067 | -57.4905 | -57.2479 | -0.13275 | 0.009155 | -0.054932 | -0.064087 | 0.37079 | 0.50354 | -0.13733 | -0.20599 | -0.004578 | -0.073242 | -0.032043 | -0.041199 | 0.077819 | 0.11902 | 0.25635 | 0.28839 | 0.21515 | -0.027466 | 0.24719 | 0.004577
1061.1719 | 122.5616 | -57.4127 | 64.7415 | -0.084763 | 0.00070078 | 0.068759 | 0.066355 | -0.0017367 | -4.7032e-05 | 0.0017518 | 0.023102 | 0.0012411 | 0.067484 | 0.063405 | 0.00024568 | -0.0041165 | -0.0055393 | 0.013733 | 64.7095 | -57.6096 | 64.7141 | -57.0328 | -57.6187 | -57.5638 | 0.16479 | 64.682 | -57.486 | 64.8285 | -57.1564 | -57.4448 | -57.2159 | -0.19684 | -0.032043 | -0.027466 | 0.004577 | 0.37994 | 0.57678 | -0.15106 | -0.20599 | 0.045776 | -0.009155 | -0.073242 | -0.059509 | 0.086975 | 0.14648 | 0.25635 | 0.32959 | 0.19684 | -0.032043 | 0.27008 | 0.041199
1062.1719 | 122.4106 | -57.4539 | 64.7644 | -0.046616 | -0.0014356 | 0.067264 | 0.066996 | 0.0027493 | 0.0010212 | 4.2808e-05 | -0.028778 | 0.00081389 | 0.06492 | 0.05956 | -0.0010361 | -0.0019803 | -0.0048983 | 0.013733 | 64.7552 | -57.605 | 64.6317 | -57.1564 | -57.7148 | -57.6736 | 0.013733 | 64.7552 | -57.5226 | 64.7415 | -57.2891 | -57.5958 | -57.3624 | -0.15106 | -0.009155 | -0.13275 | -0.1236 | 0.29755 | 0.44861 | -0.21973 | -0.26093 | -0.068665 | -0.10986 | -0.068665 | -0.009155 | -0.022888 | -0.013733 | 0.16479 | 0.23346 | 0.091552 | -0.14191 | 0.16022 | -0.073242
1063.1719 | 122.5388 | -57.399 | 64.7461 | -0.084763 | -0.0016492 | 0.066836 | 0.067424 | -0.0021641 | -0.00026063 | -0.00081179 | -0.0013123 | 0.0014547 | 0.064706 | 0.057638 | -0.0010361 | -0.0045437 | 0.00044212 | 0.027466 | 64.7003 | -57.6004 | 64.7415 | -57.0053 | -57.5958 | -57.5455 | 0.14648 | 64.7049 | -57.4814 | 64.8331 | -57.1747 | -57.2342 | -57.0602 | -0.20142 | -0.045777 | -0.004578 | 0.041199 | 0.39368 | 0.59509 | -0.14648 | -0.19684 | 0.054932 | 0.004578 | -0.082397 | -0.041199 | 0.086975 | 0.12817 | 0.2243 | 0.3067 | 0.33874 | 0.16479 | 0.42114 | 0.24719
1064.1719 | 122.5067 | -57.4127 | 64.7324 | -0.084763 | -0.0025036 | 0.067904 | 0.069346 | 0.0014675 | -0.0015424 | 4.2808e-05 | 0.016998 | 0.0012411 | 0.06492 | 0.060628 | -0.0016769 | -0.0045437 | 0.00086952 | 0.013733 | 64.7095 | -57.5867 | 64.7232 | -57.019 | -57.6141 | -57.5455 | 0.16022 | 64.6957 | -57.4677 | 64.8376 | -57.1518 | -57.4356 | -57.225 | -0.17395 | -0.022888 | -0.009155 | 0.013733 | 0.39368 | 0.56763 | -0.13275 | -0.20142 | 0.041199 | -0.027466 | -0.054931 | -0.036621 | 0.10529 | 0.14191 | 0.26093 | 0.31586 | 0.18768 | -0.022888 | 0.24261 | 0.032043
1065.1719 | 122.4792 | -57.4127 | 64.769 | -0.061874 | -0.00079462 | 0.068332 | 0.068705 | -0.0023777 | -0.00090143 | -0.0010254 | -0.03183 | -0.0011087 | 0.06492 | 0.060201 | -0.0012497 | -0.0049709 | 0.0019375 | 0.032043 | 64.7415 | -57.5684 | 64.6637 | -56.7398 | -57.3944 | -57.7606 | 0.23346 | 64.7324 | -57.4722 | 64.8331 | -57.1426 | -57.4539 | -57.2205 | -0.15564 | -0.027466 | -0.10529 | -0.07782 | 0.67291 | 0.82855 | -0.3479 | 0.018311 | -0.19226 | 0.17395 | -0.059509 | -0.036621 | 0.064087 | 0.10071 | 0.27008 | 0.32959 | 0.19226 | -0.041198 | 0.25177 | 0.018311
1066.1719 | 122.5159 | -57.4081 | 64.769 | -0.054245 | -0.00036742 | 0.068973 | 0.068705 | -0.00088229 | -0.00068783 | 0.0023926 | 0.0078429 | 0.0016685 | 0.064493 | 0.06661 | -0.00082252 | -0.0041165 | -0.0027621 | -0.013733 | 64.7095 | -57.5775 | 64.7278 | -57.0511 | -57.6324 | -57.55 | 0.11444 | 64.7232 | -57.4677 | 64.8651 | -57.1426 | -57.4265 | -57.2021 | -0.16937 | -0.059509 | -0.041199 | 0.01831 | 0.35706 | 0.52643 | -0.14191 | -0.2243 | 0.027466 | -0.054931 | -0.059509 | -0.045776 | 0.09613 | 0.14191 | 0.2655 | 0.32501 | 0.20599 | -0.018311 | 0.2655 | 0.041198
1067.1719 | 122.4289 | -57.4493 | 64.7964 | -0.031357 | -0.0014356 | 0.067691 | 0.069346 | -0.0034457 | 0.00059397 | -0.00081179 | -0.068451 | 0.0031637 | 0.06492 | 0.066396 | -0.00082252 | -0.0039029 | -0.0042575 | 0.004578 | 64.7964 | -57.5867 | 64.6408 | -57.1609 | -57.7057 | -57.6553 | 0 | 64.7873 | -57.4997 | 64.7461 | -57.2525 | -57.5638 | -57.2983 | -0.13733 | 0 | -0.15564 | -0.15564 | 0.28839 | 0.42572 | -0.20599 | -0.25635 | -0.068664 | -0.11902 | -0.050354 | -0.009156 | -0.050354 | -0.041198 | 0.19684 | 0.24719 | 0.15106 | -0.11444 | 0.20142 | -0.064087
1068.1719 | 122.5113 | -57.4036 | 64.7644 | -0.13817 | -0.00036742 | 0.06705 | 0.066996 | 0.00018591 | -0.00090143 | 0.00047001 | 0.010895 | 0.0014547 | 0.065134 | 0.061483 | -0.0014633 | -0.0049709 | -0.0014805 | -0.041199 | 64.7461 | -57.5546 | 64.6912 | -57.0648 | -57.637 | -57.5867 | 0.045776 | 64.7278 | -57.4585 | 64.8514 | -57.1609 | -57.4677 | -57.2159 | -0.15106 | -0.01831 | -0.073242 | -0.054932 | 0.33875 | 0.48981 | -0.18311 | -0.23346 | -0.032044 | -0.082398 | -0.054932 | -0.036621 | 0.086975 | 0.1236 | 0.24261 | 0.29755 | 0.18768 | -0.064087 | 0.24261 | -0.009155
1069.1719 | 122.5113 | -57.399 | 64.7507 | -0.13817 | -0.00036742 | 0.067477 | 0.068278 | 0.0018949 | 0.00016657 | 0.00068361 | -0.0043641 | -0.00025411 | 0.063852 | 0.065969 | 3.2078e-05 | -0.0026211 | 0.0017239 | -0.009155 | 64.7552 | -57.5546 | 64.7369 | -57.0053 | -57.5912 | -57.5317 | 0.1236 | 64.7461 | -57.4448 | 64.8788 | -57.1152 | -57.4173 | -57.1976 | -0.15564 | 0.004578 | -0.013733 | -0.018311 | 0.39368 | 0.54932 | -0.13275 | -0.19226 | 0.022888 | -0.036622 | -0.045776 | -0.004577 | 0.12817 | 0.13275 | 0.28381 | 0.32959 | 0.20142 | -0.01831 | 0.24719 | 0.027466
1070.1719 | 122.3969 | -57.4219 | 64.8285 | -0.069504 | -0.0016492 | 0.068973 | 0.069133 | 0.00018591 | 0.00080757 | -0.0010254 | -0.089813 | 0.00060029 | 0.065134 | 0.067464 | -0.00060892 | -0.0030483 | -0.00083948 | -0.018311 | 64.8102 | -57.5775 | 64.6637 | -57.1335 | -57.7057 | -57.637 | 0 | 64.7964 | -57.4722 | 64.8056 | -57.225 | -57.518 | -57.28 | -0.15564 | -0.01831 | -0.16479 | -0.14648 | 0.28839 | 0.44403 | -0.21515 | -0.28381 | -0.059509 | -0.12817 | -0.050354 | -0.032043 | -0.022888 | 0.009155 | 0.19684 | 0.24719 | 0.14191 | -0.09613 | 0.19226 | -0.045776
1071.1719 | 122.4655 | -57.3853 | 64.7827 | -0.077133 | -0.0014356 | 0.068545 | 0.066142 | -0.0019503 | 0.0020892 | -0.0027344 | 0.0047913 | 0.0016685 | 0.064066 | 0.058492 | 3.2078e-05 | -0.0045437 | 0.00065592 | -0.013733 | 64.7598 | -57.5546 | 64.7186 | -57.019 | -57.5958 | -57.55 | 0.041199 | 64.7324 | -57.4722 | 64.8651 | -57.1243 | -57.431 | -57.1884 | -0.16937 | -0.022888 | -0.064087 | -0.041199 | 0.36621 | 0.53558 | -0.16479 | -0.21057 | 0.004577 | -0.041199 | -0.086975 | -0.050354 | 0.082397 | 0.13275 | 0.26093 | 0.3479 | 0.19684 | -0.045776 | 0.28381 | 0.041199
1072.1719 | 122.374 | -57.4036 | 64.8422 | -0.11528 | 0.00091438 | 0.068759 | 0.069133 | -0.00024129 | 0.0020892 | 0.00047001 | -0.034882 | 0.00038669 | 0.063425 | 0.06191 | 0.0013139 | -0.0030483 | 0.0023647 | -0.018311 | 64.8102 | -57.5409 | 64.682 | -57.1106 | -57.6691 | -57.6004 | 0.013733 | 64.8102 | -57.4585 | 64.7736 | -57.2296 | -57.4997 | -57.2662 | -0.13733 | -0.032043 | -0.16022 | -0.12817 | 0.29297 | 0.4303 | -0.19684 | -0.2655 | -0.059509 | -0.12817 | -0.054932 | -0.032043 | -0.068664 | -0.036621 | 0.17395 | 0.22888 | 0.13733 | -0.096131 | 0.19226 | -0.041199
1073.1719 | 122.4289 | -57.3898 | 64.8102 | -0.13054 | -0.00015382 | 0.067264 | 0.066142 | -0.00088229 | 0.0025166 | 0.0017518 | 0.010895 | -0.00046791 | 0.065988 | 0.060842 | 0.00024568 | -0.0047573 | 0.0025785 | -0.045776 | 64.7598 | -57.5363 | 64.7141 | -57.0511 | -57.6416 | -57.5455 | 0 | 64.7507 | -57.4448 | 64.8422 | -57.1472 | -57.4539 | -57.1884 | -0.14648 | -0.050354 | -0.096131 | -0.045777 | 0.33875 | 0.48523 | -0.15564 | -0.25177 | -0.009155 | -0.10529 | -0.054931 | -0.05951 | 0.032043 | 0.091553 | 0.24262 | 0.29755 | 0.20142 | -0.064086 | 0.25635 | -0.009155
1074.1719 | 122.4792 | -57.3853 | 64.7919 | -0.12291 | 0.00070078 | 0.070041 | 0.069987 | 0.0012539 | -0.00068783 | 0.0023926 | -0.019623 | -0.00025411 | 0.06492 | 0.060415 | -0.00060892 | -0.0036893 | -0.0025485 | -0.059509 | 64.7781 | -57.5409 | 64.7049 | -57.1152 | -57.6187 | -57.5317 | -0.013733 | 64.7736 | -57.4356 | 64.856 | -57.1564 | -57.4814 | -57.2159 | -0.15564 | -0.013733 | -0.086975 | -0.073242 | 0.27008 | 0.42572 | -0.14648 | -0.23346 | 0.009156 | -0.077819 | -0.050354 | -0.01831 | 0.064087 | 0.082397 | 0.22888 | 0.27924 | 0.16937 | -0.09613 | 0.21973 | -0.045776
1075.1875 | 122.4014 | -57.3853 | 64.8148 | -0.0084687 | -0.00058102 | 0.068759 | 0.072337 | -0.00024129 | 0.00016657 | 0.0043152 | -0.12338 | 0.0010275 | 0.064706 | 0.072805 | -0.0018905 | -0.0015529 | 0.0023647 | 0.013733 | 64.8102 | -57.5821 | 64.7049 | -57.0511 | -57.6553 | -57.605 | 0.022888 | 64.7964 | -57.4356 | 64.8605 | -57.1793 | -57.5134 | -57.2296 | -0.19684 | -0.004577 | -0.10986 | -0.10529 | 0.33417 | 0.53101 | -0.21973 | -0.27008 | -0.022888 | -0.073242 | -0.050354 | -0.01831 | 0.045777 | 0.064087 | 0.20599 | 0.25635 | 0.15564 | -0.12817 | 0.20599 | -0.07782
1076.1719 | 122.5021 | -57.3669 | 64.7873 | -0.15343 | -0.0027172 | 0.06705 | 0.067637 | -0.0019503 | 0.0018756 | 0.00047001 | 0.010895 | 0.0014547 | 0.065561 | 0.058706 | 0.0011001 | -0.0036893 | 0.0021511 | -0.064087 | 64.8102 | -57.55 | 64.7141 | -57.0328 | -57.6279 | -57.5363 | -0.009155 | 64.7781 | -57.4356 | 64.8605 | -57.1381 | -57.4356 | -57.1884 | -0.18311 | 0.022889 | -0.073242 | -0.096131 | 0.33417 | 0.51727 | -0.16937 | -0.26093 | 0.013733 | -0.07782 | -0.068665 | -0.009155 | 0.073243 | 0.082398 | 0.22888 | 0.29755 | 0.17853 | -0.068665 | 0.24719 | 0
1077.1719 | 122.5204 | -57.3578 | 64.7827 | -0.061874 | -0.00058102 | 0.069186 | 0.067637 | -0.0010959 | 0.0012348 | 0.00047001 | -0.0013123 | -0.00025411 | 0.065988 | 0.060201 | -0.00060892 | -0.0056119 | -0.00041228 | -0.022888 | 64.8056 | -57.518 | 64.7369 | -56.9916 | -57.5729 | -57.5089 | 0.059509 | 64.7736 | -57.4081 | 64.9017 | -57.0969 | -57.3669 | -57.1335 | -0.16022 | 0.022888 | -0.045777 | -0.068665 | 0.36621 | 0.52643 | -0.15106 | -0.21515 | 0.009155 | -0.054932 | -0.050354 | -0.009155 | 0.11902 | 0.12817 | 0.26093 | 0.31128 | 0.2243 | -0.009155 | 0.27466 | 0.041199
1078.1719 | 122.4747 | -57.3624 | 64.7964 | -0.10765 | -0.00015382 | 0.069186 | 0.070201 | 0.00018591 | 0.00016657 | -0.00059819 | -0.016571 | 0.00017309 | 0.066202 | 0.062124 | -0.00039512 | -0.0043301 | 0.0017239 | -0.073242 | 64.7827 | -57.5363 | 64.7141 | -57.0648 | -57.6324 | -57.5455 | -0.064087 | 64.769 | -57.4402 | 64.8514 | -57.1472 | -57.4722 | -57.2159 | -0.17395 | -0.013733 | -0.082398 | -0.068665 | 0.29755 | 0.4715 | -0.1831 | -0.27008 | -0.009155 | -0.09613 | -0.07782 | -0.027466 | 0.054931 | 0.082397 | 0.21515 | 0.29297 | 0.14648 | -0.10986 | 0.2243 | -0.032043
1079.1719 | 122.4884 | -57.3624 | 64.8102 | -0.12291 | 0.0015552 | 0.068973 | 0.067851 | 0.0025357 | 0.0010212 | -0.0014526 | -0.016571 | 0.00081389 | 0.063638 | 0.063405 | -0.00018152 | -0.0026211 | 0.00086952 | -0.082397 | 64.8102 | -57.5317 | 64.7278 | -57.0511 | -57.6187 | -57.5546 | -0.018311 | 64.7873 | -57.4081 | 64.856 | -57.1564 | -57.4448 | -57.1793 | -0.16937 | 0 | -0.082398 | -0.082398 | 0.31128 | 0.48065 | -0.19226 | -0.25635 | -0.022888 | -0.086975 | -0.045776 | -0.022889 | 0.045776 | 0.068665 | 0.20599 | 0.25177 | 0.18311 | -0.082397 | 0.22888 | -0.036621
1080.1719 | 122.3419 | -57.3441 | 64.8605 | -0.069504 | -0.0010082 | 0.0694 | 0.066996 | -0.0013095 | -0.00026063 | -0.00038459 | -0.077606 | 0.0020957 | 0.06492 | 0.06896 | -0.0025315 | -0.0013393 | -0.0027621 | -0.022888 | 64.856 | -57.5226 | 64.6866 | -57.0694 | -57.6324 | -57.5821 | 0 | 64.8468 | -57.4219 | 64.8102 | -57.1976 | -57.2296 | -57.1014 | -0.17853 | -0.004578 | -0.17395 | -0.16937 | 0.27466 | 0.45319 | -0.23804 | -0.28839 | -0.059509 | -0.10986 | -0.07782 | -0.013733 | -0.050354 | -0.036621 | 0.14648 | 0.2243 | 0.24262 | 0.11444 | 0.32044 | 0.19226
1081.1719 | 122.4609 | -57.3761 | 64.8102 | -0.13054 | 0.00070078 | 0.067904 | 0.066569 | 0.00018591 | 0.0025166 | 0.00025641 | -0.037933 | -0.00068151 | 0.064279 | 0.064046 | -0.00039512 | -0.0017667 | 0.0015103 | -0.082397 | 64.7781 | -57.5226 | 64.7461 | -57.019 | -57.6462 | -57.5592 | -0.027466 | 64.8056 | -57.4127 | 64.8651 | -57.1381 | -57.431 | -57.1884 | -0.14648 | -0.032044 | -0.064087 | -0.032043 | 0.35706 | 0.50354 | -0.1831 | -0.27008 | -0.036621 | -0.1236 | -0.036621 | -0.004578 | 0.054931 | 0.059509 | 0.23804 | 0.27466 | 0.18768 | -0.054931 | 0.2243 | -0.01831
1082.1719 | 122.5616 | -57.3624 | 64.8148 | -0.1458 | 0.00091438 | 0.068332 | 0.066569 | -0.0028049 | -0.001756 | -0.0018798 | -0.0013123 | -4.0508e-05 | 0.063425 | 0.063192 | -0.0025315 | -0.0060391 | 0.00065592 | 0.036621 | 64.7964 | -57.4951 | 64.8148 | -56.9138 | -57.4997 | -57.4585 | 0.14191 | 64.8239 | -57.3853 | 65.1215 | -57.3074 | -57.5363 | -57.193 | -0.13275 | -0.01831 | 0 | 0.01831 | 0.44861 | 0.58136 | -0.09613 | -0.13733 | 0.036621 | -0.004578 | -0.022888 | 0.009156 | 0.3067 | 0.29755 | 0.054932 | 0.07782 | 0.16937 | -0.17395 | 0.19226 | -0.15106
1083.1719 | 122.4792 | -57.3441 | 64.8422 | -0.12291 | 0.001128 | 0.068545 | 0.067424 | -0.0023777 | 0.0010212 | 0.0011108 | -0.037933 | -0.00089511 | 0.063852 | 0.061056 | 0.00045928 | -0.0043301 | -0.0016941 | -0.073242 | 64.8102 | -57.4768 | 64.7324 | -57.0282 | -57.605 | -57.5272 | -0.032043 | 64.7873 | -57.3944 | 64.8605 | -57.1426 | -57.4402 | -57.1701 | -0.13275 | -0.032043 | -0.10986 | -0.07782 | 0.31586 | 0.44861 | -0.18311 | -0.26093 | -0.050354 | -0.12817 | -0.050354 | -0.054932 | 0.018311 | 0.073243 | 0.20142 | 0.25177 | 0.17395 | -0.096131 | 0.2243 | -0.045777
1084.1719 | 122.4518 | -57.3486 | 64.8285 | -0.10002 | -0.00015382 | 0.067264 | 0.068064 | 0.0012539 | 0.0014484 | 0.00089721 | -0.03183 | -0.0013223 | 0.06492 | 0.058706 | 0.00024568 | -0.0032619 | -0.00083948 | -0.07782 | 64.801 | -57.5043 | 64.7598 | -56.9962 | -57.605 | -57.5089 | -0.068665 | 64.7964 | -57.4127 | 64.8788 | -57.1289 | -57.4402 | -57.1838 | -0.15564 | -0.027466 | -0.068664 | -0.041198 | 0.35248 | 0.50812 | -0.16022 | -0.25635 | -0.004578 | -0.10071 | -0.064087 | -0.032043 | 0.050354 | 0.082397 | 0.21973 | 0.28381 | 0.16479 | -0.091553 | 0.22888 | -0.027466
1085.1719 | 122.4884 | -57.3486 | 64.8056 | -0.092392 | 0.00070078 | 0.068973 | 0.069773 | -0.00024129 | 0.0012348 | 0.00025641 | -0.056244 | 0.00060029 | 0.065347 | 0.062764 | 0.0015275 | -0.0039029 | -0.0010531 | -0.059509 | 64.7827 | -57.5226 | 64.7644 | -56.9916 | -57.6004 | -57.5317 | -0.032043 | 64.801 | -57.4219 | 64.8743 | -57.1381 | -57.431 | -57.1701 | -0.17395 | -0.022888 | -0.041199 | -0.018311 | 0.35706 | 0.53101 | -0.1831 | -0.25177 | -0.009155 | -0.07782 | -0.073242 | -0.004578 | 0.068665 | 0.073243 | 0.21057 | 0.28381 | 0.17853 | -0.082397 | 0.25177 | -0.009155
1086.1719 | 122.4518 | -57.3395 | 64.8376 | -0.10765 | 0.00048698 | 0.067477 | 0.068919 | -0.00088229 | 0.0020892 | 0.00047001 | -0.013519 | 0.0012411 | 0.063638 | 0.058492 | -0.0016769 | -0.0034757 | 1.4916e-05 | -0.086975 | 64.8148 | -57.4997 | 64.7461 | -57.0007 | -57.5821 | -57.5089 | -0.041199 | 64.8056 | -57.3807 | 64.8926 | -57.106 | -57.4173 | -57.1518 | -0.16022 | -0.022888 | -0.091552 | -0.068664 | 0.33875 | 0.49896 | -0.16937 | -0.24261 | -0.009155 | -0.082397 | -0.041198 | -0.032043 | 0.054932 | 0.086975 | 0.23346 | 0.27466 | 0.18768 | -0.077819 | 0.22888 | -0.036621
1087.1719 | 122.5204 | -57.312 | 64.856 | -0.077133 | -0.0010082 | 0.068118 | 0.068705 | -0.0043003 | 0.0020892 | 4.2808e-05 | -0.013519 | 0.0010275 | 0.065561 | 0.061696 | -0.00082252 | -0.0045437 | 0.0012967 | 0.018311 | 64.8285 | -57.4585 | 64.8285 | -56.8817 | -57.4722 | -57.4219 | 0.14648 | 64.7964 | -57.3715 | 64.9796 | -56.987 | -57.312 | -57.0877 | -0.14648 | -0.027466 | -0.027466 | 0 | 0.4303 | 0.57678 | -0.10986 | -0.16022 | 0.036621 | -0.013733 | -0.059509 | -0.059509 | 0.1236 | 0.1831 | 0.32501 | 0.38452 | 0.2243 | 0 | 0.28381 | 0.059509
1088.1719 | 122.4655 | -57.3166 | 64.8376 | -0.023727 | 5.9776e-05 | 0.070895 | 0.070414 | 0.00018591 | 0.00038017 | 0.00025641 | -0.065399 | 0.0018821 | 0.063638 | 0.062978 | -0.00039512 | -0.0032619 | -0.0027621 | -0.073242 | 64.7827 | -57.4814 | 64.8056 | -56.9595 | -57.5684 | -57.486 | 0.018311 | 64.8239 | -57.3761 | 64.9338 | -57.0831 | -57.3807 | -57.1152 | -0.16479 | -0.054931 | -0.032043 | 0.022888 | 0.35706 | 0.52185 | -0.16937 | -0.25177 | -0.004578 | -0.086975 | -0.05951 | -0.013732 | 0.096131 | 0.10986 | 0.23346 | 0.29297 | 0.20142 | -0.064087 | 0.26093 | -0.004577
1089.1719 | 122.5021 | -57.3212 | 64.801 | -0.13054 | -0.00036742 | 0.069613 | 0.06956 | 0.0010403 | -0.001756 | 0.0023926 | -0.056244 | -4.0508e-05 | 0.061929 | 0.066396 | -0.00082252 | -0.0034757 | -0.0014805 | -0.036621 | 64.8148 | -57.4722 | 64.92 | -56.9412 | -57.5958 | -57.4402 | 0.013733 | 64.7964 | -57.3853 | 64.9246 | -57.0786 | -57.3807 | -57.0969 | -0.15106 | 0.013733 | 0.11902 | 0.10529 | 0.37994 | 0.53101 | -0.11902 | -0.27466 | 0.032043 | -0.1236 | -0.064087 | -0.004577 | 0.1236 | 0.12817 | 0.24262 | 0.3067 | 0.2243 | -0.059509 | 0.28839 | 0.004578
1090.1719 | 122.4518 | -57.3624 | 64.8148 | -0.13054 | -0.00015382 | 0.070041 | 0.068705 | 0.00061311 | 0.0012348 | -0.00059819 | -0.044037 | 0.0020957 | 0.061502 | 0.063833 | -0.0012497 | -0.0026211 | 0.0015103 | -0.054932 | 64.7964 | -57.4722 | 64.8193 | -56.9183 | -57.5592 | -57.4768 | 0.013733 | 64.8239 | -57.3807 | 64.9292 | -57.0786 | -57.399 | -57.1243 | -0.10986 | -0.01831 | 0.004578 | 0.022888 | 0.44403 | 0.55389 | -0.11444 | -0.19684 | -0.004578 | -0.086975 | -0.01831 | 0.009156 | 0.11444 | 0.10528 | 0.28381 | 0.30212 | 0.23804 | -0.036621 | 0.25635 | -0.018311
1091.1719 | 122.4701 | -57.2983 | 64.8285 | -0.077133 | 0.00070078 | 0.068759 | 0.069773 | -0.00024129 | 0.0012348 | 0.0034606 | -0.068451 | 0.0016685 | 0.065775 | 0.062337 | -0.00018152 | -0.0032619 | -0.0023349 | -0.059509 | 64.7919 | -57.4677 | 64.8056 | -56.9046 | -57.5638 | -57.4722 | 0.009155 | 64.8056 | -57.3624 | 64.9384 | -57.074 | -57.3669 | -57.1106 | -0.16937 | -0.036621 | -0.022888 | 0.013733 | 0.39368 | 0.56305 | -0.17395 | -0.2655 | -0.004578 | -0.096131 | -0.064087 | -0.022888 | 0.10986 | 0.13275 | 0.2243 | 0.28839 | 0.18768 | -0.068664 | 0.25177 | -0.004577
1092.1719 | 122.5113 | -57.3074 | 64.8788 | -0.077133 | -0.00015382 | 0.067477 | 0.068064 | 0.00061311 | 0.00059397 | -0.0010254 | -0.0043641 | 0.00060029 | 0.061716 | 0.064473 | -0.00082252 | -0.0045437 | 0.00065592 | 0.036621 | 64.8605 | -57.4265 | 64.8605 | -56.8497 | -57.4539 | -57.4127 | 0.064087 | 64.856 | -57.3074 | 65.0162 | -57.0511 | -57.3624 | -57.0831 | -0.11902 | -0.01831 | -0.01831 | 0 | 0.45776 | 0.57678 | -0.10529 | -0.14648 | 0.013733 | -0.027465 | 0 | -0.022888 | 0.13733 | 0.16022 | 0.25635 | 0.25635 | 0.2243 | -0.054932 | 0.2243 | -0.054932
1093.1719 | 122.4655 | -57.2937 | 64.8468 | -0.077133 | 0.0015552 | 0.068118 | 0.068278 | 0.00039951 | -0.00068783 | -0.0010254 | -0.047089 | 0.0025229 | 0.061716 | 0.067251 | 3.2078e-05 | -0.0026211 | 0.00044212 | -0.050354 | 64.8193 | -57.4768 | 64.9521 | -56.8954 | -57.5867 | -57.4448 | 0.022888 | 64.8102 | -57.3761 | 64.975 | -57.0511 | -57.3624 | -57.0923 | -0.18311 | -0.027466 | 0.10529 | 0.13275 | 0.39825 | 0.58136 | -0.15106 | -0.29297 | 0.032044 | -0.10986 | -0.082398 | -0.036621 | 0.12817 | 0.16479 | 0.24262 | 0.32501 | 0.20142 | -0.068665 | 0.28381 | 0.013733
1094.1719 | 122.5204 | -57.28 | 64.8331 | -0.10002 | 0.00027338 | 0.0694 | 0.068492 | -0.00045509 | 0.00080757 | 0.0011108 | -0.095917 | -4.0508e-05 | 0.061929 | 0.067891 | 0.00045928 | -0.0019803 | 0.00065592 | -0.059509 | 64.8376 | -57.4493 | 64.8239 | -56.9138 | -57.5363 | -57.637 | 0.17395 | 64.8422 | -57.3349 | 64.9658 | -57.0328 | -57.3532 | -57.0877 | -0.16937 | 0.004577 | -0.009155 | -0.013732 | 0.36621 | 0.53558 | -0.35706 | -0.25635 | -0.18768 | -0.086975 | -0.054932 | 0.009155 | 0.13275 | 0.1236 | 0.24719 | 0.30212 | 0.19226 | -0.073242 | 0.24719 | -0.01831
1095.1719 | 122.4976 | -57.2754 | 64.8376 | -0.11528 | 0.00048698 | 0.068118 | 0.069987 | 0.0010403 | 0.0029438 | 0.0034606 | -0.074554 | 0.0016685 | 0.06492 | 0.06426 | -0.00082252 | -0.0034757 | -0.0042575 | -0.054932 | 64.8148 | -57.4677 | 64.8376 | -56.9 | -57.5226 | -57.4631 | -0.004578 | 64.8239 | -57.3715 | 65.0574 | -57.0969 | -57.3303 | -57.0831 | -0.19226 | -0.022888 | 0 | 0.022888 | 0.37537 | 0.56763 | -0.18768 | -0.24719 | 0.004577 | -0.054932 | -0.09613 | -0.013732 | 0.21973 | 0.23346 | 0.17853 | 0.27466 | 0.19226 | -0.054931 | 0.28839 | 0.041199
1096.1719 | 122.4838 | -57.3212 | 64.8056 | -0.19157 | 0.0032642 | 0.070254 | 0.068705 | 0.0010403 | 0.00038017 | -0.0016662 | -0.074554 | 0.00038669 | 0.065134 | 0.063619 | -0.0023177 | -0.0021939 | -0.0048983 | -0.050354 | 64.8422 | -57.4356 | 64.8239 | -56.9046 | -57.5226 | -57.4402 | 0.004578 | 64.8331 | -57.3669 | 64.9704 | -57.0236 | -57.5638 | -57.193 | -0.11444 | 0.036621 | 0.018311 | -0.01831 | 0.41657 | 0.53101 | -0.11902 | -0.20142 | -0.004578 | -0.086975 | -0.045776 | 0.027466 | 0.16479 | 0.13733 | 0.29755 | 0.34332 | 0.12817 | -0.24262 | 0.17395 | -0.19684
1097.1719 | 122.4472 | -57.28 | 64.8743 | -0.17632 | 0.00027338 | 0.066836 | 0.068919 | -0.0013095 | 0.00038017 | -0.0016662 | -0.065399 | 0.0020957 | 0.063638 | 0.068319 | -0.0010361 | -0.0034757 | 0.0017239 | 0.082397 | 64.856 | -57.431 | 64.8376 | -56.8588 | -57.518 | -57.4448 | 0.013733 | 64.8422 | -57.3441 | 64.9796 | -57.0328 | -57.3395 | -57.0831 | -0.15106 | -0.018311 | -0.036622 | -0.018311 | 0.42114 | 0.5722 | -0.1648 | -0.23804 | -0.013733 | -0.086975 | -0.064087 | -0.032044 | 0.10528 | 0.13733 | 0.24719 | 0.31128 | 0.19684 | -0.05951 | 0.26093 | 0.004577
1098.1719 | 122.5204 | -57.2571 | 64.8605 | -0.10765 | 5.9776e-05 | 0.068759 | 0.06956 | -0.0062229 | 0.00038017 | -0.00081179 | -0.037933 | 0.00038669 | 0.063852 | 0.061696 | -0.0010361 | -0.0019803 | 1.4916e-05 | -0.041199 | 64.8376 | -57.4402 | 64.8651 | -56.7169 | -57.4219 | -57.431 | 0.013733 | 64.856 | -57.3212 | 64.9979 | -57.0145 | -57.3486 | -57.074 | -0.18311 | -0.022889 | 0.004577 | 0.027466 | 0.54016 | 0.72327 | -0.17395 | -0.16479 | 0.009156 | 0.018311 | -0.064087 | -0.004578 | 0.13733 | 0.14191 | 0.24262 | 0.3067 | 0.18311 | -0.091553 | 0.24719 | -0.027466
1099.1719 | 122.4701 | -57.2708 | 64.8834 | -0.1458 | 0.001128 | 0.070895 | 0.067851 | 0.0033901 | 0.0012348 | 0.0019654 | -0.065399 | -4.0508e-05 | 0.066415 | 0.064687 | 0.00045928 | -0.0034757 | -0.0038303 | -0.041199 | 64.856 | -57.4493 | 64.856 | -56.8863 | -57.4997 | -57.4356 | -0.20142 | 64.801 | -57.3715 | 65.0116 | -57.0099 | -57.3303 | -57.0374 | -0.17853 | -0.027466 | -0.027466 | 0 | 0.38452 | 0.56305 | -0.1648 | -0.22888 | 0.013733 | -0.050354 | -0.10071 | -0.082398 | 0.12817 | 0.21057 | 0.26093 | 0.36163 | 0.23346 | -0.059509 | 0.33417 | 0.041199
1100.1719 | 122.4884 | -57.2571 | 64.8651 | -0.16106 | 0.00091438 | 0.069186 | 0.07191 | -0.0010959 | 0.0025166 | 0.00047001 | -0.028778 | -0.00046791 | 0.065347 | 0.068746 | 0.00045928 | -0.0034757 | 0.00065592 | -0.059509 | 64.856 | -57.4448 | 64.8468 | -56.8863 | -57.4997 | -57.4265 | 0.013733 | 64.8239 | -57.3441 | 64.8926 | -56.9733 | -57.1335 | -56.8771 | -0.18768 | -0.009155 | -0.01831 | -0.009155 | 0.37079 | 0.55847 | -0.16937 | -0.24262 | 0.01831 | -0.054932 | -0.086975 | -0.041198 | 0.027466 | 0.068664 | 0.28381 | 0.37079 | 0.37994 | 0.1236 | 0.46692 | 0.21057
1101.1719 | 122.4792 | -57.2571 | 64.8788 | -0.092392 | -0.0018628 | 0.068759 | 0.06721 | 0.0016813 | -0.00047423 | 0.002179 | -0.034882 | -0.00025411 | 0.061929 | 0.064901 | 0.00067288 | -0.0019803 | -0.00062588 | -0.07782 | 64.8788 | -57.4356 | 64.8331 | -56.8771 | -57.4768 | -57.4081 | 0.004578 | 64.8743 | -57.3303 | 64.9979 | -57.0236 | -57.3257 | -57.0511 | -0.17853 | 0 | -0.045776 | -0.045776 | 0.37994 | 0.55847 | -0.15106 | -0.21973 | 0.027466 | -0.041199 | -0.073242 | -0.004577 | 0.11902 | 0.1236 | 0.23346 | 0.3067 | 0.20599 | -0.068665 | 0.27924 | 0.004577
1102.1719 | 122.5204 | -57.2571 | 64.9155 | -0.13817 | -0.0014356 | 0.071536 | 0.06721 | 0.0023221 | 0.001662 | 0.00025641 | -0.034882 | 0.0010275 | 0.062784 | 0.070669 | 3.2078e-05 | -0.0011257 | 0.00086952 | -0.11444 | 64.856 | -57.486 | 64.8651 | -56.7993 | -57.4905 | -57.4219 | 0.013733 | 64.8697 | -57.3074 | 65.007 | -56.9916 | -57.3257 | -57.0328 | -0.22888 | -0.059509 | -0.050354 | 0.009155 | 0.45776 | 0.68665 | -0.16479 | -0.23346 | 0.064087 | -0.004578 | -0.050354 | -0.045776 | 0.091553 | 0.13733 | 0.2655 | 0.31586 | 0.2243 | -0.068665 | 0.27466 | -0.018311
1103.1719 | 122.4792 | -57.2617 | 64.8468 | -0.084763 | -0.00036742 | 0.067264 | 0.068492 | -0.00088229 | -0.00047423 | -0.0016662 | -0.065399 | -0.00089511 | 0.064066 | 0.066823 | 0.00088648 | -0.0024075 | 0.00065592 | -0.086975 | 64.8834 | -57.4493 | 64.8514 | -56.8909 | -57.518 | -57.4585 | -0.032043 | 64.8605 | -57.3074 | 65.0162 | -57.0007 | -57.3166 | -57.0465 | -0.18768 | 0.036621 | 0.004577 | -0.032044 | 0.37079 | 0.55847 | -0.19684 | -0.25635 | -0.009155 | -0.068664 | -0.045776 | 0.013733 | 0.16937 | 0.15564 | 0.26093 | 0.3067 | 0.21515 | -0.054931 | 0.26093 | -0.009155
1104.1719 | 122.5204 | -57.2388 | 64.888 | -0.12291 | 5.9776e-05 | 0.067477 | 0.068705 | -0.0032321 | 0.0012348 | 0.00047001 | -0.068451 | 0.0018821 | 0.063425 | 0.064473 | -0.00060892 | -0.0013393 | 0.00022852 | -0.045776 | 64.8697 | -57.4219 | 64.8834 | -56.8588 | -57.4814 | -57.4127 | 0 | 64.8834 | -57.3257 | 65.007 | -56.9824 | -57.2845 | -57.0236 | -0.1831 | -0.01831 | -0.004577 | 0.013733 | 0.37994 | 0.56305 | -0.17395 | -0.24261 | 0.009155 | -0.059509 | -0.086975 | -0.004577 | 0.11902 | 0.1236 | 0.25635 | 0.34332 | 0.21515 | -0.045776 | 0.30212 | 0.041199
1105.1719 | 122.5616 | -57.2617 | 64.8972 | -0.1458 | -0.00036742 | 0.068545 | 0.066569 | 0.0027493 | 0.0027302 | 0.0030334 | -0.077606 | 0.0014547 | 0.064066 | 0.063192 | -0.00018152 | -0.0024075 | -0.0014805 | -0.045776 | 64.9063 | -57.4081 | 64.8743 | -56.8542 | -57.4585 | -57.399 | -0.018311 | 64.856 | -57.312 | 65.0757 | -57.1381 | -57.4722 | -57.0099 | -0.14648 | 0.009155 | -0.022888 | -0.032043 | 0.40741 | 0.55389 | -0.13733 | -0.19684 | 0.009155 | -0.050354 | -0.050354 | -0.041199 | 0.17853 | 0.21973 | 0.1236 | 0.17395 | 0.25177 | -0.21057 | 0.30212 | -0.16022
1106.1719 | 122.5113 | -57.2525 | 64.9109 | -0.077133 | -0.00015382 | 0.067691 | 0.067637 | -0.00045509 | 0.00016657 | 0.002179 | -0.074554 | 0.0014547 | 0.06257 | 0.067251 | -0.0023177 | -0.0021939 | -0.0044711 | -0.045776 | 64.888 | -57.4081 | 64.8651 | -56.8542 | -57.4677 | -57.4173 | -0.022888 | 64.8788 | -57.3166 | 65.0253 | -56.9962 | -57.312 | -57.0099 | -0.15564 | -0.022889 | -0.045777 | -0.022888 | 0.39825 | 0.55389 | -0.16479 | -0.21515 | -0.009155 | -0.059509 | -0.064087 | -0.032044 | 0.11444 | 0.14648 | 0.25635 | 0.32043 | 0.24261 | -0.05951 | 0.3067 | 0.004577
1107.1719 | 122.4838 | -57.2617 | 64.9063 | -0.13817 | -0.00079462 | 0.069186 | 0.067424 | -0.0019503 | 0.001662 | -0.0016662 | 0.02005 | -0.0011087 | 0.061929 | 0.062551 | -0.00018152 | -0.0041165 | 0.0015103 | -0.068665 | 64.9292 | -57.3761 | 64.8514 | -56.8817 | -57.4722 | -57.399 | -0.009155 | 64.9246 | -57.2983 | 65.0024 | -56.9687 | -57.2937 | -57.0419 | -0.11444 | 0.022888 | -0.054932 | -0.07782 | 0.37994 | 0.49439 | -0.13733 | -0.21057 | -0.022888 | -0.09613 | -0.036621 | 0.018311 | 0.09613 | 0.077819 | 0.29297 | 0.32959 | 0.21973 | -0.032043 | 0.25635 | 0.004578
1108.1719 | 122.5159 | -57.2205 | 64.9155 | -0.13817 | 0.0013416 | 0.067691 | 0.069346 | 0.0038175 | 0.0020892 | -0.0010254 | 0.029205 | 0.00060029 | 0.066202 | 0.065328 | -0.0014633 | -0.0030483 | 0.00065592 | -0.068665 | 64.888 | -57.399 | 64.9109 | -56.9321 | -57.7515 | -57.3303 | -0.036621 | 64.8972 | -57.2937 | 65.0299 | -56.987 | -57.2891 | -57.0236 | -0.17853 | -0.027466 | -0.004577 | 0.022889 | 0.28839 | 0.46692 | -0.10986 | -0.53101 | 0.068665 | -0.35248 | -0.073242 | -0.01831 | 0.11444 | 0.13275 | 0.23346 | 0.3067 | 0.19684 | -0.068665 | 0.27008 | 0.004577
1109.1719 | 122.493 | -57.2525 | 64.92 | -0.1458 | 0.00091438 | 0.068118 | 0.067424 | 0.0016813 | 0.0031574 | -0.00038459 | -0.065399 | 0.0014547 | 0.063211 | 0.065328 | -0.0012497 | -0.0034757 | 0.0021511 | -0.10986 | 64.9338 | -57.4081 | 64.8331 | -56.8863 | -57.5043 | -57.4356 | -0.027466 | 64.9155 | -57.2937 | 65.0391 | -56.9916 | -57.3074 | -57.019 | -0.15564 | 0.013733 | -0.086975 | -0.10071 | 0.36621 | 0.52185 | -0.18311 | -0.25177 | -0.027466 | -0.09613 | -0.041199 | -0.004578 | 0.11902 | 0.1236 | 0.26093 | 0.30212 | 0.23346 | -0.054932 | 0.27466 | -0.013733
1110.1719 | 122.5067 | -57.2342 | 64.9155 | -0.12291 | 0.0015552 | 0.068545 | 0.068705 | -0.0017367 | 0.00016657 | 0.00025641 | -0.077606 | 0.0010275 | 0.064493 | 0.068532 | -0.0010361 | -0.0015529 | -0.0027621 | -0.11902 | 64.9521 | -57.4265 | 64.8331 | -56.9046 | -57.518 | -57.4585 | -0.068665 | 64.9063 | -57.3212 | 65.0436 | -56.9916 | -57.3257 | -57.019 | -0.19226 | 0.036621 | -0.082397 | -0.11902 | 0.32959 | 0.52185 | -0.2243 | -0.28381 | -0.032043 | -0.091552 | -0.086975 | -0.009155 | 0.12817 | 0.13733 | 0.24262 | 0.32959 | 0.21515 | -0.091553 | 0.30212 | -0.004578
1111.1719 | 122.5113 | -57.2342 | 64.8788 | -0.084763 | 0.00091438 | 0.070468 | 0.069346 | 0.0025357 | -4.7032e-05 | -0.0016662 | -0.056244 | 0.0023093 | 0.063638 | 0.064473 | 0.00024568 | -0.0017667 | -0.0025485 | -0.064087 | 64.92 | -57.3761 | 64.888 | -56.8222 | -57.4631 | -57.3761 | -0.009155 | 64.8697 | -57.2754 | 65.0436 | -56.955 | -57.28 | -56.9962 | -0.14191 | 0.041199 | 0.009155 | -0.032044 | 0.41199 | 0.55389 | -0.14191 | -0.22888 | 0 | -0.086975 | -0.041199 | -0.009155 | 0.16479 | 0.17395 | 0.27924 | 0.32043 | 0.23804 | -0.045776 | 0.27924 | -0.004577
1112.1719 | 122.493 | -57.2205 | 64.9292 | -0.11528 | -0.00079462 | 0.07175 | 0.069346 | 0.0029629 | 0.0037982 | -0.00017099 | -0.025726 | -0.0021769 | 0.06492 | 0.068746 | -0.0014633 | -0.0028347 | 0.0034329 | -0.24719 | 64.8926 | -57.4036 | 64.9292 | -56.7673 | -57.4631 | -57.3669 | 0 | 64.9063 | -57.2891 | 65.0436 | -56.955 | -57.28 | -57.0007 | -0.18311 | -0.036621 | 0 | 0.036621 | 0.45319 | 0.63629 | -0.14648 | -0.24262 | 0.036621 | -0.05951 | -0.068665 | -0.022888 | 0.11444 | 0.13733 | 0.2655 | 0.33417 | 0.21973 | -0.059509 | 0.28839 | 0.009156
1113.1719 | 122.5021 | -57.2205 | 64.9246 | -0.15343 | 5.9776e-05 | 0.069613 | 0.071055 | 0.0038175 | 0.001662 | 0.0013244 | -0.062347 | 0.00081389 | 0.063852 | 0.072378 | -0.0014633 | -0.0026211 | -0.00041228 | -0.082397 | 64.9246 | -57.3715 | 64.8834 | -56.8588 | -57.4585 | -57.3761 | -0.027466 | 64.9155 | -57.2617 | 65.0528 | -56.9412 | -57.2708 | -56.9778 | -0.15106 | 0 | -0.041199 | -0.041199 | 0.36163 | 0.5127 | -0.15564 | -0.23804 | -0.004578 | -0.086975 | -0.041199 | -0.009156 | 0.12817 | 0.13733 | 0.27924 | 0.32043 | 0.24262 | -0.050354 | 0.28381 | -0.009155
1114.1719 | 122.493 | -57.1976 | 64.9155 | -0.11528 | 0.00027338 | 0.070468 | 0.069773 | 0.0018949 | 0.001662 | 0.0013244 | -0.080658 | 0.00017309 | 0.065988 | 0.064046 | -0.0014633 | -0.0021939 | -0.0016941 | -0.10071 | 64.9292 | -57.399 | 64.8651 | -56.8359 | -57.4448 | -57.399 | -0.054932 | 64.9246 | -57.2662 | 65.0711 | -56.9366 | -57.28 | -56.987 | -0.20142 | 0.013733 | -0.050354 | -0.064087 | 0.36163 | 0.56305 | -0.20142 | -0.24719 | 0 | -0.045776 | -0.068664 | 0.009156 | 0.15564 | 0.14648 | 0.26093 | 0.32959 | 0.21057 | -0.082397 | 0.27923 | -0.013733
1115.1719 | 122.4792 | -57.2067 | 64.9612 | -0.092392 | 0.00048698 | 0.067904 | 0.070628 | -0.0053683 | 0.0031574 | -0.0014526 | -0.05014 | 0.0014547 | 0.064493 | 0.06426 | -0.0016769 | -0.0019803 | -0.0014805 | -0.11444 | 64.9612 | -57.3715 | 64.8514 | -56.8771 | -57.4722 | -57.4127 | -0.068665 | 64.9292 | -57.2525 | 65.0482 | -56.9458 | -57.2479 | -56.9778 | -0.16479 | 0 | -0.10986 | -0.10986 | 0.32959 | 0.49439 | -0.20599 | -0.2655 | -0.041199 | -0.10071 | -0.045776 | -0.032044 | 0.086975 | 0.11902 | 0.26093 | 0.3067 | 0.22888 | -0.041199 | 0.27466 | 0.004577
1116.1719 | 122.4655 | -57.2113 | 64.9612 | -0.11528 | 0.00048698 | 0.069827 | 0.070628 | -0.0017367 | 0.0040118 | 0.002179 | -0.056244 | 0.0018821 | 0.063425 | 0.068532 | 0.0011001 | -0.0021939 | 0.0023647 | -0.11444 | 64.9933 | -57.3578 | 64.8331 | -56.8817 | -57.5226 | -57.4448 | -0.091553 | 64.9521 | -57.28 | 65.0345 | -56.9687 | -57.3029 | -57.0145 | -0.14648 | 0.032043 | -0.12817 | -0.16022 | 0.32959 | 0.47607 | -0.23346 | -0.31128 | -0.086975 | -0.16479 | -0.068664 | -0.009156 | 0.073242 | 0.082398 | 0.24262 | 0.31128 | 0.19684 | -0.091552 | 0.2655 | -0.022888
1117.1719 | 122.4609 | -57.2205 | 64.9475 | -0.11528 | 0.0015552 | 0.067477 | 0.065501 | -0.00066869 | 0.00016657 | -0.0014526 | 0.026154 | -0.00025411 | 0.066843 | 0.065542 | -0.00082252 | -0.0045437 | -0.00062588 | -0.091553 | 64.975 | -57.3395 | 64.888 | -56.8314 | -57.4448 | -57.3578 | -0.045776 | 64.9429 | -57.2388 | 65.0162 | -56.9504 | -57.2571 | -57.0053 | -0.11902 | 0.027466 | -0.05951 | -0.086976 | 0.3891 | 0.50812 | -0.13733 | -0.2243 | -0.01831 | -0.10529 | -0.018311 | -0.004578 | 0.068664 | 0.073242 | 0.27008 | 0.28839 | 0.21515 | -0.036621 | 0.23346 | -0.01831
1118.1719 | 122.4884 | -57.1976 | 64.9796 | -0.12291 | -0.00079462 | 0.068118 | 0.068705 | -0.0032321 | 0.0029438 | 0.003247 | -0.056244 | -4.0508e-05 | 0.063638 | 0.062124 | 3.2078e-05 | -0.0019803 | -0.00083948 | -0.091553 | 64.9521 | -57.3715 | 64.9109 | -56.8176 | -57.4356 | -57.3761 | -0.032043 | 64.9521 | -57.2479 | 65.0894 | -56.9321 | -57.2479 | -56.9595 | -0.17395 | -0.027466 | -0.068664 | -0.041198 | 0.37994 | 0.55389 | -0.17853 | -0.23804 | -0.004578 | -0.064087 | -0.050354 | -0.027466 | 0.10986 | 0.13733 | 0.2655 | 0.31586 | 0.23804 | -0.050354 | 0.28839 | 0
1119.1719 | 122.5113 | -57.193 | 64.975 | -0.13054 | 0.002196 | 0.0694 | 0.068705 | -0.00024129 | 0.001662 | 0.0011108 | -0.062347 | -4.0508e-05 | 0.064706 | 0.065542 | -0.00060892 | -0.0019803 | 0.00022852 | -0.11444 | 64.975 | -57.3669 | 64.8651 | -56.8497 | -57.4677 | -57.4036 | -0.059509 | 64.9567 | -57.2662 | 65.0665 | -56.9641 | -57.2845 | -56.9824 | -0.17395 | 0 | -0.10986 | -0.10986 | 0.34332 | 0.51727 | -0.21057 | -0.27466 | -0.036621 | -0.10071 | -0.073242 | -0.018311 | 0.091552 | 0.10986 | 0.22888 | 0.30212 | 0.21057 | -0.091553 | 0.28381 | -0.018311
1120.1719 | 122.4335 | -57.1838 | 64.9704 | -0.13054 | -0.00079462 | 0.0694 | 0.068492 | -0.0010959 | 0.00038017 | 0.003247 | -0.077606 | 0.00060029 | 0.066202 | 0.067464 | -0.00018152 | -0.0013393 | -0.00041228 | -0.10529 | 65.0024 | -57.3761 | 64.8743 | -56.8726 | -57.4951 | -57.4356 | -0.10986 | 64.9796 | -57.2525 | 65.0345 | -56.9504 | -57.2937 | -57.0145 | -0.19226 | 0.032043 | -0.09613 | -0.12817 | 0.31128 | 0.50354 | -0.25177 | -0.31128 | -0.059509 | -0.11902 | -0.068664 | 0.009155 | 0.064087 | 0.054932 | 0.23346 | 0.30212 | 0.16937 | -0.10986 | 0.23804 | -0.041199
1121.1719 | 122.4289 | -57.2067 | 64.975 | -0.077133 | -0.0025036 | 0.067691 | 0.07191 | -0.0025913 | 0.0025166 | 0.003247 | -0.080658 | 0.0018821 | 0.065134 | 0.061696 | 0.00067288 | -0.0036893 | 0.0021511 | -0.11444 | 64.9979 | -57.3578 | 64.856 | -56.8817 | -57.5043 | -57.431 | -0.10986 | 64.9796 | -57.2479 | 65.0299 | -56.9824 | -57.3257 | -57.0328 | -0.15106 | 0.022888 | -0.11902 | -0.14191 | 0.32501 | 0.47607 | -0.2243 | -0.29755 | -0.073242 | -0.14648 | -0.041199 | 0.004577 | 0.054931 | 0.050354 | 0.2243 | 0.2655 | 0.17395 | -0.11902 | 0.21515 | -0.07782
1122.1719 | 122.4609 | -57.1564 | 64.9841 | -0.1458 | -0.00015382 | 0.068759 | 0.068705 | 0.0012539 | 0.00038017 | -0.00081179 | -0.010467 | 0.00017309 | 0.063852 | 0.067678 | 0.0017411 | -0.0011257 | -0.00019868 | -0.059509 | 64.9521 | -57.3074 | 64.9475 | -56.749 | -57.3944 | -57.3166 | -0.013733 | 64.9796 | -57.2433 | 65.0162 | -56.749 | -57.0923 | -56.9229 | -0.15106 | -0.032044 | -0.036621 | -0.004577 | 0.40741 | 0.55847 | -0.16022 | -0.23804 | -0.009155 | -0.086975 | -0.086975 | -0.004578 | 0.032043 | 0.036621 | 0.40741 | 0.49439 | 0.23346 | 0.064087 | 0.32043 | 0.15106
1123.1719 | 122.3877 | -57.1884 | 65.0116 | -0.13054 | 0.00027338 | 0.069613 | 0.069987 | 0.0014675 | -0.00068783 | -0.00081179 | -0.08371 | 0.0025229 | 0.063638 | 0.063833 | 0.0017411 | -0.0032619 | 0.0017239 | -0.10986 | 65.0116 | -57.3395 | 64.9521 | -56.987 | -57.4997 | -57.3669 | -0.13275 | 65.0162 | -57.2388 | 65.007 | -57.0053 | -57.3486 | -57.0419 | -0.15106 | 0 | -0.05951 | -0.05951 | 0.20142 | 0.35248 | -0.17853 | -0.31128 | -0.027465 | -0.16022 | -0.050354 | 0.004577 | -0.004578 | -0.009155 | 0.18311 | 0.23346 | 0.14648 | -0.16022 | 0.19684 | -0.10986
1124.1719 | 122.406 | -57.1793 | 65.0208 | -0.092392 | -0.0010082 | 0.07175 | 0.06956 | 0.0025357 | 0.00080757 | -0.00059819 | -0.047089 | 0.0029501 | 0.066629 | 0.064901 | -0.00039512 | -0.0011257 | -0.0012667 | -0.11902 | 65.0162 | -57.3486 | 64.8605 | -56.8634 | -57.4997 | -57.4127 | -0.10986 | 65.0116 | -57.2159 | 65.0482 | -56.9733 | -57.2891 | -57.0145 | -0.16937 | -0.004578 | -0.16022 | -0.15564 | 0.31586 | 0.48523 | -0.23346 | -0.32044 | -0.064087 | -0.15106 | -0.036621 | -0.009155 | 0.027466 | 0.036621 | 0.20599 | 0.24261 | 0.16479 | -0.10986 | 0.20142 | -0.073243
1125.1719 | 122.4152 | -57.1976 | 65.0253 | -0.061874 | -0.00015382 | 0.069186 | 0.068064 | -0.0032321 | -0.00068783 | 0.003888 | -0.071503 | 0.0025229 | 0.065988 | 0.066182 | -0.00082252 | -0.0028347 | -0.00041228 | -0.13733 | 65.0299 | -57.3395 | 64.856 | -56.8634 | -57.4539 | -57.4036 | -0.11444 | 65.0253 | -57.2205 | 65.0345 | -56.9824 | -57.3029 | -57.019 | -0.14191 | 0.004577 | -0.16937 | -0.17395 | 0.33417 | 0.47608 | -0.20599 | -0.25635 | -0.064086 | -0.11444 | -0.022888 | 0 | 0.009155 | 0.009155 | 0.21515 | 0.23804 | 0.17853 | -0.10528 | 0.20142 | -0.082397
1126.1719 | 122.4014 | -57.1747 | 65.0436 | -0.054245 | 0.002196 | 0.0694 | 0.070414 | -0.0032321 | 0.0018756 | -0.00059819 | -0.044037 | 0.0020957 | 0.064279 | 0.070241 | 0.00024568 | -0.0011257 | -0.0010531 | -0.37537 | 64.9887 | -57.3578 | 64.9017 | -56.868 | -57.4905 | -57.399 | -0.14648 | 65.0253 | -57.2067 | 65.0208 | -56.9778 | -57.3074 | -57.0145 | -0.1831 | -0.054932 | -0.14191 | -0.086975 | 0.3067 | 0.48981 | -0.2243 | -0.31586 | -0.041199 | -0.13275 | -0.032043 | -0.01831 | -0.022888 | -0.004578 | 0.19684 | 0.22888 | 0.16022 | -0.13275 | 0.19226 | -0.10071
1127.1719 | 122.406 | -57.1564 | 65.062 | -0.10765 | -0.001222 | 0.070041 | 0.068278 | -0.0034457 | 0.003371 | -0.0010254 | 0.078034 | 0.0031637 | 0.06727 | 0.06661 | 0.00088648 | -0.0011257 | 0.00044212 | -0.1236 | 65.0299 | -57.3303 | 64.8605 | -56.781 | -57.843 | -57.5089 | -0.18768 | 65.0482 | -57.2021 | 65.0391 | -56.9824 | -57.3486 | -57.0099 | -0.17395 | -0.032044 | -0.20142 | -0.16937 | 0.37537 | 0.54932 | -0.35248 | -0.68665 | -0.17853 | -0.5127 | -0.045776 | -0.013733 | -0.022889 | -0.009156 | 0.17395 | 0.21973 | 0.14648 | -0.19226 | 0.19226 | -0.14648
1128.1719 | 122.3831 | -57.1243 | 65.0848 | -0.13817 | 0.00070078 | 0.0694 | 0.071269 | -0.0010959 | 0.0029438 | 0.0013244 | -0.05014 | 0.00081389 | 0.066843 | 0.076009 | 0.0019547 | -0.00069854 | -0.0019077 | -0.12817 | 65.0711 | -57.3029 | 64.8514 | -56.8588 | -57.4722 | -57.4173 | -0.16479 | 65.0528 | -57.2067 | 65.0986 | -56.7581 | -57.3853 | -57.0786 | -0.17853 | -0.013733 | -0.23346 | -0.21973 | 0.2655 | 0.44403 | -0.29297 | -0.3479 | -0.11444 | -0.16937 | -0.082397 | -0.032044 | 0.013733 | 0.045777 | 0.36621 | 0.44861 | 0.045777 | -0.26093 | 0.12817 | -0.17853
1129.1719 | 122.4243 | -57.1609 | 65.0665 | 0.098343 | 0.00091438 | 0.069186 | 0.069773 | -0.0030185 | 0.00016657 | 0.00089721 | -0.047089 | 0.0014547 | 0.065988 | 0.068532 | -0.0014633 | -0.0024075 | 0.0036465 | -0.16937 | 65.0848 | -57.2845 | 64.8972 | -56.8405 | -57.4631 | -57.3853 | -0.11902 | 65.0528 | -57.2113 | 65.0436 | -56.9733 | -57.3074 | -57.0099 | -0.1236 | 0.018311 | -0.16937 | -0.18768 | 0.32043 | 0.44403 | -0.2243 | -0.30212 | -0.10071 | -0.17853 | -0.050354 | -0.013733 | -0.022888 | -0.009155 | 0.18768 | 0.23804 | 0.15106 | -0.14648 | 0.20142 | -0.09613
1130.1719 | 122.3373 | -57.1564 | 65.0803 | -0.13817 | 5.9776e-05 | 0.066836 | 0.071482 | 0.00018591 | 0.00016657 | -0.0014526 | -0.065399 | -0.00089511 | 0.064066 | 0.071523 | 0.00067288 | -0.0026211 | 0.00044212 | -0.09613 | 65.0711 | -57.2983 | 64.8834 | -56.8405 | -57.4677 | -57.399 | -0.16479 | 65.0757 | -57.2067 | 65.0299 | -56.9824 | -57.3395 | -57.0282 | -0.14191 | -0.009155 | -0.19684 | -0.18768 | 0.31586 | 0.45776 | -0.24262 | -0.31128 | -0.10071 | -0.16937 | -0.050354 | -0.004577 | -0.050354 | -0.045777 | 0.17395 | 0.2243 | 0.12817 | -0.18311 | 0.17853 | -0.13275
1131.1719 | 122.3969 | -57.1289 | 65.0574 | -0.13054 | -0.00015382 | 0.068545 | 0.069987 | -0.0028049 | -0.00026063 | -0.00017099 | -0.21494 | 0.00081389 | 0.064493 | 0.065114 | -0.00018152 | -0.0017667 | 0.0017239 | -0.10986 | 65.1031 | -57.3029 | 64.9017 | -56.8451 | -57.4585 | -57.4081 | -0.009155 | 65.0299 | -57.1884 | 65.0391 | -56.9595 | -57.2754 | -56.9962 | -0.17395 | 0.045776 | -0.15564 | -0.20142 | 0.28381 | 0.45776 | -0.27924 | -0.32959 | -0.10529 | -0.15564 | -0.05951 | -0.027466 | -0.018311 | 0.009155 | 0.16937 | 0.22888 | 0.13275 | -0.14648 | 0.19226 | -0.086975
1132.1719 | 122.4701 | -57.106 | 65.0391 | -0.13054 | 0.001128 | 0.068759 | 0.06721 | -0.0023777 | -0.001115 | -0.0038024 | 0.0078429 | 0.0020957 | 0.064279 | 0.061269 | 0.0013139 | -0.0060391 | 0.0010831 | -0.07782 | 65.0345 | -57.2296 | 65.0024 | -56.6666 | -57.3303 | -57.2571 | 0.004578 | 65.0162 | -57.1701 | 65.1306 | -56.8405 | -57.1793 | -56.9046 | -0.1236 | -0.004577 | -0.036621 | -0.032044 | 0.43945 | 0.56305 | -0.15106 | -0.2243 | -0.027466 | -0.10071 | -0.064087 | -0.022888 | 0.091553 | 0.11444 | 0.2655 | 0.32959 | 0.20142 | -0.073242 | 0.2655 | -0.009155
1133.1719 | 122.3877 | -57.1335 | 65.0803 | -0.15343 | 0.00070078 | 0.069827 | 0.069773 | -0.0013095 | 0.001662 | -0.001239 | -0.08371 | 0.0029501 | 0.065775 | 0.06896 | 0.00024568 | -0.0039029 | 0.00086952 | -0.10986 | 65.0848 | -57.2937 | 64.8972 | -56.8222 | -57.4402 | -57.3898 | -0.17395 | 65.094 | -57.1884 | 65.0528 | -56.9412 | -57.28 | -56.9916 | -0.16022 | 0.004578 | -0.1831 | -0.18768 | 0.31128 | 0.4715 | -0.25635 | -0.3067 | -0.096131 | -0.14648 | -0.054932 | 0.013733 | -0.027466 | -0.041199 | 0.19226 | 0.24719 | 0.14191 | -0.14648 | 0.19684 | -0.091552
1134.1719 | 122.3282 | -57.1518 | 65.126 | -0.0084687 | 5.9776e-05 | 0.0694 | 0.070414 | 0.00018591 | 0.00016657 | -0.00038459 | -0.089813 | -0.00068151 | 0.063211 | 0.064687 | -0.0012497 | -0.0021939 | -0.00083948 | -0.086975 | 65.1077 | -57.2479 | 64.9063 | -56.8405 | -57.4356 | -57.3853 | -0.15564 | 65.0986 | -57.1838 | 65.0436 | -56.9687 | -57.3029 | -56.987 | -0.096131 | -0.018311 | -0.21973 | -0.20142 | 0.31128 | 0.40741 | -0.23346 | -0.28381 | -0.13733 | -0.18768 | -0.032044 | -0.027466 | -0.082398 | -0.054932 | 0.18311 | 0.21515 | 0.16479 | -0.15106 | 0.19684 | -0.11902
1135.1719 | 122.3785 | -57.1152 | 65.1077 | -0.11528 | 0.00048698 | 0.068759 | 0.069987 | 0.0012539 | 0.00016657 | -0.00038459 | -0.059296 | 0.0025229 | 0.065561 | 0.069173 | -0.0025315 | -0.0021939 | 0.0070645 | -0.082397 | 65.1215 | -57.2617 | 64.9292 | -56.781 | -57.4127 | -57.3624 | -0.13275 | 65.0665 | -57.1518 | 65.062 | -56.9275 | -57.2845 | -56.9778 | -0.14648 | 0.013733 | -0.17853 | -0.19226 | 0.33417 | 0.48065 | -0.24719 | -0.29755 | -0.10071 | -0.15106 | -0.036621 | -0.041199 | -0.045776 | -0.004577 | 0.18768 | 0.2243 | 0.13733 | -0.16937 | 0.17395 | -0.13275
1136.1875 | 122.4747 | -57.0648 | 65.1215 | -0.084763 | 0.0013416 | 0.070254 | 0.066569 | 0.0053129 | 0.0035846 | -0.001239 | 0.10245 | -0.0011087 | 0.065988 | 0.060201 | -0.00018152 | -0.0041165 | 0.0027921 | 0.045776 | 65.1306 | -57.1747 | 65.0803 | -56.5247 | -57.3349 | -57.2342 | 0.10071 | 65.0665 | -57.1198 | 65.2084 | -56.7673 | -57.0923 | -56.8359 | -0.10986 | 0.009155 | -0.041199 | -0.050354 | 0.54016 | 0.65002 | -0.16937 | -0.27008 | -0.059509 | -0.16022 | -0.054932 | -0.054932 | 0.086975 | 0.14191 | 0.29755 | 0.35248 | 0.22888 | -0.027466 | 0.28381 | 0.027466
1137.1719 | 122.5021 | -57.0648 | 65.0711 | -0.10765 | 0.00091438 | 0.068973 | 0.069773 | 0.0033901 | 0.0018756 | 0.00025641 | 0.0078429 | -0.0015359 | 0.067484 | 0.063619 | -0.00039512 | -0.0030483 | 0.00022852 | -0.045776 | 65.0574 | -57.225 | 65.0436 | -56.6483 | -57.3074 | -57.2342 | -0.004578 | 65.062 | -57.1198 | 65.181 | -56.8039 | -57.1243 | -56.8451 | -0.16022 | -0.013733 | -0.027466 | -0.013733 | 0.41656 | 0.57678 | -0.16937 | -0.24262 | -0.009155 | -0.082397 | -0.054932 | -0.009155 | 0.10986 | 0.11902 | 0.26093 | 0.31586 | 0.21973 | -0.05951 | 0.27466 | -0.004578
1138.1719 | 122.3511 | -57.1152 | 65.126 | -0.092392 | 5.9776e-05 | 0.070682 | 0.068492 | -0.0010959 | 0.001662 | 0.002179 | -0.08371 | 0.0038047 | 0.065775 | 0.065114 | -0.0025315 | -0.0021939 | 0.0055691 | -0.064087 | 65.126 | -57.2433 | 64.9567 | -56.7764 | -57.4127 | -57.3349 | -0.13275 | 65.1306 | -57.1472 | 65.094 | -56.9321 | -57.2754 | -56.9733 | -0.12817 | 0 | -0.16937 | -0.16937 | 0.33874 | 0.46692 | -0.21973 | -0.29755 | -0.091553 | -0.16937 | -0.032044 | 0.004577 | -0.032044 | -0.036621 | 0.1831 | 0.21515 | 0.14191 | -0.16022 | 0.17395 | -0.12817
1139.1719 | 122.3877 | -57.0648 | 65.126 | -0.054245 | 0.001128 | 0.070682 | 0.069987 | -0.0015231 | -0.00047423 | 0.00025641 | -0.071503 | 0.0014547 | 0.062357 | 0.065328 | -0.00018152 | -0.0028347 | -0.0010531 | -0.054932 | 65.1352 | -57.2388 | 64.9887 | -56.7307 | -57.3669 | -57.312 | -0.091553 | 65.1306 | -57.1335 | 65.1123 | -56.9 | -57.2525 | -56.9412 | -0.17395 | 0.009155 | -0.13733 | -0.14648 | 0.33417 | 0.50812 | -0.24719 | -0.30212 | -0.073242 | -0.12817 | -0.068665 | 0.004577 | -0.013733 | -0.01831 | 0.16479 | 0.23346 | 0.1236 | -0.18768 | 0.19226 | -0.11902
1140.1719 | 122.3648 | -57.0877 | 65.1443 | -0.084763 | 0.0017688 | 0.068545 | 0.068705 | -0.00088229 | -0.00090143 | -0.0018798 | -0.077606 | 0.0016685 | 0.064279 | 0.066182 | 0.00088648 | -0.0028347 | 0.0015103 | -0.036621 | 65.1535 | -57.2159 | 64.9841 | -56.7169 | -57.3669 | -57.2937 | -0.050354 | 65.1398 | -57.1198 | 65.1169 | -56.8771 | -57.2388 | -56.9366 | -0.12817 | 0.009155 | -0.16022 | -0.16937 | 0.37079 | 0.49896 | -0.20599 | -0.27923 | -0.07782 | -0.15106 | -0.032043 | -0.004577 | -0.027466 | -0.022889 | 0.21057 | 0.24262 | 0.15106 | -0.15106 | 0.1831 | -0.11902
1141.1875 | 122.4564 | -57.0282 | 65.1352 | -0.084763 | -0.0010082 | 0.069613 | 0.065928 | 0.0040311 | 0.0014484 | 0.00025641 | 0.062775 | 0.0010275 | 0.06492 | 0.05956 | 0.0011001 | -0.0043301 | -0.0010531 | -0.082397 | 65.0986 | -57.193 | 64.9933 | -56.6849 | -57.2845 | -57.2021 | -0.050354 | 65.0986 | -57.1335 | 65.1398 | -56.7947 | -57.1335 | -56.868 | -0.16479 | -0.036621 | -0.14191 | -0.10529 | 0.34332 | 0.50812 | -0.17395 | -0.25635 | -0.009155 | -0.091553 | -0.10529 | -0.036621 | 0.004578 | 0.041199 | 0.23346 | 0.33875 | 0.16022 | -0.10529 | 0.2655 | 0
1142.1719 | 122.5159 | -57.0328 | 65.1215 | -0.084763 | 5.9776e-05 | 0.0694 | 0.068919 | -0.0013095 | 0.00059397 | 0.0023926 | -0.022675 | -0.00046791 | 0.064279 | 0.06426 | -0.0025315 | -0.0032619 | -0.00041228 | -0.09613 | 65.1123 | -57.2067 | 65.0482 | -56.6711 | -57.312 | -57.2388 | -0.036621 | 65.094 | -57.1198 | 65.2084 | -56.7719 | -57.1014 | -56.8176 | -0.17395 | -0.009155 | -0.073242 | -0.064087 | 0.36163 | 0.53558 | -0.20599 | -0.27924 | -0.032044 | -0.10529 | -0.086975 | -0.027466 | 0.086975 | 0.11444 | 0.26093 | 0.3479 | 0.21515 | -0.068664 | 0.30212 | 0.018311
1143.1719 | 122.4335 | -57.0465 | 65.1535 | -0.13054 | 0.00048698 | 0.069827 | 0.071482 | -0.0028049 | -0.00047423 | -0.0014526 | -0.074554 | 0.0012411 | 0.06492 | 0.0696 | -0.00082252 | -0.0030483 | 0.0010831 | -0.018311 | 65.1489 | -57.1976 | 64.9658 | -56.662 | -57.2662 | -57.2754 | 0.022888 | 65.1535 | -57.0969 | 65.1672 | -56.8222 | -57.1793 | -56.8542 | -0.15106 | -0.004577 | -0.18768 | -0.18311 | 0.38452 | 0.53558 | -0.22888 | -0.21973 | -0.07782 | -0.068664 | -0.050354 | 0 | 0.013733 | 0.013733 | 0.2243 | 0.27466 | 0.19226 | -0.13275 | 0.24262 | -0.082397
1144.1719 | 122.4106 | -57.0648 | 65.181 | -0.031357 | 0.0013416 | 0.0694 | 0.068919 | 0.00039951 | 0.0031574 | 4.2808e-05 | -0.062347 | 0.00017309 | 0.067056 | 0.068105 | 0.0025955 | -0.0034757 | -0.0027621 | 0 | 65.1901 | -57.1976 | 65.0345 | -56.6528 | -57.2891 | -57.2433 | 0.032043 | 65.1535 | -57.1014 | 65.2039 | -56.7856 | -57.1793 | -56.8634 | -0.13275 | 0.009156 | -0.14648 | -0.15564 | 0.41199 | 0.54474 | -0.17853 | -0.22431 | -0.045776 | -0.091553 | -0.036621 | -0.027466 | 0.022888 | 0.050354 | 0.27924 | 0.31586 | 0.20142 | -0.11444 | 0.23804 | -0.07782
1145.1719 | 122.4472 | -57.0694 | 65.1672 | -0.084763 | -0.0014356 | 0.069613 | 0.068064 | 0.00039951 | -0.00047423 | 0.001538 | -0.053192 | 0.0033773 | 0.066202 | 0.070455 | -0.0010361 | -0.0021939 | 0.0012967 | -0.022888 | 65.1718 | -57.193 | 65.0436 | -56.6345 | -57.2891 | -57.2617 | 0.009155 | 65.1627 | -57.0694 | 65.1535 | -56.8314 | -57.1655 | -56.8771 | -0.1236 | 0.004578 | -0.1236 | -0.12817 | 0.43488 | 0.55847 | -0.19226 | -0.21973 | -0.068665 | -0.096131 | 0 | -0.004577 | -0.013733 | -0.009156 | 0.23804 | 0.23804 | 0.19226 | -0.09613 | 0.19226 | -0.09613
1146.1875 | 122.4564 | -57.0145 | 65.1672 | -0.16106 | -0.00058102 | 0.068545 | 0.068492 | -0.0015231 | 0.0020892 | -0.0025208 | 0.068878 | -0.0017495 | 0.062997 | 0.061269 | 0.00045928 | -0.0060391 | 0.0036465 | -0.064087 | 65.1443 | -57.1655 | 65.0436 | -56.6345 | -57.2845 | -57.193 | -0.027466 | 65.1123 | -57.0786 | 65.2451 | -56.749 | -57.1106 | -56.8222 | -0.15106 | -0.022888 | -0.1236 | -0.10071 | 0.37994 | 0.53101 | -0.17853 | -0.27008 | -0.027466 | -0.11902 | -0.064087 | -0.054931 | 0.07782 | 0.13275 | 0.2655 | 0.32959 | 0.19226 | -0.096131 | 0.25635 | -0.032044
1147.1719 | 122.4014 | -57.0419 | 65.1993 | -0.084763 | 0.0013416 | 0.070468 | 0.069133 | -0.00088229 | 0.0023028 | -0.0014526 | -0.065399 | 0.0014547 | 0.06492 | 0.067464 | -0.00039512 | -0.0017667 | -0.0042575 | -0.10529 | 65.1627 | -57.2067 | 64.9841 | -56.7398 | -57.3669 | -57.3029 | -0.27466 | 65.126 | -57.1381 | 65.1123 | -56.7856 | -57.225 | -56.9138 | -0.1648 | -0.036621 | -0.21515 | -0.17853 | 0.30212 | 0.46692 | -0.26093 | -0.32501 | -0.09613 | -0.16022 | -0.096131 | -0.073242 | -0.086975 | -0.013733 | 0.25635 | 0.35248 | 0.12817 | -0.18311 | 0.22431 | -0.086975
1148.1719 | 122.438 | -57.0236 | 65.1901 | -0.023727 | 0.0015552 | 0.068759 | 0.068705 | 0.0014675 | 0.0020892 | 0.0019654 | -0.053192 | 0.0018821 | 0.065988 | 0.066396 | 0.0013139 | -0.0021939 | -0.0016941 | 0 | 65.2039 | -57.1747 | 65.062 | -56.6162 | -57.2433 | -57.2159 | 0.036621 | 65.1901 | -57.0602 | 65.1947 | -56.7627 | -57.1152 | -56.8405 | -0.15106 | 0.013732 | -0.12817 | -0.14191 | 0.40741 | 0.55847 | -0.19226 | -0.21973 | -0.041198 | -0.068664 | -0.036621 | 0 | 0.004577 | 0.004577 | 0.26093 | 0.29755 | 0.18311 | -0.091552 | 0.21973 | -0.054931
1149.1719 | 122.4472 | -57.0145 | 65.2084 | -0.016098 | -0.00058102 | 0.0694 | 0.069987 | 0.0018949 | 0.0014484 | -0.001239 | -0.010467 | 0.0018821 | 0.065561 | 0.070028 | 0.00088648 | -0.0043301 | -0.0014805 | 0 | 65.2222 | -57.1701 | 65.0665 | -56.6254 | -57.2296 | -57.193 | 0.032043 | 65.2084 | -57.0557 | 65.2039 | -56.7398 | -57.106 | -56.8405 | -0.15564 | 0.013733 | -0.14191 | -0.15564 | 0.3891 | 0.54474 | -0.17853 | -0.21515 | -0.022888 | -0.059509 | -0.041199 | 0 | -0.004578 | -0.004578 | 0.27466 | 0.31586 | 0.17395 | -0.091553 | 0.21515 | -0.050354
1150.1719 | 122.4335 | -57.0053 | 65.2222 | -0.1458 | 0.0013416 | 0.068759 | 0.069346 | 0.0018949 | 0.0010212 | -0.00017099 | -0.025726 | 0.0018821 | 0.067484 | 0.062551 | 0.0011001 | -0.0039029 | -0.00019868 | -0.009155 | 65.1947 | -57.1884 | 65.0528 | -56.6254 | -57.2388 | -57.1976 | 0.013733 | 65.2176 | -57.0602 | 65.2176 | -56.7398 | -57.0923 | -56.8268 | -0.18311 | -0.027466 | -0.16937 | -0.14191 | 0.37994 | 0.56305 | -0.19226 | -0.23346 | -0.009155 | -0.050354 | -0.054932 | -0.004578 | -0.004578 | 0 | 0.2655 | 0.32044 | 0.17853 | -0.086975 | 0.23346 | -0.032043
1151.1719 | 122.406 | -56.9916 | 65.2451 | -0.023727 | 5.9776e-05 | 0.066836 | 0.068492 | 0.00018591 | 0.00016657 | -0.00059819 | -0.0043641 | 0.0012411 | 0.063852 | 0.062337 | -0.0012497 | -0.0051845 | 0.0012967 | -0.027466 | 65.2222 | -57.1518 | 65.0528 | -56.6574 | -57.2433 | -57.2067 | -0.004578 | 65.181 | -57.0282 | 65.181 | -56.7215 | -57.0877 | -56.8405 | -0.16022 | -0.022888 | -0.19226 | -0.16937 | 0.33417 | 0.49438 | -0.21515 | -0.25177 | -0.054932 | -0.091553 | -0.036621 | -0.064087 | -0.064087 | 0 | 0.27008 | 0.3067 | 0.15106 | -0.096131 | 0.18768 | -0.05951
1152.1719 | 122.374 | -57.0236 | 65.213 | -0.12291 | 0.001128 | 0.070041 | 0.070628 | -0.0015231 | 0.001662 | -0.002948 | -0.03183 | 0.00038669 | 0.064706 | 0.061056 | -0.00018152 | -0.0021939 | 0.0023647 | -0.064087 | 65.2267 | -57.1838 | 65.0299 | -56.6757 | -57.3074 | -57.2662 | -0.12817 | 65.1901 | -57.0557 | 65.1581 | -56.8222 | -57.1976 | -56.8863 | -0.16022 | 0.013733 | -0.18311 | -0.19684 | 0.3479 | 0.50812 | -0.24261 | -0.28381 | -0.082397 | -0.1236 | -0.032043 | -0.022888 | -0.054932 | -0.032044 | 0.20142 | 0.23346 | 0.13733 | -0.17395 | 0.16937 | -0.14191
1153.1719 | 122.4243 | -56.9778 | 65.2267 | -0.092392 | -0.00015382 | 0.068118 | 0.067851 | -0.0019503 | -0.00026063 | -0.0025208 | -0.0043641 | 0.0010275 | 0.063852 | 0.062551 | 0.00088648 | -0.0034757 | 0.00065592 | -0.013733 | 65.2588 | -57.1289 | 65.0665 | -56.6345 | -57.2388 | -57.2021 | -0.004578 | 65.213 | -57.0282 | 65.2222 | -56.7307 | -57.0831 | -56.8176 | -0.15106 | 0.032043 | -0.16022 | -0.19226 | 0.34332 | 0.49439 | -0.2243 | -0.26093 | -0.073242 | -0.10986 | -0.050354 | -0.013733 | -0.004578 | 0.009155 | 0.24719 | 0.29755 | 0.16022 | -0.10529 | 0.21057 | -0.054932
1154.1719 | 122.4152 | -56.9595 | 65.2451 | -0.092392 | 0.00027338 | 0.068973 | 0.068492 | -0.0019503 | 0.001662 | -0.0014526 | -0.019623 | 0.0027365 | 0.066202 | 0.068532 | -0.00039512 | -0.0028347 | -0.00062588 | 0 | 65.2451 | -57.1152 | 65.1077 | -56.5842 | -57.2159 | -57.1655 | 0.032043 | 65.2084 | -57.0236 | 65.2405 | -56.7123 | -57.0694 | -56.8176 | -0.15564 | 0 | -0.13733 | -0.13733 | 0.37537 | 0.53101 | -0.20599 | -0.25635 | -0.050354 | -0.10071 | -0.064087 | -0.036621 | -0.004577 | 0.032044 | 0.24719 | 0.31128 | 0.14191 | -0.10986 | 0.20599 | -0.045776
1155.1719 | 122.3923 | -56.9824 | 65.2725 | -0.038986 | -0.00058102 | 0.067264 | 0.068064 | -0.0043003 | 0.0048664 | 0.00089721 | -0.019623 | 0.0016685 | 0.068124 | 0.060628 | 0.0019547 | -0.0049709 | 0.0012967 | -0.027466 | 65.2313 | -57.1014 | 64.9796 | -56.7719 | -57.225 | -57.1747 | 0.009155 | 65.2405 | -57.0145 | 65.1947 | -56.7261 | -57.0786 | -56.813 | -0.11902 | -0.041199 | -0.29297 | -0.25177 | 0.21057 | 0.32959 | -0.19226 | -0.24262 | -0.073243 | -0.1236 | -0.032043 | -0.032043 | -0.07782 | -0.045777 | 0.25635 | 0.28839 | 0.16937 | -0.09613 | 0.20142 | -0.064087
1156.1719 | 122.3694 | -56.9824 | 65.3046 | -0.084763 | 0.0015552 | 0.069613 | 0.066569 | 0.0018949 | 0.001662 | 0.0013244 | 0.02005 | 0.00060029 | 0.067911 | 0.066182 | -0.00039512 | -0.0053983 | 0.00065592 | -0.022888 | 65.2496 | -57.0923 | 65.0711 | -56.5796 | -57.2159 | -57.1655 | 0 | 65.2405 | -56.9778 | 65.2359 | -56.9046 | -57.1518 | -56.8222 | -0.10986 | -0.054931 | -0.23346 | -0.17853 | 0.40283 | 0.51269 | -0.1831 | -0.23346 | -0.073242 | -0.1236 | 0.004578 | -0.064086 | -0.068664 | -0.004578 | 0.07782 | 0.073242 | 0.16022 | -0.16937 | 0.15564 | -0.17395
1157.1719 | 122.3923 | -56.9641 | 65.3 | -0.10765 | 5.9776e-05 | 0.068973 | 0.066783 | -0.0023777 | -0.00026063 | -0.00038459 | 0.044464 | 0.00038669 | 0.064706 | 0.061056 | 0.00024568 | -0.0045437 | -0.00041228 | -0.013733 | 65.2771 | -57.0831 | 65.0986 | -56.5659 | -57.2021 | -57.1564 | 0.027466 | 65.2679 | -57.0053 | 65.2084 | -56.7307 | -57.0511 | -56.8176 | -0.11902 | -0.022888 | -0.20142 | -0.17853 | 0.39825 | 0.51727 | -0.19226 | -0.23804 | -0.073242 | -0.11902 | -0.041199 | -0.032044 | -0.091553 | -0.059509 | 0.23346 | 0.27466 | 0.14648 | -0.086975 | 0.18768 | -0.045776
1158.1719 | 122.3923 | -56.9595 | 65.2817 | -0.046616 | -0.00015382 | 0.067477 | 0.068705 | 0.0010403 | -0.004533 | -0.00017099 | 0.013947 | 0.0020957 | 0.066843 | 0.062551 | -0.00039512 | -0.0034757 | 0.00022852 | 0.013733 | 65.2451 | -57.074 | 65.1443 | -56.5292 | -57.1655 | -57.1106 | 0.045776 | 65.2542 | -56.9962 | 65.1993 | -56.7169 | -57.0602 | -56.8039 | -0.11444 | -0.036621 | -0.13733 | -0.10071 | 0.4303 | 0.54474 | -0.15106 | -0.20599 | -0.036621 | -0.091552 | -0.036621 | -0.027466 | -0.082397 | -0.054931 | 0.24262 | 0.27924 | 0.15564 | -0.10071 | 0.19226 | -0.064087
1159.1719 | 122.3923 | -56.9412 | 65.3229 | -0.10765 | 0.00091438 | 0.067691 | 0.068705 | -0.0025913 | -0.00090143 | 0.0011108 | -0.010467 | 0.0018821 | 0.066202 | 0.068532 | -0.0012497 | -0.0053983 | -0.0010531 | -0.022888 | 65.2771 | -57.0831 | 65.094 | -56.5704 | -57.2021 | -57.1426 | -0.009155 | 65.2634 | -56.9824 | 65.2084 | -56.7123 | -57.0694 | -56.8176 | -0.14191 | -0.045776 | -0.22888 | -0.18311 | 0.37079 | 0.51269 | -0.20142 | -0.26093 | -0.059509 | -0.11902 | -0.041199 | -0.059509 | -0.11444 | -0.054932 | 0.22888 | 0.27008 | 0.1236 | -0.12817 | 0.16479 | -0.086975
1160.1719 | 122.3877 | -56.9778 | 65.3137 | -0.077133 | 5.9776e-05 | 0.070041 | 0.068705 | -2.7692e-05 | 0.00038017 | -0.00081179 | 0.032257 | 0.0012411 | 0.064279 | 0.064473 | 0.00045928 | -0.0056119 | 1.4916e-05 | 0.018311 | 65.2725 | -57.0831 | 65.126 | -56.5567 | -57.1793 | -57.0923 | 0.027466 | 65.2542 | -56.9733 | 65.2496 | -56.6986 | -57.0557 | -56.8039 | -0.10529 | -0.041199 | -0.18768 | -0.14648 | 0.42114 | 0.52643 | -0.11444 | -0.20142 | -0.009155 | -0.09613 | 0.004577 | -0.05951 | -0.064087 | -0.004577 | 0.27924 | 0.27466 | 0.17395 | -0.07782 | 0.16937 | -0.082397
1161.1719 | 122.4197 | -56.9183 | 65.3137 | -0.054245 | -0.00015382 | 0.067264 | 0.067851 | 0.0029629 | -0.0021832 | 0.00047001 | 0.029205 | 0.0010275 | 0.065347 | 0.060415 | -0.0014633 | -0.0045437 | -0.00041228 | 0.018311 | 65.2588 | -57.0694 | 65.1764 | -56.488 | -57.1243 | -57.074 | 0.09613 | 65.2725 | -56.9733 | 65.2542 | -56.6849 | -57.0145 | -56.7764 | -0.15106 | -0.054932 | -0.13733 | -0.082397 | 0.4303 | 0.58136 | -0.15564 | -0.20599 | -0.004578 | -0.054932 | -0.054932 | -0.041199 | -0.05951 | -0.018311 | 0.23346 | 0.28839 | 0.14191 | -0.09613 | 0.19684 | -0.041198
1162.1719 | 122.4335 | -56.9183 | 65.2725 | -0.092392 | 0.0015552 | 0.067904 | 0.06721 | -0.0015231 | -0.00047423 | -0.0010254 | 0.016998 | 0.00038669 | 0.065134 | 0.05721 | -0.0027451 | -0.0047573 | -0.0023349 | 0.036621 | 65.2359 | -57.0511 | 65.2039 | -56.4514 | -57.0786 | -57.0419 | 0.091553 | 65.2725 | -56.9412 | 65.3 | -56.6345 | -56.987 | -56.7307 | -0.13275 | -0.036621 | -0.068665 | -0.032044 | 0.46692 | 0.59967 | -0.1236 | -0.16022 | 0.009155 | -0.027466 | -0.022888 | 0 | 0.027466 | 0.027466 | 0.28381 | 0.3067 | 0.18768 | -0.068665 | 0.21057 | -0.045777
1163.1719 | 122.5021 | -56.868 | 65.3 | -0.046616 | 0.0015552 | 0.0694 | 0.068492 | 0.0023221 | 0.00080757 | -0.0031616 | 0.010895 | -0.00025411 | 0.06727 | 0.060628 | -0.0010361 | -0.0058255 | -0.00062588 | 0.018311 | 65.2634 | -57.0419 | 65.2267 | -56.4468 | -57.0831 | -57.019 | 0.11444 | 65.2496 | -56.955 | 65.3183 | -56.5933 | -56.9366 | -56.6849 | -0.17395 | -0.036621 | -0.073242 | -0.036621 | 0.42114 | 0.59509 | -0.15106 | -0.21515 | 0.022888 | -0.041199 | -0.086975 | -0.050354 | 0.01831 | 0.068664 | 0.27466 | 0.36163 | 0.18311 | -0.068665 | 0.27008 | 0.01831
1164.1719 | 122.4747 | -56.8863 | 65.3091 | -0.077133 | -0.00015382 | 0.065555 | 0.068064 | 0.0031765 | 0.0046528 | -0.0018798 | 0.023102 | 0.00038669 | 0.066843 | 0.063192 | -0.0029587 | -0.0021939 | -0.0021213 | 0.29755 | 65.3229 | -57.0236 | 65.2176 | -56.3919 | -57.0648 | -57.019 | 0.13733 | 65.2679 | -56.9458 | 65.3549 | -56.5613 | -56.9321 | -56.6574 | -0.13733 | 0.013733 | -0.091553 | -0.10529 | 0.49438 | 0.63171 | -0.13275 | -0.17853 | 0.004578 | -0.041198 | -0.059509 | -0.041199 | 0.045776 | 0.086975 | 0.32501 | 0.38452 | 0.22888 | -0.045776 | 0.28839 | 0.013733
1165.1719 | 122.5067 | -56.8726 | 65.2954 | -0.13817 | 0.00091438 | 0.0694 | 0.067637 | -0.0034457 | 0.00038017 | -0.00081179 | 0.084137 | 0.0018821 | 0.067056 | 0.058492 | -0.00039512 | -0.0034757 | -0.0027621 | 0.041199 | 65.2588 | -57.0328 | 65.2588 | -56.3873 | -57.0465 | -56.7719 | -0.086975 | 65.2267 | -56.9412 | 65.3778 | -56.5796 | -56.8863 | -56.662 | -0.16022 | -0.036621 | -0.036621 | 0 | 0.48523 | 0.64545 | 0.10071 | -0.17395 | 0.26093 | -0.013733 | -0.068664 | -0.068664 | 0.082398 | 0.15106 | 0.29297 | 0.36163 | 0.21057 | -0.013733 | 0.27924 | 0.054931
1166.1719 | 122.4701 | -56.8909 | 65.3183 | -0.038986 | 0.00070078 | 0.067904 | 0.068919 | 0.0010403 | 0.00038017 | 4.2808e-05 | 0.0078429 | 0.00081389 | 0.066415 | 0.06896 | -0.0025315 | -0.0045437 | -0.0044711 | 0.027466 | 65.2908 | -57.0465 | 65.2634 | -56.3965 | -57.0694 | -57.019 | 0.11444 | 65.2542 | -56.9321 | 65.3366 | -56.5659 | -56.9138 | -56.6711 | -0.15564 | -0.027465 | -0.054931 | -0.027466 | 0.49439 | 0.65002 | -0.12817 | -0.17853 | 0.027466 | -0.022888 | -0.041199 | -0.064087 | 0.018311 | 0.082398 | 0.32501 | 0.36621 | 0.21973 | -0.022888 | 0.26093 | 0.018311
1167.1719 | 122.3877 | -56.868 | 65.4053 | -0.038986 | 0.00048698 | 0.069186 | 0.069987 | 0.00061311 | 0.0012348 | 0.001538 | -0.028778 | 0.00038669 | 0.064493 | 0.061056 | 3.2078e-05 | -0.0049709 | -0.00019868 | -0.009155 | 65.3412 | -57.0602 | 65.1535 | -56.5292 | -57.1564 | -57.1106 | 0.013733 | 65.3366 | -56.9321 | 65.2817 | -56.6528 | -57.019 | -56.7581 | -0.19226 | -0.064086 | -0.25177 | -0.18768 | 0.33875 | 0.53101 | -0.24262 | -0.28839 | -0.050354 | -0.09613 | -0.064087 | -0.068664 | -0.1236 | -0.054932 | 0.21515 | 0.27924 | 0.10986 | -0.15106 | 0.17395 | -0.086975
1168.1719 | 122.5067 | -56.868 | 65.332 | -0.13054 | 0.00048698 | 0.066623 | 0.06721 | -0.0034457 | 0.00038017 | 4.2808e-05 | 0.1879 | 0.00081389 | 0.068124 | 0.065755 | -0.00039512 | -0.0021939 | -0.0040439 | 0.027466 | 65.2634 | -57.0007 | 65.2954 | -56.3599 | -57.0099 | -56.9733 | 0.059509 | 65.3091 | -56.9092 | 65.3732 | -56.543 | -56.8863 | -56.6483 | -0.13275 | -0.068664 | -0.036621 | 0.032043 | 0.50812 | 0.64087 | -0.10529 | -0.14191 | 0.027465 | -0.009156 | -0.041199 | -0.022888 | 0.041199 | 0.064087 | 0.32501 | 0.36621 | 0.21973 | -0.018311 | 0.26093 | 0.022888
1169.1719 | 122.4564 | -56.8451 | 65.2954 | -0.13817 | 5.9776e-05 | 0.066836 | 0.065928 | 0.0010403 | 0.00059397 | 0.00047001 | 0.032257 | 0.0014547 | 0.065561 | 0.061483 | 0.00024568 | -0.0049709 | 0.0034329 | 0.036621 | 65.2954 | -57.0053 | 65.3183 | -56.3416 | -57.0099 | -56.9504 | 0.14191 | 65.2771 | -56.9046 | 65.419 | -56.5613 | -56.6208 | -56.4285 | -0.16022 | 0 | 0.022888 | 0.022888 | 0.50354 | 0.66376 | -0.10529 | -0.16479 | 0.054932 | -0.004578 | -0.059509 | -0.01831 | 0.1236 | 0.14191 | 0.28381 | 0.34332 | 0.41656 | 0.2243 | 0.47607 | 0.28381
1170.1719 | 122.5113 | -56.8314 | 65.3046 | -0.10765 | 0.00070078 | 0.068332 | 0.067637 | -0.0013095 | 0.0012348 | 0.00068361 | 0.035309 | -0.00068151 | 0.06492 | 0.065755 | -0.0027451 | -0.0034757 | -0.0027621 | 0.009155 | 65.3137 | -57.0145 | 65.3 | -56.337 | -56.9962 | -56.9595 | 0.16479 | 65.2908 | -56.8909 | 65.4236 | -56.5018 | -56.8634 | -56.6116 | -0.18311 | 0.009156 | -0.004577 | -0.013733 | 0.49438 | 0.67749 | -0.12817 | -0.1648 | 0.054931 | 0.01831 | -0.059509 | -0.013732 | 0.11902 | 0.13275 | 0.32959 | 0.3891 | 0.21973 | -0.032043 | 0.27924 | 0.027466
1171.1719 | 122.5479 | -56.8405 | 65.3046 | -0.069504 | 0.00070078 | 0.067477 | 0.068064 | -2.7692e-05 | 0.00016657 | 0.00089721 | 0.0047913 | 0.00038669 | 0.063211 | 0.063619 | 3.2078e-05 | -0.0039029 | 0.0027921 | 0.013733 | 65.3183 | -57.0145 | 65.3046 | -56.3644 | -57.0236 | -56.9641 | 0.14648 | 65.2863 | -56.8771 | 65.419 | -56.488 | -56.8451 | -56.6071 | -0.17395 | 0.013733 | 0 | -0.013733 | 0.47607 | 0.65002 | -0.1236 | -0.18311 | 0.050354 | -0.009156 | -0.036621 | -0.01831 | 0.11444 | 0.13275 | 0.35248 | 0.3891 | 0.23346 | -0.004578 | 0.27008 | 0.032043
1172.1719 | 122.4014 | -56.7993 | 65.3091 | -0.069504 | -0.0018628 | 0.0694 | 0.068705 | -0.0019503 | 0.0010212 | -0.0010254 | 0.023102 | 0.00038669 | 0.062997 | 0.068532 | -0.0027451 | -0.0041165 | -0.00041228 | 0.036621 | 65.3229 | -56.9595 | 65.3229 | -56.3187 | -56.9962 | -56.9412 | 0.15106 | 65.2908 | -56.8909 | 65.4373 | -56.4926 | -56.5704 | -56.3919 | -0.16022 | 0.013733 | 0.013733 | 0 | 0.48065 | 0.64087 | -0.14191 | -0.19684 | 0.018311 | -0.036621 | -0.091553 | -0.01831 | 0.12817 | 0.14648 | 0.3067 | 0.39825 | 0.40741 | 0.22888 | 0.49896 | 0.32043
1173.1719 | 122.5571 | -56.8085 | 65.3458 | -0.084763 | 5.9776e-05 | 0.068759 | 0.066355 | -0.0079319 | -0.001115 | -0.001239 | 0.013947 | 0.0016685 | 0.065347 | 0.061483 | -0.00082252 | -0.0051845 | -0.0019077 | 0.018311 | 65.3091 | -56.9824 | 65.3275 | -56.3278 | -56.9962 | -56.9275 | 0.17853 | 65.2908 | -56.8817 | 65.4556 | -56.4697 | -56.8268 | -56.5979 | -0.17395 | -0.036621 | -0.01831 | 0.018311 | 0.48065 | 0.6546 | -0.11902 | -0.18768 | 0.054932 | -0.013733 | -0.073242 | -0.054931 | 0.10986 | 0.16479 | 0.33875 | 0.41199 | 0.21057 | -0.01831 | 0.28381 | 0.054932
1174.1719 | 122.4976 | -56.8085 | 65.3275 | -0.10002 | 5.9776e-05 | 0.06705 | 0.06956 | -0.0068637 | 0.00059397 | -0.0020936 | -0.0013123 | 0.0016685 | 0.064279 | 0.062337 | -0.0010361 | -0.0032619 | -0.0010531 | 0 | 65.3137 | -56.9824 | 65.3183 | -56.3416 | -57.0007 | -56.955 | 0.14648 | 65.3183 | -56.8863 | 65.451 | -56.4606 | -56.8222 | -56.5887 | -0.17395 | -0.013733 | -0.009156 | 0.004577 | 0.46692 | 0.64087 | -0.14648 | -0.19226 | 0.027466 | -0.01831 | -0.07782 | -0.009156 | 0.1236 | 0.13275 | 0.3479 | 0.42572 | 0.21973 | -0.013733 | 0.29755 | 0.064087
1175.1719 | 122.5113 | -56.813 | 65.3412 | -0.10765 | 0.00048698 | 0.067264 | 0.068705 | 0.00061311 | 0.001662 | -0.00038459 | 0.0078429 | 0.00060029 | 0.065347 | 0.062764 | 3.2078e-05 | -0.0039029 | 0.00044212 | -0.013733 | 65.3046 | -56.9824 | 65.2863 | -56.369 | -57.0236 | -56.9504 | 0.10529 | 65.3091 | -56.8771 | 65.4419 | -56.4743 | -56.8359 | -56.5979 | -0.16937 | -0.036622 | -0.054932 | -0.01831 | 0.44403 | 0.6134 | -0.13733 | -0.21057 | 0.032044 | -0.041199 | -0.064087 | -0.032044 | 0.10071 | 0.13275 | 0.33875 | 0.40283 | 0.21515 | -0.022888 | 0.27924 | 0.041199
1176.1719 | 122.4747 | -56.813 | 65.3412 | -0.11528 | 0.00048698 | 0.068759 | 0.068278 | -0.0025913 | -0.00090143 | 0.00068361 | -0.010467 | 0.00038669 | 0.065561 | 0.063192 | -0.00060892 | -0.0041165 | 0.0019375 | -0.041199 | 65.3412 | -56.9778 | 65.2817 | -56.3828 | -57.0145 | -56.9458 | 0.041199 | 65.3366 | -56.8726 | 65.4373 | -56.4606 | -56.8451 | -56.5979 | -0.16479 | 0 | -0.05951 | -0.05951 | 0.4303 | 0.59509 | -0.13275 | -0.20142 | 0.032043 | -0.036621 | -0.05951 | -0.004578 | 0.09613 | 0.10071 | 0.35248 | 0.41199 | 0.21515 | -0.032044 | 0.27466 | 0.027466
1177.1719 | 122.3602 | -56.8268 | 65.4327 | -0.038986 | 5.9776e-05 | 0.068545 | 0.067637 | 0.00018591 | 0.00038017 | -0.0027344 | -0.074554 | 0.0014547 | 0.063852 | 0.067251 | -0.00039512 | -0.0051845 | -0.00019868 | -0.004578 | 65.3687 | -56.9824 | 65.2496 | -56.4148 | -57.0648 | -56.9916 | 0.018311 | 65.3961 | -56.8588 | 65.3549 | -56.5567 | -56.9504 | -56.6895 | -0.15564 | -0.064087 | -0.1831 | -0.11902 | 0.41199 | 0.56763 | -0.16479 | -0.23804 | -0.009155 | -0.082397 | -0.032044 | -0.036621 | -0.07782 | -0.041199 | 0.27008 | 0.30212 | 0.13733 | -0.1236 | 0.16937 | -0.091552
1178.1719 | 122.438 | -56.7947 | 65.3641 | -0.12291 | 0.001128 | 0.067904 | 0.069346 | 0.0014675 | -0.001756 | -0.0018798 | -0.013519 | -0.00025411 | 0.065775 | 0.062551 | 0.00088648 | -0.0024075 | -0.0023349 | -0.036621 | 65.3275 | -56.9687 | 65.2863 | -56.4056 | -57.0694 | -56.9824 | 0.009155 | 65.3687 | -56.8726 | 65.4099 | -56.488 | -56.8817 | -56.6254 | -0.17395 | -0.036621 | -0.07782 | -0.041199 | 0.3891 | 0.56305 | -0.18768 | -0.27466 | -0.013733 | -0.10071 | -0.07782 | 0.004577 | 0.045776 | 0.041199 | 0.3067 | 0.38452 | 0.16937 | -0.086975 | 0.24719 | -0.009155
1179.1719 | 122.4655 | -56.7993 | 65.3549 | -0.13054 | 0.00048698 | 0.0694 | 0.069346 | 0.00018591 | 0.00038017 | 4.2808e-05 | -0.013519 | -0.0021769 | 0.065134 | 0.059774 | -0.0012497 | -0.0047573 | 0.00086952 | -0.059509 | 65.3412 | -56.955 | 65.2863 | -56.3736 | -57.0557 | -56.9733 | 0.1236 | 65.3503 | -56.8451 | 65.4053 | -56.4926 | -56.8863 | -56.6162 | -0.15564 | -0.013732 | -0.068664 | -0.054932 | 0.42572 | 0.58136 | -0.17395 | -0.25635 | -0.018311 | -0.10071 | -0.045777 | -0.004577 | 0.050354 | 0.054931 | 0.3067 | 0.35248 | 0.1831 | -0.086976 | 0.22888 | -0.041199
1180.1719 | 122.4289 | -56.8085 | 65.3641 | -0.13817 | -0.0016492 | 0.069827 | 0.068919 | 0.0038175 | 0.00080757 | 0.0013244 | -0.025726 | 0.00060029 | 0.065347 | 0.065755 | 0.00024568 | -0.0032619 | 0.0038601 | -0.059509 | 65.3778 | -56.9595 | 65.3 | -56.369 | -57.0282 | -56.9641 | 0.004578 | 65.3549 | -56.8405 | 65.4465 | -56.4835 | -56.8405 | -56.5842 | -0.15106 | 0.013733 | -0.064087 | -0.07782 | 0.43945 | 0.59051 | -0.15564 | -0.21973 | -0.004577 | -0.068664 | -0.032043 | -0.009156 | 0.082397 | 0.091553 | 0.32501 | 0.35706 | 0.22431 | -0.032043 | 0.25635 | 0
1181.1719 | 122.4838 | -56.7902 | 65.3732 | -0.12291 | -0.00079462 | 0.06705 | 0.069346 | -0.0010959 | 0.00059397 | -0.00038459 | -0.034882 | 0.00060029 | 0.065134 | 0.061483 | 0.00024568 | -0.0019803 | 0.00086952 | -0.054932 | 65.3687 | -56.955 | 65.3091 | -56.3873 | -57.0419 | -56.9687 | -0.009155 | 65.3641 | -56.8405 | 65.4465 | -56.4651 | -56.8405 | -56.5887 | -0.1648 | -0.004578 | -0.064087 | -0.059509 | 0.40283 | 0.56763 | -0.17853 | -0.25177 | -0.013733 | -0.086975 | -0.050354 | -0.009155 | 0.073242 | 0.082397 | 0.32501 | 0.37537 | 0.20142 | -0.050354 | 0.25177 | 0
1182.1719 | 122.3602 | -56.8268 | 65.4282 | -0.0084687 | -0.0014356 | 0.068118 | 0.070628 | 0.00018591 | 0.0010212 | -0.0010254 | -0.077606 | 0.0031637 | 0.065347 | 0.067037 | 0.00024568 | -0.0021939 | 0.00044212 | -0.027466 | 65.419 | -56.9595 | 65.2771 | -56.3873 | -57.0465 | -56.987 | 0.009155 | 65.4465 | -56.8222 | 65.3824 | -56.5338 | -56.9092 | -56.662 | -0.13275 | -0.009156 | -0.15106 | -0.14191 | 0.43945 | 0.5722 | -0.16022 | -0.21973 | -0.027466 | -0.086975 | 0.004577 | 0.01831 | -0.045777 | -0.064087 | 0.29297 | 0.28839 | 0.16479 | -0.082398 | 0.16022 | -0.086975
1183.1719 | 122.4884 | -56.7902 | 65.4099 | -0.10765 | -0.00058102 | 0.068973 | 0.068278 | -0.0015231 | 0.00016657 | -0.0018798 | -0.010467 | -0.00068151 | 0.067484 | 0.06426 | 0.00024568 | -0.0056119 | 0.00022852 | -0.068665 | 65.3778 | -56.9321 | 65.2954 | -56.337 | -57.0282 | -56.955 | -0.027466 | 65.3687 | -56.8314 | 65.4785 | -56.4194 | -56.8542 | -56.5613 | -0.14191 | -0.032043 | -0.11444 | -0.082398 | 0.45319 | 0.59509 | -0.1648 | -0.23804 | -0.022888 | -0.09613 | -0.041199 | -0.041199 | 0.068665 | 0.10986 | 0.37079 | 0.41199 | 0.22888 | -0.064087 | 0.27008 | -0.022888
1184.1719 | 122.4426 | -56.7947 | 65.3915 | -0.12291 | 0.00027338 | 0.069186 | 0.069133 | -0.0025913 | 0.0027302 | 0.0011108 | -0.0074159 | 0.0020957 | 0.064493 | 0.062337 | 0.00067288 | -0.0036893 | 0.0015103 | -0.082397 | 65.3778 | -56.9458 | 65.3137 | -56.3507 | -57.0053 | -56.9504 | -0.041199 | 65.387 | -56.8268 | 65.4648 | -56.4743 | -56.8405 | -56.5842 | -0.15106 | -0.013733 | -0.07782 | -0.064087 | 0.44403 | 0.59509 | -0.15564 | -0.21057 | -0.004577 | -0.059509 | -0.032043 | -0.004578 | 0.073242 | 0.07782 | 0.32044 | 0.35248 | 0.21057 | -0.045776 | 0.24262 | -0.013733
1185.1719 | 122.4701 | -56.7535 | 65.4144 | -0.15343 | 0.001128 | 0.068118 | 0.069346 | 0.0014675 | 0.0014484 | -0.002948 | -0.044037 | 0.0018821 | 0.063852 | 0.065542 | -0.0014633 | -0.0043301 | -0.00062588 | -0.073242 | 65.4053 | -56.9138 | 65.3229 | -56.3278 | -57.0007 | -56.9366 | -0.041199 | 65.3732 | -56.8176 | 65.4648 | -56.4743 | -56.8268 | -56.575 | -0.16022 | -0.009156 | -0.091553 | -0.082397 | 0.42572 | 0.58594 | -0.18311 | -0.24719 | -0.022889 | -0.086975 | -0.064087 | -0.041199 | 0.050354 | 0.091553 | 0.27924 | 0.34332 | 0.17853 | -0.073242 | 0.24262 | -0.009155
1186.1719 | 122.4243 | -56.7444 | 65.4144 | -0.092392 | 0.0024096 | 0.068545 | 0.070628 | -0.0023777 | 0.0018756 | 0.00047001 | -0.022675 | 0.00060029 | 0.064279 | 0.063405 | -0.0012497 | -0.0026211 | 0.0017239 | -0.086975 | 65.4053 | -56.8954 | 65.3275 | -56.3141 | -57.0099 | -56.9275 | -0.045776 | 65.3961 | -56.813 | 65.4785 | -56.456 | -56.8314 | -56.5521 | -0.15106 | -0.009156 | -0.086975 | -0.077819 | 0.4303 | 0.58136 | -0.1831 | -0.2655 | -0.032043 | -0.11444 | -0.068664 | -0.018311 | 0.064087 | 0.082398 | 0.28839 | 0.35706 | 0.19226 | -0.086975 | 0.26093 | -0.018311
1187.1719 | 122.4655 | -56.7215 | 65.419 | -0.11528 | 0.001128 | 0.068332 | 0.068705 | -0.0019503 | -0.0019696 | 0.00089721 | 0.18484 | -0.00068151 | 0.065988 | 0.070028 | 0.00088648 | -0.0047573 | -0.0019077 | 0.027466 | 65.4327 | -56.8726 | 65.419 | -56.1996 | -56.8771 | -56.8359 | 0.10986 | 65.4327 | -56.7719 | 65.5243 | -56.3507 | -56.7581 | -56.5201 | -0.15106 | 0.013733 | 0 | -0.013733 | 0.52185 | 0.67291 | -0.11444 | -0.15564 | 0.036622 | -0.004577 | -0.050354 | 0.013733 | 0.10529 | 0.091553 | 0.37079 | 0.42114 | 0.20142 | -0.036621 | 0.25177 | 0.013733
1188.1719 | 122.4609 | -56.7078 | 65.4327 | -0.061874 | 0.0015552 | 0.067477 | 0.065074 | -0.0021641 | -0.00026063 | 0.00068361 | 0.0078429 | -0.0013223 | 0.062784 | 0.060628 | -0.0021041 | -0.0036893 | -0.0053257 | -0.054932 | 65.4327 | -56.9092 | 65.3641 | -56.2912 | -56.9504 | -56.8863 | -0.004578 | 65.4007 | -56.781 | 65.5151 | -56.4056 | -56.781 | -56.5338 | -0.20142 | 0 | -0.068664 | -0.068664 | 0.41656 | 0.61798 | -0.17853 | -0.24261 | 0.022888 | -0.041198 | -0.073242 | -0.032043 | 0.082398 | 0.11444 | 0.30212 | 0.37537 | 0.17395 | -0.073242 | 0.24719 | 0
1189.1719 | 122.4335 | -56.7398 | 65.4465 | -0.10002 | -0.00229 | 0.069186 | 0.06956 | 0.00082671 | 0.0025166 | 0.0013244 | -0.028778 | 0.0020957 | 0.065561 | 0.06661 | -0.00039512 | -0.0024075 | -0.00019868 | -0.07782 | 65.4236 | -56.868 | 65.3595 | -56.2958 | -56.987 | -56.8954 | -0.032043 | 65.4099 | -56.7902 | 65.4831 | -56.4239 | -56.7993 | -56.5292 | -0.12817 | -0.022888 | -0.086975 | -0.064087 | 0.44403 | 0.57221 | -0.15564 | -0.24719 | -0.027466 | -0.11902 | -0.050354 | -0.036621 | 0.036621 | 0.073242 | 0.31586 | 0.36621 | 0.21057 | -0.059509 | 0.26093 | -0.009155
1190.1719 | 122.4152 | -56.6986 | 65.4282 | -0.077133 | 0.00091438 | 0.068973 | 0.06956 | 0.0018949 | 0.0023028 | 0.00068361 | -0.034882 | -4.0508e-05 | 0.063211 | 0.057851 | -0.00039512 | -0.0039029 | -0.0023349 | -0.09613 | 65.4236 | -56.8726 | 65.3641 | -56.3004 | -56.9733 | -56.9092 | -0.032043 | 65.4099 | -56.7673 | 65.5014 | -56.4148 | -56.8085 | -56.5338 | -0.17395 | -0.004578 | -0.064087 | -0.059509 | 0.39825 | 0.57221 | -0.21057 | -0.27466 | -0.036621 | -0.10071 | -0.068665 | -0.018311 | 0.073242 | 0.091553 | 0.28381 | 0.35248 | 0.16479 | -0.10986 | 0.23346 | -0.041199
1191.1719 | 122.4609 | -56.7123 | 65.4419 | -0.11528 | 5.9776e-05 | 0.070468 | 0.069346 | 0.0040311 | 0.0023028 | 0.0011108 | -0.044037 | -0.0013223 | 0.066843 | 0.073659 | -0.0014633 | -0.0064663 | 0.0036465 | -0.073242 | 65.4236 | -56.8634 | 65.3732 | -56.2546 | -56.9733 | -56.9 | -0.032043 | 65.4419 | -56.781 | 65.5197 | -56.4102 | -56.781 | -56.5109 | -0.15106 | -0.018311 | -0.068665 | -0.050354 | 0.45776 | 0.60882 | -0.18768 | -0.26093 | -0.036621 | -0.10986 | -0.068665 | 0 | 0.077819 | 0.077819 | 0.30212 | 0.37079 | 0.20142 | -0.068665 | 0.27008 | 0
1192.1719 | 122.3465 | -56.7215 | 65.4922 | -0.054245 | 0.0015552 | 0.068118 | 0.067637 | -0.0030185 | -0.00090143 | 0.001538 | -0.062347 | 0.0012411 | 0.065988 | 0.070028 | 0.00045928 | -0.0034757 | -0.0025485 | 0.013733 | 65.4877 | -56.8405 | 65.3824 | -56.2637 | -56.9412 | -56.8771 | 0.036621 | 65.5151 | -56.7352 | 65.4694 | -56.4194 | -56.8085 | -56.5567 | -0.11902 | -0.004578 | -0.10986 | -0.10529 | 0.45776 | 0.57678 | -0.15564 | -0.21973 | -0.036621 | -0.10071 | -0.013732 | 0.022888 | -0.022889 | -0.045777 | 0.30212 | 0.31586 | 0.16479 | -0.086975 | 0.17853 | -0.073243
1193.1719 | 122.4701 | -56.6895 | 65.4602 | -0.10002 | -0.00058102 | 0.070895 | 0.06956 | -0.0013095 | 0.0018756 | -0.00081179 | -0.044037 | -4.0508e-05 | 0.063852 | 0.062124 | 0.00024568 | -0.0045437 | -0.0038303 | -0.050354 | 65.4282 | -56.8542 | 65.419 | -56.1951 | -56.9138 | -56.8405 | 0.018311 | 65.4419 | -56.7673 | 65.5426 | -56.3965 | -56.7535 | -56.5109 | -0.16479 | -0.032043 | -0.041199 | -0.009156 | 0.49439 | 0.65918 | -0.15106 | -0.2243 | 0.013733 | -0.059509 | -0.07782 | -0.01831 | 0.082398 | 0.10071 | 0.29297 | 0.37079 | 0.17853 | -0.064087 | 0.25635 | 0.013733
1194.1719 | 122.4289 | -56.694 | 65.4831 | -0.13054 | 0.00091438 | 0.069827 | 0.068705 | 0.0040311 | 0.0027302 | -0.0014526 | -0.028778 | 0.0014547 | 0.065988 | 0.06426 | 0.00045928 | -0.0045437 | -0.0016941 | -0.082397 | 65.4465 | -56.8497 | 65.4099 | -56.2546 | -56.9458 | -56.868 | -0.073242 | 65.4602 | -56.7444 | 65.5334 | -56.3965 | -56.7627 | -56.5018 | -0.15564 | -0.036621 | -0.073242 | -0.036621 | 0.43945 | 0.59509 | -0.17395 | -0.25177 | -0.018311 | -0.096131 | -0.050354 | -0.022888 | 0.050354 | 0.073242 | 0.29755 | 0.3479 | 0.19226 | -0.068664 | 0.24262 | -0.01831
1195.1719 | 122.4564 | -56.662 | 65.4556 | -0.084763 | -0.001222 | 0.068332 | 0.070842 | 0.0010403 | 0.0029438 | 0.00047001 | -0.028778 | 0.0010275 | 0.06727 | 0.062551 | -0.0010361 | -0.0047573 | 0.0047147 | -0.07782 | 65.4785 | -56.8405 | 65.419 | -56.2363 | -56.9229 | -56.8588 | -0.018311 | 65.451 | -56.7215 | 65.5563 | -56.3828 | -56.781 | -56.5109 | -0.17853 | 0.022889 | -0.036621 | -0.05951 | 0.42572 | 0.60425 | -0.19684 | -0.26093 | -0.018311 | -0.082398 | -0.05951 | -0.004577 | 0.10071 | 0.10529 | 0.27924 | 0.33875 | 0.15106 | -0.11902 | 0.21057 | -0.059509
1196.1719 | 122.4472 | -56.6711 | 65.4831 | -0.11528 | 0.00070078 | 0.069827 | 0.066569 | -0.00088229 | 0.00016657 | -0.00017099 | -0.059296 | 0.00017309 | 0.062784 | 0.06661 | -0.00082252 | -0.0017667 | 0.0027921 | -0.082397 | 65.4602 | -56.8222 | 65.4327 | -56.1768 | -56.8954 | -56.8359 | -0.009155 | 65.4556 | -56.7307 | 65.5563 | -56.3599 | -56.7352 | -56.4835 | -0.15106 | -0.022888 | -0.050354 | -0.027466 | 0.49439 | 0.64545 | -0.16479 | -0.2243 | -0.013732 | -0.073242 | -0.059509 | -0.027466 | 0.073242 | 0.10071 | 0.31128 | 0.37079 | 0.18768 | -0.064086 | 0.24719 | -0.004577
1197.1719 | 122.4838 | -56.6666 | 65.4877 | -0.054245 | -0.0020764 | 0.068973 | 0.067851 | -0.0045139 | -0.0019696 | 4.2808e-05 | -0.013519 | 0.00017309 | 0.064279 | 0.065542 | -0.00060892 | -0.0036893 | -0.0014805 | -0.045776 | 65.5106 | -56.8176 | 65.4465 | -56.1905 | -56.8634 | -56.8085 | 0.036621 | 65.4739 | -56.694 | 65.6021 | -56.2912 | -56.6895 | -56.4285 | -0.15106 | 0.022888 | -0.041199 | -0.064087 | 0.47607 | 0.62714 | -0.14191 | -0.19684 | 0.009155 | -0.045776 | -0.027466 | -0.013733 | 0.11444 | 0.12817 | 0.37537 | 0.40283 | 0.23804 | -0.022888 | 0.2655 | 0.004578
1198.1719 | 122.4838 | -56.6895 | 65.4739 | -0.19157 | 0.002196 | 0.067904 | 0.070414 | 0.00018591 | 0.0020892 | 0.0028198 | -0.062347 | -0.00025411 | 0.062784 | 0.06191 | -0.00060892 | -0.0034757 | 0.0017239 | -0.054932 | 65.4694 | -56.813 | 65.451 | -56.1676 | -56.8542 | -56.7902 | 0.027466 | 65.4648 | -56.7032 | 65.5975 | -56.3324 | -56.7307 | -56.4285 | -0.1236 | -0.004578 | -0.022888 | -0.01831 | 0.52185 | 0.64545 | -0.10071 | -0.16479 | 0.022888 | -0.041199 | -0.013733 | -0.009155 | 0.1236 | 0.13275 | 0.35706 | 0.37079 | 0.26093 | -0.041199 | 0.27466 | -0.027466
1199.1719 | 122.5067 | -56.662 | 65.4877 | -0.15343 | 0.0019824 | 0.067904 | 0.068705 | 0.0018949 | 0.00038017 | 4.2808e-05 | -0.044037 | 0.00081389 | 0.065347 | 0.069173 | -0.0023177 | -0.0036893 | -0.0025485 | -0.045776 | 65.4694 | -56.8176 | 65.5655 | -55.9982 | -56.8863 | -56.8222 | 0.013733 | 65.4694 | -56.7078 | 65.6113 | -56.3049 | -56.7078 | -56.4239 | -0.15564 | -0.018311 | 0.07782 | 0.096131 | 0.66376 | 0.8194 | -0.16022 | -0.2243 | -0.004578 | -0.068665 | -0.045777 | -0.018311 | 0.1236 | 0.14191 | 0.35706 | 0.40283 | 0.23804 | -0.045777 | 0.28381 | 0
1200.1719 | 122.4747 | -56.6345 | 65.4831 | -0.10765 | 0.0026234 | 0.070041 | 0.069346 | 0.0010403 | 0.00080757 | -0.0014526 | -0.08371 | 0.0010275 | 0.065347 | 0.064687 | 0.00024568 | -0.00027134 | -0.00041228 | -0.050354 | 65.506 | -56.8039 | 65.4648 | -56.163 | -56.8542 | -56.813 | 0.032043 | 65.4648 | -56.7352 | 65.593 | -56.282 | -56.6849 | -56.4285 | -0.16937 | 0.022888 | -0.01831 | -0.041198 | 0.4715 | 0.64087 | -0.17853 | -0.21973 | -0.009155 | -0.050354 | -0.10071 | -0.01831 | 0.10986 | 0.12817 | 0.35248 | 0.45319 | 0.20599 | -0.050354 | 0.3067 | 0.050354
1201.1719 | 122.4884 | -56.6574 | 65.4968 | -0.16106 | -0.00015382 | 0.070895 | 0.069133 | -0.0023777 | 0.0010212 | 0.002179 | -0.056244 | 0.00060029 | 0.063425 | 0.061269 | 3.2078e-05 | -0.0013393 | -0.00041228 | -0.036621 | 65.506 | -56.7719 | 65.4648 | -56.131 | -56.8451 | -56.781 | 0.013733 | 65.4968 | -56.6986 | 65.6158 | -56.3141 | -56.6986 | -56.4148 | -0.11444 | 0.009155 | -0.032043 | -0.041198 | 0.52643 | 0.64087 | -0.1236 | -0.18768 | -0.009155 | -0.073242 | -0.041198 | 0 | 0.11902 | 0.11902 | 0.34332 | 0.38452 | 0.24262 | -0.041198 | 0.28381 | 0
1202.1875 | 122.3648 | -56.6299 | 65.5746 | -0.054245 | 0.00091438 | 0.06705 | 0.067637 | -0.0047275 | 0.0027302 | -0.0027344 | -0.05014 | 0.0010275 | 0.066629 | 0.06426 | -0.00018152 | -0.0039029 | 0.00086952 | 0.018311 | 65.5792 | -56.7719 | 65.4694 | -56.1539 | -56.8497 | -56.7947 | 0.013733 | 65.5701 | -56.6711 | 65.5472 | -56.3461 | -56.749 | -56.4926 | -0.14191 | 0.004578 | -0.10529 | -0.10986 | 0.47607 | 0.61798 | -0.16479 | -0.21973 | -0.022888 | -0.077819 | -0.041199 | -0.004578 | -0.027466 | -0.022888 | 0.28381 | 0.32501 | 0.13733 | -0.11902 | 0.17853 | -0.077819
1203.1719 | 122.4564 | -56.6345 | 65.538 | -0.12291 | 0.001128 | 0.070254 | 0.06956 | 0.0014675 | 0.0020892 | -0.00038459 | -0.03183 | 0.00060029 | 0.066202 | 0.064687 | 0.00024568 | -0.0030483 | -0.00062588 | -0.050354 | 65.5197 | -56.749 | 65.4831 | -56.1035 | -56.8405 | -56.7673 | 0.009155 | 65.4877 | -56.6849 | 65.6158 | -56.3095 | -56.6666 | -56.4194 | -0.11444 | -0.018311 | -0.054932 | -0.036621 | 0.53101 | 0.64545 | -0.13275 | -0.20599 | -0.018311 | -0.091553 | -0.050354 | -0.050354 | 0.07782 | 0.12817 | 0.32501 | 0.37537 | 0.21515 | -0.032044 | 0.2655 | 0.01831
1204.1719 | 122.4426 | -56.6299 | 65.4694 | -0.15343 | 0.0015552 | 0.067264 | 0.070201 | -0.0021641 | 0.0014484 | 0.001538 | -0.08371 | 0.0014547 | 0.063638 | 0.061696 | -0.00082252 | -0.0017667 | -0.0036167 | -0.064087 | 65.5197 | -56.781 | 65.506 | -56.1218 | -56.8359 | -56.7535 | 0.009155 | 65.538 | -56.6528 | 65.6616 | -56.2683 | -56.662 | -56.3736 | -0.15106 | 0.050354 | 0.036621 | -0.013733 | 0.50812 | 0.65918 | -0.1236 | -0.20599 | 0.027466 | -0.054931 | -0.022888 | 0.068665 | 0.19226 | 0.1236 | 0.36163 | 0.38452 | 0.25635 | -0.032043 | 0.27924 | -0.009155
1205.1719 | 122.4792 | -56.5887 | 65.5243 | -0.1458 | -0.00058102 | 0.069186 | 0.071055 | 0.00018591 | 0.001662 | 0.0026062 | -0.068451 | 0.0018821 | 0.064706 | 0.068746 | -0.0012497 | -0.0019803 | -0.0027621 | -0.064087 | 65.5426 | -56.7673 | 65.5151 | -56.1127 | -56.8268 | -56.7535 | -0.022888 | 65.5334 | -56.6574 | 65.6662 | -56.25 | -56.6437 | -56.3736 | -0.17853 | 0.018311 | -0.009155 | -0.027466 | 0.47607 | 0.6546 | -0.16479 | -0.23804 | 0.013733 | -0.059509 | -0.068665 | 0.009155 | 0.14191 | 0.13275 | 0.33875 | 0.40741 | 0.21515 | -0.054932 | 0.28381 | 0.013733
1206.1719 | 122.4609 | -56.5659 | 65.5518 | -0.11528 | -0.00079462 | 0.070254 | 0.070628 | -0.00024129 | 0.0012348 | 0.00025641 | -0.092865 | 0.0016685 | 0.06492 | 0.06661 | -0.00060892 | -0.0030483 | 0.0025785 | -0.09613 | 65.5563 | -56.7627 | 65.5151 | -56.131 | -56.8405 | -56.7764 | -0.018311 | 65.5426 | -56.6345 | 65.6662 | -56.2454 | -56.6528 | -56.337 | -0.19684 | 0.004577 | -0.036621 | -0.041198 | 0.43488 | 0.63171 | -0.21057 | -0.27466 | -0.013733 | -0.07782 | -0.068664 | -0.009155 | 0.11444 | 0.1236 | 0.32044 | 0.3891 | 0.22888 | -0.086975 | 0.29755 | -0.018311
1207.1875 | 122.3877 | -56.6345 | 65.5792 | -0.12291 | -0.00058102 | 0.068118 | 0.071482 | -0.0015231 | 0.0018756 | -0.00081179 | -0.13254 | 0.0014547 | 0.066843 | 0.069387 | -0.00039512 | -0.0024075 | -0.0016941 | -0.032043 | 65.5975 | -56.7719 | 65.4831 | -56.1447 | -56.8542 | -56.8085 | -0.022888 | 65.6067 | -56.662 | 65.6204 | -56.3141 | -56.7352 | -56.4239 | -0.13733 | 0.01831 | -0.096131 | -0.11444 | 0.48981 | 0.62714 | -0.17395 | -0.21973 | -0.036621 | -0.082397 | -0.027466 | 0.027465 | 0.041198 | 0.013733 | 0.32043 | 0.3479 | 0.21057 | -0.10071 | 0.23804 | -0.073242
1208.1719 | 122.4564 | -56.575 | 65.5655 | -0.038986 | 0.00091438 | 0.068759 | 0.065715 | 0.0033901 | 0.0031574 | 0.0013244 | -0.056244 | 0.00081389 | 0.064279 | 0.06896 | 0.00088648 | -0.0030483 | 0.00022852 | -0.054932 | 65.5655 | -56.7307 | 65.5289 | -56.0715 | -56.7993 | -56.7215 | -0.022888 | 65.5518 | -56.6437 | 65.6982 | -56.2134 | -56.6208 | -56.3278 | -0.15564 | 0 | -0.036621 | -0.036621 | 0.50354 | 0.65918 | -0.14648 | -0.2243 | 0.009155 | -0.068664 | -0.068665 | -0.013733 | 0.13275 | 0.14648 | 0.36163 | 0.4303 | 0.24719 | -0.045777 | 0.31586 | 0.022888
1209.1719 | 122.5113 | -56.5521 | 65.5701 | -0.15343 | -0.0014356 | 0.068332 | 0.067424 | 0.0010403 | 0.00038017 | -0.0014526 | -0.025726 | 0.00060029 | 0.06492 | 0.064687 | -0.0016769 | -0.0021939 | -0.0023349 | -0.091553 | 65.5701 | -56.7169 | 65.5426 | -56.0669 | -56.7993 | -56.7352 | -0.022888 | 65.5472 | -56.6254 | 65.6799 | -56.2225 | -56.6025 | -56.3232 | -0.16479 | 0 | -0.027465 | -0.027465 | 0.48523 | 0.65002 | -0.1831 | -0.24719 | -0.01831 | -0.082397 | -0.073242 | -0.022888 | 0.10986 | 0.13275 | 0.32959 | 0.40283 | 0.22888 | -0.050354 | 0.30212 | 0.022888
1210.1719 | 122.4564 | -56.5521 | 65.5792 | -0.10002 | 0.0017688 | 0.068332 | 0.068492 | 0.0014675 | 0.00080757 | -0.0023072 | -0.062347 | 0.0014547 | 0.066202 | 0.067891 | 3.2078e-05 | -0.0030483 | 0.00022852 | -0.10071 | 65.6067 | -56.7352 | 65.5106 | -56.0852 | -56.8085 | -56.7215 | -0.054932 | 65.5838 | -56.6071 | 65.7028 | -56.2088 | -56.6071 | -56.3278 | -0.1831 | 0.027465 | -0.068665 | -0.09613 | 0.46692 | 0.65002 | -0.16937 | -0.25635 | 0.013732 | -0.073243 | -0.054932 | 0.004577 | 0.1236 | 0.11902 | 0.34332 | 0.39825 | 0.2243 | -0.054932 | 0.27924 | 0
1211.1719 | 122.5067 | -56.5384 | 65.5792 | -0.10002 | 0.002837 | 0.068545 | 0.069346 | 0.0014675 | 0.001662 | -0.00017099 | -0.065399 | 0.0012411 | 0.065134 | 0.065328 | 3.2078e-05 | -0.0026211 | 0.00022852 | -0.10986 | 65.6113 | -56.7261 | 65.506 | -56.1218 | -56.8222 | -56.7535 | -0.086975 | 65.6067 | -56.6254 | 65.7028 | -56.218 | -56.6528 | -56.3416 | -0.18768 | 0.032043 | -0.073243 | -0.10529 | 0.41656 | 0.60425 | -0.21515 | -0.28381 | -0.027466 | -0.096131 | -0.086975 | 0.027465 | 0.1236 | 0.096131 | 0.32043 | 0.40741 | 0.19684 | -0.11444 | 0.28381 | -0.027466
1212.1719 | 122.4014 | -56.5384 | 65.6937 | 0.01442 | -0.00058102 | 0.068332 | 0.067424 | 0.0025357 | 0.00059397 | -0.0020936 | 0.0078429 | 0.0014547 | 0.063852 | 0.061696 | 0.00024568 | -0.0039029 | 0.0019375 | 0.027466 | 65.6754 | -56.6345 | 65.5655 | -56.0349 | -56.7307 | -56.6757 | 0.041199 | 65.6708 | -56.575 | 65.6708 | -56.1951 | -56.6208 | -56.3919 | -0.09613 | -0.018311 | -0.12817 | -0.10986 | 0.50354 | 0.59967 | -0.13733 | -0.19226 | -0.041199 | -0.096131 | -0.036621 | -0.022889 | -0.022889 | 0 | 0.34332 | 0.37994 | 0.14648 | -0.082398 | 0.1831 | -0.045777
1213.1719 | 122.4564 | -56.5201 | 65.6708 | -0.15343 | 0.0013416 | 0.065768 | 0.070842 | -2.7692e-05 | 0.0027302 | 0.00025641 | -0.065399 | 0.0020957 | 0.065347 | 0.062764 | -0.00060892 | -0.0011257 | -0.0010531 | -0.10986 | 65.6662 | -56.7078 | 65.5106 | -56.1218 | -56.8405 | -56.7581 | -0.10529 | 65.6387 | -56.5979 | 65.6616 | -56.2408 | -56.662 | -56.369 | -0.18768 | -0.004577 | -0.16022 | -0.15564 | 0.39825 | 0.58594 | -0.23804 | -0.32043 | -0.050354 | -0.13275 | -0.077819 | -0.032043 | -0.009155 | 0.022888 | 0.27924 | 0.35706 | 0.15106 | -0.14191 | 0.22888 | -0.064087
1214.1719 | 122.4014 | -56.5292 | 65.6799 | -0.061874 | 0.00027338 | 0.068545 | 0.069346 | -0.0017367 | 0.001662 | 0.00068361 | -0.065399 | 0.0029501 | 0.065347 | 0.063833 | 0.00024568 | -0.0030483 | 0.0040737 | -0.09613 | 65.6708 | -56.6895 | 65.4922 | -56.1218 | -56.8176 | -56.7581 | -0.11444 | 65.6708 | -56.5933 | 65.6754 | -56.2592 | -56.6757 | -56.369 | -0.16022 | -0.009156 | -0.18768 | -0.17853 | 0.40741 | 0.56763 | -0.22888 | -0.28839 | -0.068665 | -0.12817 | -0.064087 | -0.009156 | -0.004578 | 0.004578 | 0.27008 | 0.33417 | 0.16022 | -0.14648 | 0.2243 | -0.082397
1215.1719 | 122.493 | -56.5018 | 65.6525 | -0.13054 | -0.00079462 | 0.070041 | 0.069987 | 0.0025357 | 0.0035846 | 0.0011108 | -0.05014 | 0.00017309 | 0.06492 | 0.063192 | 0.00024568 | -0.0047573 | 0.0015103 | -0.10986 | 65.6662 | -56.6666 | 65.5334 | -56.0577 | -56.7856 | -56.7261 | -0.091553 | 65.6387 | -56.575 | 65.7349 | -56.1859 | -56.5887 | -56.3095 | -0.16479 | 0.013733 | -0.11902 | -0.13275 | 0.44403 | 0.60883 | -0.2243 | -0.28381 | -0.059509 | -0.11902 | -0.073242 | -0.013733 | 0.082397 | 0.09613 | 0.31586 | 0.3891 | 0.19226 | -0.086975 | 0.2655 | -0.013733
1216.1719 | 122.3877 | -56.4972 | 65.6662 | -0.061874 | -0.00058102 | 0.068332 | 0.067851 | 0.0018949 | 0.0012348 | -0.00059819 | -0.065399 | 0.0010275 | 0.065988 | 0.066823 | 3.2078e-05 | -0.0030483 | 0.00065592 | -0.10529 | 65.7028 | -56.6574 | 65.5472 | -56.0806 | -56.7902 | -56.7215 | -0.10529 | 65.6845 | -56.5475 | 65.7211 | -56.1859 | -56.5933 | -56.3141 | -0.16022 | 0.036621 | -0.11902 | -0.15564 | 0.41656 | 0.57678 | -0.22431 | -0.29297 | -0.064087 | -0.13275 | -0.050354 | 0.01831 | 0.054931 | 0.036621 | 0.31128 | 0.36163 | 0.1831 | -0.096131 | 0.23346 | -0.045777
1217.1719 | 122.4426 | -56.4606 | 65.6937 | -0.12291 | 0.00027338 | 0.070041 | 0.066142 | -0.0057955 | 0.0012348 | 4.2808e-05 | 0.016998 | -0.00046791 | 0.064066 | 0.058919 | 0.0011001 | -0.0036893 | 0.0023647 | -0.10071 | 65.712 | -56.6208 | 65.5838 | -56.012 | -56.7123 | -56.6254 | -0.045776 | 65.6891 | -56.5201 | 65.7486 | -56.1493 | -56.5201 | -56.282 | -0.16022 | 0.01831 | -0.10986 | -0.12817 | 0.44861 | 0.60883 | -0.16479 | -0.25177 | -0.004577 | -0.091552 | -0.05951 | -0.004578 | 0.054931 | 0.059509 | 0.31128 | 0.37079 | 0.17853 | -0.05951 | 0.23804 | 0
1218.1719 | 122.3969 | -56.4972 | 65.7211 | -0.32127 | 0.0013416 | 0.065982 | 0.069987 | 0.0042447 | 0.00016657 | -0.0010254 | -0.062347 | 0.0018821 | 0.067484 | 0.056997 | -0.00039512 | -0.0015529 | 0.0049283 | 0.004578 | 65.7074 | -56.6254 | 65.5334 | -56.0577 | -56.7719 | -56.7169 | -0.09613 | 65.712 | -56.5247 | 65.7074 | -56.1813 | -56.6116 | -56.3187 | -0.12817 | -0.013733 | -0.18768 | -0.17395 | 0.43945 | 0.56763 | -0.21973 | -0.27466 | -0.091553 | -0.14648 | -0.027466 | -0.009155 | -0.013733 | -0.004578 | 0.31586 | 0.34332 | 0.17853 | -0.11444 | 0.20599 | -0.086975
1219.1719 | 122.4426 | -56.4651 | 65.712 | -0.077133 | -0.00229 | 0.068332 | 0.069773 | -2.7692e-05 | 0.00080757 | -0.0016662 | -0.089813 | -0.00025411 | 0.068338 | 0.065328 | -0.0014633 | -0.0028347 | -0.0016941 | -0.11902 | 65.7486 | -56.6299 | 65.5838 | -56.0211 | -56.7444 | -56.6757 | -0.10529 | 65.7303 | -56.5109 | 65.7669 | -56.1493 | -56.5613 | -56.2775 | -0.1648 | 0.036621 | -0.12817 | -0.1648 | 0.44403 | 0.60883 | -0.21057 | -0.27924 | -0.045776 | -0.11444 | -0.045776 | 0.018311 | 0.054932 | 0.036621 | 0.31586 | 0.36163 | 0.18768 | -0.09613 | 0.23346 | -0.050354
1220.1719 | 122.3328 | -56.4423 | 65.7623 | -0.092392 | -0.00015382 | 0.068545 | 0.071269 | 0.0021085 | -0.00068783 | -0.0014526 | -0.092865 | -0.00025411 | 0.064279 | 0.063405 | -0.0010361 | -0.0043301 | 1.4916e-05 | -0.1236 | 65.7532 | -56.5933 | 65.5609 | -56.0577 | -56.7627 | -56.6895 | -0.16022 | 65.7532 | -56.5063 | 65.7486 | -56.2546 | -56.5338 | -56.25 | -0.15106 | -0.009155 | -0.20142 | -0.19226 | 0.38452 | 0.53558 | -0.24719 | -0.32043 | -0.09613 | -0.16937 | -0.064087 | -0.009155 | -0.013733 | -0.004578 | 0.18768 | 0.25177 | 0.19226 | -0.091552 | 0.25635 | -0.027465
1221.1719 | 122.3785 | -56.4606 | 65.7211 | -0.084763 | -0.0010082 | 0.068332 | 0.071482 | 0.0018949 | 0.0025166 | -0.00059819 | -0.074554 | 0.0014547 | 0.065347 | 0.065542 | 0.0011001 | -0.0028347 | 0.0010831 | -0.11444 | 65.7623 | -56.5979 | 65.5609 | -56.0669 | -56.7856 | -56.7032 | -0.13275 | 65.7486 | -56.488 | 65.744 | -56.1951 | -56.6025 | -56.3004 | -0.13733 | 0.041199 | -0.16022 | -0.20142 | 0.39368 | 0.53101 | -0.24262 | -0.32501 | -0.10529 | -0.18768 | -0.027466 | 0.027466 | 0.022889 | -0.004577 | 0.2655 | 0.29297 | 0.16022 | -0.14191 | 0.18768 | -0.11444
1222.1719 | 122.3831 | -56.4514 | 65.7761 | -0.13054 | -0.0010082 | 0.070041 | 0.066783 | -0.00088229 | 0.00059397 | -0.0010254 | -0.071503 | 0.0018821 | 0.067484 | 0.06191 | -0.00082252 | -0.0024075 | 0.0053555 | -0.11444 | 65.7669 | -56.6025 | 65.6982 | -56.0898 | -56.7215 | -56.5842 | -0.1236 | 65.7761 | -56.4835 | 65.7623 | -56.1493 | -56.5704 | -56.2775 | -0.15106 | -0.009155 | -0.07782 | -0.068665 | 0.36163 | 0.5127 | -0.13275 | -0.27008 | 0.018311 | -0.11902 | -0.032043 | 0 | -0.013733 | -0.013733 | 0.30212 | 0.33417 | 0.17395 | -0.11902 | 0.20599 | -0.086976
1223.1719 | 122.4106 | -56.4239 | 65.831 | -0.084763 | -0.00229 | 0.069613 | 0.071482 | 0.0033901 | 0.00016657 | 0.003247 | -0.062347 | 0.00060029 | 0.065561 | 0.064901 | 0.0013139 | -0.0041165 | 0.0012967 | -0.11902 | 65.7623 | -56.5704 | 65.5975 | -56.0074 | -56.7444 | -56.6574 | -0.23346 | 65.7532 | -56.5109 | 65.7394 | -56.0669 | -56.5704 | -56.2775 | -0.14649 | -0.068665 | -0.23346 | -0.16479 | 0.41656 | 0.56305 | -0.23346 | -0.32044 | -0.086975 | -0.17395 | -0.086975 | -0.07782 | -0.091553 | -0.013733 | 0.35705 | 0.44403 | 0.14648 | -0.14649 | 0.23346 | -0.05951
1224.1719 | 122.4518 | -56.4056 | 65.7669 | -0.046616 | 0.001128 | 0.068759 | 0.069133 | 0.0016813 | -0.00047423 | 0.0028198 | -0.062347 | 0.0014547 | 0.064493 | 0.062978 | 0.0015275 | -0.0026211 | 0.0012967 | -0.1236 | 65.8218 | -56.5475 | 65.6113 | -56.0211 | -56.7261 | -56.6711 | -0.13275 | 65.7761 | -56.4377 | 65.7806 | -56.1447 | -56.5659 | -56.2592 | -0.14191 | 0.054931 | -0.15564 | -0.21057 | 0.38452 | 0.52643 | -0.2655 | -0.32043 | -0.1236 | -0.17853 | -0.032043 | 0.009155 | 0.013733 | 0.004578 | 0.26093 | 0.29297 | 0.14648 | -0.16022 | 0.17853 | -0.12817
1225.1719 | 122.3785 | -56.3644 | 65.8356 | -0.16869 | -0.00058102 | 0.06705 | 0.068919 | -0.0013095 | 0.0012348 | -0.0010254 | -0.062347 | 0.0025229 | 0.066415 | 0.064473 | -0.00082252 | -0.0036893 | 0.0019375 | -0.091553 | 65.8264 | -56.5384 | 65.625 | -55.9616 | -56.694 | -56.6483 | -0.13733 | 65.8173 | -56.4423 | 65.7761 | -56.1447 | -56.5704 | -56.25 | -0.17395 | -0.009155 | -0.21057 | -0.20142 | 0.40283 | 0.57678 | -0.28381 | -0.32959 | -0.10986 | -0.15564 | -0.07782 | -0.01831 | -0.059509 | -0.041199 | 0.21973 | 0.29755 | 0.11444 | -0.20599 | 0.19226 | -0.12817
1226.1719 | 122.3511 | -56.3599 | 65.8539 | -0.13817 | 0.00048698 | 0.069613 | 0.069346 | -0.0010959 | 0.00016657 | -0.00081179 | -0.08371 | 0.0018821 | 0.066202 | 0.064046 | -0.0010361 | -0.0011257 | 0.00086952 | -0.1236 | 65.8539 | -56.5338 | 65.6479 | -55.9433 | -56.6666 | -56.5933 | -0.16022 | 65.8173 | -56.4423 | 65.7623 | -56.0989 | -56.4331 | -56.1859 | -0.17395 | 0 | -0.20599 | -0.20599 | 0.41656 | 0.59052 | -0.23346 | -0.3067 | -0.05951 | -0.13275 | -0.082398 | -0.036621 | -0.091553 | -0.054932 | 0.26093 | 0.34332 | 0.17395 | -0.073242 | 0.25635 | 0.009156
1227.1719 | 122.4838 | -56.3278 | 65.8356 | -0.19157 | 0.0017688 | 0.068545 | 0.067637 | -0.0017367 | 0.00080757 | -0.00059819 | 0.013947 | 0.0018821 | 0.064279 | 0.063192 | 0.00088648 | -0.0011257 | 0.0017239 | -0.054932 | 65.8264 | -56.4651 | 65.7944 | -55.7785 | -56.5338 | -56.4606 | -0.022888 | 65.7669 | -56.3644 | 65.8127 | -55.9479 | -56.163 | -55.9341 | -0.13733 | -0.009155 | -0.041198 | -0.032043 | 0.54932 | 0.68665 | -0.13275 | -0.20599 | 0.004578 | -0.068664 | -0.036621 | -0.068664 | -0.022888 | 0.045776 | 0.37994 | 0.41656 | 0.39368 | 0.1648 | 0.4303 | 0.20142
1228.1719 | 122.374 | -56.3232 | 65.9042 | -0.046616 | -0.00036742 | 0.067264 | 0.06956 | -0.0015231 | 0.00038017 | -0.00038459 | -0.065399 | 0.0014547 | 0.065988 | 0.06896 | -0.00039512 | -0.0026211 | 0.0017239 | -0.041199 | 65.8676 | -56.4697 | 65.7486 | -55.8289 | -56.5659 | -56.5201 | -0.032043 | 65.8768 | -56.3461 | 65.8813 | -56.0303 | -56.456 | -56.163 | -0.14648 | -0.036621 | -0.15564 | -0.11902 | 0.49439 | 0.64087 | -0.19684 | -0.24262 | -0.050354 | -0.09613 | -0.022888 | -0.027466 | -0.022888 | 0.004578 | 0.29297 | 0.31586 | 0.16022 | -0.13275 | 0.18311 | -0.10986
1229.1719 | 122.3831 | -56.3095 | 65.918 | -0.10765 | 5.9776e-05 | 0.068973 | 0.070201 | -0.0015231 | 0.00038017 | 0.00068361 | -0.044037 | 0.0023093 | 0.064706 | 0.065542 | 3.2078e-05 | -0.0039029 | 0.0012967 | -0.064087 | 65.9042 | -56.4468 | 65.7303 | -55.8701 | -56.5933 | -56.543 | -0.11902 | 65.8905 | -56.3599 | 65.8493 | -56.044 | -56.488 | -56.1722 | -0.13733 | -0.013733 | -0.18768 | -0.17395 | 0.43945 | 0.57678 | -0.23346 | -0.28381 | -0.096131 | -0.14648 | -0.050354 | -0.027466 | -0.068665 | -0.041199 | 0.2655 | 0.31586 | 0.13733 | -0.17853 | 0.18768 | -0.12817
1230.1719 | 122.3465 | -56.2958 | 65.8997 | -0.069504 | 0.0013416 | 0.066409 | 0.069987 | -0.0021641 | 0.00038017 | 0.0019654 | -0.05014 | 0.0023093 | 0.066843 | 0.065328 | 0.00088648 | -0.0034757 | 0.0025785 | -0.064087 | 65.9271 | -56.4331 | 65.7486 | -55.8289 | -56.5659 | -56.488 | -0.073242 | 65.918 | -56.3507 | 65.8813 | -56.0211 | -56.4377 | -56.1584 | -0.13733 | 0.027466 | -0.15106 | -0.17853 | 0.46692 | 0.60425 | -0.19226 | -0.27008 | -0.054932 | -0.13275 | -0.054932 | 0.018311 | -0.01831 | -0.036621 | 0.27466 | 0.32959 | 0.13733 | -0.14191 | 0.19226 | -0.086975
1231.1719 | 122.3969 | -56.2637 | 65.9225 | -0.10002 | -0.0014356 | 0.068332 | 0.071482 | 0.0012539 | 0.00059397 | 0.00068361 | -0.056244 | 0.0014547 | 0.066415 | 0.063619 | -0.00039512 | -0.0013393 | -0.0010531 | -0.032043 | 65.9409 | -56.4377 | 65.8127 | -55.7831 | -56.5109 | -56.4606 | -0.041199 | 65.9317 | -56.3049 | 65.918 | -55.9799 | -56.4056 | -56.1218 | -0.17395 | 0.018311 | -0.10986 | -0.12817 | 0.48065 | 0.6546 | -0.19684 | -0.24719 | -0.022888 | -0.073242 | -0.041199 | 0.009156 | -0.004577 | -0.013733 | 0.28381 | 0.32501 | 0.14191 | -0.14191 | 0.18311 | -0.10071
1232.1719 | 122.3511 | -56.282 | 65.9775 | -0.061874 | 0.0013416 | 0.069827 | 0.072337 | -0.0025913 | 0.001662 | -0.0031616 | -0.040985 | 0.0023093 | 0.066843 | 0.059133 | 0.0013139 | -0.0024075 | -0.00062588 | -0.07782 | 65.9409 | -56.3828 | 65.7761 | -55.7968 | -56.5338 | -56.4835 | -0.10529 | 65.9546 | -56.2958 | 65.918 | -55.9937 | -56.4285 | -56.1127 | -0.10071 | -0.036621 | -0.20142 | -0.16479 | 0.48523 | 0.58594 | -0.20142 | -0.25177 | -0.10071 | -0.15106 | -0.013733 | -0.022888 | -0.059509 | -0.036621 | 0.28839 | 0.30212 | 0.16937 | -0.14648 | 0.1831 | -0.13275
1233.1719 | 122.4197 | -56.2225 | 66.0004 | -0.10765 | 0.00091438 | 0.069186 | 0.070201 | 0.0010403 | 0.0018756 | -0.0018798 | -0.040985 | 0.0029501 | 0.067056 | 0.06426 | -0.0012497 | -0.0032619 | 0.0012967 | -0.059509 | 65.9821 | -56.3965 | 65.8127 | -55.751 | -56.5018 | -56.4514 | -0.073242 | 65.9775 | -56.2637 | 65.9546 | -55.9525 | -56.3828 | -56.0898 | -0.17395 | -0.01831 | -0.18768 | -0.16937 | 0.4715 | 0.64545 | -0.22888 | -0.27924 | -0.054932 | -0.10529 | -0.041199 | -0.022888 | -0.045776 | -0.022888 | 0.27008 | 0.31128 | 0.13275 | -0.16022 | 0.17395 | -0.11902
1234.1719 | 122.4197 | -56.2088 | 66.0095 | -0.031357 | -0.00036742 | 0.069827 | 0.068492 | 0.0012539 | 0.0010212 | -0.0010254 | -0.034882 | 0.00060029 | 0.064493 | 0.06191 | 0.00024568 | -0.0028347 | 0.00086952 | -0.018311 | 66.0004 | -56.3187 | 65.8813 | -55.6503 | -56.4194 | -56.3736 | 0.018311 | 66.0141 | -56.2454 | 66.0324 | -55.8563 | -56.3049 | -56.0394 | -0.10986 | -0.009155 | -0.12817 | -0.11902 | 0.55847 | 0.66834 | -0.16479 | -0.21057 | -0.054931 | -0.10071 | -0.036621 | 0.004578 | 0.022889 | 0.018311 | 0.35248 | 0.3891 | 0.16937 | -0.096131 | 0.20599 | -0.05951
1235.1719 | 122.4426 | -56.163 | 66.069 | -0.031357 | 0.00070078 | 0.073672 | 0.06956 | -0.0021641 | 0.0010212 | 0.0013244 | -0.025726 | -0.00068151 | 0.065988 | 0.065755 | 0.00088648 | -0.0017667 | 0.00086952 | -0.013733 | 66.0507 | -56.3004 | 65.8997 | -55.6458 | -56.4011 | -56.3507 | 0.032043 | 66.0049 | -56.218 | 66.0278 | -55.9158 | -56.2363 | -56.0165 | -0.13733 | -0.018311 | -0.16937 | -0.15106 | 0.51727 | 0.6546 | -0.18768 | -0.23804 | -0.050354 | -0.10071 | -0.054932 | -0.064087 | -0.041199 | 0.022888 | 0.24719 | 0.30212 | 0.14648 | -0.073242 | 0.20142 | -0.01831
1236.1719 | 122.3969 | -56.1676 | 66.0507 | -0.1992 | 5.9776e-05 | 0.065982 | 0.068705 | -0.0013095 | 0.0027302 | 0.0019654 | -0.03183 | 0.0012411 | 0.067056 | 0.065542 | -0.00060892 | -0.00069854 | 0.0040737 | 0.004578 | 66.0553 | -56.2866 | 65.9134 | -55.6229 | -56.3828 | -56.3461 | 0.018311 | 66.0507 | -56.1859 | 66.0553 | -55.8243 | -56.2637 | -55.9753 | -0.11902 | 0.004578 | -0.13733 | -0.14191 | 0.54474 | 0.66376 | -0.17853 | -0.21515 | -0.059509 | -0.09613 | -0.01831 | 0 | 0.004578 | 0.004578 | 0.34332 | 0.36163 | 0.19226 | -0.09613 | 0.21057 | -0.07782
1237.1719 | 122.3694 | -56.131 | 66.0645 | -0.038986 | -0.0010082 | 0.070895 | 0.069773 | 0.0021085 | -4.7032e-05 | -0.00017099 | -0.040985 | 0.00017309 | 0.06727 | 0.068532 | -0.0014633 | -0.0028347 | 0.00065592 | -0.036621 | 66.0782 | -56.2958 | 65.9042 | -55.6503 | -56.3919 | -56.3644 | -0.041199 | 66.0553 | -56.1813 | 66.0278 | -55.8426 | -56.2912 | -55.9891 | -0.16479 | 0.013733 | -0.16022 | -0.17395 | 0.48065 | 0.64545 | -0.23346 | -0.26093 | -0.068665 | -0.096131 | -0.050354 | -0.009155 | -0.036621 | -0.027466 | 0.28839 | 0.33874 | 0.14191 | -0.16022 | 0.19226 | -0.10986
1238.1719 | 122.3969 | -56.1172 | 66.1331 | -0.084763 | -0.00036742 | 0.0694 | 0.06721 | -0.00088229 | 0.0014484 | 0.0013244 | -0.019623 | 0.0027365 | 0.06492 | 0.067678 | -0.0010361 | -0.0051845 | 0.0023647 | 0 | 66.1148 | -56.2683 | 65.9592 | -55.5862 | -56.3324 | -56.282 | 0.036621 | 66.0965 | -56.1447 | 66.1148 | -55.7693 | -56.1813 | -55.9021 | -0.15106 | -0.018311 | -0.17395 | -0.15564 | 0.53101 | 0.68207 | -0.16479 | -0.21515 | -0.013732 | -0.064086 | -0.027465 | -0.036621 | -0.018311 | 0.01831 | 0.3479 | 0.37537 | 0.21515 | -0.064086 | 0.24261 | -0.036621
1239.1719 | 122.3923 | -56.0761 | 66.1285 | -0.084763 | 0.0019824 | 0.069827 | 0.068278 | -0.00066869 | 0.00080757 | -0.001239 | -0.013519 | 0.0016685 | 0.066843 | 0.06661 | 0.00045928 | -0.0045437 | -0.0012667 | -0.018311 | 66.1285 | -56.2225 | 65.9912 | -55.5405 | -56.2958 | -56.25 | 0.036621 | 66.0828 | -56.1172 | 66.124 | -55.751 | -56.1768 | -55.8151 | -0.14648 | 0 | -0.13733 | -0.13733 | 0.53558 | 0.68207 | -0.17395 | -0.21973 | -0.027466 | -0.073242 | -0.041199 | -0.045776 | -0.004578 | 0.041198 | 0.32501 | 0.36621 | 0.26093 | -0.10071 | 0.30212 | -0.059509
1240.1719 | 122.3923 | -56.0623 | 66.1606 | -0.016098 | 0.001128 | 0.070254 | 0.070628 | 0.0010403 | 0.00016657 | -0.0020936 | -0.0013123 | -0.00025411 | 0.065561 | 0.058278 | -0.00018152 | -0.0034757 | 0.0010831 | -0.041199 | 66.1331 | -56.1722 | 66.0324 | -55.5176 | -56.2729 | -56.2271 | 0 | 66.1423 | -56.0944 | 66.1331 | -55.7327 | -56.1722 | -55.8975 | -0.10986 | -0.027465 | -0.12817 | -0.10071 | 0.54474 | 0.6546 | -0.16479 | -0.21057 | -0.054932 | -0.10071 | -0.032043 | -0.01831 | -0.027465 | -0.009155 | 0.32959 | 0.36163 | 0.16479 | -0.10986 | 0.19684 | -0.07782
1241.1719 | 122.4518 | -55.9845 | 66.2018 | -0.10002 | 0.0015552 | 0.070041 | 0.070842 | 0.0029629 | -0.0015424 | 0.0023926 | 0.0017395 | 0.0012411 | 0.066202 | 0.070028 | -0.0016769 | -0.0039029 | -0.0044711 | -0.004578 | 66.1789 | -56.1722 | 66.0507 | -55.4947 | -56.2225 | -56.2042 | 0.050354 | 66.1514 | -56.0806 | 66.1652 | -55.6732 | -56.131 | -55.8334 | -0.18768 | -0.022888 | -0.15106 | -0.12817 | 0.48981 | 0.67749 | -0.21973 | -0.23804 | -0.032044 | -0.050354 | -0.09613 | -0.050354 | -0.036621 | 0.013733 | 0.31128 | 0.40741 | 0.15106 | -0.14648 | 0.24719 | -0.050354
1242.1719 | 122.4106 | -56.0028 | 66.2109 | -0.11528 | -0.00036742 | 0.067691 | 0.06956 | 0.0014675 | 0.0020892 | 0.002179 | -0.077606 | 0.0020957 | 0.064493 | 0.062764 | -0.00060892 | -0.0049709 | 0.0036465 | -0.10071 | 66.2018 | -56.163 | 66.0004 | -55.5771 | -56.3232 | -56.25 | -0.3479 | 66.1148 | -56.0669 | 66.1743 | -55.7556 | -56.1768 | -55.8746 | -0.16022 | -0.009155 | -0.21057 | -0.20142 | 0.42572 | 0.58594 | -0.24719 | -0.32043 | -0.086975 | -0.16022 | -0.064087 | -0.09613 | -0.036621 | 0.059509 | 0.24719 | 0.31128 | 0.12817 | -0.17395 | 0.19226 | -0.10986
1243.1719 | 122.4701 | -55.9525 | 66.2521 | -0.069504 | 0.0017688 | 0.068332 | 0.067851 | -0.0013095 | 0.0010212 | -0.001239 | -0.016571 | 0.0010275 | 0.06492 | 0.067891 | -0.0014633 | -0.0047573 | 0.00065592 | 0.013733 | 66.2384 | -56.0989 | 66.0873 | -55.426 | -56.2088 | -56.1447 | 0.027466 | 66.2155 | -55.9982 | 66.2567 | -55.6229 | -56.0532 | -55.7831 | -0.14648 | -0.013733 | -0.16479 | -0.15106 | 0.52643 | 0.67291 | -0.19226 | -0.25635 | -0.045776 | -0.10986 | -0.045776 | -0.036621 | 0.004578 | 0.041199 | 0.32959 | 0.37537 | 0.16937 | -0.10071 | 0.21515 | -0.054932
1244.1719 | 122.4152 | -55.9341 | 66.2842 | -0.061874 | 0.00070078 | 0.069613 | 0.067637 | 0.00018591 | 0.0029438 | -0.0031616 | -0.016571 | 0.0025229 | 0.068124 | 0.073018 | -0.0021041 | -0.0045437 | 0.00086952 | 0.009155 | 66.275 | -56.0898 | 66.156 | -55.3802 | -56.163 | -56.1264 | 0.045776 | 66.243 | -55.9616 | 66.243 | -55.5908 | -56.0394 | -55.7602 | -0.15564 | -0.009156 | -0.12817 | -0.11902 | 0.55389 | 0.70953 | -0.19226 | -0.22888 | -0.036621 | -0.073242 | -0.027466 | -0.041199 | -0.041199 | 0 | 0.34332 | 0.37079 | 0.17395 | -0.10529 | 0.20142 | -0.07782
1245.1719 | 122.4335 | -55.8929 | 66.2979 | -0.038986 | -0.00058102 | 0.068973 | 0.06956 | 0.0012539 | 0.00080757 | -0.00081179 | -0.0013123 | 0.00017309 | 0.067056 | 0.065328 | -0.0012497 | -0.0039029 | 0.0025785 | -0.004578 | 66.3162 | -56.0669 | 66.156 | -55.394 | -56.1127 | -56.0806 | 0.036621 | 66.2933 | -55.9616 | 66.2933 | -55.5634 | -56.012 | -55.7373 | -0.17395 | 0.01831 | -0.14191 | -0.16022 | 0.49896 | 0.67291 | -0.18768 | -0.21973 | -0.013732 | -0.045776 | -0.068665 | -0.004578 | -0.004578 | 0 | 0.32959 | 0.39826 | 0.15564 | -0.11902 | 0.2243 | -0.050354
1246.1719 | 122.4106 | -55.8884 | 66.3712 | -0.054245 | 5.9776e-05 | 0.069827 | 0.068064 | -0.0032321 | 0.0018756 | 0.0026062 | 0.013947 | -0.00068151 | 0.065988 | 0.060415 | -0.00039512 | -0.0047573 | -0.00062588 | 0 | 66.3254 | -56.0211 | 66.1743 | -55.3574 | -56.0944 | -56.0623 | 0.018311 | 66.3025 | -55.9113 | 66.3254 | -55.5176 | -55.9799 | -55.7007 | -0.13275 | -0.045777 | -0.19684 | -0.15106 | 0.53101 | 0.66376 | -0.17395 | -0.20599 | -0.041199 | -0.073242 | -0.022888 | -0.068665 | -0.045777 | 0.022888 | 0.37079 | 0.39368 | 0.18768 | -0.091552 | 0.21057 | -0.068664
1247.1719 | 122.406 | -55.8517 | 66.3712 | -0.077133 | 0.0013416 | 0.068759 | 0.069987 | 0.00082671 | -0.00026063 | 0.0028198 | -0.028778 | -0.00089511 | 0.065134 | 0.062978 | 0.00024568 | -0.0028347 | -0.0016941 | -0.004578 | 66.3483 | -56.012 | 66.2292 | -55.2841 | -56.1035 | -56.044 | 0.032043 | 66.3345 | -55.8929 | 66.3345 | -55.5313 | -55.9662 | -55.6778 | -0.16022 | -0.022888 | -0.14191 | -0.11902 | 0.56763 | 0.72784 | -0.19226 | -0.25177 | -0.032043 | -0.091553 | -0.041198 | -0.036621 | -0.036621 | 0 | 0.32043 | 0.36163 | 0.17395 | -0.11444 | 0.21515 | -0.073243
1248.1719 | 122.4014 | -55.806 | 66.4124 | -0.038986 | -0.00036742 | 0.068332 | 0.066569 | 0.00018591 | 0.00016657 | -0.00017099 | 0.029205 | 0.0018821 | 0.066629 | 0.064901 | -0.00018152 | -0.0028347 | -0.0040439 | -0.027466 | 66.3986 | -55.9616 | 66.2201 | -55.3345 | -56.0577 | -56.0074 | -0.009155 | 66.4032 | -55.8838 | 66.3574 | -55.4764 | -55.9296 | -55.6366 | -0.15564 | -0.013733 | -0.19226 | -0.17853 | 0.4715 | 0.62714 | -0.20142 | -0.25177 | -0.045776 | -0.09613 | -0.07782 | -0.009156 | -0.054932 | -0.045776 | 0.32959 | 0.40741 | 0.16937 | -0.1236 | 0.24719 | -0.045776
1249.1719 | 122.4426 | -55.7922 | 66.4536 | -0.031357 | 0.0019824 | 0.066836 | 0.068705 | -0.00045509 | -0.00090143 | 0.001538 | -0.010467 | 0.00081389 | 0.066415 | 0.065114 | -0.00039512 | -0.0026211 | -0.00083948 | 0.004578 | 66.4261 | -55.9204 | 66.2933 | -55.2567 | -55.9708 | -55.9616 | 0.045776 | 66.4078 | -55.8334 | 66.4215 | -55.4169 | -55.8517 | -55.5908 | -0.12817 | -0.027466 | -0.16022 | -0.13275 | 0.53558 | 0.66376 | -0.16937 | -0.17853 | -0.041199 | -0.050354 | -0.041199 | -0.045776 | -0.032043 | 0.013733 | 0.37537 | 0.41656 | 0.20142 | -0.05951 | 0.24262 | -0.018311
1250.1719 | 122.4335 | -55.7098 | 66.4902 | -0.031357 | 0.001128 | 0.068332 | 0.068705 | 0.0021085 | -4.7032e-05 | 0.00047001 | 0.032257 | 0.0040183 | 0.065347 | 0.062124 | 0.0017411 | -0.0036893 | 0.00065592 | -0.013733 | 66.481 | -55.8746 | 66.3208 | -55.2338 | -55.9799 | -55.9296 | 0.004578 | 66.4764 | -55.7831 | 66.4536 | -55.3894 | -55.8289 | -55.5496 | -0.16479 | -0.009155 | -0.16937 | -0.16022 | 0.47607 | 0.64087 | -0.21973 | -0.27008 | -0.054931 | -0.10529 | -0.073242 | -0.013733 | -0.036621 | -0.022888 | 0.32044 | 0.39368 | 0.16022 | -0.11902 | 0.23346 | -0.045776
1251.1719 | 122.3923 | -55.7098 | 66.5314 | -0.031357 | -0.0010082 | 0.067477 | 0.068492 | 0.00082671 | -0.001756 | -0.001239 | -0.0013123 | 0.0020957 | 0.065988 | 0.059987 | -0.0012497 | -0.0034757 | -0.0023349 | -0.022888 | 66.5131 | -55.8426 | 66.33 | -55.1971 | -55.9525 | -55.8975 | 0 | 66.4993 | -55.7648 | 66.4719 | -55.3482 | -55.7968 | -55.5267 | -0.13275 | -0.01831 | -0.20142 | -0.18311 | 0.5127 | 0.64545 | -0.18768 | -0.24262 | -0.054932 | -0.10986 | -0.054932 | -0.032043 | -0.059509 | -0.027466 | 0.36163 | 0.41657 | 0.18311 | -0.086975 | 0.23804 | -0.032043
1252.1719 | 122.4518 | -55.6549 | 66.5634 | -0.069504 | 5.9776e-05 | 0.06705 | 0.07191 | -0.00024129 | 0.0023028 | -0.00017099 | -0.037933 | 0.00060029 | 0.064279 | 0.069387 | -0.00039512 | -0.0053983 | -0.00062588 | -0.004578 | 66.5268 | -55.8151 | 66.4078 | -55.0964 | -55.8701 | -55.8563 | 0.032043 | 66.5543 | -55.719 | 66.5543 | -55.2979 | -55.8105 | -55.4718 | -0.16022 | -0.036622 | -0.15564 | -0.11902 | 0.55847 | 0.71869 | -0.20142 | -0.21515 | -0.041198 | -0.054931 | -0.064087 | -0.009156 | -0.009156 | 0 | 0.35706 | 0.42114 | 0.18311 | -0.15564 | 0.24719 | -0.091553
1253.1719 | 122.4152 | -55.632 | 66.6046 | -0.077133 | 0.00027338 | 0.068759 | 0.068705 | 0.00018591 | 0.0018756 | -0.0027344 | 0.029205 | 0.0025229 | 0.065775 | 0.059347 | -0.00039512 | -0.0049709 | 0.0027921 | -0.018311 | 66.5955 | -55.7556 | 66.4124 | -55.0873 | -55.9433 | -55.8426 | -0.027466 | 66.5863 | -55.6915 | 66.5131 | -55.2979 | -55.7373 | -55.4947 | -0.1236 | -0.009155 | -0.19226 | -0.18311 | 0.54474 | 0.66833 | -0.21057 | -0.31128 | -0.086975 | -0.18768 | -0.059509 | -0.01831 | -0.091552 | -0.073242 | 0.33417 | 0.39368 | 0.13733 | -0.10529 | 0.19684 | -0.045777
1254.1719 | 122.3923 | -55.5771 | 66.6595 | -0.061874 | -0.0014356 | 0.069186 | 0.065928 | 0.00018591 | 0.00080757 | -0.0027344 | 0.13297 | 0.0012411 | 0.067484 | 0.062978 | -0.0016769 | -0.0026211 | 0.0023647 | -0.041199 | 66.6458 | -55.7236 | 66.4444 | -55.069 | -55.8517 | -55.7831 | 0.045776 | 66.6641 | -55.6137 | 66.5543 | -55.2612 | -55.7098 | -55.4718 | -0.14648 | -0.013733 | -0.21515 | -0.20142 | 0.50812 | 0.6546 | -0.20599 | -0.27466 | -0.059509 | -0.12817 | -0.036621 | 0.004578 | -0.10529 | -0.10986 | 0.31586 | 0.35248 | 0.10529 | -0.13275 | 0.14191 | -0.096131
1255.1719 | 122.319 | -55.5588 | 66.7007 | -0.077133 | -0.00079462 | 0.068332 | 0.067637 | 0.0021085 | 0.0010212 | -0.0016662 | 0.013947 | 0.00038669 | 0.065347 | 0.062978 | 0.00045928 | -0.0041165 | 0.0059963 | -0.041199 | 66.6962 | -55.6824 | 66.4993 | -55.0095 | -55.8014 | -55.751 | 0.013733 | 66.6641 | -55.5954 | 66.5405 | -55.0003 | -55.5588 | -55.4489 | -0.1236 | -0.004578 | -0.20142 | -0.19684 | 0.54932 | 0.67291 | -0.19226 | -0.24262 | -0.068665 | -0.11902 | -0.036621 | -0.036621 | -0.16022 | -0.1236 | 0.55847 | 0.59509 | 0.10986 | 0 | 0.14648 | 0.036621
1256.1719 | 122.3694 | -55.513 | 66.7557 | -0.061874 | 0.00091438 | 0.070254 | 0.067637 | -0.00066869 | 0.00038017 | -0.0020936 | 0.029205 | 0.0023093 | 0.067911 | 0.06661 | 3.2078e-05 | -0.0039029 | 0.0045011 | -0.022888 | 66.6824 | -55.6641 | 66.5268 | -54.9637 | -55.7556 | -55.7053 | 0.018311 | 66.7007 | -55.5542 | 66.6321 | -55.1834 | -55.6366 | -55.3848 | -0.15106 | -0.073242 | -0.22888 | -0.15564 | 0.54932 | 0.70038 | -0.19226 | -0.24262 | -0.041199 | -0.091553 | -0.041199 | -0.054931 | -0.1236 | -0.068665 | 0.32959 | 0.37079 | 0.12817 | -0.1236 | 0.16937 | -0.082398
1257.1719 | 122.4335 | -55.4672 | 66.774 | -0.069504 | 0.00027338 | 0.06705 | 0.067637 | 0.0027493 | -0.001115 | -0.0010254 | -0.010467 | 0.0016685 | 0.066415 | 0.067251 | -0.0018905 | -0.0039029 | -0.0029757 | -0.013733 | 66.7328 | -55.6183 | 66.5405 | -54.9774 | -55.7236 | -55.6915 | -0.004578 | 66.7328 | -55.5542 | 66.7191 | -55.1331 | -55.5862 | -55.3207 | -0.15106 | -0.041199 | -0.23346 | -0.19226 | 0.48981 | 0.64087 | -0.2243 | -0.25635 | -0.073242 | -0.10529 | -0.086975 | -0.041199 | -0.054932 | -0.013733 | 0.33417 | 0.42114 | 0.14648 | -0.11902 | 0.23346 | -0.032044
1258.1719 | 122.3373 | -55.4031 | 66.8106 | -0.023727 | 0.00027338 | 0.069186 | 0.068919 | -2.7692e-05 | 0.00038017 | -0.00059819 | 0.026154 | 0.0010275 | 0.067697 | 0.06661 | 0.0017411 | -0.0049709 | -0.00041228 | -0.022888 | 66.7923 | -55.5862 | 66.5863 | -54.895 | -55.7007 | -55.632 | 0.009155 | 66.7648 | -55.4764 | 66.7099 | -55.0278 | -55.3253 | -55.1971 | -0.18311 | -0.018311 | -0.2243 | -0.20599 | 0.50812 | 0.69122 | -0.22888 | -0.29755 | -0.045776 | -0.11444 | -0.073242 | -0.045776 | -0.10071 | -0.054932 | 0.37537 | 0.44861 | 0.20599 | 0.07782 | 0.27923 | 0.15106
1259.1719 | 122.374 | -55.4031 | 66.8518 | -0.061874 | 0.0013416 | 0.068118 | 0.067424 | -0.0010959 | 0.0010212 | -0.0020936 | 0.010895 | 0.00081389 | 0.064279 | 0.063833 | 3.2078e-05 | -0.0043301 | 0.00022852 | -0.009155 | 66.8106 | -55.5267 | 66.655 | -54.8401 | -55.632 | -55.5908 | 0.022888 | 66.8335 | -55.4306 | 66.774 | -55.0598 | -55.5222 | -55.2704 | -0.1236 | -0.041199 | -0.19684 | -0.15564 | 0.56305 | 0.68664 | -0.18768 | -0.22888 | -0.064087 | -0.10529 | -0.027466 | -0.018311 | -0.07782 | -0.059509 | 0.34332 | 0.37079 | 0.13275 | -0.11902 | 0.16022 | -0.091553
1260.1719 | 122.3969 | -55.3574 | 66.8701 | -0.0084687 | 0.0013416 | 0.071109 | 0.067637 | 0.0012539 | 0.00059397 | -0.0020936 | 0.093292 | 0.0023093 | 0.065347 | 0.06426 | -0.0014633 | -0.0045437 | -0.00041228 | -0.027466 | 66.8655 | -55.5084 | 66.6916 | -54.808 | -55.5954 | -55.4855 | -0.10071 | 66.8838 | -55.394 | 66.806 | -55.0278 | -55.4901 | -55.22 | -0.15106 | -0.004577 | -0.17853 | -0.17395 | 0.54932 | 0.70038 | -0.12817 | -0.23804 | 0.022888 | -0.086975 | -0.036621 | 0.013733 | -0.064087 | -0.07782 | 0.32959 | 0.36621 | 0.13733 | -0.13275 | 0.17395 | -0.09613
1261.1719 | 122.374 | -55.2933 | 66.9159 | -0.069504 | -0.00036742 | 0.066836 | 0.068064 | 0.00039951 | 0.00038017 | 0.00025641 | 0.02005 | 0.0027365 | 0.06727 | 0.066823 | -0.00018152 | -0.0041165 | -0.0021213 | 0.022888 | 66.9113 | -55.4626 | 66.806 | -54.6799 | -55.4626 | -55.4443 | 0.064087 | 66.9022 | -55.3391 | 66.8793 | -54.9454 | -55.394 | -55.1514 | -0.16937 | -0.004578 | -0.10986 | -0.10529 | 0.6134 | 0.78278 | -0.15106 | -0.16937 | 0.01831 | 0 | -0.045776 | -0.013733 | -0.036622 | -0.022889 | 0.3479 | 0.39368 | 0.14191 | -0.10071 | 0.18768 | -0.054932
1262.1719 | 122.4564 | -55.2475 | 66.9846 | -0.061874 | 0.001128 | 0.066195 | 0.071696 | 0.0010403 | 0.00080757 | -0.00038459 | -0.056244 | 0.0012411 | 0.065134 | 0.065328 | -0.00018152 | -0.0039029 | 0.0017239 | 0.004578 | 66.925 | -55.4352 | 66.8152 | -54.6753 | -55.4901 | -55.4581 | 0.054932 | 66.9434 | -55.3162 | 66.9662 | -54.8996 | -55.3482 | -55.0964 | -0.18768 | -0.059509 | -0.16937 | -0.10986 | 0.5722 | 0.75989 | -0.21057 | -0.24261 | -0.022888 | -0.054931 | -0.068664 | -0.041199 | -0.01831 | 0.022889 | 0.3479 | 0.41656 | 0.15106 | -0.10071 | 0.21973 | -0.032044
1263.1719 | 122.3694 | -55.2155 | 67.0303 | -0.069504 | -0.0010082 | 0.068759 | 0.06486 | 0.0059537 | 0.0010212 | 0.00025641 | 0.02005 | 0.0018821 | 0.065134 | 0.061269 | -0.00060892 | -0.0051845 | 0.0017239 | 0.022888 | 67.0029 | -55.3482 | 66.8747 | -54.6112 | -55.4535 | -55.3619 | 0.059509 | 67.0029 | -55.2704 | 66.9434 | -54.8676 | -55.3253 | -55.0735 | -0.13275 | -0.027465 | -0.15564 | -0.12817 | 0.60425 | 0.737 | -0.14648 | -0.23804 | -0.013732 | -0.10529 | -0.054932 | -0.027465 | -0.086975 | -0.05951 | 0.3479 | 0.40283 | 0.14191 | -0.10986 | 0.19684 | -0.054931
1264.1719 | 122.3923 | -55.1285 | 67.0944 | -0.077133 | 0.0013416 | 0.068973 | 0.067851 | -0.0013095 | 0.00038017 | -0.0016662 | 0.032257 | -0.00025411 | 0.065134 | 0.061483 | 3.2078e-05 | -0.0039029 | 0.00044212 | 0.032043 | 67.012 | -55.2979 | 66.9479 | -54.5059 | -55.3436 | -55.2795 | 0.07782 | 67.0532 | -55.2017 | 67.0166 | -54.808 | -55.2429 | -54.9911 | -0.16937 | -0.082397 | -0.14648 | -0.064087 | 0.62256 | 0.79193 | -0.15106 | -0.21515 | 0.018311 | -0.045776 | -0.073242 | -0.041198 | -0.077819 | -0.036621 | 0.32043 | 0.39368 | 0.13733 | -0.11444 | 0.21057 | -0.041199
1265.1719 | 122.4152 | -55.0827 | 67.1036 | -0.069504 | -0.00015382 | 0.068759 | 0.068278 | -0.0032321 | -0.00026063 | -0.0033752 | 0.013947 | 0.00017309 | 0.065347 | 0.065328 | -0.00082252 | -0.0051845 | 0.00044212 | 0.045776 | 67.0761 | -55.2338 | 67.012 | -54.4327 | -55.275 | -55.2155 | 0.09613 | 67.0715 | -55.1147 | 67.0898 | -54.7028 | -55.1559 | -54.9088 | -0.15106 | -0.027466 | -0.091553 | -0.064087 | 0.65002 | 0.80109 | -0.13275 | -0.19226 | 0.018311 | -0.041198 | -0.032043 | -0.032044 | -0.013733 | 0.018311 | 0.37994 | 0.41199 | 0.17395 | -0.073242 | 0.20599 | -0.041199
1266.1719 | 122.493 | -55.0323 | 67.1448 | -0.10002 | 0.001128 | 0.067264 | 0.067424 | 0.0025357 | 0.0029438 | 0.00089721 | 0.02005 | -0.00025411 | 0.064706 | 0.064046 | -0.0029587 | -0.0045437 | 0.0017239 | 0.045776 | 67.1127 | -55.1697 | 67.0853 | -54.3915 | -55.2109 | -55.1605 | 0.1236 | 67.0898 | -55.0873 | 67.186 | -54.6433 | -55.0644 | -54.8264 | -0.13733 | -0.032043 | -0.059509 | -0.027466 | 0.64087 | 0.7782 | -0.12817 | -0.17853 | 0.009156 | -0.041198 | -0.054931 | -0.054931 | 0.041199 | 0.09613 | 0.3891 | 0.44403 | 0.20599 | -0.032043 | 0.26093 | 0.022888
1267.1719 | 122.4701 | -54.9637 | 67.2089 | 0.067825 | 0.002837 | 0.071963 | 0.068278 | 0.00018591 | 0.0012348 | -0.002948 | 0.038361 | -0.0017495 | 0.065347 | 0.06191 | -0.00082252 | -0.0026211 | -0.0021213 | -0.12817 | 67.1539 | -55.1239 | 67.1677 | -54.2953 | -55.1651 | -55.0598 | 0.12817 | 67.1631 | -55.0186 | 67.2638 | -54.5517 | -54.9866 | -54.7623 | -0.16022 | -0.054931 | -0.041198 | 0.013733 | 0.66833 | 0.82855 | -0.09613 | -0.20142 | 0.064087 | -0.041199 | -0.054932 | -0.045776 | 0.054932 | 0.10071 | 0.41199 | 0.46692 | 0.20142 | -0.022888 | 0.25635 | 0.032044
1268.1719 | 122.5204 | -54.9179 | 67.2546 | -0.092392 | -0.00079462 | 0.070041 | 0.06721 | 0.0027493 | -0.0021832 | -0.004657 | 0.0047913 | -0.00046791 | 0.066415 | 0.057851 | -0.0021041 | -0.0041165 | -0.00041228 | 0.027466 | 67.2272 | -55.0507 | 67.2409 | -53.9749 | -55.0598 | -55.2658 | 0.20599 | 67.2043 | -54.9774 | 67.3233 | -54.4693 | -54.9408 | -54.7165 | -0.13275 | -0.027466 | -0.013733 | 0.013733 | 0.94299 | 1.0757 | -0.3479 | -0.14191 | -0.21515 | -0.009155 | -0.059509 | -0.050354 | 0.068664 | 0.11902 | 0.44861 | 0.50812 | 0.20142 | -0.022888 | 0.26093 | 0.036621
1269.1719 | 122.5159 | -54.8767 | 67.2913 | -0.046616 | 5.9776e-05 | 0.068545 | 0.067637 | 0.0014675 | 0.00038017 | -0.00038459 | 0.02005 | -0.00068151 | 0.063638 | 0.061696 | -0.00018152 | -0.0034757 | 0.0023647 | 0.032043 | 67.3096 | -55.0232 | 67.2913 | -54.1397 | -55.0003 | -54.982 | 0.16937 | 67.2684 | -54.9088 | 67.4698 | -54.4006 | -55.1514 | -54.8126 | -0.14648 | 0.01831 | 0 | -0.01831 | 0.737 | 0.88348 | -0.10529 | -0.1236 | 0.041198 | 0.022888 | -0.032043 | -0.022888 | 0.17853 | 0.20142 | 0.47607 | 0.50812 | 0.064087 | -0.27466 | 0.09613 | -0.24262
1270.1719 | 122.4884 | -54.7989 | 67.3553 | -0.092392 | -0.00079462 | 0.068973 | 0.067637 | -0.0025913 | 0.0020892 | 0.002179 | 0.016998 | -0.00046791 | 0.062997 | 0.059347 | -0.0029587 | -0.0053983 | -0.0016941 | 0.29755 | 67.3782 | -54.9454 | 67.2867 | -54.126 | -54.9728 | -54.9545 | 0.14191 | 67.3187 | -54.8447 | 67.4011 | -54.3961 | -54.8401 | -54.5792 | -0.14648 | 0.022888 | -0.068665 | -0.091553 | 0.67291 | 0.8194 | -0.15564 | -0.17395 | -0.009155 | -0.027465 | -0.045777 | -0.036621 | 0.045776 | 0.082397 | 0.40283 | 0.44861 | 0.21973 | -0.041199 | 0.2655 | 0.004578
1271.1719 | 122.4792 | -54.7714 | 67.3737 | -0.046616 | -0.00079462 | 0.069186 | 0.067851 | -0.0064365 | 0.0018756 | -0.001239 | -0.0013123 | -0.0032449 | 0.062784 | 0.06191 | 3.2078e-05 | -0.0030483 | -0.0048983 | 0.036621 | 67.3782 | -54.895 | 67.3737 | -54.0527 | -54.9088 | -54.8538 | 0.40283 | 67.4332 | -54.776 | 67.4469 | -54.3365 | -54.7668 | -54.5425 | -0.1236 | 0.004578 | 0 | -0.004578 | 0.71869 | 0.84229 | -0.082398 | -0.13733 | 0.041199 | -0.013732 | -0.004578 | 0.05951 | 0.073242 | 0.013732 | 0.43487 | 0.43945 | 0.22888 | 0.004577 | 0.23346 | 0.009155
1272.1719 | 122.4014 | -54.7302 | 67.4927 | -0.046616 | 0.00091438 | 0.070041 | 0.070628 | 0.0021085 | -0.0019696 | 0.0013244 | -0.056244 | 0.0025229 | 0.066629 | 0.06661 | -0.0014633 | -0.0041165 | -0.0014805 | -0.018311 | 67.4698 | -54.9133 | 67.3141 | -54.1763 | -54.9866 | -54.95 | -0.004578 | 67.4835 | -54.776 | 67.4561 | -54.3869 | -54.8676 | -54.5837 | -0.18311 | -0.022888 | -0.17853 | -0.15564 | 0.55389 | 0.737 | -0.21973 | -0.25635 | -0.036621 | -0.073242 | -0.045776 | -0.009155 | -0.036621 | -0.027466 | 0.34332 | 0.3891 | 0.14648 | -0.13733 | 0.19226 | -0.091553
1273.1875 | 122.3923 | -54.6844 | 67.5156 | -0.092392 | 0.00048698 | 0.070468 | 0.072337 | 0.0025357 | -0.0013288 | -0.00017099 | -0.092865 | 0.0010275 | 0.065347 | 0.070882 | -0.00060892 | -0.0036893 | 0.00065592 | -0.018311 | 67.511 | -54.863 | 67.3828 | -54.0619 | -54.9408 | -54.8904 | 0.009155 | 67.5339 | -54.744 | 67.511 | -54.3549 | -54.8401 | -54.5563 | -0.17853 | -0.004578 | -0.13275 | -0.12817 | 0.62256 | 0.80109 | -0.20599 | -0.25635 | -0.027466 | -0.07782 | -0.05951 | 0.018311 | -0.004578 | -0.022889 | 0.32959 | 0.3891 | 0.12817 | -0.15564 | 0.18768 | -0.09613
1274.1719 | 122.5159 | -54.6066 | 67.5156 | -0.077133 | -0.00058102 | 0.067477 | 0.067851 | 0.0021085 | 0.0027302 | -0.00081179 | 0.047516 | 0.00081389 | 0.066415 | 0.064046 | -0.00039512 | -0.0043301 | -0.0012667 | -0.009155 | 67.5385 | -54.7302 | 67.4789 | -53.92 | -54.7943 | -54.7302 | 0.09613 | 67.511 | -54.6753 | 67.6209 | -54.1672 | -54.6158 | -54.3823 | -0.1236 | 0.022888 | -0.036621 | -0.059509 | 0.68664 | 0.81024 | -0.1236 | -0.18768 | 0 | -0.064087 | -0.068665 | -0.004578 | 0.10529 | 0.10986 | 0.43945 | 0.50812 | 0.2243 | -0.009156 | 0.29297 | 0.059509
1275.1719 | 122.4976 | -54.5654 | 67.5842 | -0.077133 | 0.00070078 | 0.068332 | 0.066569 | 0.0014675 | -0.0013288 | 0.00047001 | 0.029205 | 0.0012411 | 0.063638 | 0.06191 | -0.00039512 | -0.0066799 | 0.0027921 | -0.018311 | 67.5659 | -54.7119 | 67.5156 | -53.9017 | -54.744 | -54.7028 | 0.068665 | 67.5568 | -54.6295 | 67.6804 | -54.1306 | -54.6112 | -54.3503 | -0.14648 | -0.018311 | -0.068665 | -0.050354 | 0.66376 | 0.81024 | -0.13733 | -0.17853 | 0.009155 | -0.032044 | -0.064087 | -0.027466 | 0.09613 | 0.1236 | 0.43488 | 0.49896 | 0.21515 | -0.045776 | 0.27924 | 0.018311
1276.1719 | 122.4884 | -54.5334 | 67.6346 | -0.069504 | 0.001128 | 0.0694 | 0.070201 | 0.0031765 | 0.0023028 | 0.001538 | 0.029205 | 0.0016685 | 0.065561 | 0.070028 | -0.0014633 | -0.0058255 | 0.00044212 | -0.045776 | 67.6392 | -54.6844 | 67.5476 | -53.8605 | -54.744 | -54.6661 | 0.013733 | 67.5934 | -54.5792 | 67.7078 | -54.0802 | -54.57 | -54.3045 | -0.15106 | 0.004577 | -0.086976 | -0.091553 | 0.67291 | 0.82397 | -0.13275 | -0.21057 | 0.01831 | -0.05951 | -0.045777 | -0.041199 | 0.073242 | 0.11444 | 0.45319 | 0.49896 | 0.22888 | -0.036621 | 0.27466 | 0.009156
1277.1719 | 122.4014 | -54.4739 | 67.7582 | -0.061874 | -0.0016492 | 0.068118 | 0.066783 | -0.00024129 | 0.00016657 | -0.0010254 | -0.028778 | 0.00038669 | 0.064279 | 0.058492 | -0.00039512 | -0.0047573 | -0.0023349 | 0.027466 | 67.7078 | -54.6021 | 67.6392 | -53.7918 | -54.657 | -54.6021 | 0.054932 | 67.717 | -54.5105 | 67.7078 | -54.0848 | -54.5837 | -54.3594 | -0.12817 | -0.050354 | -0.11902 | -0.068665 | 0.68207 | 0.81024 | -0.12817 | -0.1831 | 0 | -0.054931 | -0.036621 | -0.041199 | -0.050354 | -0.009155 | 0.3891 | 0.42572 | 0.11444 | -0.10986 | 0.15106 | -0.073242
1278.1719 | 122.438 | -54.4189 | 67.7948 | -0.023727 | -0.0010082 | 0.067477 | 0.070628 | -0.0025913 | -4.7032e-05 | 0.001538 | -0.071503 | -4.0508e-05 | 0.065347 | 0.063833 | -0.0016769 | -0.0019803 | 0.00022852 | 0 | 67.7902 | -54.6021 | 67.6483 | -53.7964 | -54.6616 | -54.5975 | 0.041199 | 67.7719 | -54.4647 | 67.7765 | -54.0344 | -54.5105 | -54.2587 | -0.18311 | -0.004578 | -0.14648 | -0.14191 | 0.62256 | 0.80566 | -0.17853 | -0.24262 | 0.004578 | -0.059509 | -0.045777 | -0.022888 | -0.018311 | 0.004577 | 0.38452 | 0.4303 | 0.16022 | -0.091553 | 0.20599 | -0.045776
1279.1719 | 122.4838 | -54.364 | 67.8085 | -0.038986 | 0.0034778 | 0.070682 | 0.067851 | -0.0017367 | 0.003371 | -0.001239 | 0.026154 | -0.00089511 | 0.063852 | 0.067037 | -0.0010361 | -0.0060391 | -0.00083948 | -0.07782 | 67.7902 | -54.5471 | 67.7124 | -53.7048 | -54.5654 | -54.5105 | -0.004578 | 67.7628 | -54.4281 | 67.8406 | -53.9383 | -54.4281 | -54.1672 | -0.18311 | -0.018311 | -0.096131 | -0.07782 | 0.65918 | 0.84229 | -0.14648 | -0.20142 | 0.036621 | -0.018311 | -0.064087 | -0.045777 | 0.032043 | 0.07782 | 0.42572 | 0.48981 | 0.19684 | -0.064087 | 0.26093 | 0
1280.1719 | 122.4564 | -54.3228 | 67.8589 | -0.11528 | -0.00015382 | 0.068332 | 0.067851 | -0.0030185 | -4.7032e-05 | -0.0023072 | 0.016998 | 0.0010275 | 0.067484 | 0.060201 | -0.0012497 | -0.0051845 | -0.00083948 | -0.064087 | 67.8406 | -54.4693 | 67.7582 | -53.6591 | -54.5242 | -54.4739 | -0.013733 | 67.8223 | -54.3869 | 67.8726 | -53.9291 | -54.3869 | -54.1214 | -0.14648 | -0.018311 | -0.10071 | -0.082397 | 0.66376 | 0.81024 | -0.15106 | -0.20142 | -0.004578 | -0.054932 | -0.064087 | -0.036621 | 0.013733 | 0.050354 | 0.39368 | 0.45776 | 0.20142 | -0.064087 | 0.2655 | 0
1281.1719 | 122.4564 | -54.2816 | 67.9001 | -0.077133 | 0.0024096 | 0.068545 | 0.070414 | -0.00024129 | -0.0021832 | 0.0026062 | 0.0078429 | -0.00025411 | 0.062997 | 0.065328 | -0.0018905 | -0.0036893 | 0.0015103 | -0.073242 | 67.8818 | -54.4006 | 67.8085 | -53.5858 | -54.4785 | -54.4144 | -0.036621 | 67.8864 | -54.2999 | 67.9321 | -53.8559 | -54.3365 | -54.0939 | -0.11902 | -0.01831 | -0.091552 | -0.073242 | 0.6958 | 0.81482 | -0.13275 | -0.19684 | -0.013733 | -0.07782 | -0.018311 | -0.013732 | 0.032044 | 0.045776 | 0.42572 | 0.44403 | 0.18768 | -0.054932 | 0.20599 | -0.036621
1282.1719 | 122.4152 | -54.1763 | 68.0008 | -0.046616 | 0.00027338 | 0.070468 | 0.067424 | -0.0023777 | 0.0012348 | -0.00059819 | -0.028778 | -0.0013223 | 0.06492 | 0.066182 | -0.00039512 | -0.0053983 | -0.00041228 | 0.027466 | 67.9779 | -54.3228 | 67.9321 | -53.4576 | -54.3594 | -54.3182 | 0.068665 | 67.9733 | -54.2313 | 67.9871 | -53.7781 | -54.3045 | -54.0527 | -0.14648 | -0.022888 | -0.068664 | -0.045776 | 0.71869 | 0.86517 | -0.14191 | -0.18311 | 0.004578 | -0.036621 | -0.054931 | -0.027465 | -0.013732 | 0.013733 | 0.39825 | 0.45319 | 0.1236 | -0.12817 | 0.17853 | -0.073242
1283.1719 | 122.438 | -54.1351 | 68.0191 | -0.10002 | 0.00048698 | 0.070254 | 0.068492 | -0.00045509 | 0.00059397 | -0.001239 | -0.047089 | -0.00025411 | 0.065988 | 0.059987 | 0.00024568 | -0.0045437 | 0.0034329 | -0.050354 | 68.0237 | -54.2999 | 67.9504 | -53.4073 | -54.3549 | -54.2816 | -0.045776 | 67.9962 | -54.1763 | 68.0923 | -53.7323 | -54.2221 | -53.9566 | -0.16479 | 0.004578 | -0.068665 | -0.073243 | 0.72784 | 0.89264 | -0.14648 | -0.21973 | 0.018311 | -0.054931 | -0.041199 | -0.022888 | 0.073242 | 0.09613 | 0.40283 | 0.44403 | 0.17853 | -0.086975 | 0.21973 | -0.045776
1284.1719 | 122.4564 | -54.0665 | 68.1107 | -0.084763 | 0.001128 | 0.069186 | 0.069773 | 0.0016813 | 0.0018756 | 0.00025641 | -0.0013123 | 0.0023093 | 0.065347 | 0.060415 | 3.2078e-05 | -0.0021939 | -0.0010531 | -0.086975 | 68.0557 | -54.2175 | 68.0054 | -53.3707 | -54.2999 | -54.213 | -0.041199 | 68.0832 | -54.1489 | 68.1381 | -53.6728 | -54.1489 | -53.8879 | -0.15106 | -0.054932 | -0.10529 | -0.050354 | 0.6958 | 0.84686 | -0.14648 | -0.23346 | 0.004577 | -0.082398 | -0.082398 | -0.027466 | 0.027466 | 0.054932 | 0.39368 | 0.47607 | 0.17853 | -0.082398 | 0.26093 | 0
1285.1719 | 122.4609 | -54.007 | 68.1519 | -0.084763 | 0.0013416 | 0.070895 | 0.067851 | -2.7692e-05 | 0.0010212 | -0.00017099 | -0.022675 | 0.00060029 | 0.065988 | 0.069173 | 0.00067288 | -0.0034757 | 0.0027921 | -0.064087 | 68.1335 | -54.158 | 68.0878 | -53.2837 | -54.2404 | -54.1718 | -0.041199 | 68.1198 | -54.0573 | 68.2068 | -53.595 | -54.0848 | -53.8422 | -0.15106 | -0.01831 | -0.064086 | -0.045776 | 0.72327 | 0.87433 | -0.1648 | -0.23346 | -0.013733 | -0.082397 | -0.050354 | -0.032043 | 0.054932 | 0.086975 | 0.41199 | 0.46234 | 0.16479 | -0.07782 | 0.21515 | -0.027466
1286.1719 | 122.4335 | -53.9474 | 68.2068 | -0.16106 | 0.00048698 | 0.070041 | 0.06956 | -0.0030185 | 0.0012348 | 0.003247 | -0.071503 | 0.0016685 | 0.06492 | 0.063192 | -0.00039512 | -0.0021939 | 1.4916e-05 | -0.07782 | 68.1976 | -54.0848 | 68.1702 | -53.183 | -54.1534 | -54.0939 | 0 | 68.1885 | -54.007 | 68.2892 | -53.5263 | -54.0024 | -53.7598 | -0.13733 | -0.009155 | -0.036621 | -0.027466 | 0.76447 | 0.90179 | -0.14648 | -0.20599 | -0.009155 | -0.068664 | -0.059509 | -0.01831 | 0.082398 | 0.10071 | 0.42114 | 0.48065 | 0.18768 | -0.054931 | 0.24719 | 0.004578
1287.1719 | 122.4838 | -53.8788 | 68.2846 | -0.084763 | 0.00070078 | 0.068973 | 0.068705 | -0.0017367 | 0.0025166 | -0.00017099 | -0.025726 | 0.00038669 | 0.067697 | 0.06191 | -0.0016769 | -0.0060391 | -0.0023349 | -0.073242 | 68.2983 | -54.0115 | 68.2114 | -53.1647 | -54.0939 | -54.0207 | -0.036621 | 68.2846 | -53.92 | 68.3212 | -53.4485 | -53.952 | -53.7186 | -0.13275 | 0.013733 | -0.073242 | -0.086975 | 0.71411 | 0.84686 | -0.14191 | -0.21515 | -0.009155 | -0.082397 | -0.041199 | 0 | 0.036621 | 0.036621 | 0.4303 | 0.4715 | 0.16022 | -0.073242 | 0.20142 | -0.032043
1288.1719 | 122.4518 | -53.7827 | 68.335 | -0.13054 | 0.00070078 | 0.070041 | 0.069773 | 0.0023221 | 0.0029438 | 0.0051696 | -0.10812 | 0.0016685 | 0.06492 | 0.070028 | 0.00067288 | -0.0024075 | -0.0019077 | -0.050354 | 68.3121 | -53.9749 | 68.1473 | -53.0823 | -53.746 | -53.8147 | 0.041199 | 68.3258 | -53.8834 | 68.4402 | -53.3661 | -53.8513 | -53.5904 | -0.19226 | -0.022888 | -0.18768 | -0.16479 | 0.70038 | 0.89264 | -0.032043 | 0.036621 | 0.16022 | 0.22888 | -0.10071 | -0.009155 | 0.10529 | 0.11444 | 0.41656 | 0.51727 | 0.19226 | -0.068664 | 0.29297 | 0.032044
1289.1719 | 122.4701 | -53.7277 | 68.4036 | -0.077133 | 0.002196 | 0.07175 | 0.072337 | -0.0013095 | -0.0019696 | 0.0023926 | -0.062347 | -0.0011087 | 0.064493 | 0.067678 | -0.00018152 | -0.00048494 | -0.0051121 | -0.036621 | 68.399 | -53.8742 | 68.3762 | -52.9678 | -53.9474 | -53.8834 | -0.009155 | 68.3945 | -53.7827 | 68.4174 | -53.1509 | -53.627 | -53.4439 | -0.14648 | -0.004577 | -0.027465 | -0.022888 | 0.75989 | 0.90637 | -0.15564 | -0.21973 | -0.009155 | -0.073242 | -0.054932 | -0.009155 | 0.013733 | 0.022888 | 0.57678 | 0.63171 | 0.28381 | 0.10071 | 0.33875 | 0.15564
1290.1719 | 122.4747 | -53.6774 | 68.4494 | -0.12291 | 0.0034778 | 0.070254 | 0.069987 | -0.00024129 | 0.00080757 | 0.0013244 | -0.074554 | -4.0508e-05 | 0.063638 | 0.068532 | -0.0021041 | -0.0013393 | 0.0019375 | -0.059509 | 68.4723 | -53.8284 | 68.4357 | -52.8992 | -53.8651 | -53.801 | -0.009155 | 68.4677 | -53.7186 | 68.6005 | -53.2196 | -53.7323 | -53.4531 | -0.15106 | 0.022888 | -0.013733 | -0.036621 | 0.7782 | 0.92926 | -0.1236 | -0.18768 | 0.027466 | -0.036621 | -0.041199 | 0.01831 | 0.15106 | 0.13275 | 0.45776 | 0.49896 | 0.2243 | -0.054932 | 0.2655 | -0.013733
1291.1719 | 122.5159 | -53.5858 | 68.541 | -0.13054 | 0.00070078 | 0.070254 | 0.06721 | -0.0015231 | 0.0018756 | 0.003247 | -0.03183 | -0.0011087 | 0.060648 | 0.06661 | -0.00082252 | -0.0021939 | 0.0030057 | -0.045776 | 68.5226 | -53.746 | 68.5043 | -52.8122 | -53.8925 | -53.9658 | 0.13275 | 68.5272 | -53.6545 | 68.66 | -53.1418 | -53.6591 | -53.3936 | -0.16022 | -0.018311 | -0.036622 | -0.018311 | 0.77362 | 0.93384 | -0.37994 | -0.3067 | -0.21973 | -0.14648 | -0.068665 | -0.013733 | 0.11902 | 0.13275 | 0.44403 | 0.5127 | 0.19226 | -0.073243 | 0.26093 | -0.004578
1292.1719 | 122.4701 | -53.5172 | 68.6188 | -0.077133 | 0.001128 | 0.070682 | 0.069773 | 0.0014675 | 0.00016657 | 0.00025641 | 0.02005 | -0.0015359 | 0.067697 | 0.066396 | -0.00039512 | -0.0043301 | 0.0015103 | -0.041199 | 68.66 | -53.6728 | 68.5547 | -52.739 | -53.7094 | -53.6499 | 0.036621 | 68.6325 | -53.5904 | 68.7332 | -53.0365 | -53.54 | -53.3066 | -0.15564 | 0.041199 | -0.064087 | -0.10529 | 0.7782 | 0.93384 | -0.13275 | -0.19226 | 0.022889 | -0.036621 | -0.073242 | 0.013733 | 0.11444 | 0.10071 | 0.48065 | 0.55389 | 0.21057 | -0.022888 | 0.28381 | 0.050354
1293.1719 | 122.493 | -53.4348 | 68.6554 | -0.13054 | 0.00027338 | 0.071963 | 0.068705 | 0.0027493 | 0.00059397 | 0.00025641 | -0.080658 | 0.0012411 | 0.065347 | 0.062551 | -0.0018905 | -0.0015529 | 0.0010831 | -0.059509 | 68.6462 | -53.627 | 68.6417 | -52.6749 | -53.6545 | -53.5904 | 0.004578 | 68.6646 | -53.5263 | 68.7927 | -53.0045 | -53.5172 | -53.2288 | -0.19226 | -0.009156 | -0.013733 | -0.004577 | 0.75989 | 0.95215 | -0.15564 | -0.21973 | 0.036621 | -0.027466 | -0.091553 | 0.009155 | 0.13733 | 0.12817 | 0.4303 | 0.52185 | 0.20599 | -0.082398 | 0.29755 | 0.009155
1294.1719 | 122.4701 | -53.3798 | 68.7607 | -0.13817 | -0.00058102 | 0.0694 | 0.069773 | 0.0031765 | 0.0014484 | 0.00025641 | -0.074554 | 0.0016685 | 0.06492 | 0.068746 | -0.00018152 | -0.0021939 | 0.0015103 | -0.068665 | 68.7424 | -53.5492 | 68.6829 | -52.6062 | -53.5858 | -53.5492 | -0.013733 | 68.7424 | -53.4439 | 68.8614 | -52.9358 | -53.4531 | -53.1647 | -0.16937 | -0.01831 | -0.07782 | -0.05951 | 0.77362 | 0.94299 | -0.16937 | -0.20599 | 0 | -0.036621 | -0.064087 | -0.01831 | 0.10071 | 0.11902 | 0.44403 | 0.50812 | 0.21515 | -0.073242 | 0.27924 | -0.009155
1295.1719 | 122.4518 | -53.3066 | 68.8431 | -0.12291 | 0.0017688 | 0.070041 | 0.068919 | -0.0028049 | 0.0010212 | -0.001239 | -0.037933 | 0.00081389 | 0.065988 | 0.060628 | -0.0014633 | -0.0017667 | 0.00086952 | -0.059509 | 68.8248 | -53.4805 | 68.7881 | -52.5009 | -53.5034 | -53.4393 | 0.004578 | 68.8293 | -53.3661 | 68.9117 | -52.8809 | -53.3707 | -53.1189 | -0.17395 | -0.018311 | -0.054932 | -0.036621 | 0.80566 | 0.97961 | -0.13275 | -0.19684 | 0.041199 | -0.022888 | -0.059509 | -0.013733 | 0.068664 | 0.082397 | 0.42572 | 0.48523 | 0.18768 | -0.064087 | 0.24719 | -0.004578
1296.1719 | 122.493 | -53.2471 | 68.898 | -0.069504 | -0.00036742 | 0.070468 | 0.067851 | 0.0014675 | 0.0014484 | 4.2808e-05 | -0.10202 | 0.0010275 | 0.065347 | 0.071523 | 0.0011001 | -0.0032619 | 0.0027921 | -0.09613 | 68.9255 | -53.421 | 68.811 | -52.5101 | -53.508 | -53.4531 | -0.082397 | 68.9026 | -53.3203 | 69.017 | -52.803 | -53.3478 | -53.0548 | -0.17395 | 0.027466 | -0.086975 | -0.11444 | 0.737 | 0.91095 | -0.20599 | -0.26093 | -0.032043 | -0.086975 | -0.073242 | 0.004578 | 0.11902 | 0.11444 | 0.44403 | 0.51727 | 0.19226 | -0.10071 | 0.2655 | -0.027466
1297.1719 | 122.4747 | -53.1784 | 69.0262 | -0.13054 | 0.0032642 | 0.068332 | 0.066783 | -0.0025913 | -0.0013288 | -0.00038459 | 0.0078429 | 0.0018821 | 0.065561 | 0.061483 | -0.0010361 | -0.0036893 | 0.0010831 | -0.045776 | 68.9896 | -53.3112 | 68.9255 | -52.3727 | -53.3752 | -53.2974 | 0 | 68.9896 | -53.215 | 69.0811 | -52.6794 | -53.1921 | -52.9541 | -0.13275 | -0.036621 | -0.10071 | -0.064087 | 0.80566 | 0.93841 | -0.11902 | -0.19684 | 0.013733 | -0.064087 | -0.036621 | -0.036621 | 0.054932 | 0.091553 | 0.49896 | 0.53558 | 0.2243 | -0.013733 | 0.26093 | 0.022888
1298.1719 | 122.5067 | -53.0914 | 69.0353 | -0.10765 | 0.0013416 | 0.069613 | 0.070201 | -0.00066869 | 0.00059397 | 4.2808e-05 | -0.071503 | 0.0012411 | 0.065134 | 0.060842 | -0.0010361 | -0.0011257 | -0.0027621 | -0.082397 | 69.0491 | -53.2608 | 69.0079 | -52.2995 | -53.3295 | -53.2471 | -0.018311 | 69.0491 | -53.1647 | 69.1589 | -52.6382 | -53.1601 | -52.8717 | -0.16937 | 0.013733 | -0.027465 | -0.041198 | 0.79193 | 0.9613 | -0.15564 | -0.23804 | 0.013733 | -0.068665 | -0.073242 | 0.013733 | 0.1236 | 0.10986 | 0.45319 | 0.52643 | 0.21973 | -0.068664 | 0.29297 | 0.004578
1299.1719 | 122.4701 | -52.9999 | 69.1269 | -0.10765 | -0.00015382 | 0.069827 | 0.067851 | 0.0031765 | 0.001662 | -0.0010254 | -0.080658 | 0.00060029 | 0.064493 | 0.068532 | -0.0027451 | -0.0026211 | 0.0023647 | -0.086975 | 69.1132 | -53.1921 | 69.0445 | -52.2125 | -53.2608 | -53.215 | -0.009155 | 69.0994 | -53.0823 | 69.2322 | -52.5558 | -53.096 | -52.803 | -0.19226 | -0.013733 | -0.082397 | -0.068664 | 0.78735 | 0.97961 | -0.21515 | -0.26093 | -0.022888 | -0.068664 | -0.082397 | -0.027466 | 0.10529 | 0.13275 | 0.44403 | 0.52643 | 0.19684 | -0.09613 | 0.27923 | -0.013733
1300.1719 | 122.5159 | -52.9678 | 69.1727 | -0.11528 | 5.9776e-05 | 0.0694 | 0.072123 | -0.00088229 | 0.0040118 | -0.001239 | -0.08371 | 0.0020957 | 0.06257 | 0.064687 | -0.0010361 | -0.0028347 | 0.0017239 | -0.10529 | 69.2093 | -53.1326 | 69.104 | -52.1713 | -53.2013 | -53.1509 | -0.068665 | 69.191 | -53.0319 | 69.2917 | -52.4826 | -53.0319 | -52.7573 | -0.16479 | 0.036622 | -0.068664 | -0.10529 | 0.79651 | 0.9613 | -0.18311 | -0.23346 | -0.018311 | -0.068665 | -0.064087 | 0.018311 | 0.11902 | 0.10071 | 0.48523 | 0.54932 | 0.21057 | -0.064087 | 0.27466 | 0
1301.1719 | 122.4335 | -52.8809 | 69.2596 | -0.15343 | 0.00070078 | 0.067691 | 0.069773 | 0.00018591 | -0.00090143 | 0.002179 | -0.092865 | 0.0012411 | 0.064066 | 0.062337 | -0.0012497 | -0.0036893 | 0.0012967 | -0.10529 | 69.3054 | -53.0502 | 69.1635 | -52.121 | -53.1555 | -53.0823 | -0.073242 | 69.2642 | -52.9541 | 69.3649 | -52.4002 | -52.9587 | -52.6978 | -0.16937 | 0.045776 | -0.096131 | -0.14191 | 0.75989 | 0.92926 | -0.20142 | -0.27466 | -0.032043 | -0.10529 | -0.073243 | 0.004577 | 0.10529 | 0.10071 | 0.48065 | 0.55389 | 0.1831 | -0.07782 | 0.25635 | -0.004577
1302.1719 | 122.4197 | -52.803 | 69.342 | -0.10002 | -0.00079462 | 0.070468 | 0.068278 | 0.0033901 | -0.0013288 | 0.00068361 | 0.0017395 | 0.00038669 | 0.066629 | 0.065328 | -0.00060892 | -0.0064663 | 0.0019375 | -0.086975 | 69.3695 | -52.9541 | 69.2642 | -51.9882 | -53.0182 | -52.9358 | -0.036621 | 69.3604 | -52.8534 | 69.4244 | -52.3499 | -52.8488 | -52.5284 | -0.15106 | 0.027466 | -0.07782 | -0.10529 | 0.81482 | 0.96588 | -0.13275 | -0.21515 | 0.018311 | -0.064086 | -0.050354 | 0.018311 | 0.082397 | 0.064086 | 0.45319 | 0.50354 | 0.27466 | -0.045776 | 0.32501 | 0.004578
1303.1719 | 122.406 | -52.7481 | 69.4382 | -0.054245 | 0.0017688 | 0.069827 | 0.070201 | 0.00039951 | 0.00016657 | 0.0030334 | -0.092865 | 0.0012411 | 0.065347 | 0.070028 | -0.0016769 | -0.0026211 | -0.00083948 | -0.10529 | 69.4519 | -52.9312 | 69.2734 | -51.9791 | -53.0319 | -52.9633 | -0.11444 | 69.4336 | -52.8076 | 69.4656 | -52.3224 | -52.8488 | -52.597 | -0.1831 | 0.013733 | -0.16479 | -0.17853 | 0.76904 | 0.95215 | -0.21515 | -0.28381 | -0.032044 | -0.10071 | -0.059509 | -0.004577 | 0.027466 | 0.032043 | 0.42572 | 0.48523 | 0.15106 | -0.10071 | 0.21057 | -0.041199
1304.1719 | 122.4335 | -52.6932 | 69.5114 | -0.084763 | 0.0017688 | 0.07175 | 0.070842 | 0.00082671 | -0.0026104 | 0.00047001 | -0.092865 | 0.0018821 | 0.065988 | 0.069814 | 0.0015275 | -0.0041165 | -0.00062588 | -0.10529 | 69.5206 | -52.8625 | 69.3466 | -51.8967 | -52.9449 | -52.89 | -0.11902 | 69.5023 | -52.7252 | 69.5251 | -52.2675 | -52.803 | -52.5192 | -0.16937 | 0.009155 | -0.16479 | -0.17395 | 0.79651 | 0.96588 | -0.19684 | -0.25177 | -0.027466 | -0.082397 | -0.032044 | -0.009156 | 0.013732 | 0.022888 | 0.42572 | 0.45776 | 0.17395 | -0.10986 | 0.20599 | -0.07782
1305.1719 | 122.4426 | -52.5421 | 69.603 | -0.084763 | -0.001222 | 0.068973 | 0.068064 | -0.0040867 | 0.00080757 | 0.00025641 | -0.019623 | 0.0010275 | 0.065347 | 0.05956 | 0.00088648 | -0.0034757 | -0.00019868 | -0.07782 | 69.5984 | -52.684 | 69.4839 | -51.7868 | -52.6657 | -52.7115 | -0.009155 | 69.5801 | -52.6154 | 69.6533 | -52.0798 | -52.6291 | -52.3819 | -0.14191 | -0.004577 | -0.11902 | -0.11444 | 0.75531 | 0.89722 | -0.16937 | -0.1236 | -0.027466 | 0.018311 | -0.073242 | -0.022888 | 0.050354 | 0.073242 | 0.46234 | 0.53558 | 0.16022 | -0.086975 | 0.23346 | -0.013733
1306.1719 | 122.4655 | -52.4963 | 69.6762 | -0.11528 | 0.0015552 | 0.071322 | 0.071269 | 0.00082671 | -0.0015424 | 0.00089721 | -0.074554 | 0.00038669 | 0.066629 | 0.068105 | -0.00060892 | -0.0019803 | 0.00086952 | -0.10986 | 69.6808 | -52.6474 | 69.5572 | -51.6632 | -52.7298 | -52.6886 | -0.082397 | 69.6899 | -52.533 | 69.7449 | -52.002 | -52.5558 | -52.2903 | -0.15106 | 0.004578 | -0.11902 | -0.1236 | 0.83313 | 0.98419 | -0.19226 | -0.23346 | -0.041199 | -0.082397 | -0.036621 | 0.013733 | 0.068665 | 0.054932 | 0.49439 | 0.53101 | 0.20599 | -0.059509 | 0.24262 | -0.022888
1307.1719 | 122.4747 | -52.3682 | 69.7952 | -0.10765 | -0.0016492 | 0.067477 | 0.068705 | -0.0034457 | 0.0014484 | -0.0016662 | 0.02005 | -4.0508e-05 | 0.066202 | 0.058492 | -0.00039512 | -0.0060391 | -0.00041228 | -0.068665 | 69.7723 | -52.5375 | 69.7037 | -51.5167 | -52.5696 | -52.5101 | -0.009155 | 69.7678 | -52.4094 | 69.8547 | -51.8692 | -52.4048 | -52.1759 | -0.16937 | -0.022888 | -0.091553 | -0.068665 | 0.85144 | 1.0208 | -0.14191 | -0.20142 | 0.027466 | -0.032043 | -0.041199 | -0.027466 | 0.059509 | 0.086975 | 0.49896 | 0.54016 | 0.19226 | -0.036621 | 0.23346 | 0.004578
1308.1719 | 122.406 | -52.3132 | 69.8868 | -0.092392 | -0.00036742 | 0.069186 | 0.072551 | 0.00018591 | 0.0042254 | -0.00059819 | -0.10812 | 0.0023093 | 0.065347 | 0.0696 | -0.0018905 | -0.0024075 | -0.00019868 | -0.1236 | 69.8776 | -52.4643 | 69.7174 | -51.503 | -52.6016 | -52.5146 | -0.11902 | 69.8639 | -52.3636 | 69.8959 | -51.8738 | -52.446 | -52.1713 | -0.15106 | -0.009155 | -0.16937 | -0.16022 | 0.81024 | 0.9613 | -0.20142 | -0.28839 | -0.050354 | -0.13733 | -0.050354 | -0.022888 | 0.009155 | 0.032043 | 0.43945 | 0.48981 | 0.14191 | -0.13275 | 0.19226 | -0.082398
1309.1719 | 122.4792 | -52.2034 | 69.9509 | -0.084763 | 0.00091438 | 0.070468 | 0.06956 | -0.00045509 | -0.00090143 | 0.0034606 | -0.089813 | -0.00046791 | 0.065988 | 0.068746 | -0.00082252 | -0.0026211 | -0.00083948 | -0.11444 | 69.9738 | -52.3499 | 69.8593 | -51.3611 | -52.4368 | -52.3682 | -0.050354 | 69.9509 | -52.2583 | 70.0653 | -51.709 | -52.2537 | -51.9836 | -0.14648 | 0.022888 | -0.091553 | -0.11444 | 0.84229 | 0.98877 | -0.16479 | -0.23346 | -0.01831 | -0.086975 | -0.054932 | 0 | 0.11444 | 0.11444 | 0.49439 | 0.54932 | 0.21973 | -0.050354 | 0.27466 | 0.004578
1310.1719 | 122.3465 | -52.1347 | 70.0882 | -0.046616 | 0.0041186 | 0.071322 | 0.070414 | -0.0023777 | -0.00090143 | -0.00038459 | -0.11423 | 0.0010275 | 0.064493 | 0.071309 | 0.00088648 | -0.0024075 | -0.00019868 | -0.29755 | 70.0333 | -52.2949 | 69.9142 | -51.3474 | -52.4002 | -52.3315 | -0.13733 | 70.0699 | -52.1759 | 70.0699 | -51.709 | -52.2629 | -51.9882 | -0.16022 | -0.054932 | -0.17395 | -0.11902 | 0.78735 | 0.94757 | -0.19684 | -0.2655 | -0.036621 | -0.10529 | -0.041198 | -0.018311 | -0.018311 | 0 | 0.42572 | 0.46692 | 0.14648 | -0.12817 | 0.18768 | -0.086975
1311.1719 | 122.3969 | -52.0294 | 70.1569 | -0.10002 | -0.00015382 | 0.070041 | 0.069133 | -0.0010959 | -0.00026063 | 0.003247 | 0.029205 | 0.0031637 | 0.066202 | 0.069173 | -0.00039512 | -0.0039029 | -0.00083948 | -0.11902 | 70.1889 | -52.1851 | 70.0104 | -51.3657 | -52.5925 | -52.1393 | -0.23804 | 70.1843 | -52.0844 | 70.1614 | -51.6312 | -52.1896 | -51.8967 | -0.15564 | 0.032044 | -0.14648 | -0.17853 | 0.66376 | 0.8194 | -0.10986 | -0.56305 | 0.045777 | -0.40741 | -0.054932 | 0.027466 | 0.004578 | -0.022888 | 0.39825 | 0.45319 | 0.13275 | -0.16022 | 0.18768 | -0.10529
1312.1719 | 122.4792 | -51.9012 | 70.2576 | -0.16106 | 0.002837 | 0.069827 | 0.069346 | 0.00082671 | 0.00059397 | 0.00047001 | -0.05014 | -4.0508e-05 | 0.063638 | 0.059133 | -0.00039512 | -0.0021939 | 0.00065592 | -0.10529 | 70.2759 | -52.0706 | 70.1523 | -51.0223 | -52.1164 | -52.0706 | -0.041199 | 70.2255 | -51.9699 | 70.3583 | -51.4069 | -51.9608 | -51.6815 | -0.16937 | 0.018311 | -0.10529 | -0.1236 | 0.87891 | 1.0483 | -0.16937 | -0.21515 | 0 | -0.045776 | -0.068665 | -0.032043 | 0.10071 | 0.13275 | 0.49439 | 0.56305 | 0.21973 | -0.059509 | 0.28839 | 0.009156
1313.1719 | 122.3969 | -51.8005 | 70.372 | 0.060196 | 0.0013416 | 0.07175 | 0.069987 | 0.00061311 | -0.0013288 | -0.0020936 | -0.077606 | 0.0012411 | 0.064066 | 0.071309 | 0.0011001 | -0.00091214 | -0.0021213 | -0.041199 | 70.4224 | -51.9974 | 70.2072 | -50.9857 | -52.1301 | -52.0523 | -0.14191 | 70.3766 | -51.9012 | 70.3583 | -51.4023 | -51.9745 | -51.6815 | -0.19684 | 0.050354 | -0.1648 | -0.21515 | 0.81482 | 1.0117 | -0.25177 | -0.32959 | -0.054932 | -0.13275 | -0.10071 | 0.004578 | -0.013733 | -0.018311 | 0.39825 | 0.49896 | 0.11902 | -0.17395 | 0.21973 | -0.073242
1314.1719 | 122.3785 | -51.7044 | 70.4681 | -0.11528 | -0.00079462 | 0.070682 | 0.072764 | 0.00082671 | 0.0029438 | -0.0016662 | -0.025726 | 0.0016685 | 0.065988 | 0.066396 | 0.0019547 | -0.0039029 | 0.00086952 | -0.10986 | 70.491 | -51.8829 | 70.4819 | -50.6653 | -52.1027 | -51.9745 | -0.14648 | 70.4773 | -51.7685 | 70.4636 | -51.3153 | -51.8738 | -51.5854 | -0.17853 | 0.022888 | 0.013733 | -0.009155 | 1.0391 | 1.2177 | -0.27008 | -0.39825 | -0.091552 | -0.21973 | -0.064087 | 0.009155 | -0.004578 | -0.013733 | 0.3891 | 0.45319 | 0.11902 | -0.16937 | 0.18311 | -0.10529
1315.1719 | 122.4426 | -51.5991 | 70.5368 | -0.10765 | 0.002837 | 0.069613 | 0.070201 | -0.00024129 | 0.001662 | -0.00017099 | 0.035309 | 0.00017309 | 0.066415 | 0.060201 | -0.00018152 | -0.0024075 | -0.0027621 | -0.073242 | 70.5597 | -51.7364 | 70.5002 | -50.6424 | -51.7914 | -51.7136 | -0.009155 | 70.5093 | -51.6403 | 70.6284 | -51.0956 | -51.6449 | -51.4114 | -0.13733 | 0.022888 | -0.036621 | -0.059509 | 0.95673 | 1.0941 | -0.11444 | -0.19226 | 0.022888 | -0.054932 | -0.041199 | -0.027466 | 0.091553 | 0.11902 | 0.50354 | 0.54474 | 0.18768 | -0.045776 | 0.22888 | -0.004577
1316.1719 | 122.4106 | -51.4847 | 70.6696 | -0.10765 | 0.00048698 | 0.071109 | 0.070628 | 0.0050991 | 0.001662 | -0.001239 | -0.065399 | 0.0012411 | 0.066415 | 0.067464 | -0.0016769 | -0.0041165 | 0.0040737 | -0.091553 | 70.7153 | -51.6495 | 70.5048 | -50.6424 | -51.7502 | -51.709 | -0.1236 | 70.6787 | -51.5442 | 70.6833 | -51.0681 | -51.6403 | -51.3565 | -0.1648 | 0.045776 | -0.16479 | -0.21057 | 0.84228 | 1.0071 | -0.2243 | -0.2655 | -0.059509 | -0.10071 | -0.059509 | 0.009155 | 0.013733 | 0.004578 | 0.41656 | 0.47607 | 0.12817 | -0.15564 | 0.18768 | -0.096131
1317.1719 | 122.4564 | -51.3748 | 70.7748 | -0.11528 | -0.00058102 | 0.070682 | 0.066142 | -0.0032321 | 0.0018756 | -0.00081179 | 0.013947 | 0.0016685 | 0.066629 | 0.060201 | 0.00067288 | -0.0036893 | -0.00083948 | -0.086975 | 70.7932 | -51.535 | 70.6879 | -50.4456 | -51.5671 | -51.5121 | -0.036621 | 70.7565 | -51.3977 | 70.8298 | -50.885 | -51.4435 | -51.1963 | -0.16022 | 0.018311 | -0.086975 | -0.10529 | 0.92926 | 1.0895 | -0.13733 | -0.19226 | 0.022888 | -0.032044 | -0.022888 | -0.01831 | 0.054932 | 0.073242 | 0.48981 | 0.5127 | 0.17853 | -0.068664 | 0.20142 | -0.045776
1318.1719 | 122.3648 | -51.3062 | 70.9351 | -0.061874 | -0.00058102 | 0.070254 | 0.069987 | -0.0013095 | 0.0040118 | 0.005597 | -0.089813 | 0.00081389 | 0.064706 | 0.063192 | -0.0021041 | -0.0021939 | 0.0010831 | -0.082397 | 70.9213 | -51.4526 | 70.7199 | -50.4547 | -51.5717 | -51.5259 | -0.13733 | 70.9076 | -51.3474 | 70.8618 | -50.8987 | -51.4938 | -51.1963 | -0.14649 | -0.013733 | -0.21515 | -0.20142 | 0.85144 | 0.99793 | -0.21973 | -0.2655 | -0.073242 | -0.11902 | -0.041199 | -0.027466 | -0.073243 | -0.045777 | 0.40741 | 0.44861 | 0.10986 | -0.18768 | 0.15106 | -0.14648
1319.1719 | 122.3602 | -51.1826 | 71.0541 | -0.061874 | 0.0019824 | 0.069827 | 0.070414 | 0.0033901 | -0.00026063 | 0.0030334 | -0.03183 | 0.0031637 | 0.066202 | 0.065114 | 0.00045928 | -0.0026211 | 0.00086952 | -0.10986 | 71.0403 | -51.3107 | 70.8435 | -50.3311 | -51.535 | -51.59 | 0.018311 | 71.0358 | -51.2421 | 71.0129 | -50.7339 | -51.3474 | -51.0544 | -0.12817 | -0.013733 | -0.21057 | -0.19684 | 0.85144 | 0.97961 | -0.40741 | -0.35248 | -0.27924 | -0.2243 | -0.059509 | -0.01831 | -0.041199 | -0.022889 | 0.44861 | 0.50812 | 0.12817 | -0.1648 | 0.18768 | -0.10529
1320.1719 | 122.3831 | -51.059 | 71.1731 | -0.10002 | 0.0017688 | 0.067904 | 0.066355 | 0.0055265 | 0.00038017 | -0.00038459 | -0.077606 | 0.00038669 | 0.067911 | 0.062337 | -0.0018905 | -0.0036893 | 0.0023647 | -0.11444 | 71.1411 | -51.21 | 70.9351 | -50.2167 | -51.3428 | -51.297 | -0.13275 | 71.1273 | -51.1414 | 71.0999 | -50.6882 | -51.123 | -50.9033 | -0.15106 | -0.032044 | -0.23804 | -0.20599 | 0.84228 | 0.99335 | -0.23804 | -0.28381 | -0.086975 | -0.13275 | -0.082397 | -0.045777 | -0.073242 | -0.027465 | 0.37079 | 0.45319 | 0.15564 | -0.064087 | 0.23804 | 0.01831
1321.1719 | 122.4792 | -50.9537 | 71.283 | -0.10765 | 0.0024096 | 0.068973 | 0.069133 | -0.00024129 | 0.001662 | 0.0036744 | -0.071503 | -4.0508e-05 | 0.065988 | 0.061483 | 0.0017411 | -0.0011257 | 0.0042875 | -0.10529 | 71.2738 | -51.1139 | 71.0449 | -50.1114 | -51.2512 | -51.1688 | -0.14191 | 71.2784 | -51.0086 | 71.3882 | -50.7477 | -51.2604 | -50.8347 | -0.16022 | -0.009155 | -0.23804 | -0.22888 | 0.84228 | 1.0025 | -0.21515 | -0.29755 | -0.054931 | -0.13733 | -0.054932 | -0.004578 | 0.10529 | 0.10986 | 0.20599 | 0.26093 | 0.11902 | -0.3067 | 0.17395 | -0.25177
1322.1719 | 122.4426 | -50.8118 | 71.3745 | -0.13817 | 0.002196 | 0.070468 | 0.068492 | -0.00024129 | -0.00068783 | -0.0014526 | 0.035309 | 0.00060029 | 0.064279 | 0.058492 | 0.00024568 | -0.0043301 | 0.0019375 | -0.09613 | 71.3379 | -50.9491 | 71.2555 | -49.8322 | -51.0086 | -50.9262 | -0.054932 | 71.3196 | -50.8713 | 71.402 | -50.2853 | -50.885 | -50.6516 | -0.13733 | -0.036621 | -0.11902 | -0.082398 | 0.97962 | 1.1169 | -0.11444 | -0.19684 | 0.022889 | -0.059509 | -0.059509 | -0.054932 | 0.027466 | 0.082398 | 0.52643 | 0.58594 | 0.16022 | -0.073242 | 0.21973 | -0.013733
1323.1719 | 122.3785 | -50.7202 | 71.5073 | -0.1458 | -0.0014356 | 0.068759 | 0.071055 | -0.0023777 | 0.00059397 | -0.0018798 | -0.053192 | 0.0020957 | 0.065988 | 0.063405 | -0.00039512 | -0.0026211 | -0.00019868 | -0.15106 | 71.4615 | -50.9125 | 71.3058 | -49.7269 | -50.9674 | -50.9262 | -0.13275 | 71.5073 | -50.7523 | 71.4752 | -50.2441 | -50.8667 | -50.5737 | -0.19226 | -0.045776 | -0.20142 | -0.15564 | 0.99335 | 1.1856 | -0.20599 | -0.24719 | -0.013732 | -0.054931 | -0.032043 | 0 | -0.032043 | -0.032043 | 0.47607 | 0.50812 | 0.14649 | -0.14648 | 0.17853 | -0.11444
1324.1719 | 122.3328 | -50.5829 | 71.6263 | -0.054245 | 5.9776e-05 | 0.071322 | 0.07191 | 0.0033901 | -0.00026063 | 0.00025641 | -0.074554 | 0.0023093 | 0.06492 | 0.06426 | -0.00060892 | -0.0024075 | -0.0019077 | -0.10529 | 71.6125 | -50.766 | 71.4386 | -49.7131 | -50.8575 | -50.8118 | -0.14648 | 71.608 | -50.6424 | 71.5622 | -50.1755 | -50.7751 | -50.5051 | -0.1831 | -0.013733 | -0.18768 | -0.17395 | 0.86975 | 1.0529 | -0.22888 | -0.27466 | -0.045777 | -0.091553 | -0.059509 | -0.018311 | -0.064087 | -0.045776 | 0.40741 | 0.46692 | 0.07782 | -0.19226 | 0.13733 | -0.13275
1325.1719 | 122.5113 | -50.441 | 71.6766 | -0.16106 | 0.002837 | 0.072391 | 0.069773 | 0.00082671 | 0.0012348 | -0.00081179 | 0.0078429 | -4.0508e-05 | 0.06492 | 0.05956 | -0.0016769 | -0.0028347 | 0.0042875 | -0.059509 | 71.6675 | -50.5829 | 71.7316 | -49.3057 | -50.6744 | -50.5875 | -0.009155 | 71.6629 | -50.5051 | 71.7957 | -49.8917 | -50.5051 | -50.2487 | -0.14191 | -0.009156 | 0.054931 | 0.064087 | 1.1353 | 1.2772 | -0.14648 | -0.23346 | -0.004577 | -0.091552 | -0.064087 | -0.013733 | 0.11902 | 0.13275 | 0.54932 | 0.6134 | 0.19226 | -0.064087 | 0.25635 | 0
1326.1719 | 122.3694 | -50.3723 | 71.8781 | -0.10765 | -0.00015382 | 0.070682 | 0.070414 | 0.0040311 | 0.0012348 | 0.0011108 | -0.098968 | 0.00017309 | 0.065347 | 0.064901 | -0.0023177 | -0.0056119 | 0.0038601 | -0.13733 | 71.8552 | -50.5051 | 71.6675 | -49.4431 | -50.6332 | -50.5692 | -0.15106 | 71.846 | -50.3906 | 71.8689 | -49.9832 | -50.4547 | -50.1846 | -0.13275 | -0.022888 | -0.21057 | -0.18768 | 0.92926 | 1.062 | -0.19684 | -0.26093 | -0.064087 | -0.12817 | -0.018311 | -0.032044 | -0.009156 | 0.022888 | 0.3891 | 0.40741 | 0.18768 | -0.082398 | 0.20599 | -0.064087
1327.1719 | 122.4518 | -50.2075 | 71.9604 | -0.11528 | -0.00036742 | 0.071109 | 0.066142 | -0.0032321 | 0.0010212 | -0.0010254 | 0.023102 | -0.00089511 | 0.067056 | 0.061056 | -0.00082252 | -0.0041165 | 0.0010831 | -0.086975 | 72.0062 | -50.3311 | 71.8597 | -49.2142 | -50.4227 | -50.3494 | -0.036621 | 71.9284 | -50.2533 | 71.9971 | -49.6719 | -50.267 | -50.0336 | -0.1236 | 0.045777 | -0.10071 | -0.14648 | 0.99335 | 1.1169 | -0.14191 | -0.21515 | -0.01831 | -0.091552 | -0.045776 | -0.032043 | 0.036621 | 0.068664 | 0.53558 | 0.58136 | 0.17395 | -0.059509 | 0.21973 | -0.013733
1328.1719 | 122.3465 | -50.1343 | 72.1207 | -0.069504 | 0.00027338 | 0.070254 | 0.070414 | -0.0010959 | 0.0020892 | 4.2808e-05 | -0.077606 | 0.0029501 | 0.067056 | 0.068319 | -0.0010361 | -0.0013393 | -0.0010531 | -0.10986 | 72.1161 | -50.2533 | 71.8964 | -49.2096 | -50.3906 | -50.3357 | -0.13733 | 72.1298 | -50.1343 | 72.0703 | -49.6628 | -50.2579 | -49.9878 | -0.11902 | -0.004578 | -0.22431 | -0.21973 | 0.92468 | 1.0437 | -0.20142 | -0.25635 | -0.082397 | -0.13733 | 0 | 0.009155 | -0.050355 | -0.05951 | 0.4715 | 0.4715 | 0.14648 | -0.1236 | 0.14648 | -0.1236
1329.1719 | 122.3694 | -49.9558 | 72.258 | -0.13054 | 0.00070078 | 0.069186 | 0.071482 | -0.00066869 | 0.0020892 | 0.0019654 | -0.080658 | 0.0012411 | 0.063425 | 0.071737 | -0.00018152 | -0.0011257 | 0.00086952 | -0.10529 | 72.2626 | -50.1297 | 72.0337 | -49.0494 | -50.2441 | -50.1892 | -0.16479 | 72.2626 | -50.0381 | 72.1848 | -49.5117 | -50.1251 | -49.8642 | -0.17395 | 0.004577 | -0.22431 | -0.22888 | 0.90637 | 1.0803 | -0.23346 | -0.28839 | -0.059509 | -0.11444 | -0.082397 | 0.004577 | -0.073243 | -0.07782 | 0.44403 | 0.52643 | 0.091553 | -0.16937 | 0.17395 | -0.086975
1330.1719 | 122.374 | -49.8276 | 72.3953 | -0.077133 | -0.00036742 | 0.071963 | 0.070201 | -0.0032321 | -0.0032514 | -0.0020936 | -0.047089 | -4.0508e-05 | 0.065561 | 0.067678 | 0.0019547 | -0.0024075 | 0.0012967 | -0.11444 | 72.3816 | -49.9832 | 72.345 | -48.9075 | -50.1709 | -49.9924 | -0.13733 | 72.3862 | -49.8917 | 72.3312 | -49.3744 | -49.9832 | -49.7086 | -0.15564 | -0.013733 | -0.050354 | -0.036621 | 0.9201 | 1.0757 | -0.16479 | -0.34332 | -0.009156 | -0.18768 | -0.064087 | -0.009156 | -0.064087 | -0.054931 | 0.45319 | 0.51727 | 0.11902 | -0.15564 | 0.18311 | -0.091552
1331.1719 | 122.3785 | -49.6948 | 72.5372 | -0.077133 | 0.0024096 | 0.0694 | 0.071055 | 0.00039951 | 0.00059397 | 0.001538 | -0.077606 | 0.0020957 | 0.063852 | 0.062337 | -0.00018152 | -0.0021939 | 0.0021511 | -0.07782 | 72.5143 | -49.8642 | 72.3175 | -48.7473 | -49.9558 | -49.9008 | -0.13733 | 72.5098 | -49.7498 | 72.4594 | -49.2279 | -49.8825 | -49.6124 | -0.16937 | -0.022888 | -0.21973 | -0.19684 | 0.94757 | 1.1169 | -0.20599 | -0.26093 | -0.036621 | -0.091553 | -0.054932 | -0.027465 | -0.077819 | -0.050354 | 0.46692 | 0.52185 | 0.082397 | -0.18768 | 0.13733 | -0.13275
1332.1719 | 122.5021 | -49.53 | 72.6196 | -0.16106 | 0.00070078 | 0.071109 | 0.071269 | -0.0021641 | 0.00059397 | -0.0010254 | 0.087189 | 0.0010275 | 0.066202 | 0.062978 | -0.00060892 | -0.0036893 | -0.00019868 | -0.041199 | 72.6013 | -49.6582 | 72.5555 | -48.5641 | -49.9924 | -49.5987 | -0.027466 | 72.583 | -49.5895 | 72.7066 | -48.967 | -49.585 | -49.3286 | -0.12817 | -0.018311 | -0.064087 | -0.045776 | 0.96588 | 1.0941 | -0.068665 | -0.46234 | 0.059509 | -0.33417 | -0.05951 | -0.036621 | 0.086975 | 0.1236 | 0.56305 | 0.62256 | 0.20142 | -0.054932 | 0.26093 | 0.004578
1333.1719 | 122.4564 | -49.4247 | 72.821 | -0.069504 | 0.00048698 | 0.069186 | 0.068064 | -0.0038729 | 0.0037982 | 0.00089721 | -0.092865 | 0.0027365 | 0.066202 | 0.061696 | 3.2078e-05 | -0.0032619 | -0.0031895 | -0.11444 | 72.7753 | -49.5667 | 72.5647 | -48.5046 | -49.6948 | -49.6307 | -0.13733 | 72.7707 | -49.4797 | 72.7936 | -49.1638 | -49.7955 | -49.4019 | -0.14191 | -0.045776 | -0.25635 | -0.21057 | 0.9201 | 1.062 | -0.20599 | -0.27008 | -0.064087 | -0.12817 | -0.054931 | -0.050354 | -0.027466 | 0.022888 | 0.26093 | 0.31586 | 0.022889 | -0.37079 | 0.07782 | -0.31586
1334.1719 | 122.4106 | -49.2737 | 72.9034 | -0.084763 | 0.002837 | 0.071963 | 0.065928 | -0.0047275 | 0.0025166 | -0.00081179 | 0.062775 | 0.0010275 | 0.063852 | 0.061483 | 0.00067288 | -0.0036893 | -0.0021213 | -0.068665 | 72.9172 | -49.3835 | 72.7982 | -48.2208 | -49.4431 | -49.3698 | 0.068665 | 72.8897 | -49.3332 | 72.963 | -48.7152 | -49.3469 | -49.0814 | -0.10986 | 0.013733 | -0.10528 | -0.11902 | 1.0529 | 1.1627 | -0.09613 | -0.16937 | 0.013733 | -0.059509 | -0.059509 | -0.013733 | 0.05951 | 0.073243 | 0.55847 | 0.61798 | 0.19226 | -0.073242 | 0.25177 | -0.013733
1335.1719 | 122.4655 | -49.118 | 73.082 | -0.16106 | 0.0019824 | 0.070682 | 0.068705 | -0.0040867 | -4.7032e-05 | 0.0030334 | 0.016998 | 0.00017309 | 0.067484 | 0.065114 | 0.00024568 | -0.0041165 | 0.00065592 | -0.07782 | 73.0545 | -49.2554 | 72.9538 | -48.0423 | -49.3057 | -49.2279 | -0.041199 | 73.0453 | -49.1547 | 73.1003 | -48.5687 | -49.1913 | -48.9349 | -0.13733 | -0.027466 | -0.12817 | -0.10071 | 1.0757 | 1.2131 | -0.10986 | -0.18768 | 0.027466 | -0.050354 | -0.036621 | -0.036621 | 0.018311 | 0.054932 | 0.54932 | 0.58594 | 0.18311 | -0.073242 | 0.21973 | -0.036621
1336.1719 | 122.3557 | -48.9853 | 73.2559 | -0.077133 | 0.0019824 | 0.073245 | 0.072764 | 0.0012539 | -0.001115 | -0.00059819 | -0.080658 | 0.00081389 | 0.063211 | 0.065755 | -0.0025315 | -0.0041165 | 0.0025785 | -0.091553 | 73.2193 | -49.1455 | 73.0316 | -48.0103 | -49.2508 | -49.1867 | -0.16479 | 73.2147 | -49.0402 | 73.1644 | -48.5275 | -49.1776 | -48.912 | -0.16022 | -0.036621 | -0.2243 | -0.18768 | 0.97504 | 1.1353 | -0.20142 | -0.2655 | -0.041199 | -0.10529 | -0.054931 | -0.041198 | -0.091552 | -0.050354 | 0.45776 | 0.5127 | 0.073243 | -0.19226 | 0.12817 | -0.13733
1337.1719 | 122.4792 | -48.7885 | 73.3704 | -0.10002 | 0.00070078 | 0.068759 | 0.066996 | -0.0025913 | 0.001662 | -0.00017099 | 0.041412 | 0.0020957 | 0.065561 | 0.063833 | -0.0010361 | -0.0043301 | 0.0038601 | -0.09613 | 73.3521 | -48.9349 | 73.2513 | -47.7402 | -48.9761 | -48.9212 | -0.041199 | 73.3383 | -48.8708 | 73.3932 | -48.2529 | -48.88 | -48.6603 | -0.14648 | -0.01831 | -0.11902 | -0.10071 | 1.0483 | 1.1948 | -0.13275 | -0.18768 | 0.013733 | -0.041198 | -0.082398 | -0.032043 | 0.022889 | 0.054932 | 0.53558 | 0.61798 | 0.12817 | -0.091553 | 0.21057 | -0.009155
1338.1719 | 122.406 | -48.6603 | 73.558 | -0.11528 | 0.0041186 | 0.071109 | 0.070414 | 0.0044583 | 0.00080757 | -0.00081179 | -0.08371 | 0.0018821 | 0.068979 | 0.067464 | -0.00039512 | 0.00015586 | -0.0046847 | -0.082397 | 73.5168 | -48.8251 | 73.3475 | -47.6578 | -48.9304 | -48.8754 | -0.11902 | 73.5352 | -48.7198 | 73.4802 | -48.1979 | -48.8617 | -48.587 | -0.1648 | -0.041198 | -0.21057 | -0.16937 | 1.0025 | 1.1673 | -0.21515 | -0.27008 | -0.050354 | -0.10529 | -0.05951 | -0.022888 | -0.077819 | -0.054931 | 0.46234 | 0.52185 | 0.073242 | -0.20142 | 0.13275 | -0.14191
1339.1719 | 122.4564 | -48.4909 | 73.6679 | -0.11528 | 5.9776e-05 | 0.071963 | 0.066355 | 0.0018949 | 0.001662 | 0.0017518 | 0.0047913 | 0.0018821 | 0.066415 | 0.057638 | 0.00067288 | -0.0047573 | 0.00022852 | -0.045776 | 73.6542 | -48.6191 | 73.6038 | -47.3785 | -48.6557 | -48.5962 | 0.009155 | 73.6496 | -48.555 | 73.7366 | -47.9324 | -48.5596 | -48.3398 | -0.12817 | -0.013733 | -0.064087 | -0.050354 | 1.1124 | 1.2405 | -0.10528 | -0.16479 | 0.022889 | -0.036621 | -0.064087 | -0.018311 | 0.068664 | 0.086975 | 0.55847 | 0.62256 | 0.15106 | -0.068664 | 0.21515 | -0.004577
1340.1719 | 122.3511 | -48.3719 | 73.8693 | -0.10765 | 0.0015552 | 0.071536 | 0.072337 | 0.0012539 | 0.0020892 | 0.0013244 | -0.089813 | 0.0010275 | 0.065134 | 0.067037 | 0.00045928 | -0.0028347 | -0.0021213 | -0.10071 | 73.851 | -48.5413 | 73.6221 | -47.3877 | -48.642 | -48.5825 | -0.16022 | 73.8327 | -48.436 | 73.8052 | -47.8821 | -48.5596 | -48.2758 | -0.16937 | -0.018311 | -0.24719 | -0.22888 | 0.98419 | 1.1536 | -0.21057 | -0.27008 | -0.041198 | -0.10071 | -0.064087 | -0.036621 | -0.064087 | -0.027466 | 0.48981 | 0.55389 | 0.09613 | -0.18768 | 0.16022 | -0.1236
1341.1719 | 122.3511 | -48.23 | 73.9975 | -0.077133 | -0.00079462 | 0.072391 | 0.070628 | 0.0018949 | 0.0012348 | 0.0013244 | -0.065399 | 0.00081389 | 0.064066 | 0.067464 | -0.00060892 | -0.0015529 | -0.0019077 | -0.073242 | 73.9975 | -48.3627 | 73.7686 | -47.1954 | -48.468 | -48.4222 | -0.15106 | 74.0112 | -48.2712 | 73.9334 | -47.7356 | -48.4177 | -48.143 | -0.13275 | 0 | -0.22888 | -0.22888 | 1.0345 | 1.1673 | -0.19226 | -0.23804 | -0.059509 | -0.10529 | -0.041199 | 0.013732 | -0.064087 | -0.077819 | 0.49438 | 0.53558 | 0.086975 | -0.18768 | 0.12817 | -0.14649
1342.1719 | 122.4838 | -48.0286 | 74.1165 | -0.16869 | -0.00058102 | 0.072391 | 0.068064 | 0.00018591 | 0.0031574 | 0.00068361 | -0.0013123 | 0.00038669 | 0.064066 | 0.065755 | -0.00060892 | -0.0034757 | 0.0012967 | -0.045776 | 74.1257 | -48.1659 | 74.0616 | -46.8842 | -48.2849 | -48.3582 | 0.13275 | 74.0799 | -48.0606 | 74.2264 | -47.4152 | -48.0881 | -47.8546 | -0.13733 | 0.009155 | -0.054932 | -0.064087 | 1.1444 | 1.2817 | -0.32959 | -0.25635 | -0.19226 | -0.11902 | -0.032044 | -0.036621 | 0.10986 | 0.14648 | 0.6134 | 0.64545 | 0.17395 | -0.05951 | 0.20599 | -0.027466
1343.1719 | 122.3785 | -47.8821 | 74.2722 | -0.11528 | -0.00036742 | 0.073459 | 0.071482 | -0.0025913 | 0.00016657 | 0.0030334 | -0.068451 | 0.0014547 | 0.066415 | 0.065755 | -0.00018152 | -0.0036893 | 0.0038601 | -0.10986 | 74.3088 | -48.0835 | 74.0707 | -46.907 | -48.1934 | -48.1247 | -0.14191 | 74.2905 | -47.9691 | 74.2676 | -47.406 | -48.0835 | -47.8134 | -0.20142 | 0.036621 | -0.20142 | -0.23804 | 0.97504 | 1.1765 | -0.24262 | -0.31128 | -0.041199 | -0.10986 | -0.086975 | 0.01831 | -0.004578 | -0.022888 | 0.47607 | 0.56305 | 0.068664 | -0.20142 | 0.15564 | -0.11444
1344.1719 | 122.4792 | -47.7081 | 74.469 | -0.12291 | -0.00079462 | 0.071963 | 0.066569 | -0.00024129 | 0.00059397 | 0.001538 | 0.068878 | 0.0016685 | 0.065988 | 0.058706 | 0.00024568 | -0.0034757 | 0.00086952 | -0.022888 | 74.4598 | -47.8592 | 74.3683 | -46.5637 | -47.8409 | -47.7722 | 0.073242 | 74.4186 | -47.7676 | 74.5239 | -47.0581 | -47.7036 | -47.525 | -0.15106 | -0.009155 | -0.10071 | -0.091553 | 1.1444 | 1.2955 | -0.064087 | -0.13275 | 0.086975 | 0.018311 | -0.059509 | -0.050354 | 0.054932 | 0.10529 | 0.65002 | 0.70953 | 0.18311 | 0.004578 | 0.24262 | 0.064087
1345.1719 | 122.438 | -47.5616 | 74.6109 | -0.11528 | -0.00058102 | 0.07175 | 0.070201 | -0.00045509 | 0.00038017 | -0.00038459 | -0.044037 | -0.00025411 | 0.068338 | 0.06661 | -0.0016769 | -0.0030483 | -0.0031895 | -0.10071 | 74.6384 | -47.7264 | 74.4507 | -46.5134 | -47.818 | -47.7539 | -0.10986 | 74.6063 | -47.6624 | 74.6292 | -47.0215 | -47.6852 | -47.4472 | -0.16479 | 0.027466 | -0.16022 | -0.18768 | 1.0483 | 1.2131 | -0.19226 | -0.25635 | -0.027466 | -0.091553 | -0.10071 | -0.004578 | 0.01831 | 0.022888 | 0.54016 | 0.64087 | 0.11444 | -0.1236 | 0.21515 | -0.022888
1346.1719 | 122.374 | -47.4335 | 74.7849 | -0.016098 | -0.00079462 | 0.073031 | 0.068919 | -0.00066869 | 0.00080757 | -0.0014526 | -0.080658 | 0.0023093 | 0.066629 | 0.066396 | 0.00024568 | -0.0026211 | -0.0010531 | -0.027466 | 74.8169 | -47.5296 | 74.6521 | -46.2891 | -47.5845 | -47.5662 | 0 | 74.7849 | -47.4564 | 74.794 | -46.8613 | -47.525 | -47.2824 | -0.09613 | 0.032044 | -0.13275 | -0.16479 | 1.1444 | 1.2405 | -0.13275 | -0.15106 | -0.036621 | -0.054932 | -0.022888 | 0 | 0.009155 | 0.009155 | 0.57221 | 0.59509 | 0.15106 | -0.091552 | 0.17395 | -0.068664
1347.1719 | 122.4426 | -47.2229 | 74.8672 | -0.10002 | 0.0030506 | 0.072604 | 0.070201 | -0.00045509 | 0.0023028 | 0.0019654 | -0.028778 | -0.00025411 | 0.063211 | 0.064473 | 3.2078e-05 | -0.0019803 | -0.00083948 | -0.036621 | 74.9084 | -47.3923 | 74.8627 | -46.0464 | -47.3785 | -47.319 | 0 | 74.9084 | -47.2504 | 75.032 | -46.5912 | -47.2778 | -46.9757 | -0.16937 | 0.041198 | -0.004578 | -0.045776 | 1.1765 | 1.3458 | -0.096131 | -0.15564 | 0.073242 | 0.013733 | -0.027466 | 0.041198 | 0.16479 | 0.1236 | 0.63171 | 0.65918 | 0.24719 | -0.054932 | 0.27466 | -0.027466
1348.1719 | 122.3465 | -47.0993 | 75.1373 | -0.10002 | 0.0017688 | 0.071963 | 0.070201 | -0.00045509 | -0.00047423 | 0.0011108 | -0.077606 | 0.00017309 | 0.06727 | 0.071096 | -0.0010361 | -0.0041165 | 0.00044212 | -0.036621 | 75.1328 | -47.2137 | 74.9725 | -45.9549 | -47.2687 | -47.2229 | -0.032043 | 75.1099 | -47.1039 | 75.1144 | -46.5134 | -47.2092 | -46.9666 | -0.11444 | -0.004578 | -0.16479 | -0.16022 | 1.1444 | 1.2589 | -0.1236 | -0.16937 | -0.009155 | -0.054932 | -0.004578 | -0.027466 | -0.022888 | 0.004578 | 0.58594 | 0.59051 | 0.13275 | -0.10986 | 0.13733 | -0.10529
1349.1719 | 122.3969 | -46.907 | 75.3067 | -0.061874 | -0.0016492 | 0.071109 | 0.071269 | 0.0016813 | 0.00059397 | -0.0016662 | -0.077606 | 0.0025229 | 0.066629 | 0.071096 | -0.00039512 | -0.0045437 | -0.0010531 | -0.022888 | 75.3387 | -47.0352 | 75.1556 | -45.7581 | -47.0856 | -47.0489 | -0.009155 | 75.293 | -46.9391 | 75.2884 | -46.3348 | -47.0261 | -46.7789 | -0.12817 | 0.032043 | -0.15106 | -0.1831 | 1.149 | 1.2772 | -0.14191 | -0.17853 | -0.013733 | -0.050354 | -0.032044 | -0.013733 | -0.018311 | -0.004578 | 0.5722 | 0.60425 | 0.12817 | -0.11902 | 0.16022 | -0.086975
1350.1719 | 122.3785 | -46.7194 | 75.4852 | -0.069504 | 0.0032642 | 0.071322 | 0.071269 | -0.00024129 | 0.00080757 | -0.00081179 | 0.013947 | 0.0018821 | 0.067056 | 0.065755 | -0.00039512 | -0.0019803 | 0.0027921 | -0.009155 | 75.4852 | -46.8704 | 75.3342 | -45.5383 | -46.9299 | -46.8109 | -0.013733 | 75.4623 | -46.7514 | 75.4807 | -46.1517 | -46.8338 | -46.6049 | -0.15106 | 0 | -0.15106 | -0.15106 | 1.181 | 1.3321 | -0.091553 | -0.21057 | 0.059509 | -0.05951 | -0.032044 | -0.022888 | -0.004577 | 0.018311 | 0.56763 | 0.59967 | 0.11444 | -0.11444 | 0.14648 | -0.082397
1351.1719 | 122.3785 | -46.5775 | 75.6775 | -0.038986 | 5.9776e-05 | 0.073245 | 0.069987 | -0.00088229 | -0.00068783 | 0.003247 | -0.077606 | -0.00025411 | 0.065561 | 0.068746 | -0.0018905 | -0.0024075 | 0.0010831 | -0.022888 | 75.6546 | -46.6965 | 75.4944 | -45.4056 | -46.7377 | -46.7056 | -0.004578 | 75.6363 | -46.5912 | 75.6409 | -45.9778 | -46.6873 | -46.4493 | -0.11902 | -0.022888 | -0.18311 | -0.16022 | 1.1719 | 1.2909 | -0.12817 | -0.16022 | -0.009155 | -0.041199 | -0.013733 | -0.041198 | -0.036621 | 0.004577 | 0.59967 | 0.6134 | 0.12817 | -0.10986 | 0.14191 | -0.09613
1352.1719 | 122.4106 | -46.376 | 75.8377 | -0.069504 | 0.00070078 | 0.072391 | 0.070628 | -0.0010959 | -0.0023968 | 0.003888 | -0.03183 | 0.0027365 | 0.065561 | 0.064046 | -0.00060892 | -0.0026211 | -0.00041228 | -0.12817 | 75.8148 | -46.5317 | 75.6271 | -45.2728 | -46.6187 | -46.5408 | -0.10529 | 75.8286 | -46.4355 | 75.8148 | -45.8176 | -46.5179 | -46.2524 | -0.15564 | -0.022889 | -0.21057 | -0.18768 | 1.1032 | 1.2589 | -0.16479 | -0.24261 | -0.009156 | -0.086975 | -0.059509 | -0.009156 | -0.022889 | -0.013733 | 0.55847 | 0.61798 | 0.1236 | -0.14191 | 0.18311 | -0.082397
1353.1719 | 122.3877 | -46.1975 | 76.0437 | -0.1458 | 0.00070078 | 0.072818 | 0.069133 | 0.0027493 | -0.00026063 | -0.0016662 | 0.023102 | 0.0029501 | 0.065988 | 0.068319 | -0.0014633 | -0.0021939 | 0.0021511 | -0.009155 | 76.0208 | -46.3303 | 75.8652 | -45.0027 | -46.4081 | -46.2616 | -0.009155 | 76.0162 | -46.2112 | 76.0162 | -45.5704 | -46.2845 | -46.0648 | -0.13275 | -0.022888 | -0.17853 | -0.15564 | 1.1948 | 1.3275 | -0.064087 | -0.21057 | 0.068664 | -0.07782 | -0.013733 | -0.027466 | -0.027466 | 0 | 0.62714 | 0.64087 | 0.13275 | -0.086975 | 0.14649 | -0.073242
1354.1719 | 122.406 | -45.9686 | 76.236 | -0.061874 | 0.001128 | 0.071963 | 0.069346 | 0.0031765 | -4.7032e-05 | -0.0018798 | -0.022675 | 0.0012411 | 0.064493 | 0.065755 | -0.0021041 | -0.0036893 | -0.0031895 | -0.013733 | 76.2177 | -46.1426 | 76.0529 | -44.8105 | -46.1746 | -46.1472 | 0.041199 | 76.2085 | -46.051 | 76.1719 | -45.2591 | -46.1014 | -45.8725 | -0.17395 | -0.018311 | -0.18311 | -0.16479 | 1.1581 | 1.3321 | -0.17853 | -0.20599 | -0.004578 | -0.032044 | -0.082397 | -0.027466 | -0.064087 | -0.036621 | 0.70953 | 0.79193 | 0.09613 | -0.13275 | 0.17853 | -0.050354
1355.1719 | 122.4655 | -45.7947 | 76.4328 | -0.12291 | 0.0024096 | 0.072177 | 0.070842 | 0.0018949 | 0.0010212 | 0.00047001 | -0.040985 | 0.0031637 | 0.06492 | 0.069387 | 3.2078e-05 | -0.0045437 | 0.0032193 | -0.10986 | 76.387 | -45.9869 | 76.1581 | -44.7556 | -46.0876 | -46.0144 | -0.15564 | 76.3779 | -45.9 | 76.4786 | -45.4193 | -46.17 | -45.845 | -0.19226 | -0.045776 | -0.27466 | -0.22888 | 1.0391 | 1.2314 | -0.21973 | -0.29297 | -0.027466 | -0.10071 | -0.10529 | -0.054931 | 0.045777 | 0.10071 | 0.37537 | 0.48065 | -0.050354 | -0.37537 | 0.054931 | -0.27008
1356.1719 | 122.4289 | -45.6161 | 76.6159 | -0.061874 | 0.0036914 | 0.072604 | 0.069773 | 0.0036039 | 0.00059397 | 0.00025641 | -0.013519 | 0.0027365 | 0.068979 | 0.06191 | 0.00088648 | -0.0028347 | 0.0010831 | -0.009155 | 76.5884 | -45.7535 | 76.4145 | -44.426 | -45.7764 | -45.7489 | 0.064087 | 76.561 | -45.6757 | 76.5839 | -45.0119 | -45.7077 | -45.5017 | -0.13733 | -0.027466 | -0.20142 | -0.17395 | 1.1902 | 1.3275 | -0.13275 | -0.16022 | 0.004578 | -0.022888 | -0.059509 | -0.054932 | -0.032044 | 0.022888 | 0.60425 | 0.66376 | 0.11444 | -0.091553 | 0.17395 | -0.032044
1357.1719 | 122.4243 | -45.4422 | 76.7578 | -0.069504 | -0.00036742 | 0.073245 | 0.069346 | 0.0012539 | 0.0037982 | 0.003888 | -0.03183 | 0.0029501 | 0.065988 | 0.063192 | 0.00024568 | -0.0045437 | -0.00062588 | -0.11444 | 76.7487 | -45.6024 | 76.5564 | -44.2932 | -45.6665 | -45.607 | -0.086975 | 76.7166 | -45.5063 | 76.767 | -44.8425 | -45.5612 | -45.3278 | -0.16022 | -0.009155 | -0.20142 | -0.19226 | 1.149 | 1.3092 | -0.16479 | -0.2243 | -0.004578 | -0.064087 | -0.064087 | -0.041198 | 0.009156 | 0.050354 | 0.59967 | 0.66376 | 0.11444 | -0.11902 | 0.17853 | -0.054931
1358.1719 | 122.4518 | -45.2225 | 76.9775 | -0.11528 | 0.0013416 | 0.072604 | 0.070842 | 0.0027493 | -0.0013288 | -0.00059819 | -0.044037 | 0.0023093 | 0.065347 | 0.063833 | -0.0016769 | -0.0013393 | -0.0012667 | 0 | 76.9592 | -45.3827 | 76.799 | -44.0414 | -45.4102 | -45.3781 | 0.032043 | 76.9501 | -45.282 | 76.9547 | -44.6182 | -45.3598 | -45.108 | -0.16022 | -0.01831 | -0.17853 | -0.16022 | 1.181 | 1.3412 | -0.15564 | -0.18768 | 0.004577 | -0.027466 | -0.059509 | -0.027466 | -0.022888 | 0.004578 | 0.60425 | 0.66376 | 0.11444 | -0.13733 | 0.17395 | -0.07782
1359.1719 | 122.3969 | -45.0668 | 77.1515 | -0.084763 | 0.0017688 | 0.071536 | 0.06956 | -0.0030185 | 0.0018756 | -0.00059819 | -0.0013123 | 0.0014547 | 0.06727 | 0.064473 | -0.00060892 | -0.0032619 | 0.0010831 | -0.022888 | 77.1744 | -45.1904 | 76.9501 | -43.9178 | -45.2454 | -45.1904 | 0 | 77.1515 | -45.0897 | 77.1286 | -44.4305 | -45.1859 | -44.9387 | -0.1236 | 0.022888 | -0.20142 | -0.2243 | 1.149 | 1.2726 | -0.1236 | -0.17853 | 0 | -0.054931 | -0.022889 | 0 | -0.022888 | -0.022888 | 0.63629 | 0.65918 | 0.12817 | -0.11902 | 0.15106 | -0.09613
1360.1719 | 122.4289 | -44.8608 | 77.3804 | -0.084763 | 0.0034778 | 0.0741 | 0.068919 | 0.0018949 | -0.00026063 | -0.001239 | -0.025726 | 0.0010275 | 0.064279 | 0.068746 | 0.0013139 | -0.0026211 | 0.0032193 | -0.004578 | 77.3712 | -44.9982 | 77.1835 | -43.6478 | -45.0302 | -44.9753 | 0.054932 | 77.3163 | -44.8883 | 77.3346 | -44.22 | -44.9615 | -44.7418 | -0.13733 | -0.009155 | -0.19684 | -0.18768 | 1.2131 | 1.3504 | -0.11444 | -0.16937 | 0.022888 | -0.032043 | -0.027466 | -0.064087 | -0.045776 | 0.018311 | 0.64087 | 0.66833 | 0.11902 | -0.10071 | 0.14648 | -0.073242
1361.1719 | 122.4152 | -44.6457 | 77.5497 | -0.084763 | 0.002196 | 0.073245 | 0.068705 | -0.0013095 | 0.0027302 | 0.0013244 | -0.0013123 | -4.0508e-05 | 0.06727 | 0.059774 | -0.0010361 | -0.0045437 | 0.0053555 | -0.013733 | 77.5635 | -44.8151 | 77.3666 | -43.483 | -44.8288 | -44.7922 | 0.027466 | 77.5314 | -44.7052 | 77.5269 | -44.0231 | -44.7647 | -44.5496 | -0.16937 | 0.013733 | -0.18311 | -0.19684 | 1.1627 | 1.3321 | -0.14648 | -0.1831 | 0.022888 | -0.013733 | -0.059509 | -0.018311 | -0.022889 | -0.004578 | 0.62256 | 0.68207 | 0.09613 | -0.11902 | 0.15564 | -0.059509
1362.1719 | 122.3969 | -44.4763 | 77.7603 | -0.19157 | 0.002196 | 0.072177 | 0.073191 | 0.0025357 | 0.00038017 | 0.00089721 | -0.025726 | 0.0031637 | 0.065988 | 0.066396 | 0.0021683 | -0.0043301 | 0.0017239 | -0.045776 | 77.7466 | -44.6228 | 77.5223 | -43.309 | -44.7052 | -44.6365 | -0.10071 | 77.71 | -44.5129 | 77.7191 | -43.8721 | -44.6136 | -44.3802 | -0.14648 | -0.013733 | -0.23804 | -0.2243 | 1.1673 | 1.3138 | -0.16022 | -0.22888 | -0.013733 | -0.082397 | -0.036621 | -0.050354 | -0.041199 | 0.009155 | 0.60425 | 0.64087 | 0.09613 | -0.13733 | 0.13275 | -0.10071
1363.1719 | 122.4518 | -44.2657 | 77.9343 | -0.077133 | 0.0017688 | 0.07175 | 0.071696 | -0.0040867 | -0.0013288 | -0.0014526 | -0.016571 | 0.0023093 | 0.065347 | 0.06426 | 3.2078e-05 | -0.0039029 | 0.0023647 | 0.004578 | 77.9755 | -44.3985 | 77.7832 | -43.0389 | -44.4031 | -44.3756 | 0.045776 | 77.948 | -44.2932 | 77.9388 | -43.6203 | -44.371 | -44.1376 | -0.13275 | 0.041199 | -0.15106 | -0.19226 | 1.2268 | 1.3596 | -0.10986 | -0.13733 | 0.022889 | -0.004577 | -0.027466 | 0.013733 | 0.004578 | -0.009155 | 0.64545 | 0.67291 | 0.12817 | -0.10529 | 0.15564 | -0.07782
1364.1719 | 122.4152 | -44.0964 | 78.1586 | -0.13054 | 0.00070078 | 0.070682 | 0.067424 | 0.0014675 | 0.0010212 | -0.0035888 | 0.010895 | 0.00060029 | 0.066415 | 0.064687 | 0.00024568 | -0.0034757 | -0.00019868 | 0 | 78.1677 | -44.2245 | 77.9343 | -42.897 | -44.2657 | -44.2245 | -0.018311 | 78.1082 | -44.1605 | 78.0533 | -43.4647 | -44.2245 | -44.0048 | -0.12817 | 0.009156 | -0.2243 | -0.23346 | 1.1993 | 1.3275 | -0.12817 | -0.16937 | 0 | -0.041199 | -0.064086 | -0.050354 | -0.10528 | -0.054931 | 0.63171 | 0.6958 | 0.091553 | -0.12817 | 0.15564 | -0.064087
1365.1719 | 122.4014 | -43.8492 | 78.3463 | -0.038986 | 0.0019824 | 0.073031 | 0.070842 | -0.0013095 | 0.00038017 | 0.00068361 | -0.047089 | 0.0012411 | 0.066202 | 0.067464 | 0.00067288 | -0.0013393 | -0.0036167 | -0.032043 | 78.3691 | -44.0002 | 78.1815 | -42.6361 | -44.0277 | -44.0002 | -0.009155 | 78.3371 | -43.9316 | 78.3142 | -43.2449 | -44.014 | -43.7759 | -0.15106 | 0.022889 | -0.16479 | -0.18768 | 1.2131 | 1.3641 | -0.15106 | -0.17853 | 0 | -0.027466 | -0.082398 | -0.009155 | -0.032043 | -0.022888 | 0.60425 | 0.68665 | 0.073242 | -0.16479 | 0.15564 | -0.082397
1366.1719 | 122.3877 | -43.6386 | 78.5477 | -0.031357 | -0.0014356 | 0.071109 | 0.068705 | 0.0010403 | 0.0014484 | -0.00081179 | 0.010895 | 0.0038047 | 0.067484 | 0.062124 | 0.00067288 | -0.0041165 | 0.0036465 | -0.018311 | 78.5706 | -43.8034 | 78.3646 | -42.4713 | -43.8309 | -43.7988 | -0.009155 | 78.5385 | -43.7164 | 78.5248 | -43.0206 | -43.7668 | -43.5562 | -0.1648 | 0.022889 | -0.1831 | -0.20599 | 1.1673 | 1.3321 | -0.16022 | -0.19226 | 0.004578 | -0.027466 | -0.07782 | -0.009155 | -0.022888 | -0.013733 | 0.61798 | 0.6958 | 0.082398 | -0.12817 | 0.16022 | -0.050354
1367.1719 | 122.3557 | -43.5104 | 78.7033 | -0.10765 | 0.00091438 | 0.07175 | 0.073832 | 0.0042447 | 0.0014484 | 0.0019654 | -0.062347 | 0.00017309 | 0.068124 | 0.067037 | -0.00018152 | -0.0036893 | -0.0044711 | -0.07782 | 78.7216 | -43.6798 | 78.4882 | -42.3523 | -43.7485 | -43.689 | -0.12817 | 78.7033 | -43.5654 | 78.6209 | -42.9428 | -43.7073 | -43.4509 | -0.16937 | 0.018311 | -0.21515 | -0.23346 | 1.1581 | 1.3275 | -0.17853 | -0.23804 | -0.009155 | -0.068664 | -0.054932 | 0 | -0.082397 | -0.082397 | 0.56763 | 0.62256 | 0.059509 | -0.19684 | 0.11444 | -0.14191
1368.1719 | 122.3465 | -43.2953 | 78.9368 | 0.12123 | -0.0016492 | 0.073886 | 0.070414 | 0.0012539 | -0.00026063 | 0.0013244 | 0.0017395 | 0.0014547 | 0.067056 | 0.063619 | -0.0010361 | -0.0045437 | 0.0015103 | -0.045776 | 79.01 | -43.4006 | 78.7125 | -42.0868 | -43.4647 | -43.4189 | -0.018311 | 78.9185 | -43.3273 | 78.8589 | -42.6407 | -43.428 | -43.2083 | -0.10529 | 0.073242 | -0.22431 | -0.29755 | 1.2085 | 1.3138 | -0.1236 | -0.16937 | -0.01831 | -0.064087 | -0.032044 | -0.018311 | -0.07782 | -0.059509 | 0.6546 | 0.68665 | 0.086975 | -0.13275 | 0.11902 | -0.10071
1369.1719 | 122.3969 | -43.0893 | 79.1519 | -0.069504 | 0.0024096 | 0.073886 | 0.071482 | -0.00088229 | 0.00059397 | -0.00059819 | -0.025726 | 0.0016685 | 0.067484 | 0.067251 | 0.00088648 | -0.0047573 | 0.00044212 | -0.036621 | 79.1565 | -43.2175 | 78.9185 | -41.629 | -43.19 | -43.364 | 0.036621 | 79.1382 | -43.1305 | 79.0604 | -42.4347 | -43.2312 | -43.0069 | -0.12817 | 0.004577 | -0.23346 | -0.23804 | 1.4603 | 1.5884 | -0.27466 | -0.10071 | -0.14649 | 0.027466 | -0.041199 | -0.013733 | -0.091553 | -0.07782 | 0.6546 | 0.6958 | 0.082397 | -0.14191 | 0.1236 | -0.10071
1370.1719 | 122.4335 | -42.8833 | 79.3488 | -0.10002 | 0.0013416 | 0.073031 | 0.071269 | -0.0025913 | -0.0019696 | -0.00038459 | -0.086761 | 0.0014547 | 0.065775 | 0.067037 | -0.0016769 | -0.0034757 | -0.00041228 | -0.004578 | 79.3167 | -43.0573 | 79.1473 | -41.6336 | -43.0481 | -42.9794 | 0.050354 | 79.3304 | -42.9016 | 79.3121 | -42.2195 | -43.0161 | -42.7826 | -0.17395 | -0.032044 | -0.20142 | -0.16937 | 1.2497 | 1.4236 | -0.09613 | -0.16479 | 0.07782 | 0.009155 | -0.01831 | -0.018311 | -0.036621 | -0.01831 | 0.66376 | 0.68207 | 0.10071 | -0.13275 | 0.11902 | -0.11444
1371.1719 | 122.3969 | -42.6727 | 79.5593 | -0.16106 | -0.0010082 | 0.07175 | 0.072337 | -0.0017367 | -0.0019696 | 0.0013244 | 0.023102 | 0.0029501 | 0.066843 | 0.06426 | 0.00045928 | -0.0053983 | -0.00083948 | 0.018311 | 79.5731 | -42.8329 | 79.3259 | -41.4597 | -42.8513 | -42.8284 | -0.022888 | 79.5273 | -42.7322 | 79.4678 | -42.0227 | -42.8101 | -42.6086 | -0.16022 | 0.013733 | -0.23346 | -0.24719 | 1.2131 | 1.3733 | -0.15564 | -0.17853 | 0.004578 | -0.01831 | -0.05951 | -0.032043 | -0.091553 | -0.05951 | 0.65002 | 0.70953 | 0.064086 | -0.13733 | 0.1236 | -0.07782
1372.1719 | 122.3694 | -42.4896 | 79.7653 | -0.10765 | 0.0030506 | 0.075168 | 0.069133 | 0.0038175 | -0.00090143 | 0.0019654 | -0.062347 | 0.0010275 | 0.063638 | 0.069814 | 0.00067288 | -0.0021939 | 0.0042875 | -0.091553 | 79.715 | -42.6453 | 79.4632 | -41.3498 | -42.7505 | -42.6682 | -0.12817 | 79.715 | -42.5858 | 79.6371 | -41.8121 | -42.9199 | -42.5491 | -0.15564 | -0.050354 | -0.30212 | -0.25177 | 1.1398 | 1.2955 | -0.17853 | -0.26093 | -0.022888 | -0.10529 | -0.09613 | -0.050354 | -0.12817 | -0.07782 | 0.67749 | 0.77362 | -0.059509 | -0.4303 | 0.036621 | -0.33417
1373.1719 | 122.3969 | -42.3203 | 79.9072 | -0.10002 | 0.0015552 | 0.073459 | 0.069346 | -0.0038729 | -0.00026063 | 0.003888 | 0.047516 | -0.00089511 | 0.067911 | 0.059774 | 0.0013139 | -0.0060391 | -0.00062588 | 0.022888 | 79.9072 | -42.4255 | 79.7241 | -41.0248 | -42.4484 | -42.3798 | 0.032043 | 79.8615 | -42.3477 | 79.8386 | -41.6336 | -42.4164 | -42.247 | -0.10529 | 0 | -0.18311 | -0.18311 | 1.2955 | 1.4008 | -0.05951 | -0.12817 | 0.045776 | -0.022888 | -0.027466 | -0.045777 | -0.068665 | -0.022888 | 0.68664 | 0.71411 | 0.073242 | -0.096131 | 0.10071 | -0.068665
1374.1719 | 122.4106 | -42.0639 | 80.2231 | -0.038986 | 0.00027338 | 0.071963 | 0.068919 | -2.7692e-05 | 0.00016657 | 0.00089721 | 0.041412 | 0.0018821 | 0.069406 | 0.06426 | 0.00067288 | -0.0058255 | 0.0021511 | -0.032043 | 80.1315 | -42.2516 | 79.9072 | -40.86 | -42.2699 | -42.2241 | 0 | 80.1086 | -42.1738 | 80.0354 | -41.4459 | -42.2424 | -42.1326 | -0.18768 | -0.091552 | -0.31586 | -0.2243 | 1.2039 | 1.3916 | -0.16022 | -0.20599 | 0.027466 | -0.01831 | -0.10986 | -0.11444 | -0.18768 | -0.073243 | 0.61798 | 0.72784 | -0.068664 | -0.17853 | 0.041199 | -0.068665
1375.1719 | 122.3785 | -41.9128 | 80.3329 | -0.077133 | 0.0019824 | 0.073031 | 0.071482 | -0.0023777 | -0.00090143 | 0.001538 | -0.028778 | 0.0014547 | 0.063638 | 0.067678 | 0.00045928 | -0.0041165 | -0.0016941 | -0.022888 | 80.3284 | -42.0547 | 80.0995 | -40.6631 | -42.0593 | -42.0456 | 0.022888 | 80.2963 | -41.9632 | 80.2917 | -41.2308 | -42.0044 | -41.8076 | -0.14191 | -0.004578 | -0.23346 | -0.22888 | 1.2497 | 1.3916 | -0.13275 | -0.14648 | 0.009156 | -0.004577 | -0.050354 | -0.036621 | -0.041199 | -0.004578 | 0.68207 | 0.73242 | 0.10529 | -0.091553 | 0.15564 | -0.041199
1376.1719 | 122.3831 | -41.716 | 80.5344 | -0.12291 | 0.00091438 | 0.072818 | 0.06721 | -0.00066869 | -4.7032e-05 | -0.0014526 | 0.044464 | 0.0016685 | 0.067484 | 0.061696 | 3.2078e-05 | -0.0053983 | -0.0021213 | 0.022888 | 80.5115 | -41.8259 | 80.3146 | -40.4251 | -41.8716 | -41.8167 | 0.027466 | 80.4794 | -41.7984 | 80.4337 | -41.0568 | -41.8625 | -41.684 | -0.10986 | -0.022888 | -0.21973 | -0.19684 | 1.2909 | 1.4008 | -0.10071 | -0.15564 | 0.009156 | -0.045776 | -0.082398 | -0.054932 | -0.10071 | -0.045776 | 0.65918 | 0.74158 | 0.032043 | -0.14648 | 0.11444 | -0.064087
1377.1719 | 122.4289 | -41.5283 | 80.7037 | -0.092392 | 0.0032642 | 0.072604 | 0.071696 | 0.0038175 | 0.00038017 | -0.0016662 | -0.037933 | 0.0018821 | 0.067911 | 0.067891 | 0.00045928 | -0.0039029 | 0.0015103 | 0.018311 | 80.6625 | -41.7023 | 80.5069 | -40.2603 | -41.684 | -41.6473 | 0.064087 | 80.6442 | -41.597 | 80.6717 | -40.8463 | -41.6519 | -41.4276 | -0.17395 | -0.041198 | -0.19684 | -0.15564 | 1.268 | 1.442 | -0.11902 | -0.15564 | 0.054932 | 0.018311 | -0.068665 | -0.059509 | -0.032043 | 0.027466 | 0.68207 | 0.75073 | 0.10071 | -0.1236 | 0.16937 | -0.054932
1378.1719 | 122.4106 | -41.3544 | 80.8868 | -0.10765 | 0.0017688 | 0.073672 | 0.068919 | -0.0013095 | 0.0012348 | 4.2808e-05 | 0.041412 | -0.00089511 | 0.067484 | 0.059133 | -0.00018152 | -0.0047573 | -0.00041228 | 0.036621 | 80.8731 | -41.4597 | 80.7312 | -39.9673 | -41.423 | -41.391 | 0.10071 | 80.8182 | -41.3864 | 80.8319 | -40.6357 | -41.4368 | -41.2582 | -0.10529 | -0.013733 | -0.15564 | -0.14191 | 1.387 | 1.4923 | -0.036621 | -0.068665 | 0.068665 | 0.036621 | -0.032044 | -0.068665 | -0.054932 | 0.013733 | 0.71869 | 0.75073 | 0.09613 | -0.082398 | 0.12817 | -0.050354
1379.1719 | 122.4564 | -41.1346 | 81.0837 | -0.092392 | 0.0017688 | 0.073672 | 0.069346 | -0.0025913 | 0.00059397 | 0.0017518 | 0.059723 | 0.0010275 | 0.066415 | 0.067037 | -0.0010361 | -0.0047573 | 0.0025785 | 0.041199 | 81.0379 | -41.2857 | 80.9509 | -39.7614 | -41.2216 | -41.1758 | 0.1236 | 81.0104 | -41.1987 | 81.0608 | -40.4114 | -41.2033 | -41.011 | -0.15106 | -0.045776 | -0.13275 | -0.086975 | 1.3733 | 1.5244 | -0.041198 | -0.086975 | 0.10986 | 0.064087 | -0.064086 | -0.073242 | -0.022888 | 0.050354 | 0.72327 | 0.78735 | 0.1236 | -0.068664 | 0.18768 | -0.004578
1380.1719 | 122.3877 | -40.9698 | 81.2714 | -0.11528 | 0.0013416 | 0.073459 | 0.070201 | 0.00039951 | 0.001662 | -0.0023072 | -0.11423 | 0.0025229 | 0.067911 | 0.07195 | -0.00039512 | -0.0028347 | -0.0023349 | 0.004578 | 81.2531 | -41.1118 | 81.0974 | -39.6286 | -41.1118 | -41.0339 | 0.041199 | 81.2851 | -40.979 | 81.2531 | -40.2603 | -41.0843 | -40.8737 | -0.14191 | -0.01831 | -0.17395 | -0.15564 | 1.3412 | 1.4832 | -0.064087 | -0.14191 | 0.077819 | 0 | -0.009155 | 0.013733 | -0.01831 | -0.032043 | 0.70953 | 0.71869 | 0.096131 | -0.11444 | 0.10529 | -0.10529
1381.1719 | 122.3877 | -40.7684 | 81.4682 | -0.092392 | 0.0019824 | 0.071963 | 0.070628 | 0.0025357 | 0.001662 | 0.00025641 | 0.041412 | 0.0018821 | 0.065775 | 0.064901 | -0.0012497 | -0.0058255 | 0.00065592 | 0.041199 | 81.4362 | -40.8875 | 81.3126 | -39.386 | -40.86 | -40.8096 | 0.082397 | 81.4316 | -40.8051 | 81.4087 | -40.0589 | -40.86 | -40.6677 | -0.11902 | -0.032044 | -0.15564 | -0.1236 | 1.3824 | 1.5015 | -0.041198 | -0.091552 | 0.07782 | 0.027466 | -0.036621 | -0.036621 | -0.05951 | -0.022889 | 0.70953 | 0.74616 | 0.10071 | -0.091552 | 0.13733 | -0.054931
1382.1719 | 122.3969 | -40.5533 | 81.665 | -0.084763 | 0.0030506 | 0.073031 | 0.072123 | 0.0031765 | -0.0015424 | 0.0019654 | -0.071503 | 0.0031637 | 0.066843 | 0.067037 | -0.0018905 | -0.0045437 | -0.0012667 | 0.027466 | 81.6147 | -40.7089 | 81.4682 | -39.2258 | -40.7135 | -40.686 | 0.045776 | 81.6193 | -40.6219 | 81.6284 | -39.8712 | -40.6815 | -40.48 | -0.15564 | -0.050354 | -0.19684 | -0.14648 | 1.3275 | 1.4832 | -0.13275 | -0.16022 | 0.022888 | -0.004578 | -0.068664 | -0.045776 | -0.036621 | 0.009155 | 0.68207 | 0.75073 | 0.073242 | -0.12817 | 0.14191 | -0.05951
1383.1719 | 122.3694 | -40.4068 | 81.8161 | -0.054245 | 0.0019824 | 0.0741 | 0.069133 | 0.0065945 | 0.0010212 | 0.00025641 | 0.02005 | 0.0020957 | 0.065988 | 0.06426 | -0.00018152 | -0.0051845 | -0.0012667 | 0.018311 | 81.8069 | -40.5075 | 81.633 | -39.0564 | -40.5167 | -40.4892 | 0.050354 | 81.8115 | -40.448 | 81.7291 | -39.6973 | -40.535 | -40.361 | -0.10071 | -0.009155 | -0.1831 | -0.17395 | 1.3504 | 1.4511 | -0.082398 | -0.10986 | 0.01831 | -0.009156 | -0.041199 | -0.004578 | -0.086975 | -0.082397 | 0.70953 | 0.75073 | 0.045776 | -0.12817 | 0.086975 | -0.086975
1384.1719 | 122.4747 | -40.2008 | 82.0084 | 0.029678 | 0.0030506 | 0.075381 | 0.066996 | -0.00088229 | 0.0037982 | -0.00017099 | 0.038361 | -0.00025411 | 0.065988 | 0.052938 | -0.0012497 | -0.0045437 | 0.0032193 | -0.027466 | 81.9626 | -40.3107 | 81.8893 | -38.7772 | -40.2603 | -40.2283 | 0.12817 | 81.9626 | -40.242 | 82.0221 | -39.4226 | -40.2191 | -40.0681 | -0.10986 | -0.045777 | -0.11902 | -0.073242 | 1.4236 | 1.5335 | -0.027465 | -0.059509 | 0.082398 | 0.050354 | -0.041198 | -0.045777 | 0.013733 | 0.05951 | 0.7782 | 0.8194 | 0.13275 | -0.01831 | 0.17395 | 0.022888
1385.1875 | 122.406 | -40.0177 | 82.1732 | -0.038986 | 5.9776e-05 | 0.073672 | 0.071696 | 0.0036039 | 0.0023028 | 0.0013244 | -0.03183 | 0.00060029 | 0.066415 | 0.067037 | -0.0010361 | -0.0045437 | -0.0012667 | -0.10071 | 82.1823 | -40.1825 | 81.9763 | -38.7085 | -40.2237 | -40.155 | -0.068665 | 82.1503 | -40.0909 | 82.1732 | -39.3173 | -40.1505 | -39.9353 | -0.1648 | 0.009155 | -0.19684 | -0.20599 | 1.3092 | 1.474 | -0.13733 | -0.20599 | 0.027466 | -0.041199 | -0.073242 | -0.022888 | 0 | 0.022888 | 0.70038 | 0.77362 | 0.082397 | -0.13275 | 0.15564 | -0.05951
1386.1719 | 122.4792 | -39.7797 | 82.3929 | -0.084763 | 0.00048698 | 0.07474 | 0.069987 | -0.0032321 | 0.0014484 | 0.0013244 | 0.0017395 | 0.0027365 | 0.065775 | 0.061696 | -0.00039512 | -0.0045437 | -0.0010531 | 0.041199 | 82.37 | -39.917 | 82.1915 | -38.6124 | -39.8575 | -39.7156 | 0.1236 | 82.3425 | -39.8438 | 82.4203 | -38.9969 | -39.8392 | -39.6561 | -0.13733 | -0.022888 | -0.20142 | -0.17853 | 1.1673 | 1.3046 | 0.064087 | -0.07782 | 0.20142 | 0.059509 | -0.064087 | -0.050354 | 0.027466 | 0.07782 | 0.78278 | 0.84686 | 0.1236 | -0.059509 | 0.18768 | 0.004578
1387.1719 | 122.4289 | -39.6378 | 82.6035 | -0.0084687 | 0.0019824 | 0.07175 | 0.070842 | 0.00039951 | -0.001756 | 0.0036744 | -0.044037 | 0.00038669 | 0.066202 | 0.066823 | -0.0023177 | -0.0041165 | 0.0015103 | -0.022888 | 82.6172 | -39.7797 | 82.3929 | -38.3148 | -39.7659 | -39.7339 | 0.027466 | 82.5623 | -39.6698 | 82.5531 | -38.9008 | -39.743 | -39.5554 | -0.14191 | 0.013732 | -0.21057 | -0.2243 | 1.3229 | 1.4648 | -0.096131 | -0.12817 | 0.045776 | 0.013733 | -0.032044 | -0.041199 | -0.050354 | -0.009155 | 0.737 | 0.76904 | 0.082397 | -0.10529 | 0.11444 | -0.073242
1388.1719 | 122.4564 | -39.4638 | 82.7225 | -0.038986 | 0.0030506 | 0.074527 | 0.071482 | -0.00045509 | 0.0010212 | 0.0017518 | 0.029205 | 0.0023093 | 0.066202 | 0.062124 | -0.0014633 | -0.0043301 | -0.0012667 | 0.032043 | 82.7087 | -39.5737 | 82.6218 | -38.031 | -39.5233 | -39.4775 | 0.11444 | 82.7179 | -39.5096 | 82.7408 | -38.681 | -39.5325 | -39.3677 | -0.10986 | -0.013733 | -0.10071 | -0.086975 | 1.4328 | 1.5427 | -0.013733 | -0.059509 | 0.09613 | 0.050354 | -0.045777 | -0.004577 | 0.018311 | 0.022888 | 0.78278 | 0.82855 | 0.09613 | -0.068665 | 0.14191 | -0.022888
1389.1719 | 122.438 | -39.3082 | 82.9147 | -0.10002 | 0.0030506 | 0.076022 | 0.068919 | -0.0025913 | -4.7032e-05 | -0.0016662 | 0.029205 | 0.0018821 | 0.065561 | 0.063405 | -0.0018905 | -0.0041165 | -0.0016941 | 0.059509 | 82.8735 | -39.4135 | 82.7957 | -37.8571 | -39.3448 | -39.3265 | 0.11444 | 82.8644 | -39.3448 | 82.8964 | -38.5025 | -39.3448 | -39.1754 | -0.10529 | -0.041199 | -0.11902 | -0.07782 | 1.4511 | 1.5564 | -0.01831 | -0.036621 | 0.086975 | 0.068664 | -0.036621 | -0.050354 | -0.018311 | 0.032043 | 0.80566 | 0.84229 | 0.13275 | -0.036621 | 0.16937 | 0
1390.1875 | 122.5113 | -39.1113 | 83.0658 | -0.15343 | 0.0036914 | 0.073886 | 0.070414 | -0.0062229 | 0.0014484 | 0.00047001 | 0.041412 | 0.00081389 | 0.066202 | 0.056356 | 0.00045928 | -0.0049709 | 0.0036465 | -0.027466 | 83.0338 | -39.2807 | 82.9193 | -37.7106 | -39.2212 | -39.1663 | 0.045776 | 83.0109 | -39.1891 | 83.0887 | -38.3286 | -39.1891 | -38.9969 | -0.16937 | -0.032044 | -0.14648 | -0.11444 | 1.4008 | 1.5701 | -0.054932 | -0.10986 | 0.11444 | 0.05951 | -0.07782 | -0.054932 | 0.022888 | 0.07782 | 0.78278 | 0.8606 | 0.11444 | -0.07782 | 0.19226 | 0
1391.1719 | 122.4609 | -39.0106 | 83.1848 | -0.069504 | 0.0030506 | 0.077731 | 0.070842 | -0.0038729 | 0.0023028 | 0.0017518 | 0.0078429 | -0.00068151 | 0.067056 | 0.06896 | 0.00067288 | -0.0043301 | 1.4916e-05 | 0.073242 | 83.1894 | -39.1251 | 83.0887 | -37.5137 | -39.0381 | -39.0106 | 0.1236 | 83.1711 | -39.0244 | 83.2352 | -38.1821 | -39.0335 | -38.8779 | -0.11444 | 0.004578 | -0.09613 | -0.10071 | 1.4969 | 1.6113 | 0 | -0.027466 | 0.11444 | 0.086975 | -0.013733 | -0.013732 | 0.050354 | 0.064086 | 0.82855 | 0.84228 | 0.13275 | -0.022888 | 0.14648 | -0.009155
1392.1719 | 122.4152 | -38.8321 | 83.3862 | -0.069504 | 5.9776e-05 | 0.07474 | 0.074473 | -0.0032321 | 0.00016657 | -0.00081179 | -0.10202 | 0.0020957 | 0.066202 | 0.065969 | -0.0016769 | -0.0036893 | -0.0019077 | -0.009155 | 83.3817 | -38.9832 | 83.2443 | -37.3489 | -38.9236 | -38.9465 | 0.027466 | 83.3954 | -38.8641 | 83.3679 | -38.0951 | -38.9557 | -38.736 | -0.15106 | -0.004577 | -0.14191 | -0.13733 | 1.4832 | 1.6342 | -0.11444 | -0.091553 | 0.036621 | 0.059509 | -0.032044 | 0.009156 | -0.01831 | -0.027466 | 0.737 | 0.76904 | 0.09613 | -0.1236 | 0.12817 | -0.091552
1393.1719 | 122.4518 | -38.6536 | 83.519 | -0.046616 | 0.00091438 | 0.076022 | 0.071055 | -0.0010959 | 0.001662 | 0.0011108 | 0.02005 | 0.00038669 | 0.06492 | 0.066182 | -0.00060892 | -0.0051845 | -0.0016941 | 0.032043 | 83.4961 | -38.7772 | 83.432 | -37.1704 | -38.7222 | -38.6948 | 0.1236 | 83.4961 | -38.7131 | 83.551 | -37.8387 | -38.6948 | -38.5345 | -0.1236 | -0.022888 | -0.086975 | -0.064087 | 1.4832 | 1.6068 | -0.041199 | -0.068665 | 0.082398 | 0.054932 | -0.05951 | -0.022888 | 0.032043 | 0.054931 | 0.81482 | 0.87433 | 0.11902 | -0.041199 | 0.17853 | 0.018311
1394.1719 | 122.3328 | -38.5254 | 83.7524 | -0.092392 | 0.0043322 | 0.075381 | 0.070414 | -0.0013095 | 0.00016657 | 0.00047001 | -0.013519 | -4.0508e-05 | 0.066843 | 0.061483 | 0.0013139 | -0.0049709 | 0.00022852 | 0.032043 | 83.6975 | -38.6536 | 83.551 | -37.0605 | -38.5666 | -38.5574 | 0.050354 | 83.6884 | -38.5437 | 83.6517 | -37.7518 | -38.6398 | -38.4659 | -0.12817 | -0.054931 | -0.20142 | -0.14648 | 1.4648 | 1.593 | -0.032043 | -0.041198 | 0.09613 | 0.086975 | -0.01831 | -0.064087 | -0.10071 | -0.036621 | 0.77362 | 0.79193 | 0.05951 | -0.11444 | 0.07782 | -0.096131
1395.1875 | 122.3648 | -38.3377 | 83.8898 | -0.00083923 | 0.0026234 | 0.075595 | 0.073619 | -0.0013095 | 0.00016657 | 0.0017518 | -0.053192 | 0.0023093 | 0.06727 | 0.061269 | 0.00067288 | -0.0045437 | -0.0012667 | -0.009155 | 83.9172 | -38.4979 | 83.5785 | -37.0239 | -38.3835 | -38.4384 | -0.013733 | 83.8943 | -38.4155 | 83.8257 | -37.6007 | -38.475 | -38.2965 | -0.16022 | 0.027465 | -0.31128 | -0.33875 | 1.3138 | 1.474 | -0.10071 | -0.045776 | 0.059509 | 0.11444 | -0.077819 | 0.004577 | -0.064087 | -0.068664 | 0.737 | 0.81482 | 0.041199 | -0.13733 | 0.11902 | -0.05951
1396.1719 | 122.4518 | -38.1958 | 84.0408 | -0.023727 | 0.0034778 | 0.076236 | 0.072978 | -0.00088229 | 0.00016657 | 0.00047001 | -0.044037 | 0.0023093 | 0.067484 | 0.071309 | -0.0012497 | -0.0028347 | -0.00041228 | 0.013733 | 83.9905 | -38.3652 | 83.8257 | -36.7538 | -38.3102 | -38.2919 | 0.068665 | 83.9813 | -38.2553 | 83.9905 | -37.4405 | -38.3194 | -38.1088 | -0.16937 | -0.050354 | -0.21515 | -0.16479 | 1.442 | 1.6113 | -0.09613 | -0.11444 | 0.073242 | 0.054931 | -0.059509 | -0.05951 | -0.050354 | 0.009156 | 0.75531 | 0.81482 | 0.086975 | -0.1236 | 0.14648 | -0.064087
1397.1875 | 122.4564 | -38.0356 | 84.1644 | -0.084763 | 0.0030506 | 0.075381 | 0.069987 | 0.0033901 | 0.0023028 | 0.00089721 | -0.044037 | 0.0035911 | 0.06257 | 0.064687 | -0.00060892 | -0.0032619 | 0.0027921 | -0.091553 | 84.1873 | -38.1958 | 83.9447 | -36.6211 | -38.205 | -38.15 | -0.082397 | 84.1599 | -38.1042 | 84.1599 | -37.262 | -38.1454 | -37.9211 | -0.16022 | 0.022888 | -0.21973 | -0.24261 | 1.4145 | 1.5747 | -0.11444 | -0.16937 | 0.045777 | -0.009155 | -0.068665 | -0.004578 | -0.004578 | 0 | 0.77362 | 0.84228 | 0.11444 | -0.10986 | 0.1831 | -0.041199
1398.1719 | 122.3694 | -37.8479 | 84.3842 | -0.038986 | 0.001128 | 0.075381 | 0.069773 | 0.00082671 | 0.00038017 | -0.0010254 | -0.0043641 | 0.0023093 | 0.065561 | 0.062551 | -0.00082252 | -0.0045437 | 0.0010831 | -0.018311 | 84.3338 | -38.031 | 84.1141 | -36.4288 | -38.0127 | -37.9807 | 0.004578 | 84.3246 | -37.9395 | 84.2606 | -37.1475 | -37.7975 | -37.6694 | -0.18311 | -0.050354 | -0.27008 | -0.21973 | 1.4191 | 1.6022 | -0.13275 | -0.16479 | 0.050354 | 0.018311 | -0.091553 | -0.059509 | -0.1236 | -0.064087 | 0.70038 | 0.79193 | 0.17853 | 0.050354 | 0.27008 | 0.14191
1399.1719 | 122.3511 | -37.7884 | 84.4849 | -0.038986 | 0.003905 | 0.07474 | 0.070628 | -0.0013095 | 0.0010212 | -0.0020936 | -0.016571 | 0.0012411 | 0.06492 | 0.066396 | 0.0015275 | -0.0036893 | -0.00041228 | -0.027466 | 84.4757 | -37.8616 | 84.2468 | -36.2915 | -37.8708 | -37.8342 | 0.004578 | 84.4757 | -37.8113 | 84.407 | -36.9965 | -37.8799 | -37.706 | -0.073242 | -0.009155 | -0.23804 | -0.22888 | 1.4969 | 1.5701 | -0.045776 | -0.082398 | 0.027466 | -0.009156 | -0.022888 | -0.009155 | -0.07782 | -0.068665 | 0.79193 | 0.81482 | 0.082397 | -0.091553 | 0.10529 | -0.068665
1400.2031 | 122.374 | -37.5961 | 84.6359 | -0.092392 | 0.0026234 | 0.076449 | 0.074473 | -0.00066869 | 0.0029438 | 0.002179 | -0.0074159 | 0.0048727 | 0.066629 | 0.064473 | -0.00082252 | -0.00048494 | -0.0029757 | -0.086975 | 84.5901 | -37.7792 | 84.4528 | -36.3831 | -37.8616 | -37.4771 | -0.20142 | 84.5947 | -37.6785 | 84.5581 | -36.8866 | -37.7747 | -37.5549 | -0.18311 | -0.045776 | -0.1831 | -0.13733 | 1.2131 | 1.3962 | 0.11902 | -0.2655 | 0.30212 | -0.082397 | -0.082398 | -0.041198 | -0.07782 | -0.036622 | 0.70953 | 0.79193 | 0.041198 | -0.17853 | 0.1236 | -0.09613
1401.1719 | 122.3694 | -37.4634 | 84.7504 | -0.077133 | 0.0015552 | 0.074313 | 0.06956 | 0.0038175 | 0.003371 | 0.00025641 | -0.11118 | 0.0033773 | 0.067911 | 0.06191 | -0.0010361 | -0.0034757 | 0.0038601 | -0.018311 | 84.7641 | -37.5961 | 84.4849 | -35.7651 | -37.3444 | -37.6556 | 0.050354 | 84.7687 | -37.5183 | 84.6863 | -36.6898 | -37.5961 | -37.4039 | -0.13275 | 0.013733 | -0.2655 | -0.27924 | 1.6983 | 1.8311 | -0.19226 | 0.11902 | -0.05951 | 0.25177 | -0.054932 | 0.018311 | -0.064087 | -0.082398 | 0.77362 | 0.82855 | 0.059509 | -0.13275 | 0.11444 | -0.077819
1402.1719 | 122.3923 | -37.3169 | 84.9106 | -0.077133 | 0.0047596 | 0.077731 | 0.073405 | 0.0044583 | 0.001662 | 0.0077332 | -0.059296 | 0.0031637 | 0.066843 | 0.063619 | -0.00082252 | -0.0032619 | 0.0015103 | 0 | 84.8785 | -37.468 | 84.8053 | -35.7788 | -37.2665 | -37.3489 | 0.036621 | 84.9106 | -37.3398 | 84.8648 | -36.5295 | -37.4176 | -37.2391 | -0.15106 | -0.032043 | -0.10529 | -0.073242 | 1.5381 | 1.6891 | -0.032043 | 0.050354 | 0.11902 | 0.20142 | -0.022888 | 0 | -0.045776 | -0.045776 | 0.78735 | 0.81024 | 0.07782 | -0.10071 | 0.10071 | -0.07782
1403.1719 | 122.406 | -37.1704 | 85.0342 | -0.10002 | 0.003905 | 0.075381 | 0.06956 | 0.00039951 | -0.00068783 | -0.0010254 | 0.023102 | 0.0029501 | 0.065347 | 0.060842 | 0.0019547 | -0.0030483 | -0.00019868 | 0.036621 | 85.0204 | -37.3032 | 84.8694 | -35.5865 | -37.2162 | -37.1887 | 0.064087 | 85.0021 | -37.2391 | 84.9792 | -36.3831 | -37.3077 | -37.1201 | -0.13275 | -0.013733 | -0.1648 | -0.15106 | 1.5839 | 1.7166 | -0.018311 | -0.045777 | 0.11444 | 0.086975 | -0.068665 | -0.032044 | -0.054932 | -0.022888 | 0.78735 | 0.85602 | 0.050354 | -0.13733 | 0.11902 | -0.068664
1404.1719 | 122.3785 | -36.9919 | 85.2173 | -0.12291 | 0.0026234 | 0.077945 | 0.071482 | 0.0050991 | -0.00026063 | 0.003247 | -0.059296 | 0.0050863 | 0.066202 | 0.066396 | -0.0027451 | -0.0017667 | -0.0014805 | 0.004578 | 85.1624 | -37.2208 | 84.9884 | -35.5042 | -37.1338 | -37.1155 | 0.050354 | 85.1532 | -37.1017 | 85.1532 | -36.2686 | -37.1704 | -36.8912 | -0.22888 | -0.054931 | -0.22888 | -0.17395 | 1.4877 | 1.7166 | -0.1236 | -0.14191 | 0.10529 | 0.086975 | -0.10986 | -0.064087 | -0.064087 | 0 | 0.72327 | 0.83313 | 0.10071 | -0.17853 | 0.21057 | -0.068664
1405.1875 | 122.3419 | -36.8637 | 85.3867 | -0.092392 | 0.0030506 | 0.0741 | 0.072764 | -2.7692e-05 | -4.7032e-05 | -0.00081179 | -0.034882 | 0.0029501 | 0.068979 | 0.070028 | -0.0016769 | -0.0030483 | 0.0036465 | -0.041199 | 85.3958 | -36.969 | 85.1395 | -35.321 | -36.937 | -36.9415 | 0.004578 | 85.3638 | -36.9324 | 85.2997 | -36.0535 | -36.9827 | -36.795 | -0.10529 | 0.009155 | -0.24719 | -0.25635 | 1.5427 | 1.6479 | -0.07782 | -0.073243 | 0.027466 | 0.032043 | -0.068665 | -0.022888 | -0.086975 | -0.064087 | 0.81024 | 0.87891 | 0.068664 | -0.11902 | 0.13733 | -0.050354
1406.1719 | 122.4243 | -36.7676 | 85.4782 | -0.054245 | 0.0024096 | 0.075168 | 0.072978 | 0.0048855 | 0.00059397 | 0.001538 | -0.037933 | 0.0016685 | 0.068765 | 0.06661 | -0.0016769 | -0.0024075 | 0.0012967 | -0.013733 | 85.4416 | -36.9507 | 85.199 | -35.2432 | -36.8866 | -36.85 | -0.018311 | 85.437 | -36.8546 | 85.3729 | -35.994 | -36.9186 | -36.7081 | -0.18311 | -0.036621 | -0.27923 | -0.24261 | 1.5244 | 1.7075 | -0.082398 | -0.11902 | 0.10071 | 0.064087 | -0.086975 | -0.041198 | -0.10529 | -0.064087 | 0.77362 | 0.8606 | 0.059509 | -0.15106 | 0.14648 | -0.064087
1407.1719 | 122.4014 | -36.6348 | 85.6293 | -0.054245 | 0.0032642 | 0.076022 | 0.072978 | -2.7692e-05 | 0.0023028 | 0.0028198 | -0.044037 | 0.0033773 | 0.065561 | 0.059133 | -0.0010361 | -0.0019803 | 0.0017239 | 0.004578 | 85.6155 | -36.7722 | 85.405 | -35.0418 | -36.6669 | -36.6623 | 0.045776 | 85.5927 | -36.6623 | 85.5927 | -35.7697 | -36.7172 | -36.5204 | -0.13733 | -0.013732 | -0.2243 | -0.21057 | 1.593 | 1.7303 | -0.027465 | -0.032043 | 0.10986 | 0.10529 | -0.027465 | -0.036621 | -0.036621 | 0 | 0.86517 | 0.89264 | 0.11444 | -0.082397 | 0.14191 | -0.054932
1408.1719 | 122.406 | -36.5524 | 85.6842 | -0.069504 | 0.0015552 | 0.074954 | 0.069773 | -0.00066869 | 0.00080757 | 0.0011108 | -0.025726 | 0.0016685 | 0.064493 | 0.059987 | 0.00067288 | -0.0043301 | 0.0012967 | -0.009155 | 85.6705 | -36.6943 | 85.4599 | -34.9823 | -36.6165 | -36.5936 | 0 | 85.6567 | -36.6211 | 85.6384 | -35.7147 | -36.6577 | -36.4517 | -0.14191 | -0.013733 | -0.2243 | -0.21057 | 1.5701 | 1.712 | -0.041199 | -0.064087 | 0.10071 | 0.07782 | -0.068665 | -0.027466 | -0.045776 | -0.01831 | 0.83771 | 0.90637 | 0.10071 | -0.10529 | 0.16937 | -0.036621
1409.1719 | 122.3969 | -36.4334 | 85.8215 | -0.016098 | 0.0034778 | 0.07709 | 0.071055 | 0.0010403 | 0.00038017 | -0.0010254 | -0.025726 | 0.0035911 | 0.063425 | 0.062978 | -0.0012497 | -0.0039029 | -0.00041228 | -0.013733 | 85.8032 | -36.5479 | 85.5789 | -34.8633 | -36.5067 | -36.4746 | -0.013733 | 85.7758 | -36.4746 | 85.7391 | -35.6094 | -36.5295 | -36.3464 | -0.11444 | -0.01831 | -0.24262 | -0.22431 | 1.5701 | 1.6846 | -0.041198 | -0.073242 | 0.073243 | 0.041199 | -0.041198 | -0.045776 | -0.082397 | -0.036621 | 0.82397 | 0.86517 | 0.086975 | -0.09613 | 0.12817 | -0.054932
1410.1875 | 122.4518 | -36.264 | 85.9085 | -0.16106 | 0.003905 | 0.074313 | 0.071482 | -0.00045509 | 0.0040118 | -0.001239 | 0.074982 | 0.0035911 | 0.066202 | 0.061696 | 3.2078e-05 | -0.0053983 | -0.00019868 | -0.036621 | 85.8627 | -36.4426 | 85.7483 | -34.6207 | -36.3373 | -36.2686 | 0.018311 | 85.8719 | -36.319 | 85.9268 | -35.3943 | -36.3235 | -36.145 | -0.17853 | -0.045776 | -0.16022 | -0.11444 | 1.6434 | 1.8219 | -0.004578 | -0.073242 | 0.17395 | 0.10529 | -0.054932 | -0.036621 | 0.018311 | 0.054932 | 0.86975 | 0.92468 | 0.11902 | -0.059509 | 0.17395 | -0.004577
1411.1719 | 122.5067 | -36.1633 | 86.0229 | -0.023727 | 0.005614 | 0.074954 | 0.068064 | -0.0010959 | 0.0010212 | 0.00089721 | 0.065826 | 0.0016685 | 0.066629 | 0.065969 | -0.0021041 | -0.0034757 | 0.0030057 | 0.022888 | 86.0046 | -36.2778 | 85.9268 | -34.4238 | -36.1588 | -35.9482 | -0.045776 | 85.9451 | -36.2366 | 86.0825 | -35.2753 | -36.1816 | -36.0077 | -0.11444 | -0.01831 | -0.09613 | -0.07782 | 1.7395 | 1.8539 | 0.21515 | 0.004578 | 0.32959 | 0.11902 | -0.073242 | -0.07782 | 0.059509 | 0.13733 | 0.88806 | 0.9613 | 0.15564 | -0.018311 | 0.22888 | 0.054931
1412.1719 | 122.3694 | -36.0352 | 86.2061 | -0.084763 | 0.002837 | 0.075809 | 0.072764 | 0.00061311 | 0.0020892 | 0.0023926 | 0.0017395 | 0.0040183 | 0.067911 | 0.066823 | 3.2078e-05 | -0.0041165 | 0.0023647 | 0.027466 | 86.2106 | -36.1404 | 86.0001 | -34.3735 | -36.0718 | -36.0443 | 0.068665 | 86.1877 | -36.0901 | 86.1008 | -35.2066 | -36.1588 | -35.9528 | -0.10529 | 0.004577 | -0.20599 | -0.21057 | 1.6617 | 1.767 | -0.009156 | -0.036621 | 0.09613 | 0.068665 | -0.054932 | -0.018311 | -0.10529 | -0.086975 | 0.82855 | 0.88348 | 0.082397 | -0.1236 | 0.13733 | -0.068664
1413.1719 | 122.3602 | -35.9253 | 86.2976 | -0.069504 | 0.0041186 | 0.076022 | 0.071055 | -0.0043003 | 0.00059397 | 0.003888 | -0.016571 | 0.0025229 | 0.065988 | 0.061269 | -0.0018905 | -0.0034757 | 0.0040737 | 0.050354 | 86.2701 | -36.0352 | 86.1328 | -34.1949 | -35.9299 | -35.8978 | 0.082397 | 86.2518 | -35.9711 | 86.2289 | -35.0693 | -36.0397 | -35.8521 | -0.10986 | -0.027465 | -0.16479 | -0.13733 | 1.7303 | 1.8402 | 0.027466 | -0.004578 | 0.13733 | 0.10529 | -0.045776 | -0.045776 | -0.068664 | -0.022888 | 0.85602 | 0.90179 | 0.073242 | -0.11444 | 0.11902 | -0.068665
1414.1719 | 122.4701 | -35.7971 | 86.3709 | -0.092392 | 0.0030506 | 0.07474 | 0.070628 | 0.0027493 | -0.00068783 | -0.00038459 | -0.071503 | 0.0010275 | 0.067056 | 0.061483 | -0.0010361 | -0.0039029 | 0.0015103 | 0.041199 | 86.3663 | -35.939 | 86.2839 | -34.0256 | -35.5545 | -35.6552 | 0.11444 | 86.3113 | -35.8658 | 86.4258 | -34.8862 | -35.8109 | -35.6598 | -0.14191 | -0.004578 | -0.086975 | -0.082397 | 1.7715 | 1.9135 | 0.14191 | 0.24262 | 0.28381 | 0.38452 | -0.068665 | -0.05951 | 0.054931 | 0.11444 | 0.91095 | 0.97961 | 0.13733 | -0.013733 | 0.20599 | 0.054932
1415.1719 | 122.4014 | -35.7422 | 86.5128 | -0.031357 | 0.0036914 | 0.075595 | 0.071482 | 0.0021085 | -0.00026063 | 0.001538 | -0.028778 | 0.0016685 | 0.065988 | 0.072805 | 0.00088648 | -0.0051845 | 0.0030057 | -0.027466 | 86.5036 | -35.9024 | 86.261 | -34.0714 | -35.8109 | -35.7742 | 0.21973 | 86.5494 | -35.7697 | 86.4075 | -34.8953 | -35.8429 | -35.6689 | -0.16022 | -0.009155 | -0.25177 | -0.24262 | 1.6708 | 1.8311 | -0.032044 | -0.068665 | 0.12817 | 0.091553 | -0.027466 | 0.036621 | -0.10528 | -0.14191 | 0.84686 | 0.87433 | 0.073242 | -0.10071 | 0.10071 | -0.073243
1416.1719 | 122.3648 | -35.6461 | 86.5952 | -0.092392 | 0.004546 | 0.076877 | 0.069987 | -0.0010959 | 0.0018756 | 0.0011108 | -0.010467 | 0.0020957 | 0.067911 | 0.064687 | -0.0012497 | -0.0039029 | 0.0023647 | 0.036621 | 86.5677 | -35.7376 | 86.4304 | -33.8608 | -35.6369 | -35.5774 | 0.059509 | 86.5677 | -35.6781 | 86.5219 | -34.7672 | -35.7056 | -35.5408 | -0.091553 | -0.027466 | -0.16479 | -0.13733 | 1.7853 | 1.8768 | 0.068664 | 0.009155 | 0.16022 | 0.10071 | -0.032044 | -0.027466 | -0.073242 | -0.045776 | 0.87891 | 0.91095 | 0.10529 | -0.059509 | 0.13733 | -0.027465
1417.1875 | 122.3694 | -35.5316 | 86.6867 | -0.092392 | 0.002837 | 0.078586 | 0.072764 | 0.0018949 | 0.00038017 | 0.0034606 | -0.098968 | 0.0044455 | 0.065988 | 0.06896 | -0.0021041 | -0.0030483 | -0.0014805 | -0.032043 | 86.6821 | -35.6644 | 86.4944 | -33.783 | -35.5911 | -35.5453 | -0.013733 | 86.6547 | -35.6186 | 86.6638 | -34.6985 | -35.6644 | -35.463 | -0.13275 | -0.004578 | -0.19226 | -0.18768 | 1.7487 | 1.8814 | -0.013733 | -0.059509 | 0.11902 | 0.073243 | -0.086975 | -0.032044 | -0.022889 | 0.009155 | 0.83313 | 0.9201 | 0.068664 | -0.13275 | 0.15564 | -0.045777
1418.1719 | 122.4838 | -35.4126 | 86.7462 | -0.11528 | 0.0024096 | 0.074527 | 0.069773 | -0.00088229 | 0.0014484 | 0.002179 | 0.044464 | 0.00081389 | 0.067697 | 0.056783 | 0.0025955 | -0.0036893 | -0.00041228 | 0.041199 | 86.7508 | -35.5362 | 86.6684 | -33.5358 | -35.3577 | -35.3119 | 0.14191 | 86.7325 | -35.4721 | 86.8378 | -34.4147 | -35.5087 | -35.2844 | -0.1236 | 0.004577 | -0.07782 | -0.082397 | 1.8768 | 2.0004 | 0.10071 | 0.054932 | 0.2243 | 0.17853 | -0.059509 | -0.013733 | 0.091553 | 0.10529 | 0.99793 | 1.0574 | 0.12817 | -0.09613 | 0.18768 | -0.036621
1419.1719 | 122.4747 | -35.2798 | 86.8881 | -0.061874 | 0.0041186 | 0.076022 | 0.07191 | 0.0016813 | 0.0025166 | -0.00038459 | 0.041412 | 0.0025229 | 0.06727 | 0.066396 | 0.00067288 | -0.0049709 | -0.00062588 | 0.022888 | 86.8561 | -35.4492 | 86.8011 | -33.4076 | -35.2386 | -35.1883 | 0.14191 | 86.8469 | -35.3439 | 86.9293 | -34.3277 | -35.3165 | -35.1517 | -0.16937 | -0.032044 | -0.086976 | -0.054932 | 1.8723 | 2.0416 | 0.091553 | 0.041199 | 0.26093 | 0.21057 | -0.064087 | -0.041199 | 0.041198 | 0.082397 | 0.95215 | 1.0162 | 0.12817 | -0.036621 | 0.19226 | 0.027466
1420.1875 | 122.3557 | -35.2615 | 87.0026 | -0.13054 | 0.0026234 | 0.076449 | 0.07426 | -0.0019503 | 0.0029438 | 0.0011108 | -0.098968 | 0.0035911 | 0.067484 | 0.066823 | -0.00082252 | -0.0036893 | -0.00041228 | -0.064087 | 86.9843 | -35.4126 | 86.7599 | -33.4763 | -35.3302 | -35.2844 | -0.1236 | 86.9522 | -35.3073 | 86.911 | -34.4147 | -35.408 | -35.1791 | -0.15106 | -0.01831 | -0.24261 | -0.2243 | 1.7853 | 1.9363 | -0.022888 | -0.068664 | 0.12817 | 0.082398 | -0.045776 | -0.050354 | -0.091552 | -0.041198 | 0.84686 | 0.89264 | 0.082398 | -0.14648 | 0.12817 | -0.10071
1421.1719 | 122.3785 | -35.1562 | 86.9751 | -0.15343 | 0.0030506 | 0.077945 | 0.07191 | -0.0030185 | 0.00059397 | 0.00089721 | 0.035309 | 0.0020957 | 0.06962 | 0.061483 | -0.0012497 | -0.0056119 | 0.00022852 | -0.073242 | 87.0621 | -35.2524 | 86.9019 | -33.2886 | -35.1379 | -35.0647 | -0.036621 | 87.0529 | -35.1608 | 87.0621 | -34.3185 | -35.2203 | -34.9823 | -0.09613 | 0.086975 | -0.073243 | -0.16022 | 1.8677 | 1.9638 | 0.091553 | 0.018311 | 0.18768 | 0.11444 | -0.004578 | 0.077819 | 0.086975 | 0.009156 | 0.83771 | 0.84229 | 0.17395 | -0.064087 | 0.17853 | -0.059509
1422.1719 | 122.4655 | -35.0555 | 87.1124 | -0.092392 | 0.0024096 | 0.075809 | 0.074046 | -0.0023777 | 0.0035846 | -0.0014526 | 0.029205 | 0.0020957 | 0.067697 | 0.066182 | 0.00024568 | -0.0036893 | -0.0036167 | -0.091553 | 87.0941 | -35.2432 | 86.9705 | -33.1879 | -35.0739 | -35.0006 | 0 | 87.117 | -35.1242 | 87.1445 | -34.1217 | -35.1105 | -34.9045 | -0.18768 | -0.018311 | -0.14191 | -0.1236 | 1.8677 | 2.0554 | 0.054932 | -0.018311 | 0.24262 | 0.16937 | -0.068665 | 0.004577 | 0.032043 | 0.027466 | 0.93384 | 1.0025 | 0.15106 | -0.054932 | 0.21973 | 0.013733
1423.1719 | 122.3785 | -34.9777 | 87.2589 | -0.077133 | 0.0036914 | 0.074954 | 0.072123 | 0.0014675 | 0.00038017 | 0.0017518 | -0.034882 | 0.0027365 | 0.069406 | 0.068319 | -0.00060892 | -0.0041165 | 0.0038601 | -0.013733 | 87.2498 | -35.1013 | 87.0209 | -33.1329 | -35.0281 | -34.9731 | 0.013733 | 87.2314 | -35.0327 | 87.1582 | -34.1034 | -35.1242 | -34.9228 | -0.1236 | -0.009155 | -0.23804 | -0.22888 | 1.8448 | 1.9684 | 0.004577 | -0.050354 | 0.12817 | 0.073242 | -0.054932 | -0.027466 | -0.10071 | -0.073242 | 0.87433 | 0.92926 | 0.054931 | -0.14648 | 0.10986 | -0.091553
1424.1719 | 122.406 | -34.8679 | 87.355 | -0.092392 | 0.002196 | 0.078158 | 0.074046 | -2.7692e-05 | 0.00016657 | 0.0026062 | 0.013947 | 0.0010275 | 0.06727 | 0.06426 | -0.00060892 | -0.0056119 | 0.0012967 | 0.027466 | 87.323 | -34.996 | 87.1948 | -32.9315 | -34.8312 | -34.79 | 0.10529 | 87.3001 | -34.9228 | 87.3138 | -33.9386 | -34.9228 | -34.7443 | -0.12817 | -0.032044 | -0.16022 | -0.12817 | 1.9363 | 2.0645 | 0.07782 | 0.036621 | 0.20599 | 0.16479 | -0.054932 | -0.054932 | -0.041199 | 0.013733 | 0.92926 | 0.98419 | 0.1236 | -0.054932 | 0.17853 | 0
1425.1719 | 122.4884 | -34.79 | 87.4008 | -0.1458 | 0.0024096 | 0.076236 | 0.071482 | 0.00018591 | 0.0012348 | 0.0028198 | 0.013947 | 0.0012411 | 0.066415 | 0.063833 | 0.00024568 | -0.0041165 | 0.0025785 | -0.004578 | 87.3779 | -34.9182 | 87.2955 | -32.8125 | -34.7397 | -34.6802 | 0.10071 | 87.3642 | -34.8404 | 87.4512 | -33.7921 | -34.7717 | -34.6161 | -0.12817 | -0.022888 | -0.10529 | -0.082398 | 1.9775 | 2.1057 | 0.10986 | 0.050354 | 0.23804 | 0.17853 | -0.050354 | -0.036621 | 0.050354 | 0.086975 | 0.99793 | 1.0483 | 0.17395 | 0.01831 | 0.2243 | 0.068664
1426.1719 | 122.3694 | -34.7305 | 87.4878 | -0.061874 | 0.003905 | 0.074954 | 0.073832 | -0.0010959 | 0.0010212 | 0.0045288 | -0.013519 | 0.0033773 | 0.068338 | 0.066823 | -0.0014633 | -0.0028347 | 1.4916e-05 | 0.027466 | 87.4649 | -34.8587 | 87.3367 | -32.7576 | -34.6939 | -34.6481 | 0.059509 | 87.4786 | -34.7946 | 87.442 | -33.8058 | -34.8129 | -34.6252 | -0.12817 | -0.022888 | -0.15106 | -0.12817 | 1.973 | 2.1011 | 0.082398 | 0.036621 | 0.21057 | 0.1648 | -0.064087 | -0.009155 | -0.045776 | -0.036621 | 0.92468 | 0.98877 | 0.10529 | -0.082397 | 0.16937 | -0.01831
1427.1719 | 122.4609 | -34.6252 | 87.5153 | -0.11528 | 0.0034778 | 0.078372 | 0.074473 | 0.0025357 | -4.7032e-05 | 0.003247 | 0.0078429 | 0.0031637 | 0.069192 | 0.060415 | -0.00082252 | -0.0034757 | 0.0032193 | -0.022888 | 87.5153 | -34.7763 | 87.4008 | -32.6614 | -34.6252 | -34.5566 | 0.027466 | 87.5153 | -34.7305 | 87.5839 | -33.6639 | -34.6481 | -34.4788 | -0.15106 | 0 | -0.11444 | -0.11444 | 1.9638 | 2.1149 | 0.068664 | 0 | 0.21973 | 0.15106 | -0.10529 | 0 | 0.068664 | 0.068664 | 0.9613 | 1.0666 | 0.14648 | -0.022888 | 0.25177 | 0.082398
1428.1719 | 122.4518 | -34.5612 | 87.6114 | -0.069504 | 0.0019824 | 0.077731 | 0.072551 | -0.0036593 | 0.0027302 | 0.0028198 | 0.02005 | 0.0014547 | 0.066629 | 0.064473 | -0.0018905 | -0.0034757 | 0.00086952 | -0.004578 | 87.616 | -34.7031 | 87.5198 | -32.5058 | -34.5062 | -34.4421 | 0.1236 | 87.5977 | -34.6207 | 87.6892 | -33.5312 | -34.5474 | -34.3735 | -0.14191 | 0.004578 | -0.091553 | -0.096131 | 2.0554 | 2.1973 | 0.11902 | 0.054931 | 0.26093 | 0.19684 | -0.05951 | -0.013733 | 0.07782 | 0.091553 | 1.03 | 1.0895 | 0.18768 | 0.013733 | 0.24719 | 0.073243
1429.1719 | 122.3511 | -34.4742 | 87.7579 | -0.023727 | 0.0032642 | 0.076449 | 0.075755 | 0.0014675 | 0.0020892 | 0.0026062 | -0.034882 | 0.0031637 | 0.069192 | 0.067251 | -0.0014633 | -0.0028347 | 1.4916e-05 | -0.032043 | 87.7075 | -34.671 | 87.4832 | -32.5241 | -34.5383 | -34.4879 | -0.004578 | 87.7213 | -34.5657 | 87.648 | -33.6044 | -34.3369 | -34.2682 | -0.19684 | -0.050354 | -0.27466 | -0.22431 | 1.9501 | 2.1469 | -0.013733 | -0.064087 | 0.18311 | 0.13275 | -0.091553 | -0.036622 | -0.10986 | -0.073242 | 0.86975 | 0.9613 | 0.20599 | 0.13733 | 0.29755 | 0.22888
1430.1719 | 122.4792 | -34.4055 | 87.7487 | -0.11528 | 0.002196 | 0.075809 | 0.072978 | -0.0015231 | 0.0010212 | -0.00017099 | 0.0078429 | 0.0033773 | 0.068124 | 0.065542 | -0.00082252 | -0.0045437 | 0.0017239 | -0.032043 | 87.7625 | -34.5566 | 87.6297 | -32.3593 | -34.3872 | -34.3231 | 0.018311 | 87.7441 | -34.4696 | 87.7945 | -33.403 | -34.4192 | -34.2453 | -0.15106 | 0.013733 | -0.11902 | -0.13275 | 2.0462 | 2.1973 | 0.082398 | 0.018311 | 0.23346 | 0.16937 | -0.064086 | -0.004577 | 0.045777 | 0.050354 | 1.0025 | 1.0666 | 0.16022 | -0.013732 | 0.2243 | 0.050354
1431.1719 | 122.4426 | -34.3643 | 87.8082 | -0.11528 | 0.002196 | 0.077731 | 0.072123 | -0.00045509 | 0.0014484 | 0.003247 | 0.0017395 | 0.0027365 | 0.068338 | 0.065755 | -0.00060892 | -0.0034757 | 1.4916e-05 | -0.07782 | 87.8174 | -34.5016 | 87.6801 | -32.2998 | -34.346 | -34.2728 | -0.027466 | 87.8265 | -34.4238 | 87.8265 | -33.3755 | -34.4192 | -34.2087 | -0.13733 | 0.009155 | -0.12817 | -0.13733 | 2.0645 | 2.2018 | 0.091553 | 0.018311 | 0.22888 | 0.15564 | -0.059509 | 0.01831 | 0.01831 | 0 | 0.98877 | 1.0483 | 0.15564 | -0.054931 | 0.21515 | 0.004578
1432.1875 | 122.4243 | -34.3094 | 87.8906 | -0.031357 | 0.0032642 | 0.078799 | 0.0749 | -0.0040867 | 0.00016657 | 0.005597 | -0.11118 | 0.0046591 | 0.066202 | 0.068105 | -0.0021041 | -0.0036893 | -0.00062588 | 0 | 87.9181 | -34.4421 | 87.735 | -32.2311 | -34.2865 | -34.2636 | 0.041199 | 87.9044 | -34.3643 | 87.9135 | -33.3252 | -34.3964 | -34.1812 | -0.13275 | 0.027466 | -0.15564 | -0.18311 | 2.0782 | 2.211 | 0.045776 | 0.022888 | 0.17853 | 0.15564 | -0.054932 | 0.013733 | 0.022888 | 0.009155 | 0.98419 | 1.0391 | 0.12817 | -0.086975 | 0.18311 | -0.032043
1433.1719 | 122.4701 | -34.227 | 87.9318 | -0.11528 | 0.003905 | 0.078372 | 0.075114 | -0.0019503 | 0.0010212 | 0.0026062 | -0.016571 | 0.0018821 | 0.067056 | 0.060415 | -0.0016769 | -0.0047573 | 0.0042875 | -0.036621 | 87.9318 | -34.3735 | 87.822 | -32.1167 | -34.1995 | -34.1309 | 0.045776 | 87.8998 | -34.3002 | 87.9913 | -33.1879 | -34.2316 | -34.0668 | -0.14648 | 0 | -0.10986 | -0.10986 | 2.1103 | 2.2568 | 0.096131 | 0.027466 | 0.24262 | 0.17395 | -0.073242 | -0.032044 | 0.059509 | 0.091553 | 1.0391 | 1.1124 | 0.16022 | -0.004577 | 0.23346 | 0.068665
1434.1719 | 122.4197 | -34.1949 | 87.9913 | -0.092392 | 0.0024096 | 0.076022 | 0.072978 | -0.0017367 | 0.00059397 | 0.0017518 | -0.013519 | 0.0018821 | 0.066629 | 0.063405 | -0.00039512 | -0.0049709 | -0.0010531 | -0.045776 | 87.9776 | -34.3414 | 87.854 | -32.0526 | -34.1492 | -34.0759 | 0.009155 | 87.9959 | -34.2545 | 88.0188 | -33.1512 | -34.1995 | -34.0118 | -0.14648 | -0.013733 | -0.13733 | -0.1236 | 2.1423 | 2.2888 | 0.11902 | 0.045776 | 0.2655 | 0.19226 | -0.05951 | 0.004578 | 0.027466 | 0.022888 | 1.0437 | 1.1032 | 0.18311 | -0.004578 | 0.24262 | 0.054932
1435.1719 | 122.4014 | -34.1537 | 88.0508 | -0.046616 | 0.0043322 | 0.07709 | 0.074687 | 0.00018591 | 0.00038017 | 0.0034606 | -0.10507 | 0.0042319 | 0.067697 | 0.067464 | -0.0012497 | -0.0024075 | 0.0015103 | 0.009155 | 88.1058 | -34.2911 | 87.8677 | -32.0389 | -34.1171 | -34.0942 | 0.036621 | 88.0554 | -34.1904 | 88.0554 | -33.1421 | -34.2041 | -34.0118 | -0.13733 | 0.054932 | -0.1831 | -0.23804 | 2.1149 | 2.2522 | 0.05951 | 0.036622 | 0.19684 | 0.17395 | -0.036621 | 0.004578 | 0.004578 | 0 | 1.0117 | 1.0483 | 0.14191 | -0.050354 | 0.17853 | -0.013733
1436.1719 | 122.4289 | -34.0118 | 88.1516 | -0.16106 | 0.0030506 | 0.079654 | 0.073832 | -0.0010959 | 0.00059397 | 0.0026062 | -0.010467 | 0.0031637 | 0.068338 | 0.06661 | 0.0015275 | -0.0028347 | 0.0030057 | -0.068665 | 88.1699 | -34.1354 | 88.0234 | -31.8466 | -33.989 | -33.9066 | -0.068665 | 88.1653 | -34.0576 | 88.179 | -32.991 | -34.0485 | -33.8699 | -0.1236 | 0.018311 | -0.12817 | -0.14648 | 2.1652 | 2.2888 | 0.10529 | 0.022888 | 0.22888 | 0.14648 | -0.045776 | 0.013733 | 0.027466 | 0.013733 | 1.0208 | 1.0666 | 0.14191 | -0.036621 | 0.18768 | 0.009155
1437.1719 | 122.438 | -33.9432 | 88.2294 | -0.1458 | 0.0047596 | 0.078372 | 0.073832 | 0.0063809 | 0.0012348 | 0.0026062 | -0.025726 | 0.0033773 | 0.068338 | 0.062978 | -0.00060892 | -0.0028347 | 0.00044212 | -0.086975 | 88.2431 | -34.0714 | 88.1012 | -31.7963 | -33.902 | -33.8196 | -0.036621 | 88.2339 | -33.9844 | 88.2248 | -32.9269 | -33.9523 | -33.7738 | -0.12817 | 0.013733 | -0.12817 | -0.14191 | 2.1469 | 2.2751 | 0.1236 | 0.041198 | 0.25177 | 0.16937 | -0.041199 | 0.004578 | -0.004578 | -0.009156 | 1.0162 | 1.0574 | 0.16937 | -0.009156 | 0.21057 | 0.032043
1438.1875 | 122.4106 | -33.902 | 88.3301 | -0.061874 | 0.0017688 | 0.079654 | 0.075328 | -0.0010959 | 0.00080757 | 0.00089721 | -0.077606 | 0.0059409 | 0.069192 | 0.06896 | -0.0029587 | -0.0043301 | -0.0042575 | -0.032043 | 88.3209 | -34.0576 | 88.0829 | -31.8008 | -33.8974 | -33.8654 | -0.009155 | 88.3072 | -33.9615 | 88.2843 | -32.8857 | -33.9844 | -33.7738 | -0.15564 | -0.009155 | -0.24719 | -0.23804 | 2.1011 | 2.2568 | 0.036622 | 0.004578 | 0.19226 | 0.16022 | -0.059509 | -0.022888 | -0.045776 | -0.022888 | 1.0162 | 1.0757 | 0.12817 | -0.082397 | 0.18768 | -0.022888
1439.1719 | 122.4243 | -33.8791 | 88.3163 | -0.12291 | 0.0049732 | 0.078372 | 0.073405 | -0.0021641 | 0.0010212 | 0.004956 | -0.013519 | 0.0020957 | 0.068338 | 0.05956 | 0.0011001 | -0.0032619 | 0.0010831 | -0.054932 | 88.3026 | -33.9798 | 88.2065 | -31.6269 | -33.815 | -33.7372 | 0.009155 | 88.2889 | -33.9249 | 88.3438 | -32.8079 | -33.8882 | -33.6777 | -0.10071 | -0.013733 | -0.10986 | -0.09613 | 2.2522 | 2.3529 | 0.14191 | 0.064087 | 0.24261 | 0.16479 | -0.045777 | -0.027466 | 0.027466 | 0.054932 | 1.0712 | 1.1169 | 0.20142 | -0.009156 | 0.24719 | 0.036621
1440.1719 | 122.4426 | -33.8608 | 88.3163 | -0.11528 | 0.0049732 | 0.078586 | 0.072978 | -0.0013095 | 0.0020892 | 0.0026062 | -0.0013123 | 0.0029501 | 0.068124 | 0.059987 | 0.0015275 | -0.0043301 | -0.00019868 | -0.073242 | 88.3301 | -33.9844 | 88.0692 | -31.6681 | -33.9569 | -33.7601 | -0.036621 | 88.2935 | -33.9294 | 88.353 | -32.7713 | -33.8928 | -33.6868 | -0.1236 | 0.013733 | -0.24719 | -0.26093 | 2.1927 | 2.3163 | 0.10071 | -0.09613 | 0.2243 | 0.027466 | -0.068664 | -0.022888 | 0.036621 | 0.059509 | 1.0895 | 1.1581 | 0.17395 | -0.032043 | 0.24261 | 0.036621
1441.1719 | 122.4564 | -33.8333 | 88.3347 | -0.092392 | 0.0034778 | 0.078586 | 0.07426 | -0.0013095 | 0.0029438 | 0.0023926 | -0.016571 | 0.0020957 | 0.066415 | 0.066396 | -0.00039512 | -0.0017667 | 1.4916e-05 | -0.068665 | 88.3209 | -33.9478 | 88.1973 | -31.5948 | -33.7784 | -33.7326 | -0.036621 | 88.3209 | -33.8882 | 88.3713 | -32.7713 | -33.8699 | -33.6777 | -0.11444 | -0.013733 | -0.13733 | -0.1236 | 2.2385 | 2.3529 | 0.10071 | 0.054932 | 0.21515 | 0.16937 | -0.054932 | -0.013733 | 0.036621 | 0.050354 | 1.062 | 1.1169 | 0.15564 | -0.036621 | 0.21057 | 0.018311
1442.1719 | 122.4426 | -33.8058 | 88.3347 | -0.11528 | 0.0051868 | 0.079867 | 0.073619 | 0.00018591 | 0.0014484 | 0.00068361 | 0.0078429 | 0.0010275 | 0.067484 | 0.061696 | -0.0018905 | -0.0045437 | -0.00062588 | -0.022888 | 88.3484 | -33.9478 | 88.2111 | -31.5811 | -33.7646 | -33.696 | -0.009155 | 88.3301 | -33.8699 | 88.4216 | -32.7026 | -33.8287 | -33.6365 | -0.14191 | 0.013733 | -0.1236 | -0.13733 | 2.2247 | 2.3666 | 0.10986 | 0.041199 | 0.25177 | 0.18311 | -0.064087 | -0.004578 | 0.086975 | 0.091553 | 1.1032 | 1.1673 | 0.16937 | -0.022888 | 0.23346 | 0.041199
1443.1719 | 122.438 | -33.8196 | 88.3575 | -0.16106 | 0.0047596 | 0.078158 | 0.0749 | 0.0010403 | 0.0012348 | 0.003888 | -0.022675 | 0.0042319 | 0.065988 | 0.068319 | -0.00082252 | -0.0045437 | 0.0021511 | -0.086975 | 88.3575 | -33.9661 | 88.2156 | -31.5674 | -33.7601 | -33.6823 | -0.036621 | 88.3301 | -33.8837 | 88.3896 | -32.7118 | -33.8425 | -33.6594 | -0.14648 | 0 | -0.14191 | -0.14191 | 2.2522 | 2.3987 | 0.13733 | 0.059509 | 0.28381 | 0.20599 | -0.064087 | -0.027466 | 0.032043 | 0.059509 | 1.1078 | 1.1719 | 0.16022 | -0.022888 | 0.2243 | 0.041199
1444.1719 | 122.5067 | -33.7555 | 88.3942 | -0.092392 | 0.0047596 | 0.078158 | 0.073832 | 0.00082671 | 0.0014484 | 0.00089721 | -0.013519 | 0.0018821 | 0.068124 | 0.064687 | -0.00060892 | -0.0047573 | 0.0023647 | 0.018311 | 88.3987 | -33.8882 | 88.3438 | -31.366 | -33.6182 | -33.5678 | 0.13275 | 88.3667 | -33.7921 | 88.4903 | -32.5928 | -33.728 | -33.5495 | -0.13275 | 0.004578 | -0.050354 | -0.054932 | 2.3895 | 2.5223 | 0.18768 | 0.13733 | 0.32043 | 0.27008 | -0.036621 | -0.027466 | 0.09613 | 0.1236 | 1.1627 | 1.1993 | 0.20599 | 0.027466 | 0.24261 | 0.064087
1445.1719 | 122.4518 | -33.7601 | 88.3804 | -0.1458 | 0.006255 | 0.078372 | 0.076823 | -2.7692e-05 | 0.0023028 | 0.0023926 | -0.022675 | 0.0035911 | 0.068338 | 0.067464 | -0.00039512 | -0.0034757 | 0.0017239 | -0.07782 | 88.4125 | -33.8791 | 88.2751 | -31.4575 | -33.7189 | -33.6319 | -0.050354 | 88.4125 | -33.8242 | 88.4308 | -32.6569 | -33.7921 | -33.5953 | -0.11902 | 0.032044 | -0.10529 | -0.13733 | 2.3026 | 2.4216 | 0.12817 | 0.041199 | 0.24719 | 0.16022 | -0.064087 | 0.032044 | 0.050354 | 0.01831 | 1.1032 | 1.1673 | 0.1648 | -0.032043 | 0.22888 | 0.032044
1446.1719 | 122.4472 | -33.7372 | 88.4354 | -0.1458 | 0.004546 | 0.081149 | 0.076823 | -0.0010959 | 0.003371 | 0.0043152 | -0.059296 | 0.0016685 | 0.068338 | 0.06191 | 0.0011001 | -0.0013393 | -0.00019868 | -0.054932 | 88.4125 | -33.8562 | 88.3347 | -31.3522 | -33.6502 | -33.5724 | 0.004578 | 88.4125 | -33.7692 | 88.5086 | -32.5882 | -33.7418 | -33.5449 | -0.11902 | -0.022888 | -0.10071 | -0.07782 | 2.3849 | 2.504 | 0.1648 | 0.086975 | 0.28381 | 0.20599 | -0.032043 | -0.022888 | 0.073242 | 0.09613 | 1.149 | 1.181 | 0.19226 | -0.004577 | 0.2243 | 0.027466
1447.1719 | 122.4472 | -33.728 | 88.4308 | -0.10002 | 0.0041186 | 0.078799 | 0.076182 | -0.00066869 | 0.0012348 | 0.0013244 | -0.028778 | 0.0020957 | 0.066202 | 0.072378 | 3.2078e-05 | -0.0013393 | -0.00041228 | -0.059509 | 88.4583 | -33.8516 | 88.3163 | -31.3522 | -33.6456 | -33.5861 | -0.032043 | 88.4216 | -33.7692 | 88.4995 | -32.5974 | -33.7372 | -33.5358 | -0.1236 | 0.027466 | -0.11444 | -0.14191 | 2.3758 | 2.4994 | 0.14191 | 0.082397 | 0.2655 | 0.20599 | -0.041199 | -0.009155 | 0.068665 | 0.07782 | 1.1307 | 1.1719 | 0.19226 | -0.009156 | 0.23346 | 0.032043
1448.1719 | 122.4014 | -33.7418 | 88.4354 | -0.13817 | 0.0051868 | 0.07944 | 0.075969 | 0.00039951 | 0.0029438 | -0.00081179 | -0.03183 | 0.0012411 | 0.069406 | 0.06426 | -0.00082252 | -0.0030483 | -0.0012667 | -0.082397 | 88.4903 | -33.8333 | 88.3392 | -31.3339 | -33.6136 | -33.5541 | -0.036621 | 88.4766 | -33.7463 | 88.5223 | -32.6431 | -33.7509 | -33.5861 | -0.091553 | 0.054931 | -0.096131 | -0.15106 | 2.4078 | 2.4994 | 0.18768 | 0.12817 | 0.27924 | 0.21973 | -0.004578 | 0.041198 | 0.086975 | 0.045777 | 1.0986 | 1.1032 | 0.15564 | -0.009156 | 0.16022 | -0.004578
1449.1719 | 122.4701 | -33.6823 | 88.4949 | -0.084763 | 0.0068958 | 0.080295 | 0.076609 | -0.0034457 | 0.0018756 | 0.0028198 | -0.059296 | 0.0035911 | 0.067056 | 0.068746 | -0.00060892 | -0.00027134 | 0.00044212 | -0.045776 | 88.4628 | -33.815 | 88.3575 | -31.2927 | -33.5083 | -33.4854 | 0.013733 | 88.4949 | -33.7418 | 88.5727 | -32.4875 | -33.6777 | -33.4579 | -0.13275 | -0.032043 | -0.13733 | -0.10529 | 2.3895 | 2.5223 | 0.19684 | 0.17395 | 0.32959 | 0.3067 | -0.059509 | 0 | 0.07782 | 0.07782 | 1.1948 | 1.2543 | 0.2243 | 0.004578 | 0.28381 | 0.064087
1450.1719 | 122.5342 | -33.6182 | 88.5452 | -0.084763 | 0.002837 | 0.079654 | 0.07426 | 0.00061311 | 0.0029438 | 4.2808e-05 | -0.013519 | 0.0033773 | 0.068979 | 0.067464 | -0.0012497 | -0.0030483 | 0.0015103 | 0.032043 | 88.5315 | -33.7601 | 88.4491 | -31.1142 | -33.4534 | -33.4122 | 0.13275 | 88.5132 | -33.7189 | 88.6002 | -32.4142 | -33.6044 | -33.5312 | -0.14191 | -0.013733 | -0.09613 | -0.082397 | 2.504 | 2.6459 | 0.20599 | 0.16479 | 0.3479 | 0.3067 | -0.10071 | -0.032043 | 0.054932 | 0.086975 | 1.2039 | 1.3046 | 0.086975 | 0.013733 | 0.18768 | 0.11444
1451.1719 | 122.4472 | -33.6273 | 88.5132 | -0.13817 | 0.0051868 | 0.081363 | 0.076823 | 0.00082671 | 0.003371 | 0.0019654 | -0.089813 | 0.0040183 | 0.067911 | 0.064687 | -0.00082252 | -0.0032619 | 0.0030057 | -0.032043 | 88.5178 | -33.7646 | 88.4354 | -31.1325 | -33.5175 | -33.4534 | 0.004578 | 88.5086 | -33.6914 | 88.6002 | -32.46 | -33.6411 | -33.4351 | -0.13733 | 0.004577 | -0.07782 | -0.082397 | 2.4948 | 2.6321 | 0.17395 | 0.10986 | 0.31128 | 0.24719 | -0.064087 | -0.004578 | 0.086975 | 0.091553 | 1.1673 | 1.2314 | 0.19226 | -0.013733 | 0.25635 | 0.050354
1452.1719 | 122.406 | -33.6182 | 88.5544 | -0.11528 | 0.0060412 | 0.080508 | 0.073191 | 0.0025357 | 0.0018756 | 0.00047001 | -0.037933 | 0.0053001 | 0.071329 | 0.070669 | -0.00060892 | -0.0015529 | -0.00019868 | -0.07782 | 88.5315 | -33.7509 | 88.4079 | -31.1508 | -33.5358 | -33.4579 | -0.068665 | 88.5315 | -33.7006 | 88.5773 | -32.4371 | -33.6319 | -33.4351 | -0.13275 | -0.022888 | -0.14648 | -0.1236 | 2.4673 | 2.6001 | 0.16022 | 0.082397 | 0.29297 | 0.21515 | -0.082398 | -0.022888 | 0.022889 | 0.045777 | 1.181 | 1.2634 | 0.1831 | -0.013733 | 0.2655 | 0.068665
1453.1719 | 122.4838 | -33.609 | 88.5498 | -0.023727 | 0.0049732 | 0.079654 | 0.075114 | -0.0021641 | 0.0025166 | -0.00038459 | -0.16 | 0.0035911 | 0.067056 | 0.072805 | 0.0011001 | -0.0021939 | -0.00062588 | 0.027466 | 88.5498 | -33.7234 | 88.4674 | -31.0867 | -33.1146 | -33.3618 | 0.1831 | 88.5269 | -33.6731 | 88.6139 | -32.3547 | -33.5449 | -33.3801 | -0.11444 | 0 | -0.082398 | -0.082398 | 2.5223 | 2.6367 | 0.24719 | 0.49439 | 0.36163 | 0.60883 | -0.064087 | -0.022888 | 0.064087 | 0.086975 | 1.2543 | 1.3184 | 0.22888 | 0.064087 | 0.29297 | 0.12817
1454.1719 | 122.4792 | -33.6136 | 88.5498 | -0.30601 | 0.0075366 | 0.082217 | 0.078532 | -0.00088229 | 0.0025166 | 0.0043152 | -0.074554 | 0.0046591 | 0.070474 | 0.071309 | 0.0015275 | -0.0017667 | -0.00062588 | -0.041199 | 88.4811 | -33.7967 | 88.4399 | -31.0593 | -33.4579 | -33.4167 | 0.013733 | 88.5269 | -33.6868 | 88.623 | -32.3868 | -33.6136 | -33.3984 | -0.18311 | -0.068665 | -0.10986 | -0.041199 | 2.5543 | 2.7374 | 0.19684 | 0.15564 | 0.37994 | 0.33875 | -0.073243 | -0.022888 | 0.073242 | 0.09613 | 1.2268 | 1.3 | 0.21515 | 0 | 0.28839 | 0.073243
1455.1719 | 122.4243 | -33.6136 | 88.5864 | -0.11528 | 0.0054004 | 0.079226 | 0.078105 | -0.0047275 | 0.003371 | 0.001538 | -0.040985 | 0.0046591 | 0.06962 | 0.06426 | -0.0014633 | -0.0039029 | -0.00083948 | -0.054932 | 88.5406 | -33.7463 | 88.4354 | -30.9998 | -33.49 | -33.4213 | -0.004578 | 88.5544 | -33.696 | 88.5956 | -32.3914 | -33.6182 | -33.4213 | -0.13275 | -0.045777 | -0.15106 | -0.10528 | 2.6138 | 2.7466 | 0.19226 | 0.1236 | 0.32501 | 0.25635 | -0.082398 | -0.032044 | 0.009155 | 0.041199 | 1.2222 | 1.3046 | 0.19226 | -0.004578 | 0.27466 | 0.07782
1456.1875 | 122.3099 | -33.5999 | 88.6139 | -0.061874 | 0.0047596 | 0.079654 | 0.07725 | -0.0023777 | 0.00059397 | 0.00068361 | -0.062347 | 0.0025229 | 0.070901 | 0.06896 | 3.2078e-05 | -0.0030483 | 0.0036465 | -0.027466 | 88.5773 | -33.7601 | 88.353 | -31.1142 | -33.5403 | -33.4991 | 0.009155 | 88.6093 | -33.696 | 88.5132 | -32.46 | -33.3939 | -33.316 | -0.16022 | -0.036621 | -0.26093 | -0.22431 | 2.4857 | 2.6459 | 0.10071 | 0.05951 | 0.26093 | 0.21973 | -0.09613 | -0.004578 | -0.10071 | -0.09613 | 1.1398 | 1.236 | 0.28381 | 0.20599 | 0.37994 | 0.30212
1457.1719 | 122.4701 | -33.6411 | 88.5269 | -0.13817 | 0.005614 | 0.080935 | 0.076396 | 0.0044583 | 0.0042254 | 0.0047424 | -0.071503 | -0.00025411 | 0.067484 | 0.066396 | -0.00018152 | -0.0021939 | -0.00019868 | -0.027466 | 88.5223 | -33.7555 | 88.4445 | -30.9814 | -33.4808 | -33.4122 | 0.004578 | 88.5223 | -33.6823 | 88.504 | -32.4051 | -33.4167 | -33.2703 | -0.11444 | -0.004578 | -0.082398 | -0.07782 | 2.6596 | 2.774 | 0.22888 | 0.16022 | 0.34332 | 0.27466 | -0.041199 | -0.004578 | -0.022889 | -0.018311 | 1.236 | 1.2772 | 0.37079 | 0.2243 | 0.41199 | 0.2655
1458.1719 | 122.4518 | -33.5907 | 88.5269 | -0.1458 | 0.003905 | 0.080935 | 0.075541 | -0.0038729 | 0.0027302 | 0.005597 | -0.010467 | 0.0025229 | 0.069833 | 0.063192 | -0.00060892 | -0.0019803 | 1.4916e-05 | -0.032043 | 88.559 | -33.7509 | 88.4033 | -30.9906 | -33.4579 | -33.4122 | 0.004578 | 88.5452 | -33.6731 | 88.6047 | -32.3227 | -33.5678 | -33.3801 | -0.16022 | 0.032043 | -0.1236 | -0.15564 | 2.6001 | 2.7603 | 0.17853 | 0.13275 | 0.33875 | 0.29297 | -0.082398 | 0.01831 | 0.077819 | 0.059509 | 1.268 | 1.3504 | 0.21057 | 0.022888 | 0.29297 | 0.10529
1459.1719 | 122.4472 | -33.609 | 88.5635 | -0.13054 | 0.0071094 | 0.079654 | 0.078105 | -0.0034457 | 0.0023028 | 0.003247 | -0.013519 | 0.0038047 | 0.069833 | 0.059987 | 0.0013139 | -0.0026211 | 0.0012967 | -0.032043 | 88.5269 | -33.8013 | 88.4033 | -30.9769 | -33.4671 | -33.4122 | -0.004578 | 88.5406 | -33.6685 | 88.591 | -32.3318 | -33.5953 | -33.3984 | -0.19226 | -0.036621 | -0.16022 | -0.1236 | 2.6321 | 2.8244 | 0.19684 | 0.14191 | 0.3891 | 0.33417 | -0.059509 | -0.022889 | 0.027465 | 0.050354 | 1.2772 | 1.3367 | 0.21057 | 0.013733 | 0.27008 | 0.073242
1460.1719 | 122.4518 | -33.5861 | 88.5635 | -0.13054 | 0.006255 | 0.081149 | 0.076823 | -0.0028049 | 0.0044392 | -0.00038459 | -0.071503 | 0.0038047 | 0.069406 | 0.068319 | 0.00088648 | -0.0039029 | -0.00083948 | -0.032043 | 88.5498 | -33.7463 | 88.4628 | -30.8716 | -33.4396 | -33.3618 | 0.004578 | 88.5406 | -33.6411 | 88.6368 | -32.3135 | -33.5861 | -33.3755 | -0.16022 | -0.013733 | -0.10071 | -0.086975 | 2.7145 | 2.8748 | 0.22431 | 0.14648 | 0.38452 | 0.3067 | -0.054931 | -0.022889 | 0.073242 | 0.096131 | 1.2726 | 1.3275 | 0.21057 | 0 | 0.2655 | 0.054931
1461.1719 | 122.4518 | -33.577 | 88.5361 | -0.13054 | 0.006255 | 0.080722 | 0.076182 | -2.7692e-05 | 0.0052936 | 0.0028198 | -0.016571 | 0.0031637 | 0.066415 | 0.069173 | -0.00039512 | -0.0032619 | 0.0038601 | -0.036621 | 88.5544 | -33.7463 | 88.4308 | -30.8853 | -33.4213 | -33.3618 | 0.050354 | 88.5406 | -33.6868 | 88.6368 | -32.2403 | -33.5541 | -33.3481 | -0.16937 | 0.01831 | -0.10529 | -0.1236 | 2.6917 | 2.861 | 0.21515 | 0.15564 | 0.38452 | 0.32501 | -0.10986 | 0.004577 | 0.10071 | 0.096131 | 1.3367 | 1.4465 | 0.22888 | 0.022888 | 0.33875 | 0.13275
1462.1719 | 122.438 | -33.6456 | 88.5223 | -0.10765 | 0.0071094 | 0.080722 | 0.079814 | -0.0015231 | 0.0057208 | 0.001538 | -0.013519 | 0.0025229 | 0.069192 | 0.066182 | 0.0011001 | -0.0043301 | 0.0027921 | -0.059509 | 88.5635 | -33.7646 | 88.3987 | -30.9036 | -33.4442 | -33.3984 | -0.064087 | 88.5269 | -33.7189 | 88.5452 | -32.3135 | -33.5953 | -33.3939 | -0.11902 | 0.041199 | -0.1236 | -0.16479 | 2.742 | 2.861 | 0.24719 | 0.20142 | 0.36621 | 0.32043 | -0.073242 | 0.004578 | 0.022888 | 0.01831 | 1.3321 | 1.4053 | 0.25177 | 0.050354 | 0.32501 | 0.1236
1463.1719 | 122.4701 | -33.6456 | 88.5132 | -0.12291 | 0.0066822 | 0.082431 | 0.077037 | 0.00061311 | 0.0046528 | 0.0030334 | -0.086761 | 0.0031637 | 0.069192 | 0.065328 | -0.00039512 | -0.0030483 | 0.0036465 | -0.036621 | 88.5178 | -33.783 | 88.4399 | -30.8624 | -33.4305 | -33.3847 | 0.004578 | 88.4995 | -33.696 | 88.623 | -32.2586 | -33.5861 | -33.3618 | -0.13733 | 0.004577 | -0.073243 | -0.07782 | 2.7832 | 2.9205 | 0.26093 | 0.21515 | 0.39825 | 0.35248 | -0.050354 | -0.013733 | 0.10986 | 0.1236 | 1.387 | 1.4374 | 0.28381 | 0.059509 | 0.33417 | 0.10986
1464.1719 | 122.4884 | -33.6639 | 88.4811 | -0.11528 | 0.0071094 | 0.083499 | 0.080882 | -0.0023777 | 0.0040118 | 0.0028198 | -0.068451 | 0.0038047 | 0.070047 | 0.065542 | 0.00024568 | -0.0030483 | 0.00086952 | -0.24719 | 88.4125 | -33.8287 | 88.4308 | -30.8258 | -33.4763 | -33.3984 | 0.027466 | 88.4857 | -33.7418 | 88.5864 | -32.2723 | -33.5907 | -33.3847 | -0.1648 | -0.068664 | -0.050354 | 0.01831 | 2.8381 | 3.0029 | 0.2655 | 0.18768 | 0.4303 | 0.35248 | -0.07782 | 0.004578 | 0.10529 | 0.10071 | 1.3916 | 1.4694 | 0.27923 | 0.073242 | 0.35706 | 0.15106
1465.1719 | 122.438 | -33.7234 | 88.4399 | -0.13817 | 0.007323 | 0.083713 | 0.0796 | 0.0033901 | 0.0012348 | 0.0017518 | -0.056244 | 0.0031637 | 0.068979 | 0.067891 | -0.0010361 | -0.00091214 | -0.00083948 | -0.045776 | 88.4262 | -33.8791 | 88.3987 | -30.9174 | -33.5312 | -33.4076 | 0.018311 | 88.4216 | -33.7784 | 88.5269 | -32.2952 | -33.6456 | -33.4396 | -0.15564 | -0.013733 | -0.041198 | -0.027465 | 2.8061 | 2.9617 | 0.31586 | 0.19226 | 0.4715 | 0.3479 | -0.054931 | -0.01831 | 0.086976 | 0.10529 | 1.4282 | 1.4832 | 0.28381 | 0.07782 | 0.33875 | 0.13275
1466.1875 | 122.374 | -33.7784 | 88.4583 | -0.061874 | 0.0071094 | 0.082217 | 0.076609 | 0.00018591 | 0.0020892 | 0.0047424 | -0.040985 | 0.0035911 | 0.069406 | 0.065542 | -0.00060892 | -0.0013393 | -0.00019868 | 0.018311 | 88.4308 | -33.8516 | 88.266 | -30.9128 | -33.5403 | -33.4946 | 0.036621 | 88.4399 | -33.7967 | 88.4033 | -32.3776 | -33.7509 | -33.5632 | -0.073243 | -0.027466 | -0.19226 | -0.16479 | 2.8656 | 2.9388 | 0.28381 | 0.23804 | 0.35706 | 0.31128 | -0.018311 | -0.018311 | -0.054932 | -0.036621 | 1.4008 | 1.4191 | 0.21515 | 0.027465 | 0.23346 | 0.045776
1467.1719 | 122.4472 | -33.7418 | 88.3987 | -0.16869 | 0.0064686 | 0.083072 | 0.079173 | 0.0042447 | 0.0020892 | 0.0017518 | 0.010895 | 0.0029501 | 0.07261 | 0.064473 | -0.0021041 | -0.0028347 | 0.0012967 | -0.041199 | 88.3987 | -33.8516 | 88.3072 | -30.8899 | -33.8654 | -33.5037 | -0.027466 | 88.353 | -33.815 | 88.504 | -32.2998 | -33.696 | -33.4671 | -0.10986 | 0 | -0.091553 | -0.091553 | 2.8519 | 2.9617 | 0.23804 | -0.1236 | 0.3479 | -0.013732 | -0.073242 | -0.045777 | 0.10529 | 0.15106 | 1.442 | 1.5152 | 0.27466 | 0.045776 | 0.3479 | 0.11902
1468.1719 | 122.4243 | -33.783 | 88.3759 | -0.16106 | 0.0077502 | 0.082858 | 0.078959 | -0.0036593 | 0.0052936 | 0.0034606 | -0.047089 | 0.0033773 | 0.070047 | 0.062337 | 0.00088648 | -0.0019803 | -0.00062588 | -0.059509 | 88.385 | -33.8928 | 88.2843 | -30.8441 | -33.522 | -33.4808 | 0.009155 | 88.385 | -33.8287 | 88.4537 | -32.341 | -33.7143 | -33.5129 | -0.10986 | 0.009156 | -0.091552 | -0.10071 | 2.9388 | 3.0487 | 0.30212 | 0.26093 | 0.41199 | 0.37079 | -0.045776 | 0.009156 | 0.07782 | 0.068664 | 1.442 | 1.4877 | 0.27008 | 0.068665 | 0.31586 | 0.11444
1469.1719 | 122.4518 | -33.7875 | 88.3438 | -0.069504 | 0.0068958 | 0.074527 | 0.0796 | 0.00082671 | 0.00080757 | 0.0026062 | -0.022675 | 0.0027365 | 0.070261 | 0.064901 | 3.2078e-05 | -0.0024075 | 0.0027921 | 0.10529 | 88.3896 | -33.9661 | 88.2614 | -30.8807 | -33.4488 | -33.4763 | 0.13733 | 88.3804 | -33.8425 | 88.4262 | -32.2998 | -33.7189 | -33.5175 | -0.17853 | 0.045776 | -0.082397 | -0.12817 | 2.9068 | 3.0853 | 0.31128 | 0.33875 | 0.48981 | 0.51727 | -0.054931 | 0.036621 | 0.082397 | 0.045776 | 1.4877 | 1.5427 | 0.27008 | 0.068665 | 0.32501 | 0.1236
1470.1719 | 122.4197 | -33.8242 | 88.3575 | -0.10765 | 0.0086048 | 0.082217 | 0.081095 | -2.7692e-05 | 0.0059344 | 0.0045288 | -0.037933 | 0.0035911 | 0.073465 | 0.067891 | -0.00060892 | -0.0036893 | 0.0045011 | -0.07782 | 88.3484 | -33.9432 | 88.1973 | -30.8807 | -33.5999 | -33.5403 | -0.054932 | 88.3255 | -33.8791 | 88.3713 | -32.3547 | -33.7875 | -33.5678 | -0.11902 | -0.009155 | -0.16022 | -0.15106 | 2.9434 | 3.0624 | 0.28381 | 0.2243 | 0.40283 | 0.34332 | -0.054931 | -0.032044 | 0.013733 | 0.045777 | 1.4694 | 1.5244 | 0.25635 | 0.036621 | 0.31128 | 0.091552
1471.1719 | 122.4838 | -33.8104 | 88.3392 | -0.10002 | 0.0071094 | 0.081149 | 0.078959 | 0.00018591 | 0.0018756 | 0.003247 | 0.0047913 | 0.0046591 | 0.072397 | 0.069173 | -0.00082252 | -0.0039029 | 0.0012967 | -0.032043 | 88.3667 | -33.9523 | 88.2111 | -30.8166 | -33.5403 | -33.5037 | 0.041199 | 88.3438 | -33.8882 | 88.4033 | -32.2861 | -33.7189 | -33.5175 | -0.14191 | 0.027466 | -0.12817 | -0.15564 | 2.9938 | 3.1357 | 0.3067 | 0.27008 | 0.44861 | 0.41199 | -0.07782 | 0.004578 | 0.064087 | 0.059509 | 1.5244 | 1.6022 | 0.29297 | 0.091553 | 0.37079 | 0.16937
1472.1719 | 122.406 | -33.8379 | 88.3484 | -0.11528 | 0.0075366 | 0.081363 | 0.080241 | 0.0012539 | 0.0023028 | 4.2808e-05 | -0.025726 | 0.0042319 | 0.072397 | 0.070455 | -0.0014633 | -0.0051845 | 0.00022852 | -0.068665 | 88.3255 | -33.9523 | 88.1836 | -30.8121 | -33.5861 | -33.5403 | -0.036621 | 88.3118 | -33.8745 | 88.3621 | -32.3318 | -33.7784 | -33.5724 | -0.11444 | -0.022889 | -0.16479 | -0.14191 | 3.0258 | 3.1403 | 0.29755 | 0.25177 | 0.41199 | 0.36621 | -0.036621 | -0.036621 | 0.013733 | 0.050354 | 1.506 | 1.5427 | 0.2655 | 0.05951 | 0.30212 | 0.096131
1473.1719 | 122.4564 | -33.8654 | 88.2614 | -0.031357 | 0.0068958 | 0.082004 | 0.080027 | 0.00039951 | 0.0042254 | 0.003888 | -0.056244 | 0.0020957 | 0.072397 | 0.067464 | -0.0010361 | -0.00027134 | 0.0032193 | -0.059509 | 88.2431 | -34.0393 | 88.1882 | -30.8029 | -33.609 | -33.5495 | 0.013733 | 88.266 | -33.9523 | 88.3575 | -32.3273 | -33.783 | -33.5724 | -0.17395 | -0.018311 | -0.073243 | -0.054932 | 3.0624 | 3.2364 | 0.31586 | 0.25635 | 0.48981 | 0.4303 | -0.086976 | 0.004577 | 0.09613 | 0.091553 | 1.5381 | 1.6251 | 0.29297 | 0.082397 | 0.37994 | 0.16937
1474.1719 | 122.4426 | -33.8791 | 88.2568 | -0.1458 | 0.0066822 | 0.081576 | 0.078746 | -0.00066869 | 0.0037982 | 0.003888 | -0.0043641 | 0.0025229 | 0.07197 | 0.068532 | 3.2078e-05 | -0.0021939 | -0.00041228 | -0.045776 | 88.2523 | -34.0668 | 88.1012 | -30.8304 | -33.609 | -33.5861 | 0 | 88.266 | -33.9523 | 88.3163 | -32.3364 | -33.8104 | -33.6273 | -0.18768 | -0.004578 | -0.15564 | -0.15106 | 3.0487 | 3.2364 | 0.29297 | 0.27008 | 0.48065 | 0.45776 | -0.073243 | 0.009155 | 0.059509 | 0.050354 | 1.5427 | 1.6159 | 0.25177 | 0.068664 | 0.32501 | 0.14191
1475.1719 | 122.4197 | -33.9386 | 88.2339 | -0.15343 | 0.0060412 | 0.082431 | 0.080668 | -0.0025913 | 0.0031574 | 0.0023926 | -0.040985 | 0.0033773 | 0.073465 | 0.069387 | -0.00060892 | -0.0028347 | -0.0016941 | -0.10529 | 88.2523 | -34.053 | 88.1012 | -30.835 | -33.6456 | -33.609 | -0.045776 | 88.2294 | -34.0027 | 88.2797 | -32.3593 | -33.8562 | -33.6502 | -0.11444 | 0.01831 | -0.13275 | -0.15106 | 3.1036 | 3.2181 | 0.32959 | 0.29297 | 0.44403 | 0.40741 | -0.064087 | -0.004578 | 0.045776 | 0.050354 | 1.5793 | 1.6434 | 0.28839 | 0.082398 | 0.35248 | 0.14648
1476.1875 | 122.3969 | -33.9798 | 88.2339 | -0.15343 | 0.0071094 | 0.080508 | 0.082164 | 0.0012539 | 0.00508 | 0.003888 | -0.086761 | 0.0063681 | 0.074533 | 0.071309 | -0.00082252 | -0.0030483 | 0.0040737 | -0.10529 | 88.2339 | -34.1263 | 88.1424 | -30.6839 | -33.7967 | -33.783 | -0.12817 | 88.2294 | -34.0347 | 88.2156 | -32.4463 | -33.9752 | -33.7326 | -0.14649 | 0 | -0.091553 | -0.091553 | 3.2959 | 3.4424 | 0.19684 | 0.1831 | 0.34332 | 0.32959 | -0.054932 | -0.004578 | -0.018311 | -0.013733 | 1.5335 | 1.5884 | 0.24719 | 0.004577 | 0.30212 | 0.059509
1477.1719 | 122.438 | -33.9752 | 88.1836 | -0.12291 | 0.0077502 | 0.083072 | 0.078959 | -0.0030185 | 0.0048664 | 0.0030334 | -0.077606 | 0.0031637 | 0.07197 | 0.069814 | 0.00045928 | -0.00091214 | 0.0017239 | -0.032043 | 88.1882 | -34.0759 | 88.1104 | -30.7617 | -33.6502 | -33.6136 | -0.24261 | 88.1195 | -34.0714 | 88.2797 | -32.2952 | -33.8745 | -33.6456 | -0.10071 | 0.004577 | -0.073242 | -0.077819 | 3.2135 | 3.3142 | 0.36163 | 0.32501 | 0.46234 | 0.42572 | -0.09613 | -0.064087 | 0.09613 | 0.16022 | 1.68 | 1.7761 | 0.32959 | 0.10071 | 0.42572 | 0.19684
1478.1719 | 122.4197 | -34.0256 | 88.1378 | -0.12291 | 0.010527 | 0.08179 | 0.080241 | -0.0017367 | 0.0020892 | 0.003247 | -0.028778 | 0.0033773 | 0.073465 | 0.070241 | 3.2078e-05 | -0.0026211 | 0.0038601 | -0.086975 | 88.179 | -34.1034 | 88.0188 | -30.8029 | -33.7006 | -33.6594 | -0.050354 | 88.1653 | -34.0759 | 88.1836 | -32.3776 | -33.9157 | -33.7143 | -0.07782 | 0.041199 | -0.11902 | -0.16022 | 3.2227 | 3.3005 | 0.36621 | 0.32501 | 0.44403 | 0.40283 | -0.050354 | 0.027466 | 0.045777 | 0.018311 | 1.6479 | 1.6983 | 0.31128 | 0.10986 | 0.36163 | 0.16022
1479.1719 | 122.4655 | -34.0302 | 88.1241 | -0.1458 | 0.0058276 | 0.083072 | 0.079173 | -0.00024129 | 0.0023028 | 0.0034606 | -0.0013123 | 0.0046591 | 0.070261 | 0.066823 | -0.00060892 | -0.0017667 | 0.0012967 | 0.004578 | 88.1516 | -34.14 | 88.0417 | -30.7114 | -33.6182 | -33.6182 | 0.14191 | 88.1241 | -34.0622 | 88.2294 | -32.2861 | -33.8516 | -33.6777 | -0.10986 | 0.027466 | -0.082397 | -0.10986 | 3.3188 | 3.4286 | 0.41199 | 0.41199 | 0.52185 | 0.52185 | -0.032044 | 0 | 0.10529 | 0.10529 | 1.7441 | 1.7761 | 0.35248 | 0.17853 | 0.38452 | 0.21057
1480.1719 | 122.4152 | -34.1675 | 88.028 | -0.13054 | 0.0086048 | 0.083713 | 0.078105 | -0.00088229 | 0.0048664 | 0.003247 | -0.034882 | 0.0035911 | 0.073038 | 0.064046 | -0.00018152 | -0.0021939 | -0.0014805 | -0.07782 | 88.0051 | -34.2636 | 87.8952 | -30.8395 | -33.7921 | -33.7555 | -0.045776 | 88.0051 | -34.1858 | 88.06 | -32.4875 | -34.053 | -33.8425 | -0.096131 | -0.022888 | -0.13275 | -0.10986 | 3.3279 | 3.4241 | 0.41199 | 0.37537 | 0.50812 | 0.4715 | -0.018311 | -0.022888 | 0.032044 | 0.054932 | 1.68 | 1.6983 | 0.32501 | 0.11444 | 0.34332 | 0.13275
1481.1875 | 122.3419 | -34.1995 | 88.0234 | -0.084763 | 0.009032 | 0.082644 | 0.083445 | 0.00082671 | 0.00508 | 0.003888 | -0.12949 | 0.0046591 | 0.07261 | 0.073873 | -0.0012497 | -0.0017667 | 0.0015103 | -0.045776 | 88.0142 | -34.3369 | 87.7991 | -30.9219 | -33.8928 | -33.8837 | -0.059509 | 88.028 | -34.2636 | 87.973 | -32.6157 | -34.1858 | -33.9615 | -0.13733 | -0.009155 | -0.2243 | -0.21515 | 3.2776 | 3.4149 | 0.31586 | 0.3067 | 0.45319 | 0.44403 | -0.064087 | 0.004578 | -0.050354 | -0.054932 | 1.5839 | 1.6479 | 0.23804 | 0.013733 | 0.30212 | 0.07782
1482.1719 | 122.4701 | -34.2499 | 87.9456 | -0.1992 | 0.0096728 | 0.084567 | 0.078532 | -0.00024129 | 0.0048664 | 0.0043152 | 0.013947 | 0.0040183 | 0.073679 | 0.066396 | 0.00088648 | -0.0024075 | 0.00086952 | -0.059509 | 87.8906 | -34.4238 | 87.7533 | -30.9036 | -33.8654 | -33.8608 | 0 | 87.8952 | -34.3094 | 87.973 | -32.5104 | -34.1171 | -33.9203 | -0.17395 | -0.054932 | -0.19226 | -0.13733 | 3.3463 | 3.5202 | 0.3891 | 0.38452 | 0.56305 | 0.55847 | -0.059509 | -0.050354 | 0.027465 | 0.077819 | 1.7395 | 1.799 | 0.32959 | 0.13275 | 0.3891 | 0.19226
1483.1719 | 122.4152 | -34.2636 | 87.8952 | -0.13817 | 0.0083912 | 0.083499 | 0.080027 | -0.0010959 | 0.0037982 | 0.0026062 | -0.0074159 | 0.0048727 | 0.073892 | 0.06661 | -0.0023177 | -0.0026211 | 0.0012967 | -0.09613 | 87.8998 | -34.3826 | 87.7579 | -30.8899 | -33.8791 | -33.8608 | -0.032043 | 87.8815 | -34.3185 | 87.9181 | -32.5607 | -34.1812 | -33.9706 | -0.11902 | 0.004577 | -0.13733 | -0.14191 | 3.3737 | 3.4927 | 0.40283 | 0.38452 | 0.52185 | 0.50354 | -0.054931 | -0.013733 | 0.022888 | 0.036621 | 1.7029 | 1.7578 | 0.29297 | 0.082398 | 0.3479 | 0.13733
1484.1719 | 122.4655 | -34.2957 | 87.8174 | -0.13054 | 0.0079638 | 0.081363 | 0.080455 | -0.0034457 | 0.0029438 | 0.0028198 | -0.0013123 | 0.0025229 | 0.07261 | 0.06896 | -0.0016769 | -0.0017667 | 0.00086952 | -0.027466 | 87.8448 | -34.433 | 87.7441 | -30.8578 | -33.8745 | -33.8745 | 0.041199 | 87.8357 | -34.3643 | 87.9181 | -32.5195 | -34.1354 | -33.9569 | -0.13733 | 0.027466 | -0.073242 | -0.10071 | 3.4378 | 3.5751 | 0.42114 | 0.42114 | 0.55847 | 0.55847 | -0.068665 | 0.01831 | 0.10071 | 0.082398 | 1.7761 | 1.8448 | 0.33874 | 0.16022 | 0.40741 | 0.22888
1485.1719 | 122.4243 | -34.3781 | 87.7716 | -0.15343 | 0.009032 | 0.084353 | 0.082591 | -0.0013095 | 0.0059344 | 0.0047424 | -0.047089 | 0.0040183 | 0.074319 | 0.069814 | -0.00018152 | -0.00048494 | 0.0042875 | -0.054932 | 87.7533 | -34.4925 | 87.6663 | -30.867 | -33.9661 | -33.9615 | 0.018311 | 87.7579 | -34.433 | 87.8357 | -32.6111 | -34.2773 | -34.053 | -0.11444 | -0.01831 | -0.10529 | -0.086975 | 3.511 | 3.6255 | 0.41656 | 0.41199 | 0.53101 | 0.52643 | -0.054931 | -0.013732 | 0.064087 | 0.077819 | 1.767 | 1.8219 | 0.32501 | 0.10071 | 0.37994 | 0.15564
1486.1719 | 122.5067 | -34.4284 | 87.7487 | -0.12291 | 0.009032 | 0.083499 | 0.0796 | -0.00024129 | 0.00508 | 0.0081604 | -0.0013123 | 0.0042319 | 0.071542 | 0.071523 | -0.0012497 | -0.00091214 | 0.00065592 | 0.013733 | 87.7441 | -34.5474 | 87.6389 | -30.8258 | -33.9386 | -33.9615 | 0.14191 | 87.6984 | -34.4696 | 87.822 | -32.5974 | -34.2682 | -34.0851 | -0.11902 | -0.004577 | -0.10986 | -0.10529 | 3.6026 | 3.7216 | 0.46692 | 0.48981 | 0.58594 | 0.60882 | -0.041198 | -0.050354 | 0.073242 | 0.1236 | 1.8311 | 1.8723 | 0.34332 | 0.16022 | 0.38452 | 0.20142
1487.1719 | 122.4426 | -34.4833 | 87.6572 | -0.1458 | 0.010314 | 0.084353 | 0.080882 | -0.0068637 | 0.0042254 | 0.0041016 | -0.059296 | 0.0035911 | 0.076456 | 0.069814 | 0.00045928 | -0.0017667 | 0.0030057 | -0.064087 | 87.6343 | -34.6161 | 87.5885 | -30.8899 | -34.0439 | -34.0347 | 0.022888 | 87.6526 | -34.52 | 87.7396 | -32.6477 | -34.3506 | -34.1354 | -0.13275 | -0.022889 | -0.068665 | -0.045776 | 3.5934 | 3.7262 | 0.44861 | 0.43945 | 0.58136 | 0.57221 | -0.036621 | -0.004578 | 0.082397 | 0.086975 | 1.8356 | 1.8723 | 0.3479 | 0.13275 | 0.38452 | 0.16937
1488.1719 | 122.4472 | -34.5337 | 87.6068 | -0.092392 | 0.0086048 | 0.083499 | 0.08195 | 0.00082671 | 0.0044392 | 0.0030334 | -0.019623 | 0.0035911 | 0.075815 | 0.06896 | -0.0014633 | -0.0034757 | 0.0012967 | -0.07782 | 87.6434 | -34.639 | 87.5015 | -30.9219 | -34.0759 | -34.0805 | -0.041199 | 87.6343 | -34.5749 | 87.6663 | -32.7209 | -34.4147 | -34.2041 | -0.10529 | 0.036621 | -0.10529 | -0.14191 | 3.6118 | 3.717 | 0.45319 | 0.45776 | 0.55847 | 0.56305 | -0.041199 | 0.027465 | 0.059509 | 0.032044 | 1.8127 | 1.8539 | 0.32959 | 0.11902 | 0.37079 | 0.16022
1489.1719 | 122.4106 | -34.6115 | 87.5565 | -0.12291 | 0.0094592 | 0.082858 | 0.08195 | -0.0028049 | 0.0044392 | 0.0030334 | -0.0074159 | 0.0055137 | 0.07261 | 0.067891 | -0.0018905 | -0.0021939 | 0.0010831 | -0.07782 | 87.616 | -34.6939 | 87.4329 | -30.9631 | -34.0988 | -34.1171 | -0.050354 | 87.5702 | -34.6161 | 87.4649 | -32.7667 | -34.5383 | -34.2361 | -0.082398 | 0.059509 | -0.1236 | -0.18311 | 3.6484 | 3.7308 | 0.49439 | 0.5127 | 0.57678 | 0.59509 | -0.004578 | 0.013732 | -0.091553 | -0.10528 | 1.8448 | 1.8494 | 0.37537 | 0.073242 | 0.37994 | 0.07782
1490.1719 | 122.4655 | -34.6252 | 87.4969 | -0.11528 | 0.0096728 | 0.081149 | 0.08195 | -0.00066869 | 0.0048664 | 0.0053832 | -0.0074159 | 0.0038047 | 0.075388 | 0.070241 | 0.00067288 | -0.0039029 | -0.0010531 | -0.009155 | 87.5381 | -34.7443 | 87.4191 | -30.9219 | -34.1034 | -34.1354 | 0.09613 | 87.5015 | -34.6985 | 87.6068 | -32.7026 | -34.4467 | -34.259 | -0.11902 | 0.041199 | -0.07782 | -0.11902 | 3.7033 | 3.8223 | 0.48981 | 0.52185 | 0.60883 | 0.64087 | -0.073242 | 0.004578 | 0.10986 | 0.10529 | 1.9226 | 1.9958 | 0.36621 | 0.17853 | 0.43945 | 0.25177
1491.1719 | 122.4518 | -34.6939 | 87.4695 | -0.15343 | 0.0086048 | 0.082644 | 0.081736 | 0.0029629 | 0.00508 | 0.0047424 | -0.040985 | 0.0020957 | 0.074106 | 0.070669 | -0.0040267 | -0.0039029 | 0.0042875 | 0.059509 | 87.4557 | -34.8175 | 87.3642 | -30.9357 | -34.14 | -34.2041 | 0.10071 | 87.4695 | -34.726 | 87.5381 | -32.7621 | -34.5337 | -34.346 | -0.1236 | -0.013732 | -0.10528 | -0.091553 | 3.7582 | 3.8818 | 0.48981 | 0.55389 | 0.6134 | 0.67749 | -0.032043 | 0 | 0.068665 | 0.068665 | 1.9318 | 1.9638 | 0.3479 | 0.16022 | 0.37994 | 0.19226
1492.1719 | 122.4518 | -34.7122 | 87.4191 | -0.10002 | 0.009032 | 0.083285 | 0.081309 | -0.0023777 | 0.0044392 | 0.0030334 | -0.016571 | 0.0048727 | 0.074106 | 0.066396 | 0.00024568 | -0.0019803 | 0.0053555 | -0.091553 | 87.442 | -34.8541 | 87.2772 | -31.0089 | -34.2453 | -34.2636 | -0.041199 | 87.4191 | -34.8129 | 87.4557 | -32.8171 | -34.5978 | -34.4009 | -0.14191 | 0.022889 | -0.14191 | -0.1648 | 3.7033 | 3.8452 | 0.44861 | 0.46692 | 0.59052 | 0.60883 | -0.10071 | 0 | 0.036622 | 0.036622 | 1.8951 | 1.9958 | 0.31128 | 0.11444 | 0.41199 | 0.21515
1493.1719 | 122.4564 | -34.758 | 87.3962 | -0.10765 | 0.0101 | 0.086703 | 0.082164 | -2.7692e-05 | 0.006148 | 0.0047424 | -0.037933 | 0.0053001 | 0.074533 | 0.068105 | 3.2078e-05 | -0.0019803 | -0.0019077 | -0.045776 | 87.3962 | -34.8953 | 87.291 | -30.9494 | -34.2453 | -34.2682 | 0.013733 | 87.3871 | -34.7992 | 87.4695 | -32.8079 | -34.6115 | -34.3918 | -0.13733 | 0 | -0.10529 | -0.10529 | 3.8086 | 3.9459 | 0.48981 | 0.5127 | 0.62714 | 0.65002 | -0.041198 | -0.009155 | 0.073242 | 0.082397 | 1.9501 | 1.9913 | 0.36621 | 0.14648 | 0.40741 | 0.18768
1494.1875 | 122.3877 | -34.8267 | 87.3825 | 0.098343 | 0.0096728 | 0.08414 | 0.085368 | 0.00018591 | 0.0018756 | 0.0019654 | -0.086761 | 0.0046591 | 0.074319 | 0.068746 | -0.00039512 | -0.00027134 | -0.0019077 | -0.091553 | 87.4466 | -34.9274 | 87.1628 | -31.073 | -34.3277 | -34.3735 | -0.022888 | 87.4237 | -34.877 | 87.3413 | -32.9086 | -34.7214 | -34.5016 | -0.10071 | 0.064087 | -0.21973 | -0.28381 | 3.7537 | 3.8544 | 0.45319 | 0.49896 | 0.55389 | 0.59967 | -0.050354 | 0.041199 | -0.041198 | -0.082397 | 1.918 | 1.9684 | 0.32501 | 0.10528 | 0.37537 | 0.15564
1495.1719 | 122.4152 | -34.8312 | 87.291 | -0.12291 | 0.0098864 | 0.083285 | 0.08195 | -0.00088229 | 0.0037982 | 0.0064514 | -0.040985 | 0.0050863 | 0.07496 | 0.067678 | -0.0010361 | -0.0011257 | 0.0038601 | -0.036621 | 87.291 | -34.9686 | 87.1994 | -30.9402 | -34.2911 | -34.3277 | 0.022888 | 87.2818 | -34.8999 | 87.3825 | -32.8491 | -34.6664 | -34.4604 | -0.13733 | 0 | -0.091553 | -0.091553 | 3.891 | 4.0283 | 0.50354 | 0.54016 | 0.64087 | 0.67749 | -0.068664 | -0.009156 | 0.091552 | 0.10071 | 1.9821 | 2.0508 | 0.37079 | 0.16479 | 0.43945 | 0.23346
1496.1719 | 122.3877 | -34.9274 | 87.3001 | -0.054245 | 0.011382 | 0.082217 | 0.084941 | -0.0017367 | 0.003371 | 0.0030334 | -0.092865 | 0.0055137 | 0.075388 | 0.076009 | -0.0010361 | -0.0032619 | 0.0025785 | -0.022888 | 87.3276 | -35.0372 | 87.085 | -31.073 | -34.3597 | -34.4467 | -0.018311 | 87.291 | -34.964 | 87.2589 | -32.9498 | -34.7992 | -34.5749 | -0.10986 | 0.027466 | -0.21515 | -0.24261 | 3.8544 | 3.9642 | 0.48065 | 0.56763 | 0.59051 | 0.67749 | -0.036621 | -0.009155 | -0.041199 | -0.032044 | 1.9775 | 2.0142 | 0.35248 | 0.12817 | 0.3891 | 0.16479
1497.1719 | 122.4243 | -34.8953 | 87.2223 | -0.16106 | 0.013518 | 0.08649 | 0.0843 | -0.00088229 | 0.0042254 | 0.0041016 | -0.044037 | 0.0046591 | 0.075815 | 0.073446 | -0.0012497 | -0.0024075 | 0.0012967 | -0.22888 | 87.1399 | -35.083 | 87.1445 | -30.9494 | -34.3094 | -34.3552 | 0.022888 | 87.2406 | -34.9731 | 87.3184 | -32.872 | -34.7443 | -34.5108 | -0.18768 | -0.082397 | -0.07782 | 0.004577 | 3.9459 | 4.1336 | 0.54016 | 0.58594 | 0.72784 | 0.77362 | -0.07782 | 0.018311 | 0.09613 | 0.077819 | 2.0233 | 2.1011 | 0.38452 | 0.15106 | 0.46234 | 0.22888
1498.1719 | 122.4243 | -34.9319 | 87.1948 | -0.19157 | 0.010955 | 0.083499 | 0.082591 | -0.0030185 | 0.0052936 | 0.0028198 | 0.016998 | 0.0050863 | 0.075174 | 0.071523 | 0.0011001 | -0.0019803 | 0.0027921 | -0.036621 | 87.2086 | -35.083 | 87.0621 | -30.986 | -34.346 | -34.3689 | -0.022888 | 87.1948 | -34.9915 | 87.2726 | -32.8857 | -34.7534 | -34.5337 | -0.15106 | 0.013733 | -0.13275 | -0.14648 | 3.9459 | 4.097 | 0.56305 | 0.58594 | 0.71411 | 0.737 | -0.059509 | 0 | 0.07782 | 0.07782 | 2.0462 | 2.1057 | 0.39826 | 0.17853 | 0.45776 | 0.23804
1499.1719 | 122.5937 | -34.9869 | 87.1902 | -0.10765 | 0.010741 | 0.083713 | 0.079814 | -0.0028049 | 0.0044392 | 0.0064514 | 0.02005 | 0.0033773 | 0.073465 | 0.069814 | 0.00045928 | -0.0019803 | 0.0021511 | 0.022888 | 87.1582 | -35.0967 | 87.0987 | -30.8853 | -34.2636 | -34.3323 | 0.11444 | 87.1536 | -35.0098 | 87.2589 | -32.8812 | -34.7351 | -34.6481 | -0.10986 | -0.032044 | -0.091553 | -0.059509 | 4.1016 | 4.2114 | 0.6546 | 0.72327 | 0.76447 | 0.83313 | -0.022889 | -0.036622 | 0.068664 | 0.10529 | 2.1057 | 2.1286 | 0.33875 | 0.25177 | 0.36163 | 0.27466
1500.1719 | 122.406 | -34.9915 | 87.1857 | -0.12291 | 0.010741 | 0.082004 | 0.083659 | -0.0025913 | 0.0059344 | 0.0047424 | -0.03183 | 0.0040183 | 0.076028 | 0.073232 | -0.00082252 | -0.0043301 | 0.0038601 | -0.082397 | 87.1445 | -35.1334 | 87.0209 | -31.0043 | -34.4009 | -34.4467 | -0.036621 | 87.1399 | -35.0601 | 87.1719 | -32.959 | -34.845 | -34.6207 | -0.14191 | -0.041199 | -0.16479 | -0.1236 | 3.9871 | 4.129 | 0.54474 | 0.59052 | 0.68665 | 0.73242 | -0.068665 | -0.045776 | -0.013733 | 0.032043 | 2.0325 | 2.1011 | 0.37079 | 0.14648 | 0.43945 | 0.21515
1501.1719 | 122.4701 | -35.0281 | 87.1353 | -0.16869 | 0.0098864 | 0.081363 | 0.084086 | -0.0038729 | 0.0055072 | 0.0045288 | -0.022675 | 0.0035911 | 0.076242 | 0.072164 | -0.0014633 | -0.0028347 | 0.00086952 | 0 | 87.1124 | -35.1379 | 87.0392 | -30.8716 | -34.3048 | -34.3918 | 0.15106 | 87.117 | -35.0693 | 87.2269 | -32.8949 | -34.79 | -34.5932 | -0.10986 | -0.022888 | -0.09613 | -0.073242 | 4.1565 | 4.2664 | 0.63629 | 0.72327 | 0.74615 | 0.83313 | -0.041199 | -0.018311 | 0.091553 | 0.10986 | 2.1332 | 2.1744 | 0.43487 | 0.23804 | 0.47607 | 0.27924
1502.1719 | 122.4289 | -35.051 | 87.085 | -0.13817 | 0.012877 | 0.084353 | 0.083873 | -0.00088229 | 0.0046528 | 0.0041016 | -0.013519 | 0.0038047 | 0.074319 | 0.065755 | -0.0010361 | -0.0019803 | 0.0025785 | -0.036621 | 87.1078 | -35.2066 | 86.9705 | -30.9494 | -34.3781 | -34.4421 | 0.032043 | 87.085 | -35.1196 | 87.1399 | -32.9498 | -34.8404 | -34.6252 | -0.15564 | 0.022888 | -0.11444 | -0.13733 | 4.1016 | 4.2572 | 0.60883 | 0.67291 | 0.76447 | 0.82855 | -0.068665 | 0 | 0.054932 | 0.054932 | 2.1011 | 2.1698 | 0.42572 | 0.21057 | 0.49439 | 0.27924
1503.1719 | 122.4152 | -35.1242 | 87.0392 | -0.13817 | 0.011382 | 0.084567 | 0.084727 | -0.0010959 | 0.006148 | 0.0053832 | -0.025726 | 0.0057273 | 0.076242 | 0.066823 | -0.0040267 | 0.00015586 | 0.0019375 | -0.068665 | 87.0392 | -35.2386 | 86.9156 | -30.9906 | -34.4376 | -34.5154 | -0.036621 | 87.0483 | -35.1562 | 87.0895 | -33.0002 | -34.9274 | -34.7122 | -0.11444 | 0 | -0.1236 | -0.1236 | 4.1336 | 4.248 | 0.60883 | 0.68665 | 0.72327 | 0.80109 | -0.032043 | 0.009155 | 0.050354 | 0.041199 | 2.124 | 2.1561 | 0.41199 | 0.19684 | 0.44403 | 0.22888
1504.1719 | 122.4426 | -35.1379 | 87.0712 | -0.084763 | 0.01245 | 0.083926 | 0.085368 | -0.0021641 | 0.0048664 | 0.0053832 | -0.034882 | 0.0046591 | 0.073465 | 0.066182 | -0.00018152 | -0.0056119 | 0.0047147 | 0.041199 | 87.0392 | -35.2478 | 86.9476 | -30.9128 | -34.3826 | -34.4742 | 0.10986 | 87.0346 | -35.1654 | 87.0758 | -32.9773 | -34.9045 | -34.6848 | -0.10986 | -0.032043 | -0.1236 | -0.091553 | 4.2252 | 4.335 | 0.66376 | 0.75531 | 0.77362 | 0.86517 | -0.027466 | -0.036621 | 0.004578 | 0.041199 | 2.1606 | 2.1881 | 0.45319 | 0.23346 | 0.48065 | 0.26093
1505.1719 | 122.4472 | -35.1334 | 86.998 | -0.18394 | 0.013091 | 0.085849 | 0.081736 | -0.0013095 | 0.0040118 | 0.0068786 | -0.080658 | 0.0048727 | 0.075815 | 0.075368 | -0.0018905 | -0.0015529 | 0.0036465 | -0.041199 | 86.9751 | -35.2936 | 86.8698 | -30.9357 | -34.3689 | -34.5062 | 0.032043 | 86.9934 | -35.1929 | 87.0941 | -32.9819 | -34.9274 | -34.671 | -0.16022 | -0.022888 | -0.12817 | -0.10529 | 4.1977 | 4.3579 | 0.62714 | 0.76447 | 0.78735 | 0.92468 | -0.059509 | -0.004578 | 0.09613 | 0.10071 | 2.1515 | 2.211 | 0.46234 | 0.20599 | 0.52185 | 0.2655
1506.1719 | 122.4197 | -35.17 | 87.0026 | -0.13054 | 0.011595 | 0.082858 | 0.084086 | -0.0023777 | 0.0044392 | 0.004956 | -0.010467 | 0.0038047 | 0.076242 | 0.070241 | 0.00024568 | -0.0021939 | 0.0045011 | 0.073242 | 86.9888 | -35.3119 | 86.8835 | -30.867 | -34.4101 | -34.5154 | 0.12817 | 86.9431 | -35.2203 | 87.0483 | -32.9681 | -34.9594 | -34.7214 | -0.14191 | -0.013732 | -0.11902 | -0.10529 | 4.303 | 4.4449 | 0.6546 | 0.75989 | 0.79651 | 0.90179 | -0.050354 | -0.059509 | 0.045777 | 0.10529 | 2.2018 | 2.2522 | 0.44861 | 0.21057 | 0.49896 | 0.26093
1507.1719 | 122.4152 | -35.2158 | 86.9202 | -0.069504 | 0.014159 | 0.083285 | 0.085368 | -0.0013095 | 0.0055072 | 0.0053832 | -0.022675 | 0.0057273 | 0.076242 | 0.070882 | 0.0015275 | -0.0019803 | 0.0023647 | -0.059509 | 86.9659 | -35.3394 | 86.7966 | -30.9402 | -34.4971 | -34.6024 | -0.068665 | 86.9247 | -35.289 | 86.9476 | -33.0688 | -35.0464 | -34.8083 | -0.1236 | 0.045776 | -0.1236 | -0.16937 | 4.2755 | 4.3991 | 0.6134 | 0.71869 | 0.737 | 0.84228 | -0.073242 | 0.004578 | 0.027466 | 0.022888 | 2.1469 | 2.2202 | 0.40741 | 0.16937 | 0.48065 | 0.24261
1508.1719 | 122.406 | -35.2203 | 86.8927 | -0.10765 | 0.013091 | 0.08649 | 0.087077 | -0.0021641 | 0.0089252 | 0.0079468 | -0.068451 | 0.0040183 | 0.074747 | 0.070241 | -0.0010361 | 0.00015586 | 0.0015103 | -0.036621 | 86.8469 | -35.376 | 86.8103 | -30.8945 | -34.4971 | -34.584 | 0.013733 | 86.9064 | -35.2982 | 86.9797 | -33.0368 | -35.0235 | -34.7855 | -0.15564 | -0.045776 | -0.082397 | -0.036621 | 4.3259 | 4.4815 | 0.63629 | 0.72327 | 0.79193 | 0.87891 | -0.07782 | 0.013733 | 0.086975 | 0.073242 | 2.1835 | 2.2614 | 0.43488 | 0.19684 | 0.5127 | 0.27466
1509.1719 | 122.4747 | -35.2203 | 86.8927 | -0.10765 | 0.012023 | 0.08414 | 0.085368 | -0.0019503 | 0.0042254 | 0.0030334 | -0.0043641 | 0.0067953 | 0.07731 | 0.075155 | -0.00039512 | -0.0041165 | 0.0034329 | -0.018311 | 86.9293 | -35.3577 | 86.8149 | -30.8258 | -34.4238 | -34.5383 | 0.09613 | 86.911 | -35.2844 | 87.0071 | -32.9544 | -34.964 | -34.7397 | -0.13733 | 0.036621 | -0.07782 | -0.11444 | 4.3945 | 4.5319 | 0.68207 | 0.79651 | 0.8194 | 0.93384 | -0.064087 | 0.018311 | 0.11444 | 0.09613 | 2.2659 | 2.33 | 0.48065 | 0.25635 | 0.54474 | 0.32044
1510.1719 | 122.4426 | -35.2432 | 86.9019 | -0.15343 | 0.014159 | 0.084781 | 0.083445 | 0.0023221 | 0.0097796 | 0.0051696 | 0.0047913 | 0.0042319 | 0.076883 | 0.069814 | -0.0014633 | -0.0045437 | 0.0034329 | -0.07782 | 86.9339 | -35.3577 | 86.7279 | -30.8807 | -34.5154 | -34.6069 | 0.004578 | 86.8698 | -35.3073 | 86.9431 | -33.0322 | -35.0372 | -34.7946 | -0.11444 | 0.032044 | -0.17395 | -0.20599 | 4.3625 | 4.4769 | 0.63629 | 0.72784 | 0.75073 | 0.84229 | -0.064087 | -0.032043 | 0.041199 | 0.073242 | 2.211 | 2.2751 | 0.44861 | 0.20599 | 0.5127 | 0.27008
1511.1719 | 122.4564 | -35.2341 | 86.9064 | -0.11528 | 0.011382 | 0.08649 | 0.082804 | -0.0025913 | 0.0072162 | 0.0064514 | -0.059296 | 0.0040183 | 0.076883 | 0.072805 | -0.0018905 | -0.0028347 | 0.0015103 | 0.022888 | 86.911 | -35.3806 | 86.673 | -30.8395 | -34.2682 | -34.5428 | 0.13733 | 86.8607 | -35.2936 | 86.9705 | -32.9727 | -34.996 | -34.7351 | -0.14648 | 0.004578 | -0.23346 | -0.23804 | 4.3945 | 4.541 | 0.69122 | 0.96588 | 0.83771 | 1.1124 | -0.059509 | -0.045776 | 0.064087 | 0.10986 | 2.2614 | 2.3209 | 0.49896 | 0.23804 | 0.55847 | 0.29755
1512.1719 | 122.4335 | -35.2661 | 86.8835 | -0.092392 | 0.012877 | 0.083499 | 0.0843 | 0.00039951 | 0.0089252 | 0.0041016 | -0.053192 | 0.0055137 | 0.076883 | 0.06661 | -0.0016769 | -0.0026211 | 0.0047147 | -0.027466 | 86.8286 | -35.3943 | 86.8011 | -30.7571 | -34.465 | -34.5383 | 0.027466 | 86.8561 | -35.3165 | 86.9843 | -32.991 | -35.0372 | -34.758 | -0.12817 | -0.054932 | -0.082398 | -0.027466 | 4.509 | 4.6371 | 0.72784 | 0.80109 | 0.85602 | 0.92926 | -0.050354 | -0.027466 | 0.10071 | 0.12817 | 2.2751 | 2.3254 | 0.50812 | 0.22888 | 0.55847 | 0.27924
1513.1719 | 122.4518 | -35.1791 | 86.9385 | -0.092392 | 0.013304 | 0.084994 | 0.087291 | -0.0010959 | 0.010207 | 0.0053832 | -0.077606 | 0.0040183 | 0.076883 | 0.068746 | 0.0013139 | -0.0026211 | 0.0055691 | -0.045776 | 86.9202 | -35.3394 | 86.7737 | -30.5603 | -34.2819 | -34.552 | 0.041199 | 86.8835 | -35.289 | 87.0163 | -32.9727 | -34.9731 | -34.7122 | -0.16022 | -0.018311 | -0.1648 | -0.14648 | 4.6188 | 4.7791 | 0.62714 | 0.89722 | 0.78735 | 1.0574 | -0.10986 | -0.054932 | 0.077819 | 0.13275 | 2.2064 | 2.3163 | 0.46692 | 0.20599 | 0.57678 | 0.31586
1514.1719 | 122.438 | -35.2203 | 86.9293 | -0.069504 | 0.015013 | 0.083713 | 0.086009 | 0.00018591 | 0.0048664 | 0.002179 | 0.0017395 | 0.0059409 | 0.076028 | 0.073873 | -0.0010361 | -0.0019803 | 0.0012967 | 0.027466 | 86.9431 | -35.3394 | 86.911 | -30.629 | -34.4284 | -34.4604 | 0.10529 | 86.9522 | -35.2661 | 87.0255 | -32.959 | -34.9365 | -34.6939 | -0.11902 | 0.013733 | -0.01831 | -0.032043 | 4.5914 | 4.7104 | 0.75989 | 0.79193 | 0.87891 | 0.91095 | -0.045776 | 0.022888 | 0.096131 | 0.073243 | 2.2614 | 2.3071 | 0.52643 | 0.28381 | 0.5722 | 0.32959
1515.1719 | 122.4289 | -35.2386 | 86.8744 | -0.092392 | 0.014159 | 0.084994 | 0.087077 | 0.0029629 | 0.0072162 | 0.0068786 | -0.077606 | 0.0057273 | 0.07731 | 0.075582 | -0.00018152 | -0.0019803 | 0.0027921 | -0.068665 | 86.8607 | -35.3714 | 86.792 | -30.6335 | -34.4009 | -34.52 | 0.013733 | 86.8835 | -35.3165 | 86.9797 | -32.9819 | -34.996 | -34.7397 | -0.13275 | -0.013733 | -0.082398 | -0.068665 | 4.6051 | 4.7379 | 0.71869 | 0.83771 | 0.85144 | 0.97046 | -0.07782 | 0.009155 | 0.10528 | 0.09613 | 2.2568 | 2.3346 | 0.49896 | 0.24261 | 0.57678 | 0.32043
1516.1719 | 122.3648 | -35.3027 | 86.8423 | -0.16869 | 0.014159 | 0.08414 | 0.084513 | -0.00045509 | 0.0089252 | 0.0062378 | 0.013947 | 0.0057273 | 0.078165 | 0.065114 | -0.0031723 | -0.0034757 | 0.0010831 | 0.027466 | 86.8927 | -35.3714 | 86.8286 | -30.5511 | -34.3277 | -34.4879 | 0.1236 | 86.879 | -35.3302 | 86.9614 | -33.0597 | -34.7946 | -34.6664 | -0.068665 | 0.050354 | -0.013733 | -0.064087 | 4.7516 | 4.8203 | 0.81482 | 0.97504 | 0.88348 | 1.0437 | -0.027466 | 0.036621 | 0.11902 | 0.082398 | 2.243 | 2.2705 | 0.63629 | 0.50812 | 0.66376 | 0.53558
1517.1719 | 122.4426 | -35.321 | 86.824 | -0.11528 | 0.014373 | 0.085208 | 0.087931 | -0.0015231 | 0.0076434 | 0.0070922 | -0.10812 | 0.0065817 | 0.074533 | 0.072591 | 0.0011001 | -0.00091214 | 0.00044212 | -0.045776 | 86.8286 | -35.4309 | 86.7508 | -30.6335 | -34.4101 | -34.5566 | 0.013733 | 86.8515 | -35.3394 | 86.9247 | -32.8949 | -35.0967 | -34.8083 | -0.10986 | 0.004577 | -0.073243 | -0.07782 | 4.6875 | 4.7974 | 0.76447 | 0.91095 | 0.87433 | 1.0208 | -0.01831 | 0.027465 | 0.10071 | 0.073243 | 2.4261 | 2.4445 | 0.5127 | 0.2243 | 0.53101 | 0.24261
1518.1719 | 122.4701 | -35.2615 | 86.8515 | -0.10002 | 0.016936 | 0.087771 | 0.085368 | -0.0015231 | 0.0046528 | 0.004956 | -0.092865 | 0.0053001 | 0.077951 | 0.076009 | 0.00088648 | -5.7742e-05 | 0.0045011 | -0.059509 | 86.8607 | -35.4263 | 86.7599 | -30.5878 | -34.3826 | -34.5383 | -0.004578 | 86.8423 | -35.3348 | 86.9705 | -32.9819 | -35.0327 | -34.7397 | -0.16479 | 0.009156 | -0.091552 | -0.10071 | 4.6738 | 4.8386 | 0.72327 | 0.87891 | 0.88806 | 1.0437 | -0.073242 | -0.009155 | 0.11902 | 0.12817 | 2.2797 | 2.3529 | 0.52185 | 0.22888 | 0.59509 | 0.30212
1519.1719 | 122.4472 | -35.3165 | 86.8103 | -0.11528 | 0.012877 | 0.086703 | 0.085368 | -0.0023777 | 0.0076434 | 0.0062378 | -0.013519 | 0.0055137 | 0.076883 | 0.067037 | 0.00067288 | -0.0011257 | 0.0012967 | -0.059509 | 86.8561 | -35.463 | 86.6043 | -30.6427 | -34.433 | -34.5978 | 0.009155 | 86.7966 | -35.4218 | 86.8515 | -33.0322 | -35.0922 | -34.8267 | -0.14648 | 0.045776 | -0.20599 | -0.25177 | 4.6738 | 4.8203 | 0.71869 | 0.88348 | 0.86517 | 1.03 | -0.10529 | -0.013733 | 0.041198 | 0.054931 | 2.2842 | 2.3895 | 0.48981 | 0.2243 | 0.59509 | 0.32959
1520.1719 | 122.4564 | -35.3394 | 86.7142 | -0.16869 | 0.014586 | 0.086276 | 0.086222 | 0.00082671 | 0.0082844 | 0.0075196 | -0.062347 | 0.0059409 | 0.077951 | 0.073659 | -0.00018152 | -0.00091214 | 0.0021511 | -0.050354 | 86.7508 | -35.5087 | 86.6684 | -30.5832 | -34.4604 | -34.5932 | 0.004578 | 86.7416 | -35.4263 | 86.8378 | -33.0276 | -35.1334 | -34.8267 | -0.16937 | 0.036621 | -0.045776 | -0.082397 | 4.7562 | 4.9255 | 0.74615 | 0.87891 | 0.91553 | 1.0483 | -0.086976 | 0.027466 | 0.1236 | 0.096131 | 2.3117 | 2.3987 | 0.5127 | 0.20599 | 0.59967 | 0.29297
1521.1719 | 122.4243 | -35.4401 | 86.6959 | -0.16869 | 0.013732 | 0.082644 | 0.087291 | -0.0010959 | 0.0091388 | 0.0068786 | -0.0074159 | 0.0057273 | 0.078592 | 0.073018 | -0.0021041 | -0.0030483 | 0.0015103 | -0.09613 | 86.7325 | -35.5865 | 86.5402 | -30.661 | -34.52 | -34.6664 | -0.036621 | 86.6959 | -35.495 | 86.7142 | -33.1055 | -35.2112 | -34.9274 | -0.14648 | 0.036621 | -0.15564 | -0.19226 | 4.7791 | 4.9255 | 0.77362 | 0.9201 | 0.9201 | 1.0666 | -0.054932 | 0 | 0.01831 | 0.01831 | 2.3346 | 2.3895 | 0.5127 | 0.22888 | 0.56763 | 0.28381
1522.1719 | 122.4197 | -35.4996 | 86.5677 | -0.15343 | 0.013091 | 0.085422 | 0.087931 | -0.0010959 | 0.0082844 | 0.006665 | -0.062347 | 0.0042319 | 0.077097 | 0.074727 | -0.0014633 | -0.0013393 | 0.0066373 | -0.032043 | 86.6821 | -35.5637 | 86.586 | -30.5923 | -34.4925 | -34.6619 | -0.009155 | 86.6776 | -35.5087 | 86.7737 | -33.0597 | -35.1791 | -34.8358 | -0.064087 | 0.11444 | 0.018311 | -0.09613 | 4.9072 | 4.9713 | 0.83771 | 1.0071 | 0.90179 | 1.0712 | -0.009155 | 0.10986 | 0.20599 | 0.096131 | 2.4399 | 2.449 | 0.66376 | 0.32043 | 0.67291 | 0.32959
1523.1719 | 122.3694 | -35.5042 | 86.6592 | -0.1458 | 0.011168 | 0.083926 | 0.090281 | 0.0016813 | 0.0099934 | 0.0081604 | -0.095917 | 0.0085043 | 0.079874 | 0.0743 | 0.00024568 | -0.00048494 | 0.0038601 | -0.11902 | 86.673 | -35.6598 | 86.4578 | -30.7114 | -34.6115 | -34.7763 | -0.10529 | 86.6638 | -35.5774 | 86.6867 | -33.1604 | -35.3119 | -34.9823 | -0.15564 | 0.013733 | -0.20142 | -0.21515 | 4.7928 | 4.9484 | 0.72784 | 0.89264 | 0.88348 | 1.0483 | -0.073243 | 0.004577 | 0.027466 | 0.022889 | 2.3438 | 2.417 | 0.52185 | 0.19226 | 0.59509 | 0.2655
1524.1719 | 122.4243 | -35.5087 | 86.6455 | -0.1458 | 0.015654 | 0.085635 | 0.087504 | 0.00018591 | 0.0089252 | 0.007306 | 0.0017395 | 0.0048727 | 0.076242 | 0.068105 | 3.2078e-05 | -0.0013393 | 0.0021511 | -0.073242 | 86.6638 | -35.6323 | 86.5036 | -30.5923 | -34.5474 | -34.7031 | -0.041199 | 86.6318 | -35.5774 | 86.6684 | -33.1146 | -35.2478 | -34.964 | -0.1236 | 0.01831 | -0.14191 | -0.16022 | 4.9164 | 5.04 | 0.80566 | 0.9613 | 0.92926 | 1.0849 | -0.068665 | -0.013733 | 0.022888 | 0.036621 | 2.3941 | 2.4628 | 0.54474 | 0.26093 | 0.6134 | 0.32959
1525.1719 | 122.3053 | -35.5682 | 86.6089 | -0.12291 | 0.011595 | 0.086062 | 0.088145 | 0.00039951 | 0.0063618 | 0.0058106 | -0.11118 | 0.0070089 | 0.076883 | 0.073232 | -0.00018152 | -5.7742e-05 | 0.0040737 | -0.068665 | 86.6272 | -35.7147 | 86.3892 | -30.6839 | -34.6436 | -34.8404 | -0.1236 | 86.6089 | -35.6277 | 86.5814 | -33.2199 | -35.3989 | -35.083 | -0.14648 | 0.01831 | -0.21973 | -0.23804 | 4.8843 | 5.0308 | 0.72784 | 0.92468 | 0.87433 | 1.0712 | -0.05951 | 0 | -0.027466 | -0.027466 | 2.3483 | 2.4078 | 0.48523 | 0.16937 | 0.54474 | 0.22888
1526.1719 | 122.4289 | -35.5408 | 86.618 | -0.12291 | 0.01715 | 0.085422 | 0.087077 | -0.0021641 | 0.0076434 | 0.006665 | -0.016571 | 0.0059409 | 0.078165 | 0.073018 | 0.0013139 | -0.0015529 | 0.0025785 | -0.054932 | 86.5906 | -35.6918 | 86.5082 | -30.5008 | -34.5062 | -34.6802 | 0.022888 | 86.6043 | -35.5911 | 86.6684 | -33.0414 | -35.2295 | -34.9503 | -0.15106 | -0.027466 | -0.10986 | -0.082397 | 5.04 | 5.191 | 0.86059 | 1.0345 | 1.0117 | 1.1856 | -0.050354 | -0.013733 | 0.050354 | 0.064087 | 2.4994 | 2.5497 | 0.59051 | 0.31128 | 0.64087 | 0.36163
1527.1719 | 122.3282 | -35.5957 | 86.6089 | -0.12291 | 0.013945 | 0.085849 | 0.088145 | 0.0016813 | 0.006789 | 0.0075196 | -0.13254 | 0.0091453 | 0.079019 | 0.073873 | 3.2078e-05 | -0.00069854 | 0.0042875 | -0.09613 | 86.6135 | -35.7468 | 86.3617 | -30.6427 | -34.6527 | -34.8221 | -0.27466 | 86.618 | -35.6369 | 86.5952 | -33.2062 | -35.3897 | -35.0647 | -0.15106 | 0.004577 | -0.24719 | -0.25177 | 4.953 | 5.1041 | 0.77362 | 0.94299 | 0.92468 | 1.0941 | -0.041199 | 0.009155 | -0.013733 | -0.022888 | 2.3895 | 2.4307 | 0.53101 | 0.20599 | 0.5722 | 0.24719
1528.1719 | 122.3419 | -35.6094 | 86.5723 | -0.21446 | 0.013945 | 0.084994 | 0.088145 | -0.0010959 | 0.0087116 | 0.0062378 | -0.10507 | 0.0070089 | 0.080515 | 0.075368 | 3.2078e-05 | 0.0018649 | 0.0042875 | -0.054932 | 86.5768 | -35.7513 | 86.3113 | -30.629 | -34.6436 | -34.8404 | -0.11444 | 86.586 | -35.6827 | 86.5265 | -33.1696 | -35.4126 | -35.0693 | -0.14191 | 0.004577 | -0.26093 | -0.2655 | 4.9805 | 5.1224 | 0.76904 | 0.96588 | 0.91095 | 1.1078 | -0.073242 | 0.013733 | -0.045777 | -0.05951 | 2.4399 | 2.5131 | 0.54016 | 0.19684 | 0.6134 | 0.27008
1529.1719 | 122.438 | -35.6094 | 86.4853 | -0.16106 | 0.018004 | 0.085422 | 0.086436 | 0.0016813 | 0.011702 | 0.0064514 | -0.059296 | 0.0059409 | 0.079019 | 0.071737 | 0.00088648 | -0.0013393 | 0.0015103 | -0.041199 | 86.5265 | -35.7376 | 86.3937 | -30.4825 | -34.5428 | -34.7626 | -0.018311 | 86.499 | -35.701 | 86.6135 | -33.0597 | -35.3119 | -34.9915 | -0.12817 | 0.041198 | -0.091553 | -0.13275 | 5.127 | 5.2551 | 0.84686 | 1.0666 | 0.97504 | 1.1948 | -0.091553 | 0.013732 | 0.12817 | 0.11444 | 2.5497 | 2.6413 | 0.61798 | 0.29755 | 0.70953 | 0.3891
1530.1719 | 122.3557 | -35.6964 | 86.4944 | -0.21446 | 0.015654 | 0.085635 | 0.087291 | -0.0015231 | 0.0089252 | 0.0088012 | -0.080658 | 0.0087179 | 0.079874 | 0.0743 | -0.00082252 | -0.0026211 | 0.0055691 | -0.10986 | 86.5036 | -35.8475 | 86.2747 | -30.5832 | -34.6756 | -34.877 | -0.10529 | 86.4899 | -35.7422 | 86.4807 | -33.1604 | -35.4446 | -35.1242 | -0.15106 | 0.009155 | -0.21973 | -0.22888 | 5.1132 | 5.2643 | 0.8194 | 1.0208 | 0.97046 | 1.1719 | -0.045776 | -0.004578 | -0.013733 | -0.009155 | 2.536 | 2.5818 | 0.5722 | 0.25177 | 0.61798 | 0.29755
1531.1719 | 122.4197 | -35.701 | 86.4166 | -0.077133 | 0.018431 | 0.086917 | 0.085795 | 0.00082671 | 0.0087116 | 0.0058106 | -0.022675 | 0.0089315 | 0.078165 | 0.071523 | 0.00045928 | -0.0019803 | 0.0074917 | -0.059509 | 86.4441 | -35.82 | 86.3251 | -30.4504 | -34.584 | -34.7809 | 0.013733 | 86.3937 | -35.733 | 86.5173 | -33.0734 | -35.3622 | -35.0647 | -0.11902 | 0.027466 | -0.091553 | -0.11902 | 5.2506 | 5.3696 | 0.9201 | 1.1169 | 1.0391 | 1.236 | -0.032043 | -0.022888 | 0.10071 | 0.1236 | 2.6276 | 2.6596 | 0.63629 | 0.33875 | 0.66833 | 0.37079
1532.1875 | 122.406 | -35.7468 | 86.4029 | -0.19157 | 0.015227 | 0.086062 | 0.087077 | -0.00045509 | 0.0074298 | 0.0077332 | 0.023102 | 0.0053001 | 0.079233 | 0.070028 | 3.2078e-05 | -0.0017667 | 0.0047147 | -0.036621 | 86.3434 | -35.9116 | 86.2152 | -30.5283 | -34.639 | -34.8587 | -0.004578 | 86.3571 | -35.82 | 86.4258 | -33.1238 | -35.4126 | -35.1471 | -0.16479 | -0.059509 | -0.18768 | -0.12817 | 5.2185 | 5.3833 | 0.88806 | 1.1078 | 1.0529 | 1.2726 | -0.073242 | -0.045776 | 0.022888 | 0.068664 | 2.623 | 2.6962 | 0.59967 | 0.33417 | 0.67291 | 0.40741
1533.1719 | 122.3648 | -35.788 | 86.3846 | -0.10002 | 0.014373 | 0.083926 | 0.087077 | -2.7692e-05 | 0.0087116 | 0.0075196 | -0.095917 | 0.0070089 | 0.081796 | 0.081136 | -0.0012497 | -0.0011257 | 0.0045011 | 0.009155 | 86.3709 | -35.9299 | 86.1923 | -30.5328 | -34.7076 | -34.9274 | 0.050354 | 86.3663 | -35.8566 | 86.3663 | -33.1696 | -35.5042 | -35.202 | -0.14191 | -0.013733 | -0.19226 | -0.17853 | 5.2551 | 5.397 | 0.8606 | 1.0803 | 1.0025 | 1.2222 | -0.068664 | -0.018311 | -0.018311 | 0 | 2.6184 | 2.6871 | 0.58594 | 0.28381 | 0.6546 | 0.35248
1534.1719 | 122.3923 | -35.8246 | 86.2976 | -0.16869 | 0.019286 | 0.088199 | 0.087077 | -0.0025913 | 0.0093524 | 0.0094422 | -0.056244 | 0.0067953 | 0.07966 | 0.077077 | -0.0027451 | 0.0012241 | 0.0030057 | -0.041199 | 86.3846 | -35.8887 | 86.2701 | -30.4276 | -34.6344 | -34.845 | -0.004578 | 86.3388 | -35.8154 | 86.4532 | -33.078 | -35.4126 | -35.1517 | -0.064087 | 0.086976 | -0.027465 | -0.11444 | 5.397 | 5.4611 | 0.97961 | 1.1902 | 1.0437 | 1.2543 | 0.009155 | 0.041199 | 0.15564 | 0.11444 | 2.7466 | 2.7374 | 0.67291 | 0.41199 | 0.66376 | 0.40283
1535.1719 | 122.3465 | -35.8566 | 86.3388 | -0.10002 | 0.014586 | 0.083926 | 0.088145 | -0.0013095 | 0.0093524 | 0.0081604 | -0.08371 | 0.0076499 | 0.080515 | 0.074087 | -0.00039512 | -5.7742e-05 | 0.0053555 | -0.041199 | 86.3388 | -35.994 | 86.1237 | -30.5191 | -34.7305 | -34.9823 | -0.013733 | 86.3434 | -35.9253 | 86.2839 | -33.2199 | -35.5728 | -35.2615 | -0.13733 | 0 | -0.21515 | -0.21515 | 5.3375 | 5.4749 | 0.87433 | 1.1261 | 1.0117 | 1.2634 | -0.068665 | 0.004578 | -0.054931 | -0.059509 | 2.6367 | 2.7054 | 0.59509 | 0.28381 | 0.66376 | 0.35248
1536.1719 | 122.4747 | -35.7788 | 86.3342 | -0.17632 | 0.016295 | 0.08414 | 0.087931 | -0.0036593 | 0.010634 | 0.0085876 | -0.0013123 | 0.0076499 | 0.079874 | 0.073873 | 0.0019547 | -0.00091214 | 0.0023647 | -0.059509 | 86.3251 | -35.9299 | 86.2335 | -30.3772 | -34.6115 | -34.8358 | -0.19684 | 86.2427 | -35.907 | 86.4166 | -33.0826 | -35.4309 | -35.1151 | -0.15106 | -0.009156 | -0.10071 | -0.091552 | 5.4016 | 5.5527 | 0.94299 | 1.1673 | 1.0941 | 1.3184 | -0.12817 | -0.091553 | 0.082397 | 0.17395 | 2.6962 | 2.8244 | 0.66376 | 0.3479 | 0.79193 | 0.47607
1537.1719 | 122.4106 | -35.8612 | 86.2701 | -0.069504 | 0.015654 | 0.08414 | 0.087504 | -0.0034457 | 0.008498 | 0.0090148 | -0.071503 | 0.0070089 | 0.079446 | 0.078145 | -0.00018152 | -0.0021939 | 0.0038601 | -0.11444 | 86.3113 | -36.0306 | 86.1053 | -30.4871 | -34.7443 | -34.9731 | -0.073242 | 86.2793 | -35.9253 | 86.3205 | -33.1192 | -35.5087 | -35.2112 | -0.16937 | 0.041198 | -0.1648 | -0.20599 | 5.3741 | 5.5435 | 0.88806 | 1.1169 | 1.0574 | 1.2863 | -0.064087 | 0.009155 | 0.050354 | 0.041199 | 2.742 | 2.8061 | 0.65002 | 0.35248 | 0.71411 | 0.41657
1538.1719 | 122.3373 | -35.9253 | 86.2427 | -0.13817 | 0.01715 | 0.088412 | 0.090068 | -0.0019503 | 0.0074298 | 0.0081604 | -0.10812 | 0.0097861 | 0.080515 | 0.073873 | -0.0016769 | 0.00058326 | 0.0059963 | -0.050354 | 86.2656 | -36.0626 | 86.0367 | -30.4962 | -34.79 | -35.0418 | -0.027466 | 86.2289 | -35.9848 | 86.2198 | -33.2336 | -35.6506 | -35.3302 | -0.13733 | 0.022888 | -0.20599 | -0.22888 | 5.4291 | 5.5664 | 0.88348 | 1.1353 | 1.0208 | 1.2726 | -0.059509 | -0.013733 | -0.022888 | -0.009155 | 2.6917 | 2.7512 | 0.59509 | 0.27466 | 0.6546 | 0.33417
1539.1719 | 122.4564 | -35.9619 | 86.142 | -0.11528 | 0.018645 | 0.089053 | 0.087504 | -0.00024129 | 0.007857 | 0.0064514 | -0.040985 | 0.0063681 | 0.079233 | 0.075368 | 0.00045928 | 0.00079686 | 0.0038601 | -0.10529 | 86.1694 | -36.1084 | 86.0184 | -30.455 | -34.758 | -35.0235 | -0.045776 | 86.1465 | -36.0352 | 86.2381 | -33.1375 | -35.5911 | -35.2707 | -0.14648 | 0.027466 | -0.1236 | -0.15106 | 5.5069 | 5.6534 | 0.93841 | 1.2039 | 1.0849 | 1.3504 | -0.073242 | 0.004577 | 0.09613 | 0.091553 | 2.8244 | 2.8976 | 0.69122 | 0.37079 | 0.76447 | 0.44403
1540.1719 | 122.3328 | -36.0947 | 86.1191 | -0.13817 | 0.016509 | 0.08884 | 0.088572 | 0.00082671 | 0.0089252 | 0.0060242 | -0.080658 | 0.0078635 | 0.079874 | 0.080068 | 0.0015275 | 0.00079686 | 0.0030057 | -0.059509 | 86.1008 | -36.2045 | 85.8902 | -30.5878 | -34.9091 | -35.1791 | -0.059509 | 86.1145 | -36.1542 | 86.0504 | -33.3023 | -35.788 | -35.4584 | -0.10986 | -0.018311 | -0.22888 | -0.21057 | 5.5069 | 5.6168 | 0.91553 | 1.1856 | 1.0254 | 1.2955 | -0.059509 | -0.004578 | -0.068665 | -0.064087 | 2.7924 | 2.8519 | 0.63629 | 0.3067 | 0.6958 | 0.36621
1541.1719 | 122.3282 | -36.1496 | 86.0733 | -0.10002 | 0.0195 | 0.086276 | 0.086863 | 0.0038175 | 0.010421 | 0.008374 | -0.080658 | 0.0080771 | 0.080301 | 0.073873 | 0.00088648 | -0.00048494 | 0.0079189 | -0.091553 | 86.0825 | -36.2595 | 85.8032 | -30.6244 | -34.9777 | -35.2432 | -0.11902 | 86.0596 | -36.2091 | 85.9818 | -33.371 | -35.8521 | -35.5179 | -0.10986 | 0.009155 | -0.27008 | -0.27923 | 5.5252 | 5.6351 | 0.90637 | 1.1719 | 1.0162 | 1.2817 | -0.059509 | -0.013733 | -0.091553 | -0.07782 | 2.7786 | 2.8381 | 0.63171 | 0.29755 | 0.69122 | 0.35705
1542.1875 | 122.4152 | -36.0947 | 85.9863 | -0.16869 | 0.018431 | 0.087771 | 0.084941 | 0.0010403 | 0.012557 | 0.0094422 | 0.029205 | 0.0076499 | 0.080942 | 0.077077 | 0.0013139 | -5.7742e-05 | 0.0055691 | -0.041199 | 86.0184 | -36.2595 | 85.9589 | -30.4367 | -34.8495 | -35.083 | 0.032043 | 85.9863 | -36.1816 | 86.0596 | -33.2153 | -35.7193 | -35.4126 | -0.16479 | 0.032044 | -0.027466 | -0.05951 | 5.658 | 5.8228 | 1.0117 | 1.2451 | 1.1765 | 1.4099 | -0.086975 | 0 | 0.073242 | 0.073242 | 2.8793 | 2.9663 | 0.68207 | 0.37537 | 0.76904 | 0.46234
1543.1719 | 122.3648 | -36.1542 | 86.0138 | -0.11528 | 0.018859 | 0.085635 | 0.087718 | -0.0021641 | 0.011061 | 0.005597 | -0.074554 | 0.0055137 | 0.08201 | 0.071737 | -0.00060892 | 0.0014377 | 0.0040737 | 0.027466 | 86.0275 | -36.3098 | 85.8536 | -30.4367 | -34.8816 | -35.1196 | 0.073242 | 86.0092 | -36.2137 | 86.0184 | -33.2565 | -35.7742 | -35.4263 | -0.15564 | 0.013733 | -0.16022 | -0.17395 | 5.7175 | 5.8731 | 1.0345 | 1.2726 | 1.1902 | 1.4282 | -0.059509 | -0.004578 | 0.004578 | 0.009156 | 2.8976 | 2.9572 | 0.72784 | 0.37994 | 0.78735 | 0.43945
1544.1719 | 122.3923 | -36.1816 | 86.0092 | -0.084763 | 0.018859 | 0.084994 | 0.089427 | -0.0023777 | 0.0089252 | 0.006665 | -0.059296 | 0.0078635 | 0.078165 | 0.075582 | 3.2078e-05 | -5.7742e-05 | 0.0034329 | 0.009155 | 85.9589 | -36.3373 | 85.7895 | -30.4733 | -34.9182 | -35.2295 | 0.11902 | 86.0321 | -36.1496 | 85.9863 | -33.2703 | -35.8017 | -35.4996 | -0.15564 | -0.050354 | -0.21973 | -0.16937 | 5.7083 | 5.864 | 0.95215 | 1.2634 | 1.1078 | 1.4191 | 0.032044 | 0.022888 | -0.022888 | -0.045776 | 2.9114 | 2.8793 | 0.68207 | 0.37994 | 0.65002 | 0.3479
1545.1719 | 122.3694 | -36.2503 | 85.9634 | -0.077133 | 0.017363 | 0.086276 | 0.088786 | -0.0025913 | 0.01277 | 0.0062378 | -0.013519 | 0.0063681 | 0.082651 | 0.071096 | 0.00024568 | -0.0015529 | 0.0045011 | -0.022888 | 85.9543 | -36.3785 | 85.7254 | -30.5237 | -34.9548 | -35.2524 | 0.036621 | 85.9268 | -36.3144 | 85.9177 | -33.2748 | -35.8246 | -35.5179 | -0.12817 | -0.009155 | -0.23804 | -0.22888 | 5.7266 | 5.8548 | 0.99792 | 1.2955 | 1.1261 | 1.4236 | -0.064087 | -0.036621 | -0.045776 | -0.009155 | 2.9755 | 3.0396 | 0.73242 | 0.42572 | 0.79651 | 0.48981
1546.1719 | 122.3465 | -36.3419 | 85.8856 | -0.13817 | 0.018218 | 0.085422 | 0.09199 | 0.0018949 | 0.008498 | 0.0060242 | -0.056244 | 0.0089315 | 0.081583 | 0.0743 | 0.0015275 | 0.0010105 | 0.0062101 | -0.036621 | 85.8902 | -36.4334 | 85.6338 | -30.5557 | -35.0327 | -35.3348 | -0.07782 | 85.8856 | -36.3647 | 85.8078 | -33.3664 | -35.9665 | -35.6323 | -0.091553 | 0.004578 | -0.25177 | -0.25635 | 5.7861 | 5.8777 | 1.0071 | 1.3092 | 1.0986 | 1.4008 | -0.022888 | 0 | -0.07782 | -0.07782 | 2.9755 | 2.9984 | 0.70953 | 0.37537 | 0.73242 | 0.39825
1547.1875 | 122.4289 | -36.3052 | 85.7986 | -0.18394 | 0.019713 | 0.08649 | 0.087504 | -0.0019503 | 0.011489 | 0.0064514 | -0.0074159 | 0.0082907 | 0.081155 | 0.068319 | 0.00045928 | -0.0013393 | 0.0057827 | -0.054932 | 85.8032 | -36.4471 | 85.6934 | -30.4413 | -34.964 | -35.2615 | 0.009155 | 85.7941 | -36.3647 | 85.8765 | -33.2428 | -35.8429 | -35.5408 | -0.14191 | 0.004578 | -0.10529 | -0.10986 | 5.864 | 6.0059 | 1.0437 | 1.3412 | 1.1856 | 1.4832 | -0.059509 | -0.004578 | 0.07782 | 0.082398 | 3.0624 | 3.1219 | 0.76447 | 0.46234 | 0.82397 | 0.52185
1548.1719 | 122.3877 | -36.3876 | 85.7941 | -0.10002 | 0.019927 | 0.087985 | 0.090495 | 0.00039951 | 0.010848 | 0.0045288 | -0.022675 | 0.0082907 | 0.081796 | 0.075155 | -0.0010361 | -0.0015529 | 0.0027921 | 0 | 85.8078 | -36.5204 | 85.5743 | -30.542 | -35.0327 | -35.3577 | 0.036621 | 85.7712 | -36.4426 | 85.7712 | -33.3023 | -35.9299 | -35.6415 | -0.13275 | 0.013733 | -0.21973 | -0.23346 | 5.8456 | 5.9784 | 1.03 | 1.355 | 1.1627 | 1.4877 | -0.054932 | -0.022888 | -0.022888 | 0 | 3.0853 | 3.1403 | 0.74616 | 0.45776 | 0.80109 | 0.5127
1549.1719 | 122.3465 | -36.4471 | 85.7346 | -0.11528 | 0.017791 | 0.086703 | 0.087504 | 0.0023221 | 0.010848 | 0.0077332 | -0.065399 | 0.0078635 | 0.082864 | 0.076436 | 0.0011001 | 0.00015586 | 0.0027921 | -0.004578 | 85.7941 | -36.5021 | 85.5515 | -30.5145 | -35.0784 | -35.4034 | -0.009155 | 85.73 | -36.4929 | 85.73 | -33.3572 | -36.0031 | -35.6873 | -0.054931 | 0.059509 | -0.18311 | -0.24261 | 5.9326 | 5.9875 | 1.0437 | 1.3687 | 1.0986 | 1.4236 | -0.045776 | -0.004578 | -0.004578 | 0 | 3.0899 | 3.1357 | 0.75989 | 0.44403 | 0.80566 | 0.48981
1550.1719 | 122.3557 | -36.5021 | 85.6934 | -0.054245 | 0.01715 | 0.083926 | 0.089 | 0.0033901 | 0.010207 | 0.01051 | -0.022675 | 0.0097861 | 0.080087 | 0.075155 | 0.0011001 | 0.0018649 | 0.0072781 | -0.032043 | 85.7254 | -36.6348 | 85.4507 | -30.5878 | -35.1288 | -35.4675 | 0 | 85.6842 | -36.557 | 85.643 | -33.3527 | -36.026 | -35.7376 | -0.13275 | 0.032044 | -0.24261 | -0.27466 | 5.9143 | 6.0471 | 1.0345 | 1.3733 | 1.1673 | 1.506 | -0.054932 | -0.009155 | -0.050354 | -0.041199 | 3.1494 | 3.2043 | 0.76447 | 0.47607 | 0.8194 | 0.53101
1551.1719 | 122.3969 | -36.4883 | 85.6934 | -0.13054 | 0.016936 | 0.084994 | 0.088572 | 0.0044583 | 0.009566 | 0.0062378 | -0.019623 | 0.0070089 | 0.079874 | 0.077291 | -0.0010361 | 0.0012241 | 0.0057827 | 0.018311 | 85.7162 | -36.6211 | 85.4965 | -30.4962 | -35.1242 | -35.4538 | 0.054932 | 85.675 | -36.5524 | 85.6659 | -33.3069 | -36.0031 | -35.7147 | -0.13275 | 0.022889 | -0.19684 | -0.21973 | 5.9921 | 6.1249 | 1.0345 | 1.3641 | 1.1673 | 1.4969 | -0.064087 | -0.01831 | -0.027465 | -0.009155 | 3.1815 | 3.2455 | 0.77362 | 0.48523 | 0.83771 | 0.54932
1552.1719 | 122.374 | -36.5021 | 85.6796 | -0.11528 | 0.021209 | 0.087558 | 0.088359 | -0.00066869 | 0.01213 | 0.0085876 | -0.034882 | 0.0091453 | 0.080728 | 0.072805 | -0.0018905 | 0.0012241 | 0.0062101 | -0.009155 | 85.6888 | -36.644 | 85.4691 | -30.5191 | -35.1242 | -35.4584 | 0.045776 | 85.6842 | -36.5799 | 85.6567 | -33.2703 | -36.0031 | -35.7056 | -0.14191 | 0.009156 | -0.21057 | -0.21973 | 5.983 | 6.1249 | 1.0437 | 1.3779 | 1.1856 | 1.5198 | -0.07782 | 0.004578 | -0.022888 | -0.027466 | 3.2318 | 3.3096 | 0.79651 | 0.49896 | 0.87433 | 0.57678
1553.1719 | 122.3236 | -36.5936 | 85.6201 | -0.10765 | 0.019286 | 0.083926 | 0.086863 | -0.0010959 | 0.011061 | 0.0096558 | 0.044464 | 0.0074363 | 0.083292 | 0.071737 | -0.00039512 | -0.00091214 | 0.0038601 | -0.004578 | 85.611 | -36.7081 | 85.3683 | -30.574 | -35.2066 | -35.5042 | 0.013733 | 85.6155 | -36.6348 | 85.5194 | -33.371 | -36.0901 | -35.8246 | -0.11444 | -0.009155 | -0.25177 | -0.24262 | 6.0196 | 6.134 | 1.0895 | 1.387 | 1.2039 | 1.5015 | -0.041199 | -0.004577 | -0.10071 | -0.096131 | 3.2227 | 3.2639 | 0.76904 | 0.50354 | 0.81024 | 0.54474
1554.1719 | 122.3328 | -36.6028 | 85.6064 | -0.084763 | 0.018218 | 0.084567 | 0.088359 | -0.0021641 | 0.011489 | 0.008374 | -0.019623 | 0.0097861 | 0.081583 | 0.073446 | -0.0012497 | -0.0021939 | 0.0055691 | -0.027466 | 85.5927 | -36.7401 | 85.3638 | -30.6061 | -35.2112 | -35.5545 | -0.004578 | 85.5881 | -36.6714 | 85.5286 | -33.3252 | -36.1084 | -35.8337 | -0.13733 | -0.013733 | -0.24261 | -0.22888 | 5.9967 | 6.134 | 1.0483 | 1.3916 | 1.1856 | 1.5289 | -0.068665 | -0.01831 | -0.07782 | -0.05951 | 3.2776 | 3.3463 | 0.76904 | 0.49439 | 0.83771 | 0.56305
1555.1719 | 122.3007 | -36.6165 | 85.6018 | -0.10002 | 0.016082 | 0.083926 | 0.086436 | 0.0010403 | 0.011489 | 0.0081604 | 0.010895 | 0.0097861 | 0.083292 | 0.071737 | 0.00067288 | -0.0015529 | 0.0032193 | -0.027466 | 85.5652 | -36.7676 | 85.2905 | -30.6107 | -35.2432 | -35.5591 | -0.004578 | 85.5469 | -36.6943 | 85.4507 | -33.3664 | -35.9024 | -35.6918 | -0.15106 | -0.036621 | -0.31128 | -0.27466 | 6.0059 | 6.1569 | 1.0574 | 1.3733 | 1.2085 | 1.5244 | -0.07782 | -0.054932 | -0.15106 | -0.09613 | 3.2501 | 3.3279 | 0.92468 | 0.71411 | 1.0025 | 0.79193
1556.1719 | 122.3831 | -36.7584 | 85.4599 | -0.092392 | 0.016082 | 0.085422 | 0.088572 | -0.00088229 | 0.0074298 | 0.0092286 | -0.016571 | 0.0095725 | 0.083505 | 0.077718 | -0.00039512 | -0.00027134 | 0.0064237 | -0.013733 | 85.4507 | -36.8912 | 85.2173 | -30.6793 | -35.3302 | -35.6689 | 0.041199 | 85.4141 | -36.8088 | 85.4095 | -33.3801 | -36.2045 | -35.9207 | -0.13275 | -0.009155 | -0.24262 | -0.23346 | 6.0791 | 6.2119 | 1.0895 | 1.4282 | 1.2222 | 1.561 | -0.050354 | -0.045776 | -0.050354 | -0.004578 | 3.3783 | 3.4286 | 0.83771 | 0.55389 | 0.88806 | 0.60425
1557.1719 | 122.3373 | -36.8271 | 85.4141 | -0.10002 | 0.018645 | 0.085422 | 0.086436 | -0.0030185 | 0.011489 | 0.0085876 | 0.041412 | 0.0095725 | 0.079019 | 0.077718 | 0.0013139 | -0.0015529 | 0.0074917 | 0.009155 | 85.3912 | -36.9003 | 85.1486 | -30.6885 | -35.3806 | -35.7056 | 0.022888 | 85.3638 | -36.8546 | 85.3043 | -33.4122 | -36.2823 | -36.0214 | -0.073243 | -0.022889 | -0.2655 | -0.24261 | 6.1386 | 6.2119 | 1.1215 | 1.4465 | 1.1948 | 1.5198 | -0.027466 | -0.050354 | -0.10986 | -0.05951 | 3.4149 | 3.4424 | 0.80566 | 0.54474 | 0.83313 | 0.5722
1558.1719 | 122.3465 | -36.8591 | 85.3867 | -0.10002 | 0.018859 | 0.086703 | 0.090068 | -0.0017367 | 0.0080708 | 0.0079468 | -0.019623 | 0.0072227 | 0.080728 | 0.074514 | -0.0012497 | 0.0016513 | 0.0057827 | -0.009155 | 85.3638 | -36.9644 | 85.0983 | -30.7434 | -35.3989 | -35.7513 | -0.009155 | 85.35 | -36.8912 | 85.2859 | -33.403 | -36.2778 | -36.0214 | -0.10529 | -0.022888 | -0.28839 | -0.2655 | 6.1157 | 6.221 | 1.1078 | 1.4603 | 1.2131 | 1.5656 | -0.032043 | -0.036621 | -0.10071 | -0.064087 | 3.4561 | 3.4882 | 0.83771 | 0.58136 | 0.86975 | 0.6134
1559.1719 | 122.3419 | -36.8362 | 85.318 | -0.084763 | 0.018004 | 0.085208 | 0.087077 | 0.0036039 | 0.010207 | 0.0064514 | -0.071503 | 0.010854 | 0.083078 | 0.076436 | -0.00039512 | -0.00091214 | 0.0057827 | -0.059509 | 85.3409 | -36.9873 | 85.1303 | -30.6747 | -35.3897 | -35.7651 | -0.018311 | 85.3317 | -36.8912 | 85.2905 | -33.4259 | -36.2961 | -36.0168 | -0.15106 | 0.022888 | -0.18768 | -0.21057 | 6.1615 | 6.3126 | 1.0712 | 1.4465 | 1.2222 | 1.5976 | -0.054931 | 0.013733 | -0.027466 | -0.041199 | 3.4103 | 3.4653 | 0.8194 | 0.54016 | 0.87433 | 0.59509
1560.1719 | 122.3282 | -36.937 | 85.2905 | -0.046616 | 0.019286 | 0.084994 | 0.089427 | 0.00018591 | 0.011702 | 0.006665 | -0.0043641 | 0.0097861 | 0.080728 | 0.072164 | -0.00018152 | 0.0014377 | 0.0077053 | -0.018311 | 85.2814 | -37.0148 | 85.0616 | -30.7617 | -35.4996 | -35.82 | -0.018311 | 85.2722 | -36.969 | 85.1807 | -33.4442 | -36.3464 | -36.0809 | -0.07782 | -0.009155 | -0.22888 | -0.21973 | 6.1752 | 6.2531 | 1.1169 | 1.4374 | 1.1948 | 1.5152 | -0.032043 | -0.01831 | -0.10986 | -0.091553 | 3.4927 | 3.5248 | 0.85602 | 0.59052 | 0.88806 | 0.62256
1561.1719 | 122.3099 | -36.9873 | 85.2081 | -0.12291 | 0.019072 | 0.086062 | 0.090709 | 0.00082671 | 0.011061 | 0.0075196 | -0.068451 | 0.0093589 | 0.082223 | 0.073018 | 3.2078e-05 | -0.00048494 | 0.0027921 | -0.054932 | 85.2127 | -37.1201 | 84.9838 | -30.7755 | -35.4996 | -35.8933 | -0.050354 | 85.1807 | -37.0468 | 85.1349 | -33.49 | -36.4243 | -36.1542 | -0.13275 | 0.004578 | -0.2243 | -0.22888 | 6.2119 | 6.3446 | 1.0941 | 1.4877 | 1.2268 | 1.6205 | -0.059509 | -0.027466 | -0.073242 | -0.045776 | 3.4973 | 3.5568 | 0.83313 | 0.56305 | 0.89264 | 0.62256
1562.1719 | 122.3465 | -37.0377 | 85.1532 | -0.092392 | 0.016722 | 0.086276 | 0.086009 | -0.0021641 | 0.0097796 | 0.011151 | 0.041412 | 0.0063681 | 0.082864 | 0.081777 | 0.00067288 | 0.00036966 | 0.0042875 | 0.022888 | 85.1669 | -37.1109 | 84.9792 | -30.7251 | -35.4492 | -35.8383 | 0.059509 | 85.1349 | -37.0697 | 85.0708 | -33.4305 | -36.3831 | -36.1542 | -0.073242 | 0.013733 | -0.17395 | -0.18768 | 6.3126 | 6.3858 | 1.1993 | 1.5884 | 1.2726 | 1.6617 | -0.032043 | -0.01831 | -0.082397 | -0.064087 | 3.6072 | 3.6392 | 0.88348 | 0.6546 | 0.91553 | 0.68664
1563.1719 | 122.3099 | -36.9507 | 85.2493 | -0.13817 | 0.02014 | 0.086276 | 0.087931 | -0.0013095 | 0.014052 | 0.0081604 | 0.013947 | 0.0085043 | 0.08201 | 0.07665 | 0.00088648 | -0.00027134 | 0.0021511 | -0.013733 | 85.2676 | -37.056 | 85.0021 | -30.7114 | -35.4218 | -35.8246 | 0.013733 | 85.2493 | -37.001 | 85.1624 | -33.371 | -36.3281 | -36.0992 | -0.10528 | 0.01831 | -0.24719 | -0.2655 | 6.2393 | 6.3446 | 1.1261 | 1.5289 | 1.2314 | 1.6342 | -0.050354 | 0 | -0.086975 | -0.086975 | 3.5797 | 3.6301 | 0.85144 | 0.62256 | 0.90179 | 0.67291
1564.1719 | 122.3373 | -36.9598 | 85.2493 | -0.12291 | 0.019286 | 0.084994 | 0.092204 | 0.0031765 | 0.010207 | 0.011365 | -0.044037 | 0.010427 | 0.083078 | 0.078573 | -0.0014633 | -5.7742e-05 | 0.0062101 | -0.013733 | 85.2081 | -37.0834 | 85.0616 | -30.629 | -35.3851 | -35.788 | 0.013733 | 85.2356 | -36.9553 | 85.1944 | -33.3435 | -36.3144 | -36.0535 | -0.1236 | -0.041199 | -0.18768 | -0.14648 | 6.3309 | 6.4545 | 1.1719 | 1.5747 | 1.2955 | 1.6983 | 0.004578 | -0.013733 | -0.054932 | -0.041199 | 3.6163 | 3.6118 | 0.90637 | 0.64545 | 0.90179 | 0.64087
1565.1719 | 122.2687 | -37.0285 | 85.1349 | -0.10002 | 0.019713 | 0.084781 | 0.087077 | 0.0021085 | 0.010207 | 0.010083 | -0.0013123 | 0.0089315 | 0.084146 | 0.074941 | -0.00060892 | -0.0017667 | 0.0059963 | -0.041199 | 85.1349 | -37.1887 | 84.8602 | -30.7846 | -35.527 | -35.9299 | 0.004578 | 85.1166 | -37.1429 | 85.0113 | -33.3664 | -36.1038 | -36.0535 | -0.16022 | 0 | -0.27466 | -0.27466 | 6.2439 | 6.4041 | 1.0986 | 1.5015 | 1.2589 | 1.6617 | -0.11444 | -0.018311 | -0.1236 | -0.10529 | 3.6621 | 3.7766 | 0.97504 | 0.92468 | 1.0895 | 1.0391
1566.1719 | 122.3511 | -37.1475 | 85.0388 | -0.092392 | 0.019927 | 0.085849 | 0.08964 | 0.00039951 | 0.011275 | 0.0092286 | 0.0017395 | 0.011068 | 0.084146 | 0.08434 | -0.0010361 | 0.00058326 | 0.0032193 | -0.009155 | 85.0296 | -37.2986 | 84.7961 | -30.8533 | -35.5957 | -36.0168 | 0.018311 | 85.0159 | -37.2253 | 84.9747 | -33.4351 | -36.4655 | -36.232 | -0.15106 | -0.009155 | -0.24261 | -0.23346 | 6.2943 | 6.4453 | 1.1307 | 1.5518 | 1.2817 | 1.7029 | -0.07782 | -0.022888 | -0.064087 | -0.041199 | 3.7125 | 3.7903 | 0.91553 | 0.68207 | 0.99335 | 0.75989
1567.1719 | 122.4747 | -37.2574 | 84.8785 | -0.092392 | 0.019713 | 0.083285 | 0.086222 | -0.00066869 | 0.013198 | 0.007306 | 0.065826 | 0.0065817 | 0.083505 | 0.072591 | -0.0033859 | -0.00027134 | 0.011123 | 0.018311 | 84.9335 | -37.3169 | 84.8465 | -30.7571 | -35.527 | -35.9528 | 0.13733 | 84.8831 | -37.2574 | 84.9655 | -33.3664 | -36.3373 | -36.2091 | -0.05951 | 0.054932 | -0.032043 | -0.086975 | 6.5002 | 6.5598 | 1.3046 | 1.7303 | 1.3641 | 1.7899 | 0 | 0.004578 | 0.086975 | 0.082397 | 3.891 | 3.891 | 1.0483 | 0.9201 | 1.0483 | 0.9201
1568.1719 | 122.3236 | -37.3352 | 84.8877 | -0.084763 | 0.018431 | 0.084567 | 0.086863 | -0.0025913 | 0.011275 | 0.0090148 | 0.016998 | 0.011495 | 0.083505 | 0.076436 | -0.00018152 | 0.00036966 | 0.0079189 | -0.022888 | 84.8602 | -37.4634 | 84.6176 | -30.9998 | -35.7468 | -36.145 | 0.018311 | 84.8511 | -37.3672 | 84.7641 | -33.5541 | -36.6348 | -36.4105 | -0.12817 | -0.027466 | -0.27008 | -0.24261 | 6.3354 | 6.4636 | 1.1902 | 1.5884 | 1.3184 | 1.7166 | -0.032044 | -0.036621 | -0.1236 | -0.086975 | 3.7811 | 3.8132 | 0.92468 | 0.70038 | 0.95673 | 0.73242
1569.1719 | 122.319 | -37.4222 | 84.8099 | -0.11528 | 0.018859 | 0.083499 | 0.087504 | 0.00039951 | 0.0093524 | 0.0068786 | 0.032257 | 0.011281 | 0.085641 | 0.073232 | -0.00018152 | 0.00036966 | 0.0079189 | -0.022888 | 84.7824 | -37.5275 | 84.5444 | -31.1874 | -35.8017 | -36.2411 | 0.022888 | 84.8511 | -37.4176 | 84.7046 | -33.5861 | -36.6989 | -36.4975 | -0.10529 | -0.027465 | -0.2655 | -0.23804 | 6.2347 | 6.34 | 1.181 | 1.6205 | 1.2863 | 1.7258 | 0.004577 | 0.041199 | -0.10529 | -0.14648 | 3.8361 | 3.8315 | 0.92468 | 0.72327 | 0.9201 | 0.71869
1570.1719 | 122.4243 | -37.4268 | 84.6634 | -0.11528 | 0.018645 | 0.082431 | 0.088145 | 0.00039951 | 0.01277 | 0.011578 | 0.0017395 | 0.0099997 | 0.082437 | 0.080495 | 0.00024568 | 0.00015586 | 0.011123 | 0.1236 | 84.6771 | -37.6007 | 84.5352 | -31.1417 | -35.7147 | -36.1908 | 0.13733 | 84.6497 | -37.5092 | 84.7046 | -33.5037 | -36.6211 | -36.4563 | -0.17395 | 0.013733 | -0.12817 | -0.14191 | 6.2851 | 6.459 | 1.236 | 1.712 | 1.4099 | 1.886 | -0.082397 | -0.013733 | 0.041199 | 0.054932 | 3.923 | 4.0054 | 0.97046 | 0.80566 | 1.0529 | 0.88806
1571.1719 | 122.4701 | -37.5275 | 84.5627 | -0.1458 | 0.018859 | 0.086276 | 0.086009 | -0.00066869 | 0.012984 | 0.0070922 | 0.041412 | 0.0091453 | 0.081583 | 0.076436 | -0.00018152 | 0.00079686 | 0.0010831 | 0.009155 | 84.581 | -37.6739 | 84.4849 | -31.192 | -35.8246 | -36.2823 | 0.1236 | 84.5581 | -37.619 | 84.6497 | -33.5358 | -36.6898 | -36.525 | -0.14648 | 0.018311 | -0.07782 | -0.096131 | 6.3354 | 6.4819 | 1.2451 | 1.7029 | 1.3916 | 1.8494 | -0.091553 | -0.004578 | 0.086975 | 0.091553 | 3.9917 | 4.0833 | 1.0025 | 0.83771 | 1.0941 | 0.92926
1572.1719 | 122.3236 | -37.7243 | 84.4757 | -0.16106 | 0.021636 | 0.084781 | 0.089213 | -0.0023777 | 0.012343 | 0.0092286 | 0.010895 | 0.012136 | 0.08201 | 0.078145 | -0.0018905 | 0.0027195 | 0.0085599 | 0.045776 | 84.4666 | -37.8342 | 84.2834 | -31.3751 | -36.0123 | -36.4655 | 0.064087 | 84.4437 | -37.7609 | 84.4025 | -33.7189 | -36.9049 | -36.7401 | -0.10986 | -0.009155 | -0.19226 | -0.18311 | 6.3492 | 6.459 | 1.2588 | 1.712 | 1.3687 | 1.8219 | -0.036621 | -0.032043 | -0.073242 | -0.041199 | 4.0054 | 4.0421 | 0.98419 | 0.8194 | 1.0208 | 0.85602
1573.1719 | 122.4564 | -37.7289 | 84.3887 | -0.16106 | 0.020568 | 0.086062 | 0.085795 | -0.0036593 | 0.011061 | 0.0085876 | 0.026154 | 0.0097861 | 0.080942 | 0.073659 | -0.00039512 | 0.00058326 | 0.0085599 | 0.036621 | 84.343 | -37.8754 | 84.2514 | -31.2561 | -36.1588 | -36.7081 | 0.16022 | 84.3475 | -37.8113 | 84.4162 | -33.6411 | -36.8729 | -36.7126 | -0.14648 | -0.045776 | -0.13733 | -0.091553 | 6.4728 | 6.6193 | 1.0208 | 1.5701 | 1.1673 | 1.7166 | -0.082397 | -0.041199 | 0.027466 | 0.068665 | 4.0878 | 4.1702 | 1.0162 | 0.85602 | 1.0986 | 0.93841
1574.1875 | 122.319 | -37.9395 | 84.2743 | -0.12291 | 0.017577 | 0.086703 | 0.091563 | 0.0014675 | 0.011916 | 0.0079468 | -0.10812 | 0.011922 | 0.083292 | 0.088186 | 0.0013139 | 0.0022923 | 0.0055691 | -0.041199 | 84.2606 | -38.0219 | 84.05 | -31.7093 | -36.2778 | -36.7447 | -0.045776 | 84.2651 | -37.9532 | 84.2056 | -33.8837 | -37.1475 | -36.9415 | -0.082398 | -0.013733 | -0.2243 | -0.21057 | 6.2302 | 6.3126 | 1.1948 | 1.6617 | 1.2772 | 1.7441 | -0.013733 | -0.009155 | -0.068665 | -0.05951 | 4.0558 | 4.0695 | 0.99793 | 0.79193 | 1.0117 | 0.80566
1575.1719 | 122.438 | -37.8937 | 84.2239 | -0.13817 | 0.019927 | 0.085208 | 0.087291 | -0.0021641 | 0.01277 | 0.0077332 | 0.016998 | 0.0089315 | 0.083932 | 0.082418 | 0.0019547 | 0.0012241 | 0.0049283 | -0.022888 | 84.2377 | -38.031 | 84.1141 | -31.6635 | -36.1404 | -36.6211 | 0.064087 | 84.2102 | -37.9715 | 84.2834 | -33.7006 | -36.9644 | -36.8134 | -0.13733 | 0.013733 | -0.10986 | -0.1236 | 6.2302 | 6.3675 | 1.2726 | 1.7532 | 1.4099 | 1.8906 | -0.07782 | -0.013733 | 0.059509 | 0.073242 | 4.1931 | 4.2709 | 1.0803 | 0.92926 | 1.1581 | 1.0071
1576.1719 | 122.3328 | -37.9898 | 84.2148 | -0.031357 | 0.0195 | 0.08649 | 0.089427 | -0.0017367 | 0.012557 | 0.012646 | -0.053192 | 0.011495 | 0.085214 | 0.083059 | 0.0015275 | 0.0014377 | 0.0049283 | -0.022888 | 84.2194 | -38.0997 | 83.9127 | -31.8146 | -36.1862 | -36.7996 | -0.009155 | 84.2194 | -38.0264 | 84.1324 | -33.8287 | -37.1338 | -36.9598 | -0.10986 | 0.004577 | -0.30212 | -0.3067 | 6.1752 | 6.2851 | 1.1902 | 1.8036 | 1.3 | 1.9135 | -0.036621 | 0.004577 | -0.082398 | -0.086975 | 4.1611 | 4.1977 | 1.03 | 0.85602 | 1.0666 | 0.89264
1577.1719 | 122.4838 | -38.031 | 84.0637 | -0.11528 | 0.021209 | 0.084567 | 0.085795 | -0.0032321 | 0.01277 | 0.0090148 | 0.023102 | 0.011709 | 0.083505 | 0.076223 | 3.2078e-05 | 0.0022923 | 0.0049283 | 0.027466 | 84.1095 | -38.1454 | 84.0271 | -31.6956 | -36.2228 | -36.7218 | 0.13275 | 84.0637 | -38.0722 | 84.1599 | -33.7234 | -37.0926 | -36.9781 | -0.11444 | 0.045776 | -0.036621 | -0.082397 | 6.3354 | 6.4499 | 1.3092 | 1.8082 | 1.4236 | 1.9226 | -0.041199 | 0 | 0.09613 | 0.09613 | 4.3076 | 4.3488 | 1.0529 | 0.93842 | 1.0941 | 0.97961
1578.1719 | 122.438 | -38.0905 | 84.0317 | -0.10765 | 0.019072 | 0.084994 | 0.087718 | 0.00082671 | 0.010421 | 0.0096558 | 0.023102 | 0.0085043 | 0.085001 | 0.078573 | -0.0016769 | 0.0016513 | 0.0068509 | -0.036621 | 84.0042 | -38.2233 | 83.8806 | -31.8466 | -36.3556 | -36.8408 | 0.036621 | 83.9996 | -38.1546 | 84.0683 | -33.7784 | -37.1475 | -36.9965 | -0.13275 | -0.027466 | -0.15106 | -0.1236 | 6.2439 | 6.3766 | 1.2497 | 1.7349 | 1.3824 | 1.8677 | -0.064087 | -0.032043 | 0.036621 | 0.068664 | 4.3121 | 4.3762 | 1.0941 | 0.94299 | 1.1581 | 1.0071
1579.1719 | 122.319 | -38.2553 | 83.9355 | -0.092392 | 0.019286 | 0.08414 | 0.089 | -0.0013095 | 0.013198 | 0.0090148 | 0.050568 | 0.011281 | 0.083932 | 0.082204 | 0.0011001 | 0.0025059 | 0.0070645 | -0.027466 | 83.9264 | -38.3698 | 83.6975 | -32.0663 | -36.5616 | -36.9553 | -0.11444 | 83.9905 | -38.269 | 83.8669 | -33.9523 | -37.3718 | -37.2025 | -0.11444 | -0.009155 | -0.23804 | -0.22888 | 6.189 | 6.3034 | 1.3 | 1.6937 | 1.4145 | 1.8082 | -0.013733 | 0.054932 | -0.068665 | -0.1236 | 4.303 | 4.3167 | 1.0529 | 0.88348 | 1.0666 | 0.89722
1580.1719 | 122.4655 | -38.2278 | 83.8486 | -0.13054 | 0.0195 | 0.084567 | 0.086863 | -0.0023777 | 0.013625 | 0.0075196 | 0.010895 | 0.012136 | 0.082651 | 0.079213 | 3.2078e-05 | -0.00069854 | 0.0047147 | 0.004578 | 83.844 | -38.3514 | 83.7845 | -31.9199 | -36.4288 | -36.9049 | 0.16022 | 83.8486 | -38.3057 | 83.9447 | -33.8104 | -37.2299 | -37.1155 | -0.1236 | -0.004578 | -0.064087 | -0.059509 | 6.308 | 6.4316 | 1.3229 | 1.799 | 1.4465 | 1.9226 | -0.07782 | 0 | 0.09613 | 0.09613 | 4.4174 | 4.4952 | 1.1124 | 0.99792 | 1.1902 | 1.0757
1581.1719 | 122.4152 | -38.3102 | 83.8074 | -0.12291 | 0.0195 | 0.084353 | 0.088145 | 0.0027493 | 0.011275 | 0.010297 | 0.016998 | 0.012777 | 0.083292 | 0.078573 | -0.0018905 | 0.0020785 | 0.0034329 | -0.041199 | 83.7845 | -38.4796 | 83.6563 | -32.0984 | -36.5616 | -37.0377 | 0.013733 | 83.7845 | -38.3835 | 83.8394 | -33.8791 | -37.3123 | -37.175 | -0.16937 | -0.022888 | -0.15106 | -0.12817 | 6.2119 | 6.3812 | 1.2726 | 1.7487 | 1.442 | 1.918 | -0.073242 | -0.022888 | 0.032044 | 0.054932 | 4.4312 | 4.5044 | 1.1353 | 0.99793 | 1.2085 | 1.0712
1582.1719 | 122.4197 | -38.2095 | 83.8806 | -0.1458 | 0.020568 | 0.087344 | 0.089427 | -0.00024129 | 0.011702 | 0.008374 | -0.028778 | 0.011709 | 0.085214 | 0.082631 | -0.0018905 | 0.0035739 | 0.010055 | -0.082397 | 83.8898 | -38.3377 | 83.7524 | -32.0068 | -36.4746 | -36.9644 | -0.041199 | 83.876 | -38.2965 | 83.9035 | -33.7875 | -37.2208 | -37.0972 | -0.12817 | 0.009156 | -0.12817 | -0.13733 | 6.2027 | 6.3309 | 1.2451 | 1.7349 | 1.3733 | 1.8631 | -0.086975 | -0.004577 | 0.022888 | 0.027465 | 4.422 | 4.509 | 1.1124 | 0.98877 | 1.1993 | 1.0757
1583.1719 | 122.3877 | -38.2141 | 83.9035 | -0.13817 | 0.018645 | 0.08414 | 0.088572 | -0.0017367 | 0.011275 | 0.0098694 | -0.044037 | 0.011281 | 0.081583 | 0.077932 | -0.00039512 | 0.0016513 | 0.0030057 | -0.059509 | 83.8989 | -38.3377 | 83.7891 | -31.9748 | -36.438 | -36.9186 | -0.045776 | 83.8989 | -38.2645 | 83.9355 | -33.7463 | -37.1979 | -37.0789 | -0.1236 | -0.004577 | -0.11444 | -0.10986 | 6.2393 | 6.3629 | 1.2955 | 1.7761 | 1.4191 | 1.8997 | -0.050354 | -0.004577 | 0.032044 | 0.036621 | 4.4678 | 4.5181 | 1.1353 | 1.0162 | 1.1856 | 1.0666
1584.1719 | 122.3328 | -38.2828 | 83.8806 | -0.13054 | 0.019713 | 0.085208 | 0.086436 | -0.0030185 | 0.012343 | 0.007306 | -0.05014 | 0.012777 | 0.085428 | 0.085195 | -0.0012497 | 0.0035739 | 0.0087735 | 0.013733 | 83.8715 | -38.3881 | 83.7204 | -32.0389 | -36.4609 | -36.9873 | 0.036621 | 83.8715 | -38.3148 | 83.844 | -33.7784 | -37.2574 | -37.1567 | -0.10529 | -0.009155 | -0.16022 | -0.15106 | 6.2439 | 6.3492 | 1.2955 | 1.8219 | 1.4008 | 1.9272 | -0.032043 | -0.009155 | -0.036621 | -0.027466 | 4.5044 | 4.5364 | 1.1261 | 1.0254 | 1.1581 | 1.0574
1585.1719 | 122.4197 | -38.3972 | 83.7524 | -0.18394 | 0.019286 | 0.084567 | 0.091136 | -0.00045509 | 0.01277 | 0.011151 | -0.03183 | 0.012563 | 0.082864 | 0.077932 | 0.00024568 | 0.00015586 | 0.0059963 | -0.064087 | 83.7341 | -38.5391 | 83.5922 | -32.135 | -36.5936 | -37.2482 | -0.022888 | 83.6746 | -38.5025 | 83.7204 | -33.8287 | -37.3306 | -37.2162 | -0.14191 | -0.01831 | -0.16022 | -0.14191 | 6.2622 | 6.4041 | 1.149 | 1.8036 | 1.2909 | 1.9455 | -0.10529 | -0.077819 | -0.032043 | 0.045776 | 4.5685 | 4.6738 | 1.181 | 1.0666 | 1.2863 | 1.1719
1586.1719 | 122.4289 | -38.4384 | 83.6426 | -0.16869 | 0.019713 | 0.084781 | 0.090068 | -0.0028049 | 0.013839 | 0.0094422 | -0.010467 | 0.011922 | 0.08735 | 0.084127 | 3.2078e-05 | 0.0033603 | 0.0030057 | 0.032043 | 83.6563 | -38.6032 | 83.5785 | -32.1442 | -36.5662 | -37.1338 | 0.1236 | 83.638 | -38.5254 | 83.757 | -33.8104 | -37.3947 | -37.2665 | -0.16479 | 0.013733 | -0.064087 | -0.07782 | 6.2943 | 6.459 | 1.3046 | 1.8723 | 1.4694 | 2.037 | -0.086975 | -0.004578 | 0.11444 | 0.11902 | 4.628 | 4.715 | 1.1719 | 1.0437 | 1.2589 | 1.1307
1587.1719 | 122.3877 | -38.3881 | 83.7158 | -0.16869 | 0.022917 | 0.085635 | 0.087077 | 0.0021085 | 0.014052 | 0.0096558 | -0.044037 | 0.013845 | 0.081583 | 0.084554 | -0.00060892 | 0.0022923 | 0.0068509 | -0.045776 | 83.67 | -38.5437 | 83.5968 | -32.1121 | -36.5753 | -37.0972 | 0.013733 | 83.6792 | -38.4705 | 83.757 | -33.7692 | -37.3261 | -37.2025 | -0.15564 | -0.045776 | -0.11902 | -0.073242 | 6.2759 | 6.4316 | 1.2909 | 1.8127 | 1.4465 | 1.9684 | -0.082397 | -0.036621 | 0.041199 | 0.07782 | 4.6188 | 4.7012 | 1.1856 | 1.062 | 1.268 | 1.1444
1588.1719 | 122.4014 | -38.4888 | 83.638 | -0.17632 | 0.019927 | 0.085635 | 0.088786 | -0.0017367 | 0.013198 | 0.0081604 | -0.040985 | 0.013204 | 0.083078 | 0.080282 | -0.0010361 | 0.0044285 | 0.0066373 | -0.073242 | 83.6609 | -38.5941 | 83.519 | -32.2037 | -36.6302 | -37.1613 | -0.068665 | 83.638 | -38.5254 | 83.6609 | -33.7967 | -37.3489 | -37.2528 | -0.10528 | 0.022889 | -0.11902 | -0.14191 | 6.2851 | 6.3904 | 1.3275 | 1.8585 | 1.4328 | 1.9638 | -0.036621 | 0 | 0.022889 | 0.022889 | 4.6921 | 4.7287 | 1.236 | 1.1398 | 1.2726 | 1.1765
1589.1719 | 122.4197 | -38.4933 | 83.5648 | -0.13054 | 0.018645 | 0.084781 | 0.087077 | -0.0021641 | 0.015548 | 0.0090148 | -0.037933 | 0.013631 | 0.082651 | 0.08135 | -0.0033859 | 0.0037875 | 0.0047147 | 0.013733 | 83.5831 | -38.6444 | 83.5144 | -32.1899 | -36.6348 | -37.1887 | 0.28381 | 83.5785 | -38.6078 | 83.638 | -33.7418 | -37.3215 | -37.2757 | -0.15106 | 0.018311 | -0.050354 | -0.068665 | 6.3034 | 6.4545 | 1.3046 | 1.8585 | 1.4557 | 2.0096 | -0.11444 | 0.013733 | 0.073242 | 0.059509 | 4.7516 | 4.866 | 1.2177 | 1.1719 | 1.3321 | 1.2863
1590.1719 | 122.4106 | -38.5941 | 83.5419 | -0.13054 | 0.018859 | 0.085422 | 0.087291 | -0.00088229 | 0.011489 | 0.011792 | -0.074554 | 0.013631 | 0.083505 | 0.080495 | -0.0016769 | 0.0037875 | 0.0057827 | -0.054932 | 83.5007 | -38.6993 | 83.4137 | -32.2723 | -36.7218 | -37.2711 | -0.004578 | 83.5007 | -38.6398 | 83.5739 | -33.8196 | -37.468 | -37.3581 | -0.10529 | -0.041199 | -0.12817 | -0.086975 | 6.3217 | 6.427 | 1.3229 | 1.8723 | 1.4282 | 1.9775 | -0.045777 | -0.041199 | 0.032044 | 0.073243 | 4.7745 | 4.8203 | 1.236 | 1.1261 | 1.2817 | 1.1719
1591.1719 | 122.3511 | -38.6993 | 83.4045 | -0.12291 | 0.02014 | 0.087344 | 0.087931 | -0.0040867 | 0.012984 | 0.008374 | -0.077606 | 0.013418 | 0.08201 | 0.086477 | -0.00018152 | 0.0022923 | 0.0042875 | -0.059509 | 83.3771 | -38.8458 | 83.2809 | -32.3685 | -36.8134 | -37.3581 | -0.009155 | 83.3862 | -38.7543 | 83.4778 | -33.8745 | -37.5458 | -37.4496 | -0.14648 | -0.027466 | -0.1236 | -0.09613 | 6.3309 | 6.4774 | 1.3412 | 1.886 | 1.4877 | 2.0325 | -0.054931 | -0.018311 | 0.073242 | 0.091553 | 4.8248 | 4.8798 | 1.2497 | 1.1536 | 1.3046 | 1.2085
1592.1719 | 122.374 | -38.7863 | 83.313 | -0.13054 | 0.019927 | 0.083285 | 0.087718 | 0.00082671 | 0.013839 | 0.0090148 | -0.056244 | 0.010854 | 0.085001 | 0.084554 | 0.00045928 | 0.0052829 | 0.0059963 | -0.073242 | 83.3405 | -38.8733 | 83.2031 | -32.4738 | -36.8774 | -37.4588 | -0.059509 | 83.3176 | -38.8138 | 83.3633 | -33.9111 | -37.6007 | -37.5275 | -0.086975 | 0.027466 | -0.10986 | -0.13733 | 6.3126 | 6.3995 | 1.3275 | 1.9089 | 1.4145 | 1.9959 | -0.027466 | 0.004578 | 0.050354 | 0.045776 | 4.8752 | 4.9026 | 1.2589 | 1.1856 | 1.2863 | 1.2131
1593.1719 | 122.3007 | -38.8504 | 83.3405 | -0.092392 | 0.0195 | 0.086276 | 0.090709 | -0.0030185 | 0.013839 | 0.010297 | -0.074554 | 0.014913 | 0.083932 | 0.089895 | -0.00039512 | 0.0046421 | 0.0094143 | -0.027466 | 83.3405 | -39.0015 | 83.0795 | -32.6065 | -36.9827 | -37.5732 | 0.041199 | 83.3313 | -38.8412 | 83.2718 | -34.0256 | -37.7197 | -37.6236 | -0.15106 | 0 | -0.26093 | -0.26093 | 6.2439 | 6.395 | 1.2772 | 1.8677 | 1.4282 | 2.0187 | 0.009155 | -0.009155 | -0.068664 | -0.059509 | 4.8248 | 4.8157 | 1.2268 | 1.1307 | 1.2177 | 1.1215
1594.1719 | 122.406 | -38.8458 | 83.2306 | -0.10002 | 0.022063 | 0.085422 | 0.090281 | -0.0013095 | 0.012984 | 0.010937 | -0.022675 | 0.012563 | 0.08436 | 0.084981 | -0.00018152 | 0.0035739 | 0.0049283 | -0.009155 | 83.2397 | -38.9877 | 83.1528 | -32.4829 | -36.8912 | -37.4863 | 0.10071 | 83.2397 | -38.9236 | 83.3313 | -33.8333 | -37.5687 | -37.5458 | -0.14191 | 0.009155 | -0.07782 | -0.086975 | 6.3629 | 6.5048 | 1.3596 | 1.9547 | 1.5015 | 2.0966 | -0.07782 | 0.009155 | 0.10071 | 0.091553 | 5.0125 | 5.0903 | 1.3 | 1.2772 | 1.3779 | 1.355
1595.1719 | 122.4014 | -38.9328 | 83.1665 | -0.15343 | 0.020781 | 0.084994 | 0.090281 | 0.0014675 | 0.013411 | 0.007306 | -0.10202 | 0.015127 | 0.085641 | 0.087758 | -0.00018152 | 0.0044285 | 0.0053555 | -0.073242 | 83.1573 | -39.1113 | 83.0566 | -32.5882 | -37.0148 | -37.6282 | -0.013733 | 83.1345 | -39.0289 | 83.2169 | -33.9203 | -37.706 | -37.6144 | -0.17853 | -0.009155 | -0.10986 | -0.10071 | 6.3446 | 6.5231 | 1.3046 | 1.918 | 1.4832 | 2.0966 | -0.096131 | -0.032044 | 0.050354 | 0.082398 | 5.0125 | 5.1086 | 1.3184 | 1.2268 | 1.4145 | 1.3229
1596.1719 | 122.4289 | -39.0839 | 82.9834 | -0.12291 | 0.018004 | 0.085635 | 0.087291 | -0.0030185 | 0.014479 | 0.0098694 | -0.0074159 | 0.012777 | 0.085214 | 0.086263 | 0.00045928 | 0.0046421 | 0.0096279 | -0.027466 | 83.0246 | -39.2349 | 82.8873 | -32.7072 | -37.1201 | -37.7106 | 0.059509 | 82.9926 | -39.18 | 83.0841 | -33.9661 | -37.7838 | -37.738 | -0.15106 | 0.041199 | -0.09613 | -0.13733 | 6.3766 | 6.5277 | 1.3733 | 1.9638 | 1.5244 | 2.1149 | -0.096131 | 0.009156 | 0.10071 | 0.091552 | 5.1178 | 5.2139 | 1.3458 | 1.3 | 1.442 | 1.3962
1597.1719 | 122.4243 | -39.0839 | 83.0109 | -0.15343 | 0.021636 | 0.087344 | 0.086436 | -0.00045509 | 0.015334 | 0.010937 | -0.065399 | 0.014272 | 0.085001 | 0.08434 | -0.00018152 | 0.0044285 | 0.0077053 | -0.050354 | 82.9834 | -39.2441 | 82.9147 | -32.7255 | -37.1155 | -37.7243 | 0.013733 | 82.9834 | -39.1434 | 83.0795 | -33.9661 | -37.8113 | -37.7701 | -0.16022 | -0.027466 | -0.09613 | -0.068664 | 6.3583 | 6.5186 | 1.3596 | 1.9684 | 1.5198 | 2.1286 | -0.05951 | -0.027466 | 0.068665 | 0.096131 | 5.1178 | 5.1773 | 1.3138 | 1.2726 | 1.3733 | 1.3321
1598.1875 | 122.3511 | -39.0793 | 83.0978 | -0.084763 | 0.019286 | 0.083926 | 0.087718 | -0.0023777 | 0.012557 | 0.012006 | -0.053192 | 0.014058 | 0.085855 | 0.082845 | 3.2078e-05 | 0.0050693 | 0.0085599 | -0.032043 | 83.0612 | -39.2624 | 82.7957 | -32.8537 | -37.207 | -37.825 | -0.009155 | 83.0383 | -39.1891 | 82.9559 | -34.0073 | -37.9395 | -37.6694 | -0.18311 | -0.036621 | -0.30212 | -0.2655 | 6.2256 | 6.4087 | 1.2543 | 1.8723 | 1.4374 | 2.0554 | -0.10986 | -0.059509 | -0.14191 | -0.082397 | 5.072 | 5.1819 | 1.4099 | 1.1398 | 1.5198 | 1.2497
1599.1719 | 122.3969 | -39.1846 | 82.9102 | -0.13817 | 0.0195 | 0.086917 | 0.087291 | -2.7692e-05 | 0.014052 | 0.0077332 | -0.019623 | 0.016408 | 0.084787 | 0.076436 | 0.00024568 | 0.0033603 | 0.011337 | -0.018311 | 82.9239 | -39.3265 | 82.8049 | -32.8354 | -37.1979 | -37.8204 | 0.054932 | 82.8964 | -39.267 | 82.9926 | -33.9249 | -37.825 | -37.7975 | -0.14191 | 0.013733 | -0.10528 | -0.11902 | 6.3492 | 6.4911 | 1.3641 | 1.9867 | 1.506 | 2.1286 | -0.082398 | -0.013733 | 0.082398 | 0.096131 | 5.2597 | 5.3421 | 1.387 | 1.3596 | 1.4694 | 1.442
1600.1719 | 122.3694 | -39.2578 | 82.8827 | -0.13817 | 0.0195 | 0.084994 | 0.087931 | -0.0013095 | 0.01213 | 0.010083 | -0.059296 | 0.015767 | 0.085214 | 0.086263 | -0.0014633 | 0.0029331 | 0.010483 | -0.07782 | 82.869 | -39.3677 | 82.7362 | -32.8674 | -37.2528 | -37.8983 | -0.032043 | 82.8598 | -39.299 | 82.9147 | -33.9706 | -37.8937 | -37.8662 | -0.10986 | -0.013732 | -0.14648 | -0.13275 | 6.3904 | 6.5002 | 1.3596 | 2.005 | 1.4694 | 2.1149 | -0.041199 | -0.022888 | 0.032044 | 0.054932 | 5.2872 | 5.3284 | 1.3916 | 1.3641 | 1.4328 | 1.4053
1601.1719 | 122.3923 | -39.1617 | 82.9514 | -0.092392 | 0.02014 | 0.084567 | 0.089427 | -0.0021641 | 0.013198 | 0.0090148 | -0.025726 | 0.015767 | 0.08671 | 0.082631 | -0.0010361 | 0.0046421 | 0.010696 | -0.032043 | 82.933 | -39.2944 | 82.8369 | -32.8033 | -37.1429 | -37.7975 | 0.045776 | 82.9468 | -39.2212 | 83.0063 | -33.8104 | -37.7563 | -37.7426 | -0.13275 | -0.018311 | -0.11444 | -0.09613 | 6.3583 | 6.4911 | 1.3641 | 2.0187 | 1.4969 | 2.1515 | -0.059509 | -0.004578 | 0.054932 | 0.05951 | 5.3513 | 5.4108 | 1.4191 | 1.4053 | 1.4786 | 1.4648
1602.1719 | 122.3923 | -39.2761 | 82.8369 | -0.12291 | 0.019072 | 0.085208 | 0.089854 | 0.00082671 | 0.014052 | 0.01286 | -0.037933 | 0.015127 | 0.085855 | 0.080922 | 3.2078e-05 | 0.0046421 | 0.01091 | -0.068665 | 82.8461 | -39.386 | 82.6813 | -32.9132 | -37.2849 | -37.9074 | -0.050354 | 82.8232 | -39.3585 | 82.8781 | -33.9203 | -37.8754 | -37.8799 | -0.10986 | 0.009155 | -0.15564 | -0.16479 | 6.3629 | 6.4728 | 1.3687 | 1.9913 | 1.4786 | 2.1011 | -0.082398 | -0.013733 | 0.041199 | 0.054932 | 5.3558 | 5.4382 | 1.3962 | 1.4008 | 1.4786 | 1.4832
1603.1875 | 122.3236 | -39.386 | 82.8094 | -0.11528 | 0.019072 | 0.086062 | 0.089854 | 0.00018591 | 0.016402 | 0.010937 | -0.10812 | 0.018117 | 0.086496 | 0.087118 | 0.00067288 | 0.0078465 | 0.0089871 | -0.045776 | 82.7957 | -39.5187 | 82.5668 | -32.9544 | -37.4313 | -38.0814 | -0.027466 | 82.8049 | -39.4409 | 82.7316 | -33.9935 | -38.0173 | -38.0127 | -0.13275 | -0.013733 | -0.24262 | -0.22888 | 6.4316 | 6.5643 | 1.3046 | 1.9547 | 1.4374 | 2.0874 | -0.054932 | -0.004577 | -0.07782 | -0.073243 | 5.3925 | 5.4474 | 1.3733 | 1.3687 | 1.4282 | 1.4236
1604.1719 | 122.406 | -39.5142 | 82.6172 | -0.15343 | 0.019286 | 0.085208 | 0.087718 | 0.0021085 | 0.013411 | 0.010083 | -0.016571 | 0.015981 | 0.086069 | 0.083059 | -0.0012497 | 0.0052829 | 0.010483 | -0.07782 | 82.6218 | -39.6011 | 82.4844 | -33.0368 | -37.4863 | -38.0997 | -0.036621 | 82.6126 | -39.5325 | 82.6675 | -34.0576 | -38.0493 | -38.0356 | -0.086975 | 0.004578 | -0.13275 | -0.13733 | 6.4774 | 6.5643 | 1.4145 | 2.0279 | 1.5015 | 2.1149 | -0.018311 | -0.004577 | 0.050355 | 0.054932 | 5.4565 | 5.4749 | 1.4786 | 1.4648 | 1.4969 | 1.4832
1605.1719 | 122.3419 | -39.5599 | 82.4936 | -0.1992 | 0.019713 | 0.087131 | 0.089427 | -0.0015231 | 0.016188 | 0.0094422 | -0.071503 | 0.015981 | 0.084787 | 0.089681 | -0.0012497 | 0.0050693 | 0.013046 | -0.050354 | 82.4615 | -39.7339 | 82.3837 | -33.11 | -37.6144 | -38.2187 | -0.018311 | 82.5165 | -39.6606 | 82.5439 | -34.0576 | -38.1638 | -38.1729 | -0.17395 | -0.032043 | -0.10986 | -0.07782 | 6.4499 | 6.6238 | 1.3412 | 1.9455 | 1.5152 | 2.1194 | -0.10071 | 0.022888 | 0.050354 | 0.027466 | 5.5023 | 5.603 | 1.387 | 1.3962 | 1.4877 | 1.4969
1606.1719 | 122.406 | -39.6286 | 82.5027 | -0.092392 | 0.018431 | 0.083713 | 0.087718 | -0.0019503 | 0.011275 | 0.0092286 | -0.044037 | 0.017049 | 0.086496 | 0.080922 | -0.00018152 | 0.0065647 | 0.010696 | -0.045776 | 82.4478 | -39.7842 | 82.3517 | -33.0963 | -37.6144 | -38.2141 | 0 | 82.489 | -39.6652 | 82.5485 | -34.0027 | -38.1042 | -38.1454 | -0.15564 | -0.054932 | -0.15106 | -0.09613 | 6.5323 | 6.6879 | 1.4145 | 2.0142 | 1.5701 | 2.1698 | -0.036621 | -0.013733 | 0.045776 | 0.059509 | 5.6259 | 5.6625 | 1.4832 | 1.5244 | 1.5198 | 1.561
1607.1719 | 122.3511 | -39.7064 | 82.3929 | -0.19157 | 0.020568 | 0.087131 | 0.08964 | -0.00024129 | 0.014266 | 0.0096558 | -0.05014 | 0.017904 | 0.085855 | 0.087758 | -0.0016769 | 0.0074191 | 0.0053555 | -0.054932 | 82.3654 | -39.8346 | 82.283 | -33.1375 | -37.6785 | -38.3102 | -0.004578 | 82.3654 | -39.7751 | 82.4341 | -34.0485 | -38.2187 | -38.2233 | -0.12817 | -0.027466 | -0.10986 | -0.082397 | 6.5689 | 6.6971 | 1.3962 | 2.0279 | 1.5244 | 2.1561 | -0.068664 | -0.027466 | 0.041199 | 0.068665 | 5.658 | 5.7266 | 1.4832 | 1.4877 | 1.5518 | 1.5564
1608.1875 | 122.3373 | -39.8849 | 82.251 | -0.1458 | 0.02014 | 0.084994 | 0.089427 | 0.0027493 | 0.015334 | 0.010297 | -0.13254 | 0.02004 | 0.085428 | 0.087972 | -0.0021041 | 0.0069919 | 0.0059963 | -0.004578 | 82.2601 | -40.0497 | 82.0999 | -33.3572 | -37.8799 | -38.5254 | 0.14191 | 82.283 | -39.8758 | 82.2693 | -34.1812 | -38.4064 | -38.4247 | -0.16479 | 0.009155 | -0.15106 | -0.16022 | 6.5277 | 6.6925 | 1.3596 | 2.005 | 1.5244 | 2.1698 | 0.009156 | 0.032043 | 0.01831 | -0.013733 | 5.7037 | 5.6946 | 1.4603 | 1.4786 | 1.4511 | 1.4694
1609.1719 | 122.374 | -39.9261 | 82.1594 | -0.1458 | 0.022063 | 0.086062 | 0.089 | 0.00061311 | 0.01277 | 0.011578 | -0.062347 | 0.016622 | 0.085855 | 0.085836 | 0.0013139 | 0.0067783 | 0.0077053 | -0.036621 | 82.1503 | -40.0909 | 82.0541 | -33.3344 | -37.8937 | -38.5345 | 0.004578 | 82.1457 | -39.9902 | 82.2464 | -34.1446 | -38.3789 | -38.4338 | -0.16479 | -0.009155 | -0.10529 | -0.096131 | 6.5918 | 6.7566 | 1.3916 | 2.0325 | 1.5564 | 2.1973 | -0.064087 | -0.013733 | 0.086975 | 0.10071 | 5.7816 | 5.8456 | 1.4923 | 1.5472 | 1.5564 | 1.6113
1610.1719 | 122.4014 | -40.0635 | 82.0496 | -0.16869 | 0.020354 | 0.083926 | 0.085582 | 0.0018949 | 0.015334 | 0.0098694 | -0.022675 | 0.01769 | 0.085855 | 0.081563 | 0.00024568 | 0.0057101 | 0.0089871 | -0.032043 | 82.0404 | -40.2054 | 81.9214 | -33.4625 | -37.999 | -38.6398 | 0.027466 | 82.0404 | -40.1367 | 82.0999 | -34.1812 | -38.4613 | -38.4979 | -0.14191 | -0.009156 | -0.12817 | -0.11902 | 6.601 | 6.7429 | 1.4236 | 2.0645 | 1.5656 | 2.2064 | -0.073242 | -0.009156 | 0.050354 | 0.05951 | 5.8823 | 5.9555 | 1.5656 | 1.6022 | 1.6388 | 1.6754
1611.1719 | 122.4518 | -40.1688 | 81.8939 | -0.084763 | 0.019072 | 0.086703 | 0.089427 | -0.0060093 | 0.014693 | 0.0068786 | -0.0043641 | 0.017049 | 0.087564 | 0.087972 | 0.00067288 | 0.0037875 | 0.011123 | -0.018311 | 81.9168 | -40.3152 | 81.8298 | -33.5175 | -38.0951 | -38.7314 | 0.091553 | 81.9077 | -40.2283 | 81.9946 | -34.2133 | -38.5391 | -38.6032 | -0.14648 | 0.022888 | -0.064087 | -0.086975 | 6.6513 | 6.7978 | 1.4374 | 2.0737 | 1.5839 | 2.2202 | -0.059509 | 0.013733 | 0.10071 | 0.086975 | 5.9555 | 6.015 | 1.5656 | 1.6296 | 1.6251 | 1.6891
1612.1719 | 122.4106 | -40.2603 | 81.8161 | -0.12291 | 0.019713 | 0.086703 | 0.089 | -0.0036593 | 0.013839 | 0.012219 | -0.10507 | 0.019185 | 0.086923 | 0.086049 | -0.0018905 | 0.0076327 | 0.0070645 | -0.045776 | 81.7932 | -40.3839 | 81.514 | -33.7418 | -37.9715 | -38.7589 | 0.041199 | 81.7932 | -40.3152 | 81.9214 | -34.259 | -38.6215 | -38.6719 | -0.1236 | -0.022888 | -0.30212 | -0.27924 | 6.5186 | 6.6422 | 1.5015 | 2.2888 | 1.6251 | 2.4124 | -0.054932 | -0.022888 | 0.10529 | 0.12817 | 6.0013 | 6.0562 | 1.5884 | 1.6388 | 1.6434 | 1.6937
1613.1719 | 122.4106 | -40.2649 | 81.8207 | -0.21446 | 0.02014 | 0.087985 | 0.088786 | 0.00018591 | 0.01512 | 0.013074 | -0.059296 | 0.016836 | 0.083505 | 0.086904 | -0.00082252 | 0.0069919 | 0.01091 | -0.032043 | 81.8436 | -40.3885 | 81.7291 | -33.577 | -38.1866 | -38.8229 | 0.013733 | 81.839 | -40.3244 | 81.9031 | -34.2178 | -38.6307 | -38.6765 | -0.1236 | 0.022888 | -0.091553 | -0.11444 | 6.6879 | 6.8115 | 1.442 | 2.0782 | 1.5656 | 2.2018 | -0.059509 | 0.01831 | 0.082397 | 0.064087 | 6.0471 | 6.1066 | 1.5884 | 1.6342 | 1.6479 | 1.6937
1614.1719 | 122.4106 | -40.361 | 81.7062 | -0.084763 | 0.020568 | 0.087131 | 0.088786 | -0.0032321 | 0.013625 | 0.0098694 | -0.0043641 | 0.013631 | 0.085428 | 0.0837 | -0.0016769 | 0.0065647 | 0.011764 | 0.032043 | 81.7337 | -40.4892 | 81.665 | -33.6227 | -38.2141 | -38.887 | 0.28381 | 81.7245 | -40.4251 | 81.7932 | -34.2499 | -38.6307 | -38.7497 | -0.12817 | 0.027466 | -0.041199 | -0.068665 | 6.7383 | 6.8665 | 1.474 | 2.1469 | 1.6022 | 2.2751 | -0.064087 | 0.01831 | 0.086975 | 0.068665 | 6.1111 | 6.1752 | 1.6113 | 1.7303 | 1.6754 | 1.7944
1615.1719 | 122.4197 | -40.4343 | 81.7062 | -0.13054 | 0.020568 | 0.086062 | 0.090281 | 0.0023221 | 0.015975 | 0.013287 | -0.098968 | 0.018117 | 0.082864 | 0.092244 | -0.00082252 | 0.0084873 | 0.0096279 | -0.064087 | 81.6742 | -40.5533 | 81.5826 | -33.728 | -38.3423 | -38.9786 | -0.018311 | 81.6879 | -40.4617 | 81.7841 | -34.2499 | -38.7405 | -38.7726 | -0.11902 | -0.032044 | -0.1236 | -0.091552 | 6.7062 | 6.8253 | 1.4557 | 2.092 | 1.5747 | 2.211 | -0.027466 | -0.018311 | 0.07782 | 0.096131 | 6.1844 | 6.2119 | 1.6617 | 1.6937 | 1.6891 | 1.7212
1616.1719 | 122.3648 | -40.567 | 81.5231 | -0.12291 | 0.018645 | 0.087771 | 0.08964 | -0.0023777 | 0.015334 | 0.012219 | -0.028778 | 0.017263 | 0.086282 | 0.08904 | 0.0023819 | 0.0069919 | 0.0083463 | -0.091553 | 81.5369 | -40.7043 | 81.4087 | -33.8654 | -38.4796 | -39.1296 | -0.059509 | 81.5231 | -40.6311 | 81.5826 | -34.3643 | -38.8687 | -38.9603 | -0.13733 | 0.013733 | -0.11444 | -0.12817 | 6.7017 | 6.839 | 1.4374 | 2.0874 | 1.5747 | 2.2247 | -0.064087 | 0 | 0.05951 | 0.05951 | 6.2027 | 6.2668 | 1.6068 | 1.6983 | 1.6708 | 1.7624
1617.1719 | 122.406 | -40.5991 | 81.5048 | -0.12291 | 0.020995 | 0.08649 | 0.086863 | -2.7692e-05 | 0.017256 | 0.010083 | -0.062347 | 0.017904 | 0.08671 | 0.08135 | 0.0021683 | 0.0087009 | 0.012192 | -0.10529 | 81.4362 | -40.7455 | 81.3858 | -33.8699 | -38.4979 | -39.1571 | 0.018311 | 81.4774 | -40.6631 | 81.5781 | -34.3277 | -38.8458 | -38.942 | -0.14648 | -0.068665 | -0.11902 | -0.050354 | 6.7291 | 6.8756 | 1.442 | 2.1011 | 1.5884 | 2.2476 | -0.064087 | -0.027466 | 0.073242 | 0.10071 | 6.2714 | 6.3354 | 1.6571 | 1.7532 | 1.7212 | 1.8173
1618.1719 | 122.374 | -40.6128 | 81.4957 | -0.10002 | 0.018004 | 0.087985 | 0.088786 | -0.00045509 | 0.013625 | 0.013287 | -0.05014 | 0.019399 | 0.085428 | 0.089467 | -0.00082252 | 0.0091281 | 0.011764 | -0.050354 | 81.4499 | -40.7547 | 81.3904 | -33.8379 | -38.4888 | -39.1663 | 0.004578 | 81.4911 | -40.6494 | 81.5506 | -34.3185 | -38.8504 | -38.9511 | -0.14191 | -0.045777 | -0.10529 | -0.059509 | 6.7749 | 6.9168 | 1.4465 | 2.124 | 1.5884 | 2.2659 | -0.036621 | -0.004578 | 0.054931 | 0.059509 | 6.2943 | 6.3309 | 1.6617 | 1.7624 | 1.6983 | 1.799
1619.1875 | 122.4152 | -40.6631 | 81.4407 | -0.1458 | 0.0195 | 0.084994 | 0.090495 | -0.0034457 | 0.014693 | 0.0090148 | -0.025726 | 0.019399 | 0.087137 | 0.085195 | -0.0031723 | 0.0082737 | 0.0079189 | -0.018311 | 81.4133 | -40.8005 | 81.3629 | -33.9523 | -38.5025 | -39.1891 | 0.12817 | 81.4041 | -40.7272 | 81.5002 | -34.2865 | -38.8641 | -38.9877 | -0.13733 | -0.027466 | -0.07782 | -0.050354 | 6.7108 | 6.8481 | 1.474 | 2.1606 | 1.6113 | 2.298 | -0.064087 | -0.036621 | 0.059509 | 0.09613 | 6.3766 | 6.4407 | 1.6754 | 1.799 | 1.7395 | 1.8631
1620.1719 | 122.3831 | -40.9103 | 81.1935 | -0.18394 | 0.020781 | 0.085635 | 0.090922 | 0.0014675 | 0.015548 | 0.008374 | -0.037933 | 0.020467 | 0.086496 | 0.088613 | 0.0011001 | 0.0084873 | 0.010483 | -0.064087 | 81.1981 | -41.0294 | 81.0791 | -34.1217 | -38.7863 | -39.4547 | -0.045776 | 81.1844 | -40.9561 | 81.2347 | -34.4925 | -39.1251 | -39.2578 | -0.11902 | 0.004578 | -0.11444 | -0.11902 | 6.7886 | 6.9077 | 1.4557 | 2.124 | 1.5747 | 2.243 | -0.045777 | -0.009155 | 0.041199 | 0.050354 | 6.4178 | 6.4636 | 1.6525 | 1.7853 | 1.6983 | 1.8311
1621.1719 | 122.4014 | -41.0889 | 81.015 | -0.13054 | 0.020781 | 0.084781 | 0.086222 | -0.0013095 | 0.014693 | 0.010083 | -0.019623 | 0.017263 | 0.086282 | 0.088399 | -0.0014633 | 0.0076327 | 0.0092007 | -0.068665 | 80.983 | -41.217 | 80.8823 | -34.3002 | -38.942 | -39.6469 | -0.013733 | 81.0196 | -41.1438 | 81.0654 | -34.6024 | -39.2258 | -39.3768 | -0.12817 | -0.032044 | -0.13275 | -0.10071 | 6.7886 | 6.9168 | 1.442 | 2.1469 | 1.5701 | 2.2751 | -0.054932 | 0.004577 | 0.050354 | 0.045777 | 6.4865 | 6.5414 | 1.712 | 1.8631 | 1.767 | 1.918
1622.1719 | 122.3877 | -41.1301 | 80.9967 | -0.18394 | 0.019072 | 0.08414 | 0.090922 | 0.00018591 | 0.013411 | 0.01051 | -0.040985 | 0.01769 | 0.084787 | 0.089254 | 0.0017411 | 0.0074191 | 0.010696 | -0.073242 | 80.9326 | -41.3132 | 80.8228 | -34.3552 | -38.9877 | -39.6835 | 0.009155 | 80.9647 | -41.1896 | 81.0287 | -34.6161 | -39.2807 | -39.4135 | -0.1831 | -0.064087 | -0.17395 | -0.10986 | 6.7749 | 6.958 | 1.4465 | 2.1423 | 1.6296 | 2.3254 | -0.059509 | -0.032043 | 0.032044 | 0.064087 | 6.514 | 6.5735 | 1.7166 | 1.8494 | 1.7761 | 1.9089
1623.1719 | 122.319 | -41.3269 | 80.8044 | -0.19157 | 0.021209 | 0.084994 | 0.094126 | -0.0019503 | 0.016402 | 0.012006 | -0.12033 | 0.021535 | 0.08671 | 0.095022 | 3.2078e-05 | 0.010196 | 0.0070645 | -0.064087 | 80.8319 | -41.4505 | 80.5984 | -34.5474 | -39.2578 | -39.9673 | -0.1236 | 80.8182 | -41.3773 | 80.7999 | -34.8862 | -39.5462 | -39.6652 | -0.1236 | 0.027466 | -0.20599 | -0.23346 | 6.7795 | 6.9031 | 1.3596 | 2.0691 | 1.4832 | 2.1927 | -0.050354 | 0.013733 | -0.004577 | -0.01831 | 6.4407 | 6.4911 | 1.6617 | 1.7807 | 1.712 | 1.8311
1624.1719 | 122.3785 | -41.5604 | 80.5435 | -0.21446 | 0.020568 | 0.087131 | 0.090068 | 0.00018591 | 0.016402 | 0.0094422 | -0.03183 | 0.016195 | 0.086496 | 0.09139 | 0.0013139 | 0.0078465 | 0.0081327 | -0.032043 | 80.5115 | -41.6794 | 80.4016 | -34.6664 | -39.3997 | -40.0955 | 0.022888 | 80.5344 | -41.6336 | 80.5984 | -35.0143 | -39.6515 | -39.8163 | -0.11902 | -0.032043 | -0.14191 | -0.10986 | 6.8939 | 7.0129 | 1.4648 | 2.1606 | 1.5839 | 2.2797 | -0.073242 | -0.009155 | 0.054932 | 0.064087 | 6.546 | 6.6193 | 1.7441 | 1.9089 | 1.8173 | 1.9821
1625.1719 | 122.3877 | -42.3614 | 79.6921 | -0.16869 | 0.022277 | 0.086917 | 0.088145 | -0.00024129 | 0.012984 | 0.010297 | -0.092865 | 0.01769 | 0.083505 | 0.087972 | -0.00082252 | 0.0082737 | 0.0096279 | -0.050354 | 79.715 | -42.5171 | 79.628 | -35.321 | -40.2054 | -40.9012 | -0.013733 | 79.7379 | -42.4072 | 79.8157 | -36.1542 | -40.48 | -40.6998 | -0.15564 | 0.022888 | -0.064087 | -0.086975 | 7.0404 | 7.196 | 1.4603 | 2.1561 | 1.6159 | 2.3117 | -0.045777 | 0.045776 | 0.1236 | 0.07782 | 6.2073 | 6.2531 | 1.6617 | 1.8814 | 1.7075 | 1.9272
1626.1719 | 122.3694 | -42.6315 | 79.4495 | -0.17632 | 0.017791 | 0.08649 | 0.08665 | -0.0013095 | 0.014479 | 0.011365 | -0.071503 | 0.020467 | 0.083719 | 0.086477 | 0.00045928 | 0.0089145 | 0.0083463 | -0.045776 | 79.3808 | -42.778 | 79.3304 | -35.527 | -40.4617 | -41.1621 | -0.009155 | 79.3945 | -42.7368 | 79.4907 | -36.5479 | -40.741 | -41.011 | -0.14648 | -0.068665 | -0.11902 | -0.050354 | 7.1045 | 7.251 | 1.4694 | 2.1698 | 1.6159 | 2.3163 | -0.10528 | -0.054932 | 0.041199 | 0.096131 | 6.0837 | 6.189 | 1.6205 | 1.8906 | 1.7258 | 1.9958
1627.1719 | 122.3236 | -42.8055 | 79.3167 | -0.42809 | 0.021422 | 0.086276 | 0.091563 | -0.0036593 | 0.013625 | 0.0090148 | -0.13864 | 0.021963 | 0.083719 | 0.087118 | 0.00024568 | 0.0095555 | 0.016037 | -0.022888 | 79.3396 | -42.9337 | 79.0787 | -35.7651 | -40.6906 | -41.4185 | -0.10986 | 79.3167 | -42.8787 | 79.2892 | -36.7905 | -40.9744 | -41.2125 | -0.12817 | 0.022889 | -0.23804 | -0.26093 | 7.0404 | 7.1686 | 1.387 | 2.1149 | 1.5152 | 2.243 | -0.073242 | 0 | -0.027465 | -0.027465 | 6.015 | 6.0883 | 1.593 | 1.8311 | 1.6663 | 1.9043
1628.1719 | 122.4243 | -42.9291 | 79.1473 | -0.18394 | 0.021422 | 0.086917 | 0.091563 | 0.00039951 | 0.017256 | 0.01051 | -0.089813 | 0.018545 | 0.08671 | 0.09139 | 0.00045928 | 0.0072055 | 0.012405 | -0.054932 | 79.1428 | -43.0664 | 79.0695 | -35.7513 | -40.7318 | -41.4368 | -0.022888 | 79.1611 | -42.9794 | 79.248 | -36.85 | -40.9744 | -41.2262 | -0.13733 | -0.004578 | -0.07782 | -0.073242 | 7.1777 | 7.3151 | 1.4923 | 2.1973 | 1.6296 | 2.3346 | -0.050354 | 0.013733 | 0.10071 | 0.086975 | 6.0791 | 6.1295 | 1.7029 | 1.9547 | 1.7532 | 2.005
1629.1719 | 122.4197 | -43.0206 | 79.0741 | -0.1992 | 0.023131 | 0.087558 | 0.090495 | 0.0012539 | 0.016829 | 0.0092286 | -0.10202 | 0.019826 | 0.087137 | 0.086904 | -0.00018152 | 0.0089145 | 0.0092007 | -0.068665 | 79.0695 | -43.1442 | 78.9871 | -35.8109 | -40.8051 | -41.5558 | -0.018311 | 79.0558 | -43.0618 | 79.1702 | -37.0285 | -41.0477 | -41.3132 | -0.1236 | -0.004578 | -0.086975 | -0.082397 | 7.2098 | 7.3334 | 1.4648 | 2.2156 | 1.5884 | 2.3392 | -0.041199 | -0.018311 | 0.09613 | 0.11444 | 5.9921 | 6.0333 | 1.7075 | 1.973 | 1.7487 | 2.0142
1630.1719 | 122.3785 | -42.9153 | 79.1702 | -0.19157 | 0.020995 | 0.086062 | 0.086222 | 0.00018591 | 0.017684 | 0.014355 | -0.062347 | 0.019185 | 0.08671 | 0.089467 | 0.0013139 | 0.0084873 | 0.01091 | -0.059509 | 79.1382 | -43.0481 | 79.0695 | -35.6552 | -40.6723 | -41.4139 | 0.009155 | 79.1519 | -42.9794 | 79.2664 | -36.9827 | -40.9378 | -41.2354 | -0.13275 | -0.032043 | -0.10071 | -0.068665 | 7.2601 | 7.3929 | 1.5015 | 2.243 | 1.6342 | 2.3758 | -0.064087 | -0.01831 | 0.09613 | 0.11444 | 5.9326 | 5.9967 | 1.68 | 1.9775 | 1.7441 | 2.0416
1631.1719 | 122.374 | -42.9886 | 79.1061 | -0.11528 | 0.018645 | 0.087771 | 0.091563 | 0.0021085 | 0.01747 | 0.012006 | -0.098968 | 0.020467 | 0.083292 | 0.093526 | 0.00067288 | 0.01041 | 0.014114 | -0.10071 | 79.1107 | -43.1625 | 78.923 | -35.8292 | -40.8371 | -41.5787 | -0.09613 | 79.0878 | -43.0481 | 79.1428 | -37.1384 | -41.0614 | -41.3361 | -0.17395 | 0.004578 | -0.1831 | -0.18768 | 7.1594 | 7.3334 | 1.4099 | 2.1515 | 1.5839 | 2.3254 | -0.05951 | -0.01831 | 0.036621 | 0.054931 | 5.8502 | 5.9097 | 1.6525 | 1.9272 | 1.712 | 1.9867
1632.1719 | 122.4014 | -43.0115 | 79.0283 | -0.16869 | 0.022277 | 0.087344 | 0.091136 | 0.00018591 | 0.01747 | 0.0096558 | -0.074554 | 0.018331 | 0.08735 | 0.089467 | -0.00039512 | 0.010837 | 0.0066373 | -0.054932 | 79.0283 | -43.1717 | 78.978 | -35.7697 | -40.7822 | -41.5375 | -0.21973 | 78.9825 | -43.1534 | 79.1199 | -37.1475 | -41.0294 | -41.3269 | -0.16022 | 0 | -0.050354 | -0.050354 | 7.2418 | 7.402 | 1.474 | 2.2293 | 1.6342 | 2.3895 | -0.14191 | -0.045776 | 0.091553 | 0.13733 | 5.864 | 6.0059 | 1.6846 | 1.9821 | 1.8265 | 2.124
1633.1719 | 122.3099 | -43.222 | 78.9276 | -0.15343 | 0.018004 | 0.087131 | 0.088359 | -0.0017367 | 0.017684 | 0.011792 | -0.095917 | 0.020894 | 0.087137 | 0.093313 | 0.0015275 | 0.0087009 | 0.010055 | -0.013733 | 78.9459 | -43.2999 | 78.6896 | -35.9482 | -40.9973 | -41.8121 | -0.12817 | 78.923 | -43.2495 | 78.9093 | -37.413 | -41.2766 | -41.5512 | -0.07782 | 0.018311 | -0.23804 | -0.25635 | 7.2739 | 7.3517 | 1.4099 | 2.2247 | 1.4877 | 2.3026 | -0.027466 | -0.004577 | -0.01831 | -0.013733 | 5.809 | 5.8365 | 1.6708 | 1.9455 | 1.6983 | 1.973
1634.1719 | 122.4197 | -42.9977 | 79.0604 | -0.11528 | 0.022063 | 0.085635 | 0.08665 | -0.0021641 | 0.014052 | 0.010724 | 0.023102 | 0.018331 | 0.086069 | 0.089895 | 0.0017411 | 0.0080601 | 0.0077053 | -0.004578 | 79.1061 | -43.0801 | 79.0192 | -35.6552 | -40.6448 | -41.4459 | 0.14191 | 79.0787 | -43.0161 | 79.1931 | -37.1384 | -40.9103 | -41.2628 | -0.082397 | 0.045776 | -0.041199 | -0.086975 | 7.3425 | 7.4249 | 1.5518 | 2.3529 | 1.6342 | 2.4353 | -0.01831 | 0.01831 | 0.13275 | 0.11444 | 5.8594 | 5.8777 | 1.7349 | 2.0874 | 1.7532 | 2.1057
1635.1719 | 122.406 | -43.1671 | 78.8956 | -0.16106 | 0.02014 | 0.086917 | 0.090068 | -0.0015231 | 0.01512 | 0.010083 | -0.10507 | 0.019185 | 0.084146 | 0.093526 | -0.0016769 | 0.010196 | 0.011764 | -0.10071 | 78.9368 | -43.3365 | 78.7811 | -35.9161 | -40.9378 | -41.7114 | -0.064087 | 78.9001 | -43.2037 | 79.0192 | -37.4451 | -41.1163 | -41.4505 | -0.16937 | 0.041199 | -0.11444 | -0.15564 | 7.251 | 7.4203 | 1.4557 | 2.2293 | 1.6251 | 2.3987 | -0.036621 | 0.004577 | 0.1236 | 0.11902 | 5.722 | 5.7587 | 1.7166 | 2.0508 | 1.7532 | 2.0874
1636.1719 | 122.3236 | -43.3777 | 78.7399 | -0.13817 | 0.018431 | 0.086276 | 0.089854 | -0.00045509 | 0.017897 | 0.013074 | -0.098968 | 0.021108 | 0.084573 | 0.088186 | 0.00088648 | 0.0089145 | 0.013687 | -0.10529 | 78.7354 | -43.5242 | 78.4653 | -35.9344 | -40.9882 | -42.0502 | -0.07782 | 78.7811 | -43.4326 | 78.7125 | -37.7426 | -41.4093 | -41.7297 | -0.14648 | -0.004577 | -0.27466 | -0.27008 | 7.4432 | 7.5897 | 1.3275 | 2.3895 | 1.474 | 2.536 | -0.054931 | 0.041199 | -0.027466 | -0.068665 | 5.6351 | 5.69 | 1.6479 | 1.9684 | 1.7029 | 2.0233
1637.1719 | 122.2687 | -43.4784 | 78.6575 | -0.15343 | 0.02014 | 0.085635 | 0.092418 | -0.00024129 | 0.016188 | 0.0077332 | -0.12643 | 0.020254 | 0.087778 | 0.092031 | 0.00045928 | 0.0093417 | 0.0074917 | -0.082397 | 78.6621 | -43.5974 | 78.447 | -36.2137 | -41.272 | -42.0547 | -0.11444 | 78.6575 | -43.5745 | 78.6163 | -37.8525 | -41.5054 | -41.835 | -0.11902 | 0.004577 | -0.21057 | -0.21515 | 7.2647 | 7.3837 | 1.4236 | 2.2064 | 1.5427 | 2.3254 | -0.09613 | 0 | -0.041199 | -0.041199 | 5.6259 | 5.722 | 1.6434 | 1.973 | 1.7395 | 2.0691
1638.1719 | 122.2916 | -43.6615 | 78.479 | -0.13054 | 0.020568 | 0.086276 | 0.091136 | -0.0034457 | 0.015761 | 0.011365 | -0.098968 | 0.020254 | 0.085855 | 0.086263 | -0.00018152 | 0.010837 | 0.0096279 | -0.050354 | 78.4698 | -43.7988 | 78.2639 | -36.351 | -41.4139 | -42.215 | -0.068665 | 78.4653 | -43.7073 | 78.4561 | -38.0173 | -41.6473 | -41.9861 | -0.13733 | -0.009155 | -0.21515 | -0.20599 | 7.3105 | 7.4478 | 1.4465 | 2.2476 | 1.5839 | 2.3849 | -0.045776 | -0.013733 | -0.022888 | -0.009155 | 5.6442 | 5.69 | 1.6754 | 2.0142 | 1.7212 | 2.0599
1639.1719 | 122.4289 | -43.6157 | 78.4561 | -0.16869 | 0.020781 | 0.087558 | 0.091563 | 0.0012539 | 0.015334 | 0.011578 | -0.044037 | 0.02004 | 0.087137 | 0.085409 | -0.00039512 | 0.0095555 | 0.010696 | -0.07782 | 78.4195 | -43.7714 | 78.3646 | -36.2274 | -41.3406 | -42.1371 | -0.013733 | 78.4241 | -43.6707 | 78.5477 | -37.9166 | -41.4871 | -41.8396 | -0.15564 | -0.036621 | -0.091553 | -0.054932 | 7.3883 | 7.5439 | 1.4786 | 2.2751 | 1.6342 | 2.4307 | -0.054931 | -0.032044 | 0.091552 | 0.1236 | 5.6992 | 5.7541 | 1.7761 | 2.1286 | 1.8311 | 2.1835
1640.1719 | 122.3236 | -43.7576 | 78.3554 | -0.1458 | 0.020354 | 0.086917 | 0.090281 | -0.0060093 | 0.014907 | 0.01286 | -0.10507 | 0.022176 | 0.085214 | 0.092244 | 0.00067288 | 0.012332 | 0.0094143 | -0.022888 | 78.4241 | -43.8492 | 78.1952 | -36.3464 | -41.4642 | -42.3019 | -0.009155 | 78.3691 | -43.7988 | 78.3646 | -38.1271 | -41.6794 | -42.0456 | -0.091553 | 0.068664 | -0.16022 | -0.22888 | 7.4112 | 7.5027 | 1.4557 | 2.2934 | 1.5472 | 2.3849 | -0.041199 | 0.013733 | 0.009155 | -0.004578 | 5.6305 | 5.6717 | 1.712 | 2.0782 | 1.7532 | 2.1194
1641.1719 | 122.2733 | -43.9499 | 78.2043 | -0.10002 | 0.020781 | 0.087771 | 0.090281 | -0.0010959 | 0.012557 | 0.0098694 | -0.092865 | 0.020254 | 0.086069 | 0.090322 | 0.0017411 | 0.013187 | 0.012832 | -0.07782 | 78.186 | -44.1055 | 77.948 | -36.6119 | -41.7252 | -42.5537 | -0.15106 | 78.1815 | -44.0094 | 78.1174 | -38.3972 | -41.8854 | -42.2653 | -0.15564 | -0.018311 | -0.25635 | -0.23804 | 7.338 | 7.4936 | 1.3962 | 2.2247 | 1.5518 | 2.3804 | -0.059509 | -0.022888 | -0.086975 | -0.064087 | 5.5527 | 5.6122 | 1.6846 | 2.0645 | 1.7441 | 2.124
1642.1719 | 122.2733 | -44.0735 | 78.0807 | -0.046616 | 0.021849 | 0.08649 | 0.08964 | 0.0031765 | 0.016616 | 0.010297 | -0.068451 | 0.02004 | 0.08671 | 0.090963 | -0.00018152 | 0.0095555 | 0.010696 | -0.041199 | 78.0853 | -44.1833 | 77.8839 | -36.5982 | -41.7709 | -42.5995 | -0.022888 | 78.0762 | -44.1193 | 78.0441 | -38.4705 | -41.9449 | -42.3752 | -0.10986 | 0.004577 | -0.19684 | -0.20142 | 7.4753 | 7.5851 | 1.474 | 2.3026 | 1.5839 | 2.4124 | -0.045777 | -0.004578 | -0.036622 | -0.032044 | 5.603 | 5.6488 | 1.6983 | 2.1286 | 1.7441 | 2.1744
1643.1719 | 122.319 | -44.014 | 78.0991 | -0.1458 | 0.020354 | 0.088199 | 0.090922 | -0.0019503 | 0.016616 | 0.0085876 | -0.056244 | 0.020894 | 0.083505 | 0.088613 | 0.00024568 | 0.011692 | 0.010055 | 0.022888 | 78.0991 | -44.165 | 77.9343 | -36.4563 | -41.6977 | -42.6086 | 0.045776 | 78.1036 | -44.0552 | 78.1082 | -38.3881 | -41.8579 | -42.2836 | -0.15106 | 0 | -0.16479 | -0.16479 | 7.5577 | 7.7087 | 1.4053 | 2.3163 | 1.5564 | 2.4673 | -0.041199 | 0.004578 | 0.009155 | 0.004577 | 5.6259 | 5.6671 | 1.7303 | 2.1561 | 1.7715 | 2.1973
1644.1719 | 122.3053 | -43.927 | 78.2043 | -0.26024 | 0.021422 | 0.086917 | 0.088359 | 0.0031765 | 0.017256 | 0.010297 | -0.059296 | 0.020254 | 0.088419 | 0.088399 | 0.00024568 | 0.0091281 | 0.012619 | -0.09613 | 78.2547 | -44.0369 | 77.9755 | -36.4151 | -41.6245 | -42.5217 | -0.13275 | 78.2227 | -43.959 | 78.1723 | -38.3148 | -41.7801 | -42.1921 | -0.10986 | 0.050354 | -0.22888 | -0.27924 | 7.5119 | 7.6218 | 1.4053 | 2.3026 | 1.5152 | 2.4124 | -0.032043 | 0.01831 | -0.032044 | -0.050354 | 5.6122 | 5.6442 | 1.7349 | 2.1469 | 1.767 | 2.179
1645.1719 | 122.3557 | -43.9636 | 78.1998 | -0.069504 | 0.02014 | 0.087344 | 0.08964 | -0.0010959 | 0.017897 | 0.01051 | -0.068451 | 0.023031 | 0.087137 | 0.08904 | 0.00045928 | 0.0097691 | 0.0092007 | -0.013733 | 78.1631 | -44.1238 | 77.9846 | -36.3327 | -41.6153 | -42.5262 | 0.11902 | 78.2043 | -43.8904 | 78.1815 | -38.2782 | -41.7572 | -42.2012 | -0.16022 | -0.036621 | -0.21515 | -0.17853 | 7.6309 | 7.7911 | 1.4374 | 2.3483 | 1.5976 | 2.5085 | 0.073242 | 0.004578 | -0.01831 | -0.022888 | 5.6854 | 5.6122 | 1.7624 | 2.2064 | 1.6891 | 2.1332
1646.1719 | 122.2733 | -44.4855 | 77.655 | -0.12291 | 0.021422 | 0.087771 | 0.092631 | -0.0025913 | 0.014479 | 0.0094422 | -0.077606 | 0.022176 | 0.086282 | 0.094381 | 0.00024568 | 0.010623 | 0.010696 | -0.068665 | 77.6825 | -44.5908 | 77.3895 | -37.7747 | -42.16 | -43.1305 | -0.12817 | 77.6871 | -44.5221 | 77.6001 | -38.8321 | -42.3157 | -42.7414 | -0.10528 | 0.027466 | -0.2655 | -0.29297 | 6.7108 | 6.8161 | 1.355 | 2.3254 | 1.4603 | 2.4307 | -0.036621 | 0.032044 | -0.054931 | -0.086975 | 5.6534 | 5.69 | 1.7441 | 2.1698 | 1.7807 | 2.2064
1647.1719 | 122.3282 | -44.2017 | 77.9434 | -0.10002 | 0.02014 | 0.08884 | 0.091136 | -0.0017367 | 0.016188 | 0.010297 | -0.059296 | 0.022817 | 0.086069 | 0.086477 | 0.00045928 | 0.011051 | 0.013901 | 0.022888 | 77.9297 | -44.3481 | 77.7603 | -36.7447 | -41.8442 | -42.8329 | 0.059509 | 77.9297 | -44.22 | 77.9205 | -38.4705 | -41.9678 | -42.421 | -0.14649 | -0.013733 | -0.1831 | -0.16937 | 7.457 | 7.6035 | 1.3687 | 2.3575 | 1.5152 | 2.504 | -0.018311 | -0.013733 | -0.022888 | -0.009155 | 5.7312 | 5.7495 | 1.7807 | 2.2339 | 1.799 | 2.2522
1648.1719 | 122.2824 | -44.5221 | 77.6367 | -0.13054 | 0.021422 | 0.087985 | 0.090281 | -0.0017367 | 0.016829 | 0.011365 | -0.068451 | 0.022817 | 0.08671 | 0.089681 | 0.0013139 | 0.010196 | 0.013473 | -0.10529 | 77.6779 | -44.6411 | 77.4124 | -37.3169 | -42.1875 | -43.158 | -0.082397 | 77.6184 | -44.545 | 77.5726 | -38.7726 | -42.3065 | -42.7551 | -0.11902 | 0.041198 | -0.2243 | -0.2655 | 7.2052 | 7.3242 | 1.3641 | 2.3346 | 1.4832 | 2.4536 | -0.022888 | -0.018311 | -0.064087 | -0.045776 | 5.7495 | 5.7724 | 1.767 | 2.2156 | 1.7899 | 2.2385
1649.1719 | 122.374 | -44.5084 | 77.5864 | -0.1458 | 0.021209 | 0.087558 | 0.08964 | 0.00018591 | 0.015334 | 0.011792 | -0.040985 | 0.022603 | 0.087991 | 0.088827 | 0.00067288 | 0.013828 | 0.014541 | -0.10986 | 77.6321 | -44.664 | 77.4033 | -37.2986 | -42.1967 | -43.1625 | -0.082397 | 77.5772 | -44.5633 | 77.6047 | -38.7039 | -42.2424 | -42.6956 | -0.15564 | 0.045776 | -0.18311 | -0.22888 | 7.2098 | 7.3654 | 1.3458 | 2.3117 | 1.5015 | 2.4673 | -0.054931 | -0.009156 | 0.01831 | 0.027466 | 5.8044 | 5.8594 | 1.8127 | 2.2659 | 1.8677 | 2.3209
1650.1719 | 122.3007 | -44.5038 | 77.6367 | -0.13054 | 0.02014 | 0.088412 | 0.08964 | 0.00082671 | 0.01747 | 0.010297 | -0.12033 | 0.020467 | 0.086282 | 0.084981 | -0.0016769 | 0.012973 | 0.013046 | -0.10529 | 77.623 | -44.6457 | 77.3849 | -37.2986 | -42.2104 | -43.3044 | 0.022888 | 77.655 | -44.5358 | 77.5543 | -38.7131 | -42.2836 | -42.746 | -0.14191 | -0.013733 | -0.25177 | -0.23804 | 7.2052 | 7.3471 | 1.1993 | 2.2934 | 1.3412 | 2.4353 | -0.032044 | 0.01831 | -0.082398 | -0.10071 | 5.7907 | 5.8228 | 1.7578 | 2.2202 | 1.7899 | 2.2522
1651.1719 | 122.3373 | -44.6136 | 77.5314 | -0.1458 | 0.02249 | 0.088626 | 0.08964 | 0.0023221 | 0.01747 | 0.011365 | -0.03183 | 0.022817 | 0.085641 | 0.093953 | 0.00088648 | 0.010623 | 0.010696 | -0.013733 | 77.5726 | -44.7464 | 77.3254 | -37.3032 | -42.2104 | -43.222 | 0.032043 | 77.5131 | -44.6365 | 77.4857 | -38.7268 | -42.2928 | -42.7917 | -0.13275 | 0.041199 | -0.20599 | -0.24719 | 7.3105 | 7.4432 | 1.3916 | 2.4033 | 1.5244 | 2.536 | -0.022889 | -0.01831 | -0.045776 | -0.027466 | 5.8868 | 5.9097 | 1.8219 | 2.3209 | 1.8448 | 2.3438
1652.1719 | 122.3511 | -44.6777 | 77.4353 | -0.11528 | 0.020995 | 0.086062 | 0.090922 | 0.00082671 | 0.015761 | 0.0085876 | -0.040985 | 0.022176 | 0.085641 | 0.093313 | -0.00018152 | 0.011478 | 0.011978 | 0.009155 | 77.4445 | -44.8471 | 77.2659 | -37.3489 | -42.3111 | -43.3182 | 0.045776 | 77.4124 | -44.7327 | 77.4261 | -38.8184 | -42.3706 | -42.8741 | -0.16937 | 0.009155 | -0.16937 | -0.17853 | 7.3288 | 7.4982 | 1.3596 | 2.3666 | 1.5289 | 2.536 | -0.054932 | -0.022888 | -0.009156 | 0.013732 | 5.8594 | 5.9143 | 1.8036 | 2.3071 | 1.8585 | 2.3621
1653.1719 | 122.3145 | -44.6915 | 77.4399 | -0.11528 | 0.019713 | 0.086917 | 0.092845 | 0.0018949 | 0.016402 | 0.0085876 | -0.053192 | 0.022817 | 0.08436 | 0.090108 | 3.2078e-05 | 0.011051 | 0.0092007 | 0.009155 | 77.4673 | -44.8059 | 77.2705 | -37.3169 | -42.2836 | -43.2861 | 0.059509 | 77.4124 | -44.7006 | 77.449 | -38.7451 | -42.3386 | -42.8696 | -0.11444 | 0.027466 | -0.16937 | -0.19684 | 7.3746 | 7.489 | 1.4053 | 2.4078 | 1.5198 | 2.5223 | -0.009156 | -0.027465 | 0.009156 | 0.036621 | 5.9464 | 5.9555 | 1.8219 | 2.3529 | 1.8311 | 2.3621
1654.1719 | 122.3923 | -44.7418 | 77.3438 | -0.16106 | 0.020354 | 0.087344 | 0.089 | -0.0028049 | 0.015548 | 0.0081604 | 0.032257 | 0.023458 | 0.086282 | 0.090535 | 0.00024568 | 0.010837 | 0.010696 | -0.054932 | 77.3209 | -44.8563 | 77.2293 | -37.3032 | -42.334 | -43.2999 | 0.004578 | 77.3117 | -44.783 | 77.3804 | -38.7589 | -42.3614 | -42.8879 | -0.11444 | -0.022888 | -0.11444 | -0.091553 | 7.4387 | 7.5531 | 1.442 | 2.4078 | 1.5564 | 2.5223 | -0.041199 | -0.032043 | 0.036621 | 0.068664 | 5.983 | 6.0242 | 1.8539 | 2.3804 | 1.8951 | 2.4216
1655.1719 | 122.3282 | -44.838 | 77.2751 | -0.13817 | 0.021209 | 0.088412 | 0.090281 | -0.0030185 | 0.016829 | 0.011151 | -0.037933 | 0.02239 | 0.08735 | 0.098226 | 3.2078e-05 | 0.011692 | 0.0077053 | -0.009155 | 77.33 | -44.9341 | 77.1423 | -37.3947 | -42.3981 | -43.4418 | 0 | 77.2888 | -44.87 | 77.2842 | -38.8596 | -42.4667 | -42.9886 | -0.09613 | 0.054932 | -0.13275 | -0.18768 | 7.4432 | 7.5394 | 1.3962 | 2.4399 | 1.4923 | 2.536 | -0.032043 | 0.013733 | 0.009156 | -0.004577 | 5.9784 | 6.0104 | 1.8494 | 2.3712 | 1.8814 | 2.4033
1656.1719 | 122.3099 | -44.6686 | 77.4536 | -0.12291 | 0.020781 | 0.087771 | 0.095408 | 0.0018949 | 0.015548 | 0.010297 | -0.071503 | 0.021749 | 0.085001 | 0.094167 | -0.00039512 | 0.011478 | 0.012405 | -0.018311 | 77.449 | -44.8196 | 77.298 | -37.2803 | -42.2653 | -43.277 | 0.004578 | 77.4628 | -44.7235 | 77.4261 | -38.6993 | -42.3111 | -42.8375 | -0.15106 | -0.004577 | -0.15564 | -0.15106 | 7.3883 | 7.5394 | 1.3916 | 2.4033 | 1.5427 | 2.5543 | -0.054932 | 0.009156 | -0.027466 | -0.036622 | 5.9692 | 6.0242 | 1.8311 | 2.3575 | 1.886 | 2.4124
1657.1719 | 122.3236 | -44.7418 | 77.3987 | -0.1458 | 0.020568 | 0.087558 | 0.090281 | -0.0032321 | 0.018752 | 0.0075196 | -0.062347 | 0.021963 | 0.084787 | 0.08904 | -0.00060892 | 0.011692 | 0.011551 | -0.009155 | 77.3895 | -44.8792 | 77.2293 | -37.3306 | -42.2974 | -43.3594 | 0.036621 | 77.3987 | -44.751 | 77.3575 | -38.736 | -42.3294 | -42.8741 | -0.13733 | -0.009156 | -0.16937 | -0.16022 | 7.4112 | 7.5485 | 1.3824 | 2.4445 | 1.5198 | 2.5818 | -0.009156 | 0 | -0.041199 | -0.041199 | 6.0059 | 6.015 | 1.8677 | 2.4124 | 1.8768 | 2.4216
1658.1719 | 122.319 | -44.9661 | 77.1698 | -0.10765 | 0.019713 | 0.087131 | 0.090709 | 0.0018949 | 0.018965 | 0.011578 | -0.059296 | 0.023244 | 0.087991 | 0.093313 | -0.00039512 | 0.012332 | 0.0089871 | 0 | 77.1378 | -45.1355 | 76.9821 | -37.5229 | -42.5629 | -43.6111 | 0.054932 | 77.1423 | -45.0027 | 77.1286 | -38.9236 | -42.5491 | -43.1168 | -0.16937 | -0.032044 | -0.18768 | -0.15564 | 7.4432 | 7.6126 | 1.355 | 2.4033 | 1.5244 | 2.5726 | -0.036622 | -0.027466 | -0.041199 | -0.013733 | 6.0425 | 6.0791 | 1.8494 | 2.417 | 1.886 | 2.4536
1659.1719 | 122.3648 | -45.1218 | 76.9821 | -0.16106 | 0.02014 | 0.089267 | 0.091136 | -0.00066869 | 0.015761 | 0.007306 | -0.025726 | 0.023031 | 0.087564 | 0.092031 | 0.00067288 | 0.013401 | 0.010269 | -0.10986 | 76.9867 | -45.3003 | 76.7761 | -37.6831 | -42.7597 | -43.7759 | -0.068665 | 76.9363 | -45.1721 | 76.9821 | -39.0747 | -42.691 | -43.2358 | -0.17853 | 0.004577 | -0.20599 | -0.21057 | 7.4387 | 7.6172 | 1.3458 | 2.3621 | 1.5244 | 2.5406 | -0.050354 | -0.045777 | 0 | 0.045777 | 6.0471 | 6.0974 | 1.886 | 2.4307 | 1.9363 | 2.4811
1660.1719 | 122.3236 | -45.0531 | 77.0691 | -0.069504 | 0.021849 | 0.086703 | 0.08964 | 0.0023221 | 0.017043 | 0.0094422 | -0.040985 | 0.022603 | 0.08671 | 0.092031 | -0.0018905 | 0.012332 | 0.012405 | 0 | 77.0645 | -45.2087 | 76.8997 | -37.5504 | -42.627 | -43.5837 | 0.041199 | 77.1194 | -45.0531 | 77.0599 | -39.0015 | -42.6041 | -43.1808 | -0.15564 | -0.004578 | -0.16937 | -0.16479 | 7.5027 | 7.6584 | 1.4694 | 2.4261 | 1.6251 | 2.5818 | 0 | 0.050354 | -0.009155 | -0.059509 | 6.0516 | 6.0516 | 1.8723 | 2.449 | 1.8723 | 2.449
1661.1719 | 122.3465 | -45.2316 | 76.8723 | -0.13817 | 0.020995 | 0.088199 | 0.090281 | -0.0015231 | 0.01747 | 0.010297 | -0.065399 | 0.021108 | 0.086496 | 0.088827 | -0.0012497 | 0.013401 | 0.012405 | 0 | 76.8997 | -45.3873 | 76.7029 | -37.7014 | -42.8055 | -43.8721 | 0.009155 | 76.9592 | -45.2362 | 76.8585 | -39.1663 | -42.7872 | -43.3685 | -0.15564 | 0.027466 | -0.16937 | -0.19684 | 7.5302 | 7.6859 | 1.3596 | 2.4261 | 1.5152 | 2.5818 | -0.004578 | 0.086976 | -0.013732 | -0.10071 | 6.0654 | 6.0699 | 1.8631 | 2.4445 | 1.8677 | 2.449
1662.1719 | 122.3145 | -45.2545 | 76.8402 | -0.061874 | 0.020995 | 0.087131 | 0.088786 | -0.00024129 | 0.017684 | 0.013287 | -0.056244 | 0.021963 | 0.086923 | 0.093953 | -0.0012497 | 0.011264 | 0.01091 | 0.004578 | 76.8906 | -45.4102 | 76.6846 | -37.7518 | -42.8192 | -43.895 | 0.15106 | 76.8997 | -45.2454 | 76.8768 | -39.1434 | -42.7826 | -43.3685 | -0.15564 | 0.050354 | -0.15564 | -0.20599 | 7.5027 | 7.6584 | 1.3596 | 2.4353 | 1.5152 | 2.5909 | 0.009156 | 0.059509 | 0.036621 | -0.022888 | 6.1111 | 6.102 | 1.886 | 2.4719 | 1.8768 | 2.4628
1663.1719 | 122.287 | -45.3964 | 76.7532 | -0.11528 | 0.021422 | 0.089053 | 0.091563 | 0.0021085 | 0.015761 | 0.012219 | -0.059296 | 0.023671 | 0.085855 | 0.088186 | 0.0013139 | 0.014255 | 0.010269 | -0.032043 | 76.7303 | -45.5017 | 76.5335 | -37.8479 | -42.952 | -44.046 | -0.022888 | 76.712 | -45.4468 | 76.68 | -39.2899 | -42.9199 | -43.5287 | -0.10529 | -0.022888 | -0.21973 | -0.19684 | 7.5485 | 7.6538 | 1.3504 | 2.4445 | 1.4557 | 2.5497 | -0.050354 | -0.041199 | -0.073242 | -0.032043 | 6.1066 | 6.1569 | 1.8677 | 2.4765 | 1.918 | 2.5269
1664.1875 | 122.3831 | -45.3094 | 76.7715 | -0.0084687 | 0.018859 | 0.087985 | 0.088572 | -0.0055819 | 0.015975 | 0.010083 | 0.10245 | 0.024312 | 0.086496 | 0.090108 | 0.0015275 | 0.011478 | 0.015823 | -0.086975 | 76.7029 | -45.4834 | 76.6251 | -37.7792 | -42.8513 | -43.9087 | 0.009155 | 76.7029 | -45.369 | 76.7944 | -39.1113 | -42.7414 | -43.3914 | -0.17395 | -0.068664 | -0.14648 | -0.07782 | 7.5302 | 7.7042 | 1.4008 | 2.4582 | 1.5747 | 2.6321 | -0.05951 | -0.068664 | 0.022889 | 0.091553 | 6.1981 | 6.2576 | 1.918 | 2.5681 | 1.9775 | 2.6276
1665.1719 | 122.3053 | -45.369 | 76.799 | -0.10765 | 0.020781 | 0.085208 | 0.091777 | -0.0019503 | 0.016616 | 0.012219 | -0.05014 | 0.025594 | 0.088419 | 0.098012 | -0.00060892 | 0.012546 | 0.011764 | -0.004578 | 76.7166 | -45.5566 | 76.5152 | -37.8571 | -42.9611 | -44.0598 | 0.054932 | 76.6708 | -45.4605 | 76.6983 | -39.2212 | -42.8925 | -43.2816 | -0.18768 | -0.082397 | -0.28381 | -0.20142 | 7.5119 | 7.6996 | 1.3092 | 2.4078 | 1.4969 | 2.5955 | -0.091552 | -0.12817 | -0.10071 | 0.027466 | 6.1478 | 6.2393 | 2.0874 | 2.4765 | 2.179 | 2.5681
1666.1719 | 122.3053 | -45.2866 | 76.8677 | -0.00083923 | 0.02014 | 0.086703 | 0.089427 | 0.00039951 | 0.01747 | 0.0085876 | -0.068451 | 0.023244 | 0.086496 | 0.094381 | -0.0014633 | 0.014896 | 0.011551 | -0.054932 | 76.886 | -45.369 | 76.6708 | -37.7335 | -42.8192 | -43.927 | -0.018311 | 76.8219 | -45.3232 | 76.8036 | -39.0839 | -42.7734 | -43.4006 | -0.082398 | 0.01831 | -0.19684 | -0.21515 | 7.5531 | 7.6355 | 1.3596 | 2.4673 | 1.442 | 2.5497 | -0.036621 | -0.045777 | -0.064087 | -0.01831 | 6.2027 | 6.2393 | 1.886 | 2.5131 | 1.9226 | 2.5497
1667.1719 | 122.3099 | -45.2866 | 76.8356 | -0.13817 | 0.019927 | 0.083926 | 0.09199 | 0.0036039 | 0.015334 | 0.0098694 | -0.062347 | 0.02474 | 0.087137 | 0.092672 | -0.0010361 | 0.011264 | 0.014328 | 0.054932 | 76.8814 | -45.3781 | 76.6708 | -37.7151 | -42.8146 | -43.9133 | 0.004578 | 76.8173 | -45.3323 | 76.799 | -39.0747 | -42.7505 | -43.4006 | -0.091553 | 0.045777 | -0.16479 | -0.21057 | 7.5714 | 7.663 | 1.3733 | 2.4719 | 1.4648 | 2.5635 | -0.045776 | -0.01831 | -0.036621 | -0.018311 | 6.2119 | 6.2576 | 1.886 | 2.536 | 1.9318 | 2.5818
1668.1719 | 122.3465 | -45.2591 | 76.8723 | -0.092392 | 0.021849 | 0.086703 | 0.09199 | 0.0021085 | 0.015548 | 0.010083 | -0.08371 | 0.02239 | 0.08671 | 0.094167 | -0.00039512 | 0.012973 | 0.010055 | -0.013733 | 76.8494 | -45.401 | 76.6891 | -37.6739 | -42.7963 | -43.8812 | 0.013733 | 76.8814 | -45.2957 | 76.8402 | -38.9923 | -42.691 | -43.3548 | -0.14191 | -0.022888 | -0.1831 | -0.16022 | 7.5851 | 7.7271 | 1.3779 | 2.4628 | 1.5198 | 2.6047 | -0.036621 | 0.009156 | -0.032043 | -0.041199 | 6.2668 | 6.3034 | 1.9043 | 2.5681 | 1.9409 | 2.6047
1669.1875 | 122.2961 | -45.2957 | 76.8402 | -0.061874 | 0.021209 | 0.085635 | 0.090709 | -0.0015231 | 0.013839 | 0.0094422 | -0.025726 | 0.023885 | 0.086069 | 0.088613 | -0.00018152 | 0.0099827 | 0.012405 | -0.041199 | 76.8723 | -45.4239 | 76.5747 | -37.8708 | -42.8375 | -43.8995 | -0.022888 | 76.886 | -45.3049 | 76.7624 | -39.0015 | -42.7139 | -43.396 | -0.12817 | 0.032043 | -0.2655 | -0.29755 | 7.4249 | 7.5531 | 1.3962 | 2.4582 | 1.5244 | 2.5864 | -0.009156 | 0.045776 | -0.07782 | -0.1236 | 6.2942 | 6.3034 | 1.8997 | 2.5818 | 1.9089 | 2.5909
1670.1719 | 122.2687 | -45.3232 | 76.8311 | -0.10002 | 0.02014 | 0.087558 | 0.091136 | 0.0018949 | 0.017043 | 0.0098694 | -0.059296 | 0.022603 | 0.085428 | 0.09139 | -0.00039512 | 0.012119 | 0.012832 | -0.032043 | 76.8219 | -45.4239 | 76.6251 | -37.7426 | -42.8421 | -43.959 | -0.041199 | 76.8127 | -45.3461 | 76.7807 | -39.006 | -42.7414 | -43.4143 | -0.10071 | -0.009156 | -0.20599 | -0.19684 | 7.5806 | 7.6813 | 1.3641 | 2.4811 | 1.4648 | 2.5818 | -0.022888 | -0.018311 | -0.050354 | -0.032043 | 6.3171 | 6.34 | 1.9089 | 2.5818 | 1.9318 | 2.6047
1671.1719 | 122.3053 | -45.2637 | 76.8631 | -0.084763 | 0.019072 | 0.086062 | 0.091349 | -0.0036593 | 0.018538 | 0.011365 | -0.037933 | 0.025167 | 0.085855 | 0.09139 | 0.00088648 | 0.011692 | 0.011764 | 0 | 76.8494 | -45.401 | 76.6708 | -37.6831 | -42.8146 | -43.927 | 0.027466 | 76.8585 | -45.3278 | 76.8402 | -38.9282 | -42.627 | -43.3365 | -0.13733 | -0.013733 | -0.19226 | -0.17853 | 7.5806 | 7.7179 | 1.3367 | 2.449 | 1.474 | 2.5864 | -0.064087 | -0.004577 | -0.022888 | -0.018311 | 6.3354 | 6.3995 | 1.9272 | 2.6367 | 1.9913 | 2.7008
1672.1719 | 122.3328 | -45.1401 | 76.9501 | -0.10765 | 0.019072 | 0.087131 | 0.091136 | -2.7692e-05 | 0.017256 | 0.0070922 | -0.022675 | 0.024099 | 0.085428 | 0.090108 | -0.0014633 | 0.015537 | 0.013687 | -0.009155 | 76.9272 | -45.3049 | 76.7578 | -37.5824 | -42.6956 | -43.8126 | 0.041199 | 76.9409 | -45.2408 | 76.9135 | -38.7451 | -42.5308 | -43.2449 | -0.16479 | -0.022888 | -0.19226 | -0.16937 | 7.5577 | 7.7225 | 1.3275 | 2.4445 | 1.4923 | 2.6093 | -0.10071 | -0.009155 | -0.036621 | -0.027466 | 6.395 | 6.4957 | 1.8951 | 2.6093 | 1.9958 | 2.71
1673.1719 | 122.3099 | -45.2728 | 76.8814 | -0.16106 | 0.020781 | 0.088412 | 0.091777 | -0.0013095 | 0.016402 | 0.0092286 | -0.037933 | 0.025594 | 0.087991 | 0.094808 | -0.00039512 | 0.014896 | 0.012405 | -0.027466 | 76.8951 | -45.3827 | 76.6663 | -37.6648 | -42.8009 | -43.9178 | -0.013733 | 76.8539 | -45.314 | 76.8402 | -38.8641 | -42.6041 | -43.2999 | -0.10986 | 0.013733 | -0.21515 | -0.22888 | 7.608 | 7.7179 | 1.355 | 2.4719 | 1.4648 | 2.5818 | -0.041199 | -0.027466 | -0.041199 | -0.013733 | 6.4087 | 6.4499 | 1.973 | 2.6688 | 2.0142 | 2.71
1674.1719 | 122.4335 | -45.1355 | 76.8951 | -0.21446 | 0.0195 | 0.08884 | 0.089 | 0.0016813 | 0.018538 | 0.0090148 | 0.010895 | 0.024312 | 0.084573 | 0.08904 | -0.0021041 | 0.012119 | 0.011551 | -0.050354 | 76.9226 | -45.2683 | 76.8402 | -37.5229 | -42.627 | -43.7393 | 0.009155 | 76.9272 | -45.195 | 77.0462 | -38.7222 | -42.4118 | -43.1259 | -0.13275 | 0.027465 | -0.054932 | -0.082397 | 7.6126 | 7.7454 | 1.3962 | 2.5085 | 1.5289 | 2.6413 | -0.059509 | 0.032043 | 0.15106 | 0.11902 | 6.4133 | 6.4728 | 2.0096 | 2.7237 | 2.0691 | 2.7832
1675.1719 | 122.3282 | -45.3506 | 76.7807 | -0.11528 | 0.02014 | 0.085849 | 0.091349 | -0.00066869 | 0.018752 | 0.0064514 | -0.022675 | 0.025594 | 0.089487 | 0.094594 | -0.0012497 | 0.012546 | 0.011551 | -0.018311 | 76.799 | -45.4605 | 76.68 | -37.6831 | -42.8513 | -43.9682 | 0.004578 | 76.7899 | -45.3644 | 76.7487 | -38.887 | -42.6727 | -43.4006 | -0.10986 | 0.01831 | -0.10071 | -0.11902 | 7.6675 | 7.7774 | 1.3824 | 2.4994 | 1.4923 | 2.6093 | -0.013733 | 0.009155 | -0.032044 | -0.041199 | 6.4636 | 6.4774 | 1.9501 | 2.6779 | 1.9638 | 2.6917
1676.1719 | 122.2778 | -45.3094 | 76.8356 | -0.13054 | 0.020781 | 0.087558 | 0.090495 | 0.0023221 | 0.018538 | 0.011578 | -0.068451 | 0.025808 | 0.087137 | 0.093526 | -0.00018152 | 0.014469 | 0.013473 | -0.045776 | 76.8311 | -45.4193 | 76.6388 | -37.7655 | -42.8329 | -43.9819 | -0.036621 | 76.8082 | -45.3278 | 76.7624 | -38.8596 | -42.6361 | -43.3548 | -0.10986 | -0.004577 | -0.19684 | -0.19226 | 7.5439 | 7.6538 | 1.3275 | 2.4765 | 1.4374 | 2.5864 | -0.018311 | -0.027465 | -0.073242 | -0.045777 | 6.4499 | 6.4682 | 1.9547 | 2.6733 | 1.973 | 2.6917
1677.1719 | 122.319 | -45.5475 | 76.561 | -0.11528 | 0.020568 | 0.08948 | 0.090922 | -0.00066869 | 0.01512 | 0.0094422 | -0.065399 | 0.027517 | 0.085214 | 0.09908 | -0.0014633 | 0.011692 | 0.011978 | -0.032043 | 76.6159 | -45.6665 | 76.3916 | -38.0356 | -43.0801 | -44.2062 | -0.013733 | 76.5701 | -45.5841 | 76.5381 | -39.1159 | -42.8741 | -43.634 | -0.11902 | 0.054932 | -0.16937 | -0.2243 | 7.5119 | 7.6309 | 1.3412 | 2.4673 | 1.4603 | 2.5864 | -0.036621 | 0.009155 | -0.022888 | -0.032043 | 6.4316 | 6.4682 | 1.9135 | 2.6733 | 1.9501 | 2.71
1678.1719 | 122.3282 | -45.5566 | 76.5472 | -0.10765 | 0.021636 | 0.087985 | 0.090495 | -0.0017367 | 0.017256 | 0.0094422 | -0.022675 | 0.024099 | 0.087778 | 0.093313 | 0.00024568 | 0.012119 | 0.015823 | 0.009155 | 76.561 | -45.694 | 76.3641 | -38.1088 | -43.071 | -44.2245 | 0.036621 | 76.5381 | -45.6299 | 76.5427 | -39.0427 | -42.8192 | -43.6157 | -0.13733 | 0.013733 | -0.1831 | -0.19684 | 7.4478 | 7.5851 | 1.3321 | 2.4857 | 1.4694 | 2.623 | -0.073242 | -0.009155 | -0.004577 | 0.004578 | 6.514 | 6.5872 | 1.9409 | 2.7374 | 2.0142 | 2.8107
1679.1719 | 122.3053 | -45.7443 | 76.3962 | -0.15343 | 0.021636 | 0.088412 | 0.089427 | 0.00082671 | 0.017043 | 0.0098694 | -0.0074159 | 0.023885 | 0.086923 | 0.092458 | 0.0021683 | 0.014255 | 0.013687 | -0.013733 | 76.387 | -45.8542 | 76.1078 | -38.3789 | -43.2953 | -44.4855 | -0.10986 | 76.3824 | -45.8084 | 76.3184 | -39.2166 | -43.0481 | -43.8126 | -0.10986 | -0.009155 | -0.28839 | -0.27924 | 7.3654 | 7.4753 | 1.2588 | 2.449 | 1.3687 | 2.5589 | -0.064087 | -0.013733 | -0.07782 | -0.064087 | 6.5277 | 6.5918 | 1.9318 | 2.6962 | 1.9958 | 2.7603
1680.1719 | 122.3053 | -45.9824 | 76.0986 | -0.12291 | 0.02249 | 0.086703 | 0.090922 | -0.0015231 | 0.01747 | 0.012219 | -0.05014 | 0.023458 | 0.082223 | 0.092031 | -0.00018152 | 0.013614 | 0.01326 | -0.018311 | 76.1078 | -46.1609 | 75.9384 | -38.9099 | -43.5104 | -44.7052 | 0.036621 | 76.0666 | -46.0648 | 76.1169 | -39.3906 | -43.2541 | -44.0414 | -0.17853 | 0.009155 | -0.16022 | -0.16937 | 7.0724 | 7.251 | 1.2772 | 2.4719 | 1.4557 | 2.6505 | -0.082397 | -0.032044 | 0.01831 | 0.050354 | 6.5918 | 6.6742 | 1.9409 | 2.7283 | 2.0233 | 2.8107
1681.1719 | 122.3145 | -46.0419 | 76.0574 | -0.12291 | 0.021636 | 0.087985 | 0.090709 | 0.00039951 | 0.017897 | 0.0064514 | -0.05014 | 0.02474 | 0.08735 | 0.090322 | 3.2078e-05 | 0.012973 | 0.010483 | 0.041199 | 76.062 | -46.1929 | 75.8835 | -38.9923 | -43.5608 | -44.7464 | 0.045776 | 76.062 | -46.0968 | 76.0391 | -39.4363 | -43.277 | -44.0689 | -0.15106 | 0.004578 | -0.17395 | -0.17853 | 7.0496 | 7.2006 | 1.2955 | 2.4811 | 1.4465 | 2.6321 | -0.054932 | 0.004578 | -0.01831 | -0.022888 | 6.6055 | 6.6605 | 1.973 | 2.7649 | 2.0279 | 2.8198
1682.1719 | 122.3282 | -46.1563 | 75.943 | -0.12291 | 0.020995 | 0.087558 | 0.090709 | -0.00066869 | 0.01747 | 0.0079468 | -0.016571 | 0.025808 | 0.087137 | 0.089681 | 3.2078e-05 | 0.011692 | 0.012832 | 0.018311 | 75.943 | -46.3303 | 75.7736 | -39.1159 | -43.6478 | -44.8654 | 0.045776 | 75.9201 | -46.2067 | 75.9293 | -39.4913 | -43.3777 | -44.1696 | -0.17395 | 0 | -0.16937 | -0.16937 | 7.0404 | 7.2144 | 1.2909 | 2.5085 | 1.4648 | 2.6825 | -0.050354 | -0.022888 | -0.013733 | 0.009155 | 6.665 | 6.7154 | 1.9867 | 2.7786 | 2.037 | 2.829
1683.1719 | 122.3053 | -46.2341 | 75.8789 | -0.10002 | 0.022277 | 0.087344 | 0.093699 | -0.00024129 | 0.017684 | 0.010724 | -0.037933 | 0.025808 | 0.085214 | 0.092458 | -0.0023177 | 0.014682 | 0.01091 | -0.009155 | 75.8286 | -46.4081 | 75.6775 | -39.2487 | -43.7393 | -44.9753 | 0.036621 | 75.856 | -46.2753 | 75.8331 | -39.5554 | -43.4372 | -44.2749 | -0.17395 | -0.050354 | -0.20142 | -0.15106 | 6.9855 | 7.1594 | 1.2588 | 2.4948 | 1.4328 | 2.6688 | -0.041199 | -0.022888 | -0.045776 | -0.022888 | 6.6788 | 6.72 | 1.9592 | 2.7969 | 2.0004 | 2.8381
1684.1719 | 122.374 | -46.1197 | 75.9567 | -0.16106 | 0.021422 | 0.088412 | 0.090709 | -0.00088229 | 0.017684 | 0.010724 | -0.028778 | 0.023031 | 0.085641 | 0.086904 | -0.00060892 | 0.014896 | 0.016037 | -0.09613 | 75.9659 | -46.2753 | 75.8057 | -39.1434 | -43.6249 | -44.8242 | -0.009155 | 75.9018 | -46.1838 | 76.0025 | -39.3768 | -43.277 | -44.0781 | -0.15564 | 0.009155 | -0.15106 | -0.16022 | 6.9763 | 7.132 | 1.2955 | 2.4948 | 1.4511 | 2.6505 | -0.064087 | -0.054932 | 0.045776 | 0.10071 | 6.7429 | 6.8069 | 2.0416 | 2.8427 | 2.1057 | 2.9068
1685.1719 | 122.3145 | -46.0648 | 76.0529 | -0.10002 | 0.02014 | 0.088199 | 0.090709 | 0.0014675 | 0.016188 | 0.0098694 | -0.013519 | 0.024099 | 0.08671 | 0.098226 | -0.00039512 | 0.01276 | 0.011764 | -0.022888 | 76.062 | -46.225 | 75.824 | -39.1571 | -43.5837 | -44.8013 | -0.013733 | 76.0437 | -46.1105 | 75.9705 | -39.3768 | -43.2678 | -44.1147 | -0.16022 | 0.009156 | -0.22888 | -0.23804 | 6.9077 | 7.0679 | 1.2634 | 2.4811 | 1.4236 | 2.6413 | -0.045777 | -0.009155 | -0.082397 | -0.073242 | 6.6879 | 6.7337 | 1.9501 | 2.7969 | 1.9958 | 2.8427
1686.1719 | 122.3557 | -46.1472 | 75.9567 | -0.13817 | 0.0195 | 0.087131 | 0.090922 | 0.0018949 | 0.018752 | 0.0094422 | -0.040985 | 0.02538 | 0.086496 | 0.092458 | -0.0016769 | 0.013401 | 0.017746 | -0.022888 | 75.9705 | -46.2845 | 75.7553 | -39.212 | -43.6203 | -44.8608 | 0.041199 | 75.9384 | -46.2158 | 75.9338 | -39.386 | -43.3044 | -44.1559 | -0.13733 | 0.013733 | -0.20142 | -0.21515 | 6.9351 | 7.0724 | 1.2863 | 2.5269 | 1.4236 | 2.6642 | -0.068664 | -0.01831 | -0.022888 | -0.004578 | 6.7612 | 6.8298 | 1.9913 | 2.8427 | 2.0599 | 2.9114
1687.1719 | 122.2916 | -46.0052 | 76.0803 | -0.077133 | 0.022063 | 0.088412 | 0.092631 | -0.00066869 | 0.017256 | 0.0092286 | -0.013519 | 0.023458 | 0.087564 | 0.091176 | -0.00039512 | 0.013401 | 0.011123 | -0.018311 | 76.1536 | -46.1609 | 75.9109 | -39.1067 | -43.515 | -44.7327 | -0.004578 | 76.0986 | -46.0739 | 76.0483 | -39.2624 | -43.1488 | -43.9957 | -0.15564 | 0.073242 | -0.16937 | -0.24261 | 6.8985 | 7.0541 | 1.2726 | 2.4902 | 1.4282 | 2.6459 | -0.068665 | 0.018311 | -0.032043 | -0.050354 | 6.7429 | 6.8115 | 2.0096 | 2.8564 | 2.0782 | 2.9251
1688.1719 | 122.3053 | -46.0327 | 76.1032 | -0.1458 | 0.020995 | 0.086917 | 0.089213 | -2.7692e-05 | 0.016188 | 0.010937 | -0.016571 | 0.02538 | 0.085855 | 0.093313 | -0.0018905 | 0.013401 | 0.015396 | -0.018311 | 76.0986 | -46.17 | 75.8652 | -39.1113 | -43.515 | -44.7601 | -0.004578 | 76.0712 | -46.0785 | 76.0117 | -39.2761 | -43.158 | -44.0369 | -0.13733 | -0.004577 | -0.23804 | -0.23346 | 6.9214 | 7.0587 | 1.2726 | 2.5177 | 1.4099 | 2.655 | -0.045776 | -0.032043 | -0.091552 | -0.059509 | 6.7566 | 6.8024 | 1.9959 | 2.8748 | 2.0416 | 2.9205
1689.1719 | 122.3328 | -46.2067 | 75.8881 | -0.084763 | 0.020781 | 0.086917 | 0.091349 | -0.0015231 | 0.01747 | 0.0090148 | -0.065399 | 0.026021 | 0.086282 | 0.090322 | -0.0014633 | 0.01276 | 0.013687 | -0.032043 | 75.943 | -46.3303 | 75.7416 | -39.2258 | -43.689 | -44.9203 | 0.004578 | 75.8881 | -46.2616 | 75.8835 | -39.4043 | -43.3548 | -44.2062 | -0.1236 | 0.054931 | -0.14648 | -0.20142 | 6.9809 | 7.1045 | 1.2863 | 2.5177 | 1.4099 | 2.6413 | -0.054932 | 0 | -0.004578 | -0.004578 | 6.8024 | 6.8573 | 2.0004 | 2.8519 | 2.0554 | 2.9068
1690.1719 | 122.2961 | -45.9229 | 76.2085 | -0.069504 | 0.021849 | 0.086703 | 0.092418 | 0.0021085 | 0.016402 | 0.007306 | -0.019623 | 0.025808 | 0.087137 | 0.096303 | 0.00088648 | 0.01575 | 0.012832 | -0.027466 | 76.2405 | -46.0373 | 75.9979 | -38.9969 | -43.4097 | -44.6548 | -0.013733 | 76.1948 | -45.9778 | 76.1398 | -39.1159 | -43.0298 | -43.9224 | -0.11444 | 0.032044 | -0.21057 | -0.24262 | 6.926 | 7.0404 | 1.268 | 2.5131 | 1.3824 | 2.6276 | -0.054931 | -0.013733 | -0.068664 | -0.054931 | 6.8069 | 6.8619 | 2.0004 | 2.8931 | 2.0554 | 2.948
1691.1719 | 122.2733 | -45.9412 | 76.2314 | -0.11528 | 0.020781 | 0.087344 | 0.088572 | -0.00045509 | 0.016616 | 0.006665 | 0.0078429 | 0.023031 | 0.088205 | 0.092672 | 0.00024568 | 0.012973 | 0.013687 | -0.013733 | 76.1856 | -46.0648 | 75.9796 | -38.9877 | -43.3914 | -44.6503 | 0.018311 | 76.1627 | -45.9641 | 76.1124 | -39.1525 | -43.0344 | -43.9545 | -0.1236 | -0.045776 | -0.25177 | -0.20599 | 6.9534 | 7.077 | 1.2909 | 2.5497 | 1.4145 | 2.6733 | -0.022888 | -0.068664 | -0.11902 | -0.050354 | 6.7886 | 6.8115 | 1.9867 | 2.9068 | 2.0096 | 2.9297
1692.1719 | 122.2916 | -45.9732 | 76.1673 | -0.12291 | 0.021209 | 0.087344 | 0.089 | -0.0013095 | 0.018111 | 0.010297 | 0.041412 | 0.02474 | 0.086923 | 0.087545 | -0.00060892 | 0.014041 | 0.014969 | 0.022888 | 76.1444 | -46.0922 | 75.9796 | -38.974 | -43.396 | -44.6274 | 0.059509 | 76.1215 | -45.9915 | 76.0757 | -39.1388 | -43.0801 | -43.9774 | -0.11902 | -0.022888 | -0.18768 | -0.16479 | 6.9992 | 7.1182 | 1.3458 | 2.5772 | 1.4648 | 2.6962 | -0.01831 | -0.045776 | -0.091552 | -0.045776 | 6.8344 | 6.8527 | 1.9958 | 2.8931 | 2.0142 | 2.9114
1693.1719 | 122.2778 | -46.1884 | 75.9476 | -0.13054 | 0.024199 | 0.087131 | 0.089427 | -2.7692e-05 | 0.017256 | 0.011578 | 0.010895 | 0.024526 | 0.083932 | 0.088613 | 0.00088648 | 0.013401 | 0.014755 | 0.004578 | 75.9521 | -46.2799 | 75.7507 | -39.2029 | -43.6523 | -44.8654 | 0.027466 | 75.943 | -46.1975 | 75.8606 | -39.3448 | -43.2632 | -44.1788 | -0.091553 | 0.004577 | -0.19684 | -0.20142 | 6.9855 | 7.077 | 1.3229 | 2.536 | 1.4145 | 2.6276 | -0.009156 | -0.004578 | -0.086975 | -0.082397 | 6.8436 | 6.8527 | 2.0096 | 2.9251 | 2.0187 | 2.9343
1694.1719 | 122.2687 | -46.1197 | 75.9796 | -0.084763 | 0.020781 | 0.087558 | 0.094981 | 0.00039951 | 0.015975 | 0.0070922 | -0.080658 | 0.026876 | 0.086923 | 0.090963 | -0.00060892 | 0.015323 | 0.011978 | -0.022888 | 75.9979 | -46.2524 | 75.8102 | -39.1205 | -43.602 | -44.8608 | -0.009155 | 75.9979 | -46.1517 | 75.9338 | -39.2853 | -43.222 | -44.1513 | -0.13275 | 0.018311 | -0.16937 | -0.18768 | 6.9992 | 7.132 | 1.2588 | 2.5177 | 1.3916 | 2.6505 | -0.032043 | 0.018311 | -0.045776 | -0.064087 | 6.8344 | 6.8665 | 1.9684 | 2.8976 | 2.0004 | 2.9297
1695.1719 | 122.3648 | -46.17 | 75.9247 | -0.1458 | 0.019286 | 0.087771 | 0.089427 | -0.0013095 | 0.017684 | 0.0079468 | 0.035309 | 0.024953 | 0.085855 | 0.093526 | -0.00039512 | 0.014682 | 0.010269 | 0.032043 | 75.9109 | -46.2845 | 75.8286 | -39.0701 | -43.5562 | -44.8151 | 0.11444 | 75.8789 | -46.2158 | 75.9109 | -39.2258 | -43.1625 | -44.101 | -0.11444 | -0.013733 | -0.096131 | -0.082398 | 7.0999 | 7.2144 | 1.355 | 2.6138 | 1.4694 | 2.7283 | -0.045776 | -0.045777 | -0.013733 | 0.032044 | 6.9443 | 6.9901 | 2.0691 | 3.0075 | 2.1149 | 3.0533
1696.1719 | 122.4014 | -46.257 | 75.7919 | -0.12291 | 0.02014 | 0.088199 | 0.090709 | 0.0012539 | 0.016616 | 0.0060242 | 0.062775 | 0.024099 | 0.08735 | 0.088827 | 0.00045928 | 0.012119 | 0.0092007 | 0.018311 | 75.8011 | -46.3898 | 75.7965 | -39.3402 | -43.9087 | -44.8746 | 0.068665 | 75.769 | -46.312 | 75.8423 | -39.2899 | -43.2266 | -44.2108 | -0.13275 | 0.009155 | 0.004578 | -0.004577 | 6.9168 | 7.0496 | 1.3824 | 2.3483 | 1.5152 | 2.4811 | -0.054932 | -0.022888 | 0.050354 | 0.073242 | 6.9672 | 7.0221 | 2.0462 | 3.0304 | 2.1011 | 3.0853
1697.1719 | 122.3785 | -46.2982 | 75.7919 | -0.11528 | 0.022917 | 0.088412 | 0.090495 | 0.0018949 | 0.017684 | 0.007306 | 0.010895 | 0.024526 | 0.085428 | 0.091604 | -0.00018152 | 0.014682 | 0.014114 | 0.004578 | 75.7736 | -46.4355 | 75.7141 | -39.1663 | -43.689 | -44.9295 | 0.1236 | 75.7507 | -46.3303 | 75.8148 | -39.2899 | -43.2632 | -44.2337 | -0.13733 | -0.01831 | -0.07782 | -0.05951 | 7.132 | 7.2693 | 1.3687 | 2.6093 | 1.506 | 2.7466 | -0.032043 | -0.041199 | 0.022888 | 0.064087 | 7.0084 | 7.0404 | 2.0645 | 3.035 | 2.0966 | 3.067
1698.1719 | 122.438 | -46.1105 | 75.9293 | -0.21446 | 0.020781 | 0.086703 | 0.092204 | -0.0021641 | 0.017256 | 0.0090148 | 0.029205 | 0.02538 | 0.086923 | 0.094808 | -0.0023177 | 0.014682 | 0.016678 | -0.004578 | 75.9338 | -46.2799 | 75.8377 | -39.0106 | -43.5059 | -44.7693 | 0.13275 | 75.9018 | -46.1929 | 75.9888 | -39.1159 | -43.0618 | -44.0552 | -0.16937 | 0.004578 | -0.091552 | -0.09613 | 7.0999 | 7.2693 | 1.3412 | 2.6047 | 1.5106 | 2.774 | -0.082397 | -0.027466 | 0.05951 | 0.086976 | 6.9946 | 7.077 | 2.0554 | 3.0487 | 2.1378 | 3.1311
1699.1719 | 122.374 | -45.7306 | 76.3138 | -0.11528 | 0.021636 | 0.087131 | 0.092631 | -0.00024129 | 0.017684 | 0.0098694 | 0.013947 | 0.024099 | 0.085214 | 0.087758 | -0.00060892 | 0.014255 | 0.014755 | -0.027466 | 76.3092 | -45.8725 | 76.2131 | -38.7131 | -43.158 | -44.4077 | 0.073242 | 76.2955 | -45.8084 | 76.3641 | -38.8 | -42.6773 | -43.6844 | -0.14191 | -0.004578 | -0.10071 | -0.09613 | 7.0175 | 7.1594 | 1.3229 | 2.5726 | 1.4648 | 2.7145 | -0.07782 | -0.018311 | 0.050354 | 0.068665 | 6.9305 | 7.0084 | 2.0462 | 3.0533 | 2.124 | 3.1311
1700.1719 | 122.4243 | -46.0464 | 76.0162 | -0.13817 | 0.020568 | 0.088626 | 0.09199 | -2.7692e-05 | 0.017684 | 0.011151 | 0.026154 | 0.022817 | 0.083292 | 0.094594 | 3.2078e-05 | 0.013401 | 0.013901 | -0.013733 | 76.0208 | -46.1838 | 75.9338 | -38.9191 | -43.4235 | -44.6732 | 0.091553 | 76.0071 | -46.0785 | 76.0849 | -39.0335 | -42.9657 | -43.9682 | -0.13733 | 0.004578 | -0.082397 | -0.086975 | 7.1274 | 7.2647 | 1.3733 | 2.623 | 1.5106 | 2.7603 | -0.032043 | -0.009155 | 0.068665 | 0.07782 | 7.0129 | 7.045 | 2.0782 | 3.0807 | 2.1103 | 3.1128
1701.1719 | 122.3785 | -46.1243 | 75.9521 | -0.077133 | 0.019927 | 0.087771 | 0.091349 | 0.0023221 | 0.018111 | 0.010297 | -0.010467 | 0.023671 | 0.086496 | 0.086477 | -0.0025315 | 0.01575 | 0.0092007 | -0.013733 | 75.9293 | -46.2891 | 75.8011 | -39.061 | -43.5471 | -44.7968 | 0.045776 | 75.8926 | -46.1792 | 75.9888 | -39.1479 | -43.071 | -44.0872 | -0.16479 | -0.022888 | -0.15106 | -0.12817 | 7.0633 | 7.2281 | 1.3275 | 2.5772 | 1.4923 | 2.742 | -0.054931 | -0.059509 | 0.036622 | 0.096131 | 6.9763 | 7.0312 | 2.037 | 3.0533 | 2.092 | 3.1082
1702.1719 | 122.3785 | -46.1792 | 75.8926 | -0.13054 | 0.02014 | 0.088626 | 0.092845 | 0.00061311 | 0.01747 | 0.010297 | -0.010467 | 0.026662 | 0.084787 | 0.094594 | -0.00039512 | 0.013828 | 0.016037 | -0.050354 | 75.9018 | -46.3531 | 75.7919 | -39.1022 | -43.6203 | -44.8425 | 0.022888 | 75.8926 | -46.2204 | 75.9201 | -39.2029 | -43.1076 | -44.1376 | -0.17395 | 0.009155 | -0.10071 | -0.10986 | 7.077 | 7.251 | 1.3367 | 2.5589 | 1.5106 | 2.7328 | -0.041199 | 0 | 0.027466 | 0.027466 | 6.9763 | 7.0175 | 2.0416 | 3.0716 | 2.0828 | 3.1128
1703.1719 | 122.3694 | -46.3531 | 75.737 | -0.13817 | 0.019927 | 0.08649 | 0.091136 | -0.0019503 | 0.018752 | 0.010297 | -0.025726 | 0.024953 | 0.08671 | 0.095449 | -0.00039512 | 0.010837 | 0.011123 | -0.07782 | 75.7645 | -46.4493 | 75.6409 | -39.1983 | -43.7393 | -44.9844 | -0.032043 | 75.7416 | -46.3852 | 75.7553 | -39.4684 | -43.2816 | -44.2841 | -0.096131 | 0.027465 | -0.096131 | -0.1236 | 7.1548 | 7.251 | 1.3687 | 2.6138 | 1.4648 | 2.71 | -0.032044 | 0.004577 | 0.01831 | 0.013733 | 6.8848 | 6.9168 | 2.0691 | 3.0716 | 2.1011 | 3.1036
1704.1719 | 122.287 | -46.1609 | 75.943 | -0.10765 | 0.020781 | 0.086703 | 0.094554 | -0.00024129 | 0.018325 | 0.010937 | -0.11728 | 0.025594 | 0.085214 | 0.092672 | -0.0023177 | 0.016391 | 0.010269 | 0.013733 | 75.8789 | -46.3669 | 75.7416 | -39.0976 | -43.6432 | -44.9158 | 0.045776 | 75.8926 | -46.257 | 75.8835 | -39.3127 | -43.19 | -44.1971 | -0.20599 | -0.064087 | -0.20142 | -0.13733 | 7.0633 | 7.2693 | 1.2451 | 2.5177 | 1.4511 | 2.7237 | -0.09613 | -0.050354 | -0.059509 | -0.009155 | 6.8481 | 6.9443 | 1.9638 | 2.9709 | 2.0599 | 3.067
1705.1719 | 122.3328 | -46.0785 | 75.9567 | -0.13054 | 0.021422 | 0.089267 | 0.091777 | -0.0015231 | 0.018965 | 0.012219 | -0.05014 | 0.025167 | 0.08436 | 0.090749 | -0.0014633 | 0.013828 | 0.011551 | -0.054932 | 75.9567 | -46.2341 | 75.856 | -38.9465 | -43.5333 | -44.7556 | -0.036621 | 75.9567 | -46.1563 | 75.9979 | -39.1754 | -43.0298 | -44.0552 | -0.15564 | 0 | -0.10071 | -0.10071 | 7.132 | 7.2876 | 1.3229 | 2.5452 | 1.4786 | 2.7008 | -0.07782 | 0 | 0.041199 | 0.041199 | 6.9031 | 6.9809 | 2.0233 | 3.0487 | 2.1011 | 3.1265
1706.1719 | 122.319 | -46.1243 | 75.9293 | -0.1458 | 0.019927 | 0.089053 | 0.08964 | -0.0030185 | 0.018752 | 0.0077332 | -0.065399 | 0.023458 | 0.084787 | 0.093953 | 0.0015275 | 0.013401 | 0.019668 | -0.050354 | 75.8926 | -46.2845 | 75.8286 | -38.9465 | -43.5425 | -44.7968 | 0 | 75.9247 | -46.1975 | 75.9933 | -39.2624 | -43.0527 | -44.0964 | -0.16022 | -0.036621 | -0.10071 | -0.064087 | 7.1777 | 7.338 | 1.3275 | 2.5818 | 1.4877 | 2.742 | -0.073242 | -0.004577 | 0.064087 | 0.068664 | 6.8619 | 6.9351 | 2.0279 | 3.0716 | 2.1011 | 3.1448
1707.1719 | 122.3694 | -46.2662 | 75.769 | -0.16869 | 0.0195 | 0.088199 | 0.091563 | -0.0023777 | 0.016188 | 0.0092286 | -0.044037 | 0.026449 | 0.083078 | 0.09374 | 0.00067288 | 0.014896 | 0.0074917 | -0.073242 | 75.7919 | -46.3989 | 75.6729 | -39.1113 | -43.7073 | -44.9249 | -0.045776 | 75.8011 | -46.3394 | 75.8148 | -39.4638 | -43.1808 | -44.2383 | -0.13275 | 0.022888 | -0.09613 | -0.11902 | 7.1548 | 7.2876 | 1.3412 | 2.5589 | 1.474 | 2.6917 | -0.073243 | 0.032043 | 0.045776 | 0.013733 | 6.8024 | 6.8756 | 2.0279 | 3.0853 | 2.1011 | 3.1586
1708.1719 | 122.3328 | -46.2433 | 75.8194 | -0.13054 | 0.020568 | 0.086703 | 0.091777 | 0.0027493 | 0.017256 | 0.01051 | -0.065399 | 0.023031 | 0.085214 | 0.092458 | -0.00018152 | 0.013614 | 0.0087735 | -0.036621 | 75.8011 | -46.3715 | 75.7278 | -39.0564 | -43.6569 | -44.8975 | -0.027466 | 75.8148 | -46.2891 | 75.8789 | -39.5096 | -43.1351 | -44.2108 | -0.12817 | -0.018311 | -0.091553 | -0.073242 | 7.1869 | 7.3151 | 1.3458 | 2.5864 | 1.474 | 2.7145 | -0.045776 | -0.004578 | 0.059509 | 0.064087 | 6.7337 | 6.7795 | 2.0325 | 3.1082 | 2.0782 | 3.154
1709.1719 | 122.3557 | -46.4951 | 75.5722 | 0.15938 | 0.020781 | 0.088626 | 0.091777 | -0.00066869 | 0.016829 | 0.0081604 | -0.040985 | 0.023031 | 0.084146 | 0.092244 | -0.0021041 | 0.015537 | 0.012832 | -0.07782 | 75.5219 | -46.6599 | 75.5127 | -39.2075 | -43.8812 | -45.1126 | 0.10071 | 75.5585 | -46.5591 | 75.6226 | -39.8254 | -43.3365 | -44.4443 | -0.1648 | -0.050354 | -0.05951 | -0.009156 | 7.2876 | 7.4524 | 1.3824 | 2.6138 | 1.5472 | 2.7786 | -0.064087 | -0.013733 | 0.050354 | 0.064087 | 6.6696 | 6.7337 | 2.0508 | 3.1586 | 2.1149 | 3.2227
1710.1719 | 122.3557 | -46.3989 | 75.6912 | -0.13054 | 0.022277 | 0.088412 | 0.092204 | -0.0030185 | 0.01982 | 0.0092286 | -0.077606 | 0.025594 | 0.08436 | 0.089895 | -0.00082252 | 0.01511 | 0.011978 | -0.041199 | 75.6363 | -46.5088 | 75.5859 | -39.1251 | -43.7576 | -45.0165 | 0.009155 | 75.65 | -46.4584 | 75.7599 | -39.8895 | -43.2541 | -44.339 | -0.10986 | -0.054931 | -0.10529 | -0.050355 | 7.2739 | 7.3837 | 1.3824 | 2.6413 | 1.4923 | 2.7512 | -0.059509 | -0.041199 | 0.068665 | 0.10986 | 6.5094 | 6.5689 | 2.0599 | 3.1448 | 2.1194 | 3.2043
1711.1719 | 122.3602 | -46.6507 | 75.3983 | -0.16106 | 0.02014 | 0.088626 | 0.090922 | -0.0028049 | 0.019179 | 0.01051 | -0.077606 | 0.026449 | 0.08436 | 0.088399 | 0.0011001 | 0.017459 | 0.01326 | -0.045776 | 75.4211 | -46.7697 | 75.3296 | -39.3082 | -44.0186 | -45.2774 | -0.036621 | 75.4257 | -46.6553 | 75.4349 | -40.3107 | -43.5013 | -44.6228 | -0.11902 | 0.022889 | -0.068664 | -0.091553 | 7.3425 | 7.4615 | 1.3733 | 2.6321 | 1.4923 | 2.7512 | -0.004577 | 0.027466 | 0.036621 | 0.009155 | 6.34 | 6.3446 | 2.0279 | 3.1494 | 2.0325 | 3.154
1712.1719 | 122.3511 | -46.8384 | 75.2197 | -0.20683 | 0.021209 | 0.087985 | 0.093058 | 0.00018591 | 0.017256 | 0.0096558 | -0.13864 | 0.02538 | 0.083719 | 0.091176 | 0.00024568 | 0.01575 | 0.010269 | -0.009155 | 75.206 | -46.9574 | 75.1556 | -39.4684 | -44.2291 | -45.694 | 0.16022 | 75.1602 | -46.9254 | 75.3159 | -40.4755 | -43.6707 | -44.783 | -0.11902 | -0.013733 | -0.064087 | -0.050354 | 7.37 | 7.489 | 1.1444 | 2.6093 | 1.2634 | 2.7283 | -0.086975 | -0.05951 | 0.09613 | 0.15564 | 6.3629 | 6.4499 | 2.0554 | 3.1677 | 2.1423 | 3.2547
1713.1719 | 122.4655 | -46.8384 | 75.2151 | -0.13054 | 0.022917 | 0.08948 | 0.090922 | -0.00045509 | 0.018965 | 0.0088012 | -0.089813 | 0.026235 | 0.085001 | 0.089254 | 0.00088648 | 0.017032 | 0.015396 | -0.032043 | 75.2472 | -46.9208 | 75.2106 | -39.4363 | -44.1605 | -45.4102 | 0 | 75.2747 | -46.8246 | 75.325 | -40.8508 | -43.6523 | -44.7052 | -0.082397 | 0.032043 | -0.004578 | -0.036621 | 7.402 | 7.4844 | 1.4282 | 2.6779 | 1.5106 | 2.7603 | 0.013733 | 0.059509 | 0.10986 | 0.050354 | 5.9875 | 5.9738 | 2.1332 | 3.186 | 2.1194 | 3.1723
1714.1719 | 122.3419 | -46.8979 | 75.1236 | -0.45097 | 0.022704 | 0.087985 | 0.090709 | -0.0017367 | 0.015761 | 0.011365 | -0.071503 | 0.024312 | 0.086069 | 0.091604 | 3.2078e-05 | 0.014041 | 0.013901 | -0.032043 | 75.0961 | -47.0764 | 75.0778 | -39.743 | -44.2749 | -45.5338 | 0.022888 | 75.1007 | -46.9437 | 75.2426 | -40.8005 | -43.689 | -44.8563 | -0.17853 | -0.027466 | -0.045776 | -0.01831 | 7.1548 | 7.3334 | 1.3641 | 2.623 | 1.5427 | 2.8015 | -0.045777 | -0.022888 | 0.11902 | 0.14191 | 6.0974 | 6.1432 | 2.0416 | 3.2089 | 2.0874 | 3.2547
1715.1719 | 122.3602 | -46.8338 | 75.1556 | -0.16869 | 0.022063 | 0.088412 | 0.091777 | -0.0019503 | 0.018965 | 0.013074 | -0.12033 | 0.023885 | 0.083292 | 0.090535 | -0.0010361 | 0.016605 | 0.011978 | -0.054932 | 75.1877 | -46.9986 | 74.9084 | -39.8529 | -44.0643 | -45.4056 | 0.018311 | 75.1328 | -46.9437 | 75.2792 | -40.8188 | -43.6432 | -44.8059 | -0.16479 | 0.032043 | -0.24719 | -0.27924 | 6.9809 | 7.1457 | 1.4282 | 2.7695 | 1.593 | 2.9343 | -0.10986 | -0.022889 | 0.1236 | 0.14648 | 6.015 | 6.1249 | 2.0279 | 3.1906 | 2.1378 | 3.3005
1716.1719 | 122.3694 | -46.6415 | 75.3799 | -0.16869 | 0.020568 | 0.087985 | 0.08964 | 0.0016813 | 0.018538 | 0.010724 | -0.11118 | 0.025808 | 0.083719 | 0.094808 | -0.0014633 | 0.014682 | 0.017105 | -0.050354 | 75.412 | -46.756 | 75.3479 | -39.386 | -44.0094 | -45.2591 | 0.004578 | 75.4028 | -46.7056 | 75.4898 | -40.7272 | -43.3914 | -44.5587 | -0.11444 | 0.032043 | -0.032044 | -0.064087 | 7.2556 | 7.37 | 1.3824 | 2.6321 | 1.4969 | 2.7466 | -0.064086 | 0.022888 | 0.10986 | 0.086975 | 5.9143 | 5.9784 | 2.0828 | 3.2501 | 2.1469 | 3.3142
1717.1719 | 122.319 | -46.7331 | 75.2563 | -0.16106 | 0.021209 | 0.08948 | 0.091563 | -0.00024129 | 0.018325 | 0.011792 | -0.098968 | 0.024526 | 0.083932 | 0.090963 | -0.0010361 | 0.017032 | 0.013687 | -0.050354 | 75.3342 | -46.8338 | 75.2701 | -39.4272 | -44.0781 | -45.3506 | -0.004578 | 75.3342 | -46.7468 | 75.4395 | -40.7822 | -43.5013 | -44.8105 | -0.10071 | 0.077819 | 0.013733 | -0.064086 | 7.3059 | 7.4066 | 1.3824 | 2.655 | 1.4832 | 2.7557 | -0.013733 | 0.077819 | 0.1831 | 0.10529 | 5.9509 | 5.9647 | 1.9226 | 3.2318 | 1.9363 | 3.2455
1718.1719 | 122.3602 | -46.637 | 75.3937 | -0.18394 | 0.020781 | 0.094607 | 0.090922 | 0.0010403 | 0.01982 | 0.0085876 | -0.071503 | 0.025808 | 0.083932 | 0.091817 | 3.2078e-05 | 0.017887 | 0.0074917 | -0.27008 | 75.325 | -46.8338 | 75.3387 | -39.3402 | -44.0414 | -45.2911 | 0 | 75.3616 | -46.7102 | 75.5035 | -40.7547 | -43.4372 | -44.5999 | -0.19684 | -0.068665 | -0.054932 | 0.013733 | 7.2968 | 7.4936 | 1.3458 | 2.5955 | 1.5427 | 2.7924 | -0.073242 | -0.032044 | 0.10986 | 0.14191 | 5.8823 | 5.9555 | 2.037 | 3.1998 | 2.1103 | 3.273
1719.1719 | 122.4106 | -46.7422 | 75.2975 | -0.11528 | 0.018859 | 0.086703 | 0.090068 | -0.0019503 | 0.018538 | 0.0085876 | -0.022675 | 0.024312 | 0.084573 | 0.091604 | -0.0010361 | 0.013828 | 0.013046 | -0.018311 | 75.2655 | -46.9208 | 75.2151 | -39.4684 | -44.1147 | -45.3506 | 0.073242 | 75.2884 | -46.7972 | 75.3937 | -40.8188 | -43.4784 | -44.7006 | -0.17853 | -0.032043 | -0.082397 | -0.050354 | 7.2739 | 7.4524 | 1.3916 | 2.6276 | 1.5701 | 2.8061 | -0.054931 | -0.009155 | 0.096131 | 0.10529 | 5.9235 | 5.9784 | 2.0416 | 3.2639 | 2.0966 | 3.3188
1720.1719 | 122.3785 | -46.4401 | 75.6042 | -0.17632 | 0.020568 | 0.088199 | 0.093699 | -0.00066869 | 0.014907 | 0.012219 | -0.11118 | 0.027517 | 0.083719 | 0.10207 | -0.0018905 | 0.017673 | 0.013901 | -0.10071 | 75.6088 | -46.637 | 75.4761 | -39.2532 | -43.895 | -45.1447 | -0.041199 | 75.5859 | -46.5179 | 75.6729 | -40.5762 | -43.2632 | -44.426 | -0.19684 | 0.004578 | -0.12817 | -0.13275 | 7.1869 | 7.3837 | 1.2955 | 2.5452 | 1.4923 | 2.742 | -0.077819 | -0.018311 | 0.068665 | 0.086976 | 5.864 | 5.9418 | 2.0142 | 3.1769 | 2.092 | 3.2547
1721.1719 | 122.3328 | -46.6324 | 75.4349 | -0.10002 | 0.021209 | 0.086703 | 0.094554 | -0.00088229 | 0.017043 | 0.011151 | -0.13559 | 0.027089 | 0.083719 | 0.096731 | -0.0031723 | 0.016391 | 0.01326 | -0.10071 | 75.4578 | -46.7926 | 75.2701 | -39.3814 | -44.0918 | -45.3323 | -0.10071 | 75.4395 | -46.7011 | 75.4761 | -40.8142 | -43.4509 | -44.632 | -0.16022 | 0.022889 | -0.16479 | -0.18768 | 7.251 | 7.4112 | 1.3 | 2.5406 | 1.4603 | 2.7008 | -0.068665 | 0.004578 | 0.041199 | 0.036621 | 5.8182 | 5.8868 | 2.0004 | 3.1815 | 2.0691 | 3.2501
1722.1719 | 122.3648 | -46.8018 | 75.2151 | -0.15343 | 0.020568 | 0.086917 | 0.092204 | -0.00045509 | 0.019606 | 0.01051 | -0.092865 | 0.02538 | 0.083932 | 0.098012 | 0.0021683 | 0.017887 | 0.013473 | -0.059509 | 75.2335 | -46.9574 | 75.1419 | -39.4547 | -44.2154 | -45.4742 | -0.032043 | 75.1923 | -46.8979 | 75.3296 | -40.9698 | -43.602 | -44.7876 | -0.15564 | 0.01831 | -0.073242 | -0.091552 | 7.3471 | 7.5027 | 1.3275 | 2.5864 | 1.4832 | 2.742 | -0.09613 | -0.022888 | 0.11444 | 0.13733 | 5.8319 | 5.928 | 2.0142 | 3.1998 | 2.1103 | 3.2959
1723.1719 | 122.3602 | -46.7239 | 75.2838 | -0.1458 | 0.0195 | 0.088412 | 0.090495 | 0.0031765 | 0.019393 | 0.0096558 | -0.065399 | 0.027089 | 0.087137 | 0.095876 | -0.00082252 | 0.015323 | 0.011764 | -0.10986 | 75.3021 | -46.9162 | 75.1923 | -39.3997 | -44.1513 | -45.4102 | -0.032043 | 75.3067 | -46.8109 | 75.3754 | -40.9195 | -43.5287 | -44.6823 | -0.19226 | 0.018311 | -0.091552 | -0.10986 | 7.3242 | 7.5165 | 1.3138 | 2.5726 | 1.506 | 2.7649 | -0.086975 | 0.022889 | 0.091553 | 0.068664 | 5.8044 | 5.8914 | 2.0416 | 3.1952 | 2.1286 | 3.2822
1724.1719 | 122.3419 | -46.756 | 75.325 | -0.13054 | 0.021636 | 0.08884 | 0.088572 | -0.00024129 | 0.018538 | 0.0077332 | -0.053192 | 0.02474 | 0.085428 | 0.086477 | -0.0025315 | 0.016391 | 0.011551 | -0.082397 | 75.3708 | -46.8246 | 75.2197 | -39.3356 | -44.0964 | -45.3323 | -0.036621 | 75.3525 | -46.7834 | 75.3891 | -40.9973 | -43.5242 | -44.696 | -0.068665 | 0.045777 | -0.10528 | -0.15106 | 7.4203 | 7.489 | 1.4236 | 2.6596 | 1.4923 | 2.7283 | -0.027466 | 0.027466 | 0.064087 | 0.036621 | 5.7587 | 5.7861 | 2.0599 | 3.2318 | 2.0874 | 3.2593
1725.1719 | 122.3923 | -46.7468 | 75.3067 | -0.18394 | 0.020781 | 0.087558 | 0.091777 | 0.00061311 | 0.019606 | 0.0079468 | -0.077606 | 0.025808 | 0.083292 | 0.090108 | -0.00039512 | 0.015537 | 0.01326 | -0.068665 | 75.2747 | -46.8842 | 75.2289 | -39.3494 | -44.1376 | -45.3873 | -0.009155 | 75.2792 | -46.7926 | 75.4028 | -40.9149 | -43.5013 | -44.6869 | -0.13733 | -0.032044 | -0.07782 | -0.045776 | 7.3975 | 7.5348 | 1.3596 | 2.6093 | 1.4969 | 2.7466 | -0.045777 | -0.027466 | 0.09613 | 0.1236 | 5.8319 | 5.8777 | 2.0599 | 3.2455 | 2.1057 | 3.2913
1726.1719 | 122.3236 | -46.8063 | 75.2747 | -0.069504 | 0.022277 | 0.08649 | 0.095195 | -0.0010959 | 0.018538 | 0.014142 | -0.089813 | 0.028371 | 0.084573 | 0.095876 | -0.00039512 | 0.017459 | 0.013046 | -0.11902 | 75.2792 | -46.962 | 75.0824 | -39.4684 | -44.252 | -45.4926 | -0.10529 | 75.2975 | -46.8475 | 75.2747 | -41.0797 | -43.6478 | -44.8151 | -0.15564 | 0.004578 | -0.19226 | -0.19684 | 7.338 | 7.4936 | 1.3138 | 2.5543 | 1.4694 | 2.71 | -0.041199 | 0.022888 | 0 | -0.022888 | 5.7266 | 5.7678 | 1.9913 | 3.1586 | 2.0325 | 3.1998
1727.1719 | 122.2458 | -46.8933 | 75.1923 | -0.12291 | 0.019713 | 0.088626 | 0.090281 | 0.00018591 | 0.017684 | 0.0098694 | -0.092865 | 0.026449 | 0.086282 | 0.095235 | -0.0014633 | 0.016178 | 0.0096279 | -0.064087 | 75.2243 | -47.0261 | 74.9771 | -39.5004 | -44.2886 | -45.6024 | -0.1236 | 75.2289 | -46.9391 | 75.1831 | -41.1896 | -43.7393 | -44.9387 | -0.13275 | 0.032043 | -0.21515 | -0.24719 | 7.3929 | 7.5256 | 1.2909 | 2.6047 | 1.4236 | 2.7374 | -0.045776 | 0.036621 | -0.009156 | -0.045777 | 5.7037 | 5.7495 | 1.9547 | 3.154 | 2.0004 | 3.1998
1728.1719 | 122.3145 | -46.788 | 75.2655 | -0.092392 | 0.019713 | 0.087985 | 0.092204 | -0.00088229 | 0.01747 | 0.012219 | -0.089813 | 0.026449 | 0.083292 | 0.092031 | -0.00018152 | 0.015964 | 0.014969 | -0.082397 | 75.3754 | -46.8521 | 75.1831 | -39.3677 | -44.1559 | -45.4422 | -0.15106 | 75.2884 | -46.8567 | 75.2975 | -41.1026 | -43.602 | -44.8151 | -0.064087 | 0.10986 | -0.082398 | -0.19226 | 7.4203 | 7.4844 | 1.3458 | 2.6321 | 1.4099 | 2.6962 | -0.068664 | 0.022888 | 0.032043 | 0.009155 | 5.6854 | 5.7541 | 1.973 | 3.186 | 2.0416 | 3.2547
1729.1719 | 122.3648 | -46.8063 | 75.1785 | -0.16869 | 0.020354 | 0.089908 | 0.091349 | -0.00045509 | 0.017043 | 0.012006 | -0.0013123 | 0.025594 | 0.085641 | 0.090108 | -0.00039512 | 0.016178 | 0.0070645 | -0.059509 | 75.1923 | -46.9482 | 75.119 | -39.3402 | -44.1971 | -45.4468 | 0 | 75.1556 | -46.8933 | 75.2792 | -41.1255 | -43.6111 | -44.8105 | -0.14191 | 0.013733 | -0.059509 | -0.073242 | 7.4661 | 7.608 | 1.3596 | 2.6093 | 1.5015 | 2.7512 | -0.086976 | -0.022888 | 0.10071 | 0.1236 | 5.6808 | 5.7678 | 1.9958 | 3.1952 | 2.0828 | 3.2822
1730.1719 | 122.3099 | -47.0306 | 75.0137 | -0.13817 | 0.020354 | 0.087344 | 0.088786 | -0.0068637 | 0.017897 | 0.0088012 | 0.016998 | 0.024953 | 0.085214 | 0.090322 | 3.2078e-05 | 0.015323 | 0.01625 | -0.082397 | 75.1099 | -47.1039 | 74.9588 | -39.5142 | -44.339 | -45.5887 | -0.036621 | 75.0595 | -47.0123 | 75.119 | -41.2766 | -43.9087 | -44.9982 | -0.073242 | 0.09613 | -0.054932 | -0.15106 | 7.5165 | 7.5897 | 1.442 | 2.6917 | 1.5152 | 2.7649 | 0.018311 | 0.045776 | 0.10529 | 0.05951 | 5.7541 | 5.7358 | 2.0325 | 3.1219 | 2.0142 | 3.1036
1731.1719 | 122.3419 | -46.9894 | 75.1053 | -0.10765 | 0.019072 | 0.089053 | 0.093699 | 0.0036039 | 0.019606 | 0.01051 | -0.10202 | 0.026662 | 0.083292 | 0.095022 | 0.00088648 | 0.018314 | 0.014969 | -0.10071 | 75.1053 | -47.1588 | 74.9039 | -39.5599 | -44.4489 | -45.6848 | -0.10986 | 75.0778 | -47.0627 | 75.0916 | -41.2994 | -43.8309 | -45.0165 | -0.16937 | 0 | -0.20142 | -0.20142 | 7.4295 | 7.5989 | 1.3046 | 2.5406 | 1.474 | 2.71 | -0.073242 | -0.027466 | -0.013733 | 0.013733 | 5.69 | 5.7632 | 1.973 | 3.1586 | 2.0462 | 3.2318
1732.1719 | 122.2549 | -47.0673 | 75.0412 | -0.00083923 | 0.017791 | 0.087558 | 0.092845 | 0.00039951 | 0.017256 | 0.013287 | -0.11118 | 0.026876 | 0.083932 | 0.095022 | 3.2078e-05 | 0.019382 | 0.014114 | -0.032043 | 75.0595 | -47.1863 | 74.8306 | -39.7568 | -44.5312 | -45.7581 | -0.13275 | 75.0504 | -47.1085 | 75.0092 | -41.4047 | -43.9133 | -45.108 | -0.11902 | 0.01831 | -0.21057 | -0.22888 | 7.3105 | 7.4295 | 1.3092 | 2.536 | 1.4282 | 2.655 | -0.041198 | 0.009155 | -0.032044 | -0.041199 | 5.6625 | 5.7037 | 1.9592 | 3.154 | 2.0004 | 3.1952
1733.1719 | 122.2366 | -47.1176 | 74.968 | -0.13054 | 0.020354 | 0.087771 | 0.094554 | -0.0019503 | 0.018538 | 0.007306 | -0.12643 | 0.026876 | 0.085641 | 0.094167 | 0.0017411 | 0.0181 | 0.011123 | -0.073242 | 75.0092 | -47.2504 | 74.7757 | -39.7934 | -44.5541 | -45.8176 | -0.13275 | 74.9954 | -47.1725 | 74.9405 | -41.5054 | -43.9774 | -45.1813 | -0.13275 | 0.041198 | -0.19226 | -0.23346 | 7.3242 | 7.457 | 1.3 | 2.5635 | 1.4328 | 2.6962 | -0.054931 | 0.027465 | -0.027466 | -0.054931 | 5.6122 | 5.6671 | 1.9363 | 3.1403 | 1.9913 | 3.1952
1734.1719 | 122.3602 | -47.1817 | 74.8215 | -0.16869 | 0.021422 | 0.089908 | 0.09199 | -0.0013095 | 0.020247 | 0.011578 | -0.028778 | 0.026021 | 0.085214 | 0.085622 | -0.00018152 | 0.017246 | 0.014755 | -0.041199 | 74.8764 | -47.2961 | 74.7986 | -39.6835 | -44.5267 | -45.7764 | -0.11902 | 74.7803 | -47.2961 | 74.913 | -41.4093 | -43.9178 | -45.1584 | -0.11444 | 0.054932 | -0.022888 | -0.07782 | 7.4982 | 7.6126 | 1.4053 | 2.655 | 1.5198 | 2.7695 | -0.11444 | -0.041199 | 0.091553 | 0.13275 | 5.7724 | 5.8868 | 2.0233 | 3.2639 | 2.1378 | 3.3783
1735.1875 | 122.3785 | -47.2687 | 74.826 | -0.12291 | 0.0195 | 0.086703 | 0.090709 | -0.0030185 | 0.017256 | 0.0088012 | 0.013947 | 0.02474 | 0.084146 | 0.092458 | 0.0015275 | 0.01511 | 0.011764 | -0.054932 | 74.8123 | -47.3877 | 74.6613 | -39.7476 | -44.6182 | -45.8771 | 0.022888 | 74.7757 | -47.3373 | 74.8444 | -41.5375 | -43.9682 | -45.2408 | -0.11902 | -0.013733 | -0.16479 | -0.15106 | 7.5211 | 7.6401 | 1.3916 | 2.6505 | 1.5106 | 2.7695 | -0.068664 | -0.050354 | 0.01831 | 0.068664 | 5.7312 | 5.7999 | 2.0279 | 3.3005 | 2.0966 | 3.3691
1736.1719 | 122.2275 | -47.4655 | 74.6613 | -0.12291 | 0.022704 | 0.08884 | 0.094981 | -0.0013095 | 0.018111 | 0.011578 | -0.12338 | 0.027303 | 0.085001 | 0.096731 | -0.0033859 | 0.017887 | 0.013473 | -0.054932 | 74.6246 | -47.6028 | 74.4278 | -39.9719 | -44.902 | -46.1472 | -0.09613 | 74.6384 | -47.5204 | 74.5926 | -41.8442 | -44.3069 | -45.4926 | -0.13733 | -0.036621 | -0.23346 | -0.19684 | 7.4936 | 7.6309 | 1.3184 | 2.5635 | 1.4557 | 2.7008 | -0.054932 | -0.022888 | -0.068665 | -0.045777 | 5.6213 | 5.6763 | 1.973 | 3.1586 | 2.0279 | 3.2135
1737.1719 | 122.2504 | -47.4609 | 74.6429 | -0.1458 | 0.021849 | 0.090121 | 0.092631 | 0.00018591 | 0.017256 | 0.013287 | -0.12033 | 0.025594 | 0.084573 | 0.095662 | -0.00060892 | 0.0181 | 0.0094143 | -0.068665 | 74.6475 | -47.6303 | 74.4598 | -39.9353 | -44.8471 | -46.1426 | -0.064087 | 74.6246 | -47.5021 | 74.5972 | -41.8259 | -44.2795 | -45.4834 | -0.16937 | 0.004578 | -0.1831 | -0.18768 | 7.5256 | 7.695 | 1.3184 | 2.6138 | 1.4877 | 2.7832 | -0.041199 | -0.01831 | -0.045776 | -0.027466 | 5.6351 | 5.6763 | 1.9775 | 3.1815 | 2.0187 | 3.2227
1738.1719 | 122.2504 | -47.4289 | 74.6704 | -0.13817 | 0.022704 | 0.089694 | 0.092418 | 0.00061311 | 0.017684 | 0.0092286 | -0.092865 | 0.027089 | 0.085428 | 0.095235 | 0.00045928 | 0.018741 | 0.0087735 | -0.07782 | 74.6338 | -47.5845 | 74.4141 | -39.9628 | -44.8837 | -46.138 | -0.13733 | 74.6613 | -47.5204 | 74.6017 | -41.8396 | -44.2886 | -45.4971 | -0.15564 | -0.036621 | -0.25635 | -0.21973 | 7.4661 | 7.6218 | 1.2909 | 2.5452 | 1.4465 | 2.7008 | -0.091553 | -0.009155 | -0.068664 | -0.059509 | 5.5893 | 5.6808 | 1.9318 | 3.1403 | 2.0233 | 3.2318
1739.1719 | 122.3831 | -47.5021 | 74.4965 | -0.17632 | 0.023345 | 0.08884 | 0.091777 | -0.0021641 | 0.020461 | 0.0096558 | -0.040985 | 0.026876 | 0.083292 | 0.092244 | -0.0014633 | 0.017032 | 0.014328 | -0.054932 | 74.5102 | -47.6669 | 74.4507 | -39.8666 | -44.8608 | -46.1197 | -0.24261 | 74.4461 | -47.6212 | 74.5972 | -41.7892 | -44.2245 | -45.4697 | -0.16479 | 0.013733 | -0.045776 | -0.059509 | 7.6355 | 7.8003 | 1.3824 | 2.6413 | 1.5472 | 2.8061 | -0.11902 | -0.050354 | 0.10071 | 0.15106 | 5.7129 | 5.8319 | 2.0325 | 3.2776 | 2.1515 | 3.3966
1740.1875 | 122.3511 | -47.4884 | 74.5651 | -0.13817 | 0.020354 | 0.088199 | 0.09199 | -0.0034457 | 0.015975 | 0.0090148 | 0.0078429 | 0.025167 | 0.085641 | 0.087972 | -0.00060892 | 0.017032 | 0.010055 | -0.073242 | 74.6155 | -47.5891 | 74.4781 | -39.8575 | -44.8196 | -46.0648 | -0.050354 | 74.5789 | -47.5113 | 74.5651 | -41.7709 | -44.1788 | -45.4605 | -0.10071 | 0.050354 | -0.086976 | -0.13733 | 7.6309 | 7.7316 | 1.4236 | 2.6688 | 1.5244 | 2.7695 | -0.022889 | 0.013732 | 0 | -0.013732 | 5.7175 | 5.7404 | 2.0279 | 3.3096 | 2.0508 | 3.3325
1741.1719 | 122.287 | -47.0856 | 75.0092 | -0.084763 | 0.022277 | 0.089694 | 0.092845 | 0.00082671 | 0.019179 | 0.0088012 | -0.10202 | 0.027089 | 0.085641 | 0.099935 | -0.0025315 | 0.020664 | 0.012405 | 0.004578 | 75.0137 | -47.2275 | 74.8627 | -39.4958 | -44.4534 | -45.7169 | 0.045776 | 75.0275 | -47.1405 | 74.9771 | -41.3818 | -43.8538 | -45.1401 | -0.14191 | 0.004578 | -0.14648 | -0.15106 | 7.5897 | 7.7316 | 1.3687 | 2.6321 | 1.5106 | 2.774 | -0.054932 | 0.018311 | -0.032043 | -0.050354 | 5.7037 | 5.7587 | 1.9455 | 3.2318 | 2.0004 | 3.2867
1742.1719 | 122.3053 | -47.1588 | 74.9451 | -0.11528 | 0.020354 | 0.08948 | 0.094126 | -0.0010959 | 0.016829 | 0.012006 | -0.080658 | 0.027517 | 0.083078 | 0.096944 | 0.00045928 | 0.018314 | 0.014541 | -0.009155 | 75.0366 | -47.2504 | 74.7986 | -39.5325 | -44.5312 | -45.7901 | -0.009155 | 74.9542 | -47.1909 | 74.9268 | -41.4871 | -43.9453 | -45.1721 | -0.091553 | 0.091553 | -0.14648 | -0.23804 | 7.6263 | 7.7179 | 1.3687 | 2.6276 | 1.4603 | 2.7191 | -0.032044 | 0.009156 | -0.01831 | -0.027466 | 5.6717 | 5.7037 | 1.9867 | 3.2135 | 2.0187 | 3.2455
1743.1719 | 122.2687 | -46.9574 | 75.1328 | -0.13054 | 0.020781 | 0.087985 | 0.091777 | 0.0025357 | 0.020034 | 0.01051 | -0.095917 | 0.024526 | 0.082223 | 0.099508 | -0.0016769 | 0.018955 | 0.014755 | -0.050354 | 75.1602 | -47.1039 | 74.9359 | -39.4272 | -44.3527 | -45.6436 | -0.050354 | 75.119 | -47.0215 | 75.0916 | -41.3406 | -43.7805 | -45.0073 | -0.14648 | 0.027466 | -0.19684 | -0.2243 | 7.5302 | 7.6767 | 1.3138 | 2.6047 | 1.4603 | 2.7512 | -0.064087 | -0.013732 | -0.041198 | -0.027466 | 5.6168 | 5.6808 | 1.9501 | 3.1769 | 2.0142 | 3.241
1744.1719 | 122.2549 | -47.0627 | 75.0183 | -0.1458 | 0.020568 | 0.087985 | 0.094767 | 0.00018591 | 0.018538 | 0.01051 | 0.0017395 | 0.026235 | 0.086282 | 0.094381 | 0.00024568 | 0.018314 | 0.011123 | -0.10986 | 75.0275 | -47.2412 | 74.8169 | -39.5874 | -44.6594 | -45.8725 | -0.045776 | 75.0366 | -47.113 | 74.9634 | -41.4139 | -43.9087 | -45.1172 | -0.17853 | 0.009155 | -0.20142 | -0.21057 | 7.4753 | 7.6538 | 1.1902 | 2.4033 | 1.3687 | 2.5818 | -0.050354 | 0.01831 | -0.054932 | -0.073242 | 5.6488 | 5.6992 | 1.9455 | 3.154 | 1.9958 | 3.2043
1745.1719 | 122.2916 | -47.2275 | 74.8901 | -0.12291 | 0.02014 | 0.088199 | 0.09434 | 0.00082671 | 0.019606 | 0.0094422 | -0.077606 | 0.026876 | 0.085001 | 0.092031 | 3.2078e-05 | 0.016178 | 0.014755 | -0.022888 | 74.9405 | -47.319 | 74.675 | -39.6332 | -44.6091 | -45.8725 | -0.009155 | 74.8764 | -47.2504 | 74.8672 | -41.51 | -43.9957 | -45.2133 | -0.091553 | 0.050354 | -0.21515 | -0.2655 | 7.5943 | 7.6859 | 1.355 | 2.6184 | 1.4465 | 2.71 | -0.022888 | -0.013733 | -0.022888 | -0.009155 | 5.7175 | 5.7404 | 2.0142 | 3.2318 | 2.037 | 3.2547
1746.1719 | 122.3099 | -47.2275 | 74.8764 | -0.16869 | 0.019927 | 0.086276 | 0.091777 | -0.0021641 | 0.020888 | 0.015424 | -0.068451 | 0.02773 | 0.085428 | 0.092031 | -0.00018152 | 0.017246 | 0.015182 | -0.073242 | 74.794 | -47.4289 | 74.7665 | -39.505 | -44.5404 | -45.845 | 0.013733 | 74.8856 | -47.2458 | 74.8581 | -41.5283 | -43.9819 | -45.2408 | -0.20142 | -0.082398 | -0.10986 | -0.027465 | 7.7225 | 7.9239 | 1.3824 | 2.6871 | 1.5839 | 2.8885 | -0.018311 | 0.009155 | -0.018311 | -0.027466 | 5.6992 | 5.7175 | 1.9867 | 3.2455 | 2.005 | 3.2639
1747.1719 | 122.3099 | -47.3465 | 74.7482 | -0.13054 | 0.02014 | 0.087131 | 0.089854 | -0.0013095 | 0.017043 | 0.012646 | -0.068451 | 0.027303 | 0.085214 | 0.087118 | -0.0014633 | 0.017032 | 0.015182 | -0.013733 | 74.7757 | -47.4747 | 74.5789 | -39.7202 | -44.6823 | -45.9412 | 0.041199 | 74.7437 | -47.374 | 74.7665 | -41.5924 | -44.0781 | -45.3369 | -0.12817 | 0.027466 | -0.16937 | -0.19684 | 7.6263 | 7.7545 | 1.4053 | 2.6642 | 1.5335 | 2.7924 | -0.027465 | -0.004578 | 0.018311 | 0.022889 | 5.7541 | 5.7816 | 2.0096 | 3.2684 | 2.037 | 3.2959
1748.1719 | 122.3007 | -47.081 | 75.0458 | -0.077133 | 0.021422 | 0.088626 | 0.091777 | 0.0014675 | 0.019606 | 0.010937 | -0.053192 | 0.026449 | 0.085428 | 0.094167 | 0.0013139 | 0.015964 | 0.0083463 | -0.018311 | 75.032 | -47.2183 | 74.826 | -39.505 | -44.4305 | -45.6985 | 0.013733 | 75.0366 | -47.1085 | 75.0046 | -41.3406 | -43.8217 | -45.0531 | -0.13733 | -0.013733 | -0.21973 | -0.20599 | 7.576 | 7.7133 | 1.3824 | 2.6505 | 1.5198 | 2.7878 | -0.027465 | -0.009155 | -0.041198 | -0.032043 | 5.7404 | 5.7678 | 2.0279 | 3.2593 | 2.0554 | 3.2867
1749.1719 | 122.3099 | -47.1176 | 74.9863 | -0.084763 | 0.0195 | 0.088199 | 0.091777 | -2.7692e-05 | 0.018538 | 0.01051 | -0.065399 | 0.026449 | 0.084573 | 0.095022 | 0.00045928 | 0.0181 | 0.011764 | 0.022888 | 74.9863 | -47.2733 | 74.8123 | -39.5004 | -44.4717 | -45.7443 | 0.050354 | 74.968 | -47.1588 | 74.9771 | -41.3956 | -43.8766 | -45.1172 | -0.15564 | 0 | -0.17395 | -0.17395 | 7.6172 | 7.7728 | 1.3733 | 2.6459 | 1.5289 | 2.8015 | -0.041198 | -0.01831 | -0.009155 | 0.009155 | 5.722 | 5.7632 | 2.0004 | 3.241 | 2.0416 | 3.2822
1750.1719 | 122.287 | -47.1039 | 75.0183 | -0.10002 | 0.021209 | 0.089053 | 0.093486 | -0.0010959 | 0.019179 | 0.012433 | -0.05014 | 0.028798 | 0.086496 | 0.084768 | 0.0021683 | 0.017246 | 0.01091 | 0.009155 | 75.0046 | -47.2412 | 74.8215 | -39.4958 | -44.4489 | -45.7123 | 0.059509 | 75 | -47.1405 | 74.9725 | -41.3544 | -43.84 | -45.0943 | -0.13733 | -0.013733 | -0.19684 | -0.18311 | 7.608 | 7.7454 | 1.3916 | 2.655 | 1.5289 | 2.7924 | -0.036621 | -0.018311 | -0.045777 | -0.027466 | 5.7495 | 5.7861 | 2.0096 | 3.2639 | 2.0462 | 3.3005
1751.1719 | 122.3099 | -46.9894 | 75.1282 | -0.11528 | 0.021849 | 0.088626 | 0.093272 | -0.00066869 | 0.018965 | 0.010937 | -0.05014 | 0.027303 | 0.083078 | 0.091817 | -0.0012497 | 0.017673 | 0.011337 | 0.004578 | 75.1007 | -47.1222 | 74.9313 | -39.4089 | -44.3573 | -45.6207 | 0.041199 | 75.0916 | -47.0398 | 75.1099 | -41.2491 | -43.7485 | -44.9982 | -0.13275 | -0.027466 | -0.19684 | -0.16937 | 7.5806 | 7.7133 | 1.3687 | 2.6321 | 1.5015 | 2.7649 | -0.050354 | -0.036621 | -0.018311 | 0.01831 | 5.7404 | 5.7907 | 1.9913 | 3.241 | 2.0416 | 3.2913
1752.1719 | 122.3099 | -47.0535 | 75.0641 | -0.13054 | 0.020995 | 0.087985 | 0.092845 | -0.00024129 | 0.020034 | 0.0094422 | -0.056244 | 0.027517 | 0.083505 | 0.08669 | -0.0012497 | 0.016605 | 0.012405 | -0.004578 | 75.0366 | -47.2229 | 74.8535 | -39.4638 | -44.4305 | -45.6985 | 0.032043 | 74.9954 | -47.1268 | 75.0366 | -41.3177 | -43.84 | -45.076 | -0.16937 | -0.027466 | -0.21057 | -0.1831 | 7.5897 | 7.7591 | 1.355 | 2.623 | 1.5244 | 2.7924 | -0.073242 | -0.068665 | -0.027466 | 0.041199 | 5.7358 | 5.809 | 1.9775 | 3.2135 | 2.0508 | 3.2867
1753.1719 | 122.2687 | -47.0993 | 75 | -0.12291 | 0.019286 | 0.089053 | 0.093058 | -0.00045509 | 0.018111 | 0.0098694 | -0.047089 | 0.028585 | 0.086282 | 0.091176 | 0.00067288 | 0.016391 | 0.01326 | -0.004578 | 74.9222 | -47.3053 | 74.7437 | -39.6149 | -44.5541 | -45.813 | 0.050354 | 74.9039 | -47.2321 | 74.8901 | -41.4139 | -43.8583 | -45.0623 | -0.20599 | -0.07782 | -0.25635 | -0.17853 | 7.4844 | 7.6904 | 1.2863 | 2.5452 | 1.4923 | 2.7512 | -0.13275 | -0.09613 | -0.10986 | -0.013733 | 5.6854 | 5.8182 | 2.037 | 3.241 | 2.1698 | 3.3737
1754.1719 | 122.2778 | -46.8109 | 75.3296 | -0.13817 | 0.02014 | 0.087771 | 0.093058 | 0.0048855 | 0.015548 | 0.0098694 | -0.056244 | 0.026876 | 0.087137 | 0.089467 | -0.00060892 | 0.017459 | 0.011123 | -0.064087 | 75.325 | -46.9391 | 75.1007 | -39.299 | -44.2154 | -45.4514 | -0.068665 | 75.325 | -46.843 | 75.2609 | -41.066 | -43.5928 | -44.8151 | -0.12817 | -0.004578 | -0.22888 | -0.2243 | 7.5119 | 7.6401 | 1.3596 | 2.5955 | 1.4877 | 2.7237 | -0.032044 | -0.004578 | -0.068665 | -0.064087 | 5.7449 | 5.777 | 1.9958 | 3.2181 | 2.0279 | 3.2501
1755.1719 | 122.2824 | -46.994 | 75.1602 | -0.054245 | 0.019713 | 0.087558 | 0.093486 | 0.0014675 | 0.019606 | 0.0085876 | -0.0074159 | 0.026449 | 0.08671 | 0.090108 | -0.00082252 | 0.017459 | 0.016464 | -0.022888 | 75.174 | -47.0673 | 74.9222 | -39.4455 | -44.3161 | -45.5841 | 0 | 75.1556 | -46.9986 | 75.087 | -41.1896 | -43.689 | -44.9432 | -0.073242 | 0.013733 | -0.23804 | -0.25177 | 7.5485 | 7.6218 | 1.4099 | 2.6779 | 1.4832 | 2.7512 | -0.004577 | -0.004577 | -0.073242 | -0.068665 | 5.8044 | 5.809 | 2.0508 | 3.3051 | 2.0554 | 3.3096
1756.1719 | 122.3099 | -46.7606 | 75.3296 | -0.084763 | 0.021422 | 0.086703 | 0.09199 | 0.0012539 | 0.016829 | 0.0088012 | -0.040985 | 0.026662 | 0.086069 | 0.094594 | 0.00067288 | 0.016391 | 0.011978 | 0 | 75.3571 | -46.8658 | 75.1831 | -39.2212 | -44.1101 | -45.3461 | 0.041199 | 75.3387 | -46.7926 | 75.3387 | -41.0019 | -43.5013 | -44.7327 | -0.10529 | 0.027466 | -0.14648 | -0.17395 | 7.5394 | 7.6447 | 1.4145 | 2.6505 | 1.5198 | 2.7557 | -0.032044 | 0.009155 | 0.009155 | 0 | 5.7587 | 5.7907 | 2.0279 | 3.2593 | 2.0599 | 3.2913
1757.1719 | 122.2824 | -46.7377 | 75.3525 | -0.061874 | 0.019927 | 0.086062 | 0.091563 | 0.0010403 | 0.01982 | 0.0081604 | -0.040985 | 0.025808 | 0.085214 | 0.085836 | 3.2078e-05 | 0.017887 | 0.010483 | 0 | 75.3754 | -46.8567 | 75.1968 | -39.2258 | -44.0872 | -45.3415 | 0.036621 | 75.3387 | -46.7606 | 75.3433 | -40.9332 | -43.4738 | -44.7281 | -0.11902 | 0.022888 | -0.15564 | -0.17853 | 7.5119 | 7.6309 | 1.3962 | 2.6505 | 1.5152 | 2.7695 | -0.022888 | -0.013733 | -0.009155 | 0.004578 | 5.8044 | 5.8273 | 2.0096 | 3.2639 | 2.0325 | 3.2867
1758.1719 | 122.3099 | -46.5591 | 75.5219 | -0.10002 | 0.019072 | 0.087344 | 0.091136 | 0.0014675 | 0.018538 | 0.0085876 | -0.034882 | 0.027089 | 0.086282 | 0.09139 | -0.00039512 | 0.017246 | 0.013473 | -0.009155 | 75.5173 | -46.7239 | 75.3387 | -39.0976 | -43.9453 | -45.2179 | 0.036621 | 75.499 | -46.6278 | 75.499 | -40.7684 | -43.3044 | -44.5587 | -0.16479 | -0.004578 | -0.18311 | -0.17853 | 7.4615 | 7.6263 | 1.3412 | 2.6138 | 1.506 | 2.7786 | -0.068665 | -0.022889 | -0.022889 | 0 | 5.7907 | 5.8594 | 2.0004 | 3.2547 | 2.0691 | 3.3234
1759.1719 | 122.2641 | -46.5088 | 75.5814 | -0.21446 | 0.019713 | 0.087985 | 0.090281 | 0.00061311 | 0.017897 | 0.0079468 | -0.05014 | 0.026235 | 0.085001 | 0.084554 | -0.0023177 | 0.019382 | 0.014755 | -0.07782 | 75.6088 | -46.6553 | 75.3708 | -39.1113 | -43.9453 | -45.1813 | -0.082397 | 75.5997 | -46.5775 | 75.5081 | -40.7913 | -43.3228 | -44.5404 | -0.14648 | 0.027466 | -0.21057 | -0.23804 | 7.3975 | 7.5439 | 1.3275 | 2.5635 | 1.474 | 2.71 | -0.068665 | 0.01831 | -0.073242 | -0.091552 | 5.7175 | 5.7861 | 1.9684 | 3.186 | 2.037 | 3.2547
1760.1719 | 122.287 | -46.5042 | 75.6042 | -0.092392 | 0.021422 | 0.089694 | 0.093058 | -0.0010959 | 0.017043 | 0.010724 | -0.053192 | 0.026021 | 0.085641 | 0.094381 | 0.0013139 | 0.017459 | 0.0083463 | 0.013733 | 75.5814 | -46.6599 | 75.4028 | -39.0701 | -43.8721 | -45.1309 | 0.036621 | 75.5905 | -46.5546 | 75.5585 | -40.7272 | -43.2678 | -44.5175 | -0.15564 | -0.022888 | -0.20142 | -0.17853 | 7.4341 | 7.5897 | 1.3733 | 2.6321 | 1.5289 | 2.7878 | -0.050354 | -0.013733 | -0.045776 | -0.032043 | 5.777 | 5.8273 | 1.9867 | 3.2364 | 2.037 | 3.2867
1761.1719 | 122.2778 | -46.3531 | 75.769 | -0.069504 | 0.019286 | 0.087771 | 0.090922 | 0.0018949 | 0.01982 | 0.0081604 | -0.013519 | 0.02773 | 0.086496 | 0.08904 | -0.0023177 | 0.018314 | 0.0092007 | -0.013733 | 75.7736 | -46.4905 | 75.5447 | -38.9786 | -43.7302 | -45.0027 | -0.004578 | 75.737 | -46.3989 | 75.6729 | -40.5991 | -43.1442 | -44.3939 | -0.13733 | 0.004578 | -0.2243 | -0.22888 | 7.3746 | 7.5119 | 1.3504 | 2.623 | 1.4877 | 2.7603 | -0.045777 | -0.032043 | -0.09613 | -0.064087 | 5.7541 | 5.7999 | 1.9592 | 3.2089 | 2.005 | 3.2547
1762.1719 | 122.2824 | -46.4172 | 75.7324 | -0.054245 | 0.018218 | 0.086062 | 0.091563 | -0.0010959 | 0.018325 | 0.011578 | -0.010467 | 0.025594 | 0.083719 | 0.094167 | 0.00024568 | 0.016178 | 0.013901 | 0.027466 | 75.7782 | -46.4813 | 75.499 | -39.0427 | -43.7988 | -45.0348 | -0.013733 | 75.705 | -46.4447 | 75.6226 | -40.6265 | -43.1671 | -44.4305 | -0.064087 | 0.045776 | -0.23346 | -0.27924 | 7.3746 | 7.4387 | 1.3824 | 2.6184 | 1.4465 | 2.6825 | -0.027466 | -0.027466 | -0.10986 | -0.082397 | 5.7907 | 5.8182 | 1.9867 | 3.2501 | 2.0142 | 3.2776
1763.1719 | 122.319 | -46.582 | 75.5493 | -0.15343 | 0.020781 | 0.086062 | 0.092845 | -0.00066869 | 0.018752 | 0.0092286 | -0.019623 | 0.028158 | 0.08436 | 0.098226 | 3.2078e-05 | 0.017673 | 0.015182 | -0.018311 | 75.5402 | -46.6965 | 75.3296 | -39.2212 | -43.9453 | -45.2087 | -0.004578 | 75.531 | -46.6278 | 75.4761 | -40.7593 | -43.309 | -44.5724 | -0.11444 | -0.009155 | -0.21973 | -0.21057 | 7.3608 | 7.4753 | 1.3733 | 2.6367 | 1.4877 | 2.7512 | -0.045777 | -0.01831 | -0.073242 | -0.054932 | 5.8228 | 5.8685 | 2.0096 | 3.273 | 2.0554 | 3.3188
1764.1719 | 122.3053 | -46.7468 | 75.3662 | -0.10765 | 0.020568 | 0.08884 | 0.093272 | 0.0012539 | 0.017043 | 0.0085876 | -0.034882 | 0.027517 | 0.084787 | 0.091604 | 3.2078e-05 | 0.017887 | 0.010055 | 0 | 75.3616 | -46.9025 | 75.1877 | -39.3265 | -44.101 | -45.369 | 0.050354 | 75.3479 | -46.8018 | 75.3296 | -40.9286 | -43.4875 | -44.7327 | -0.15564 | -0.004578 | -0.17853 | -0.17395 | 7.4203 | 7.576 | 1.3779 | 2.6459 | 1.5335 | 2.8015 | -0.054932 | -0.018311 | -0.036621 | -0.01831 | 5.8182 | 5.8731 | 2.0142 | 3.2593 | 2.0691 | 3.3142
1765.1719 | 122.287 | -46.4905 | 75.6088 | -0.10765 | 0.018218 | 0.087131 | 0.091349 | -0.00024129 | 0.017043 | 0.012006 | -0.010467 | 0.02773 | 0.086496 | 0.097585 | -0.00018152 | 0.013614 | 0.01326 | 0.18768 | 75.6821 | -46.5866 | 75.3525 | -39.2349 | -43.8858 | -45.163 | -0.018311 | 75.5905 | -46.5637 | 75.531 | -40.7135 | -43.2724 | -44.5358 | -0.09613 | 0.073242 | -0.25635 | -0.32959 | 7.2556 | 7.3517 | 1.3275 | 2.6047 | 1.4236 | 2.7008 | -0.073242 | -0.018311 | -0.07782 | -0.059509 | 5.777 | 5.8502 | 1.9547 | 3.2181 | 2.0279 | 3.2913
1766.1719 | 122.2687 | -46.5591 | 75.5676 | -0.16869 | 0.019927 | 0.087344 | 0.092204 | 0.00039951 | 0.019606 | 0.0077332 | -0.016571 | 0.027517 | 0.083719 | 0.096731 | -0.0012497 | 0.016818 | 0.01091 | -0.018311 | 75.5722 | -46.7148 | 75.3525 | -39.3494 | -43.9545 | -45.195 | 0 | 75.5539 | -46.6049 | 75.499 | -40.7639 | -43.2999 | -44.545 | -0.15564 | 0.004578 | -0.21515 | -0.21973 | 7.2098 | 7.3654 | 1.3641 | 2.6047 | 1.5198 | 2.7603 | -0.045776 | -0.013733 | -0.068665 | -0.054932 | 5.7953 | 5.8411 | 2.0142 | 3.2593 | 2.0599 | 3.3051
1767.1719 | 122.2824 | -46.7651 | 75.3433 | -0.10002 | 0.020781 | 0.085422 | 0.091777 | -0.00066869 | 0.018111 | 0.010724 | -0.019623 | 0.027303 | 0.085641 | 0.090322 | -0.00082252 | 0.018955 | 0.012832 | -0.022888 | 75.3708 | -46.9162 | 75.119 | -39.6927 | -44.1467 | -45.4147 | -0.013733 | 75.2975 | -46.8338 | 75.2426 | -40.9653 | -43.5196 | -44.7784 | -0.15106 | 0.027466 | -0.2243 | -0.25177 | 7.0724 | 7.2235 | 1.3504 | 2.6184 | 1.5015 | 2.7695 | -0.068664 | -0.045777 | -0.10071 | -0.054931 | 5.7999 | 5.8685 | 1.9867 | 3.2455 | 2.0554 | 3.3142
1768.1719 | 122.2229 | -46.8338 | 75.293 | -0.084763 | 0.019713 | 0.085422 | 0.090709 | -0.0034457 | 0.017684 | 0.011151 | 0.0047913 | 0.02773 | 0.085855 | 0.09908 | 0.0011001 | 0.016818 | 0.010055 | -0.013733 | 75.2792 | -46.9482 | 75.0595 | -40.1688 | -44.22 | -45.4605 | 0 | 75.2609 | -46.8613 | 75.1968 | -41.0202 | -43.5791 | -44.8517 | -0.11444 | -0.013733 | -0.23346 | -0.21973 | 6.665 | 6.7795 | 1.3733 | 2.6138 | 1.4877 | 2.7283 | -0.027466 | -0.032044 | -0.096131 | -0.064087 | 5.8136 | 5.8411 | 1.9821 | 3.2547 | 2.0096 | 3.2822
1769.1719 | 122.2824 | -47.1176 | 74.9771 | -0.13817 | 0.022277 | 0.086917 | 0.093699 | -2.7692e-05 | 0.016829 | 0.01051 | -0.077606 | 0.026449 | 0.084573 | 0.095662 | -0.0016769 | 0.017673 | 0.012832 | -0.004578 | 74.9542 | -47.287 | 74.9359 | -40.6174 | -44.3802 | -45.7031 | 0.032043 | 74.9817 | -47.1497 | 74.9405 | -41.2903 | -43.8766 | -45.1172 | -0.16937 | -0.022888 | -0.041199 | -0.018311 | 6.5002 | 6.6696 | 1.4145 | 2.7374 | 1.5839 | 2.9068 | -0.032043 | 0.004577 | -0.036621 | -0.041198 | 5.8273 | 5.8594 | 2.0004 | 3.241 | 2.0325 | 3.273
1770.1719 | 122.2458 | -47.3648 | 74.762 | -0.092392 | 0.019713 | 0.086062 | 0.090922 | 0.00018591 | 0.018325 | 0.013074 | 0.010895 | 0.027089 | 0.082651 | 0.088827 | -0.0016769 | 0.014469 | 0.0087735 | -0.036621 | 74.7894 | -47.4747 | 74.5285 | -41.0385 | -44.7144 | -45.9869 | 0.013733 | 74.7528 | -47.3694 | 74.6796 | -41.5375 | -44.1101 | -45.3827 | -0.10986 | 0.027466 | -0.23346 | -0.26093 | 6.3263 | 6.4362 | 1.3779 | 2.6505 | 1.4877 | 2.7603 | -0.004578 | -0.009155 | -0.082398 | -0.073243 | 5.8273 | 5.8319 | 1.9821 | 3.2547 | 1.9867 | 3.2593
1771.1719 | 122.2412 | -47.3831 | 74.7482 | -0.084763 | 0.020354 | 0.08414 | 0.090068 | 0.0025357 | 0.018752 | 0.0085876 | 0.023102 | 0.02538 | 0.085855 | 0.091176 | 3.2078e-05 | 0.014896 | 0.011978 | 0.009155 | 74.7528 | -47.4701 | 74.5468 | -41.0294 | -44.7372 | -45.9915 | 0.022888 | 74.7253 | -47.4106 | 74.6475 | -41.5329 | -44.0872 | -45.3918 | -0.086975 | 0.004578 | -0.20142 | -0.20599 | 6.3538 | 6.4407 | 1.3916 | 2.6459 | 1.4786 | 2.7328 | -0.027465 | -0.022888 | -0.10071 | -0.07782 | 5.8502 | 5.8777 | 1.9913 | 3.2959 | 2.0187 | 3.3234
1772.1719 | 122.2458 | -47.7173 | 74.4003 | -0.1458 | 0.02249 | 0.086276 | 0.090922 | 0.00061311 | 0.016188 | 0.0070922 | 0.0017395 | 0.026235 | 0.084146 | 0.085836 | -0.0025315 | 0.015537 | 0.0186 | 0.041199 | 74.4324 | -47.7951 | 74.2218 | -41.3223 | -44.9982 | -46.2799 | 0.027466 | 74.3866 | -47.7402 | 74.3042 | -41.8396 | -44.4443 | -45.726 | -0.07782 | 0.032043 | -0.17853 | -0.21057 | 6.395 | 6.4728 | 1.4374 | 2.7191 | 1.5152 | 2.7969 | -0.022888 | -0.013733 | -0.096131 | -0.082398 | 5.8777 | 5.9006 | 1.9913 | 3.273 | 2.0142 | 3.2959
1773.1719 | 122.2366 | -47.7127 | 74.4049 | -0.10765 | 0.0195 | 0.08649 | 0.092418 | -0.0013095 | 0.018752 | 0.010724 | -0.010467 | 0.024312 | 0.084146 | 0.086904 | -0.0012497 | 0.017887 | 0.012832 | -0.018311 | 74.4003 | -47.8455 | 74.1806 | -41.3818 | -45.0714 | -46.3394 | 0.009155 | 74.3774 | -47.7585 | 74.3134 | -41.8396 | -44.4489 | -45.7352 | -0.13275 | -0.004577 | -0.2243 | -0.21973 | 6.3309 | 6.4636 | 1.3733 | 2.6413 | 1.506 | 2.774 | -0.045776 | -0.027466 | -0.091553 | -0.064087 | 5.8731 | 5.9189 | 1.9775 | 3.2639 | 2.0233 | 3.3096
1774.1719 | 122.2961 | -47.85 | 74.2264 | -0.092392 | 0.021422 | 0.086062 | 0.093058 | 0.0029629 | 0.019179 | 0.013715 | -0.10202 | 0.027089 | 0.08436 | 0.09374 | -0.0010361 | 0.017887 | 0.011123 | 0.022888 | 74.231 | -48.0331 | 74.0662 | -41.5192 | -45.1996 | -46.4859 | 0.059509 | 74.2218 | -47.9095 | 74.2035 | -41.9586 | -44.5862 | -45.8405 | -0.1831 | 0.004578 | -0.16022 | -0.16479 | 6.3309 | 6.514 | 1.3641 | 2.6505 | 1.5472 | 2.8336 | -0.059509 | -0.004577 | -0.022888 | -0.018311 | 5.8914 | 5.9509 | 2.0096 | 3.2639 | 2.0691 | 3.3234
1775.1719 | 122.3419 | -47.8729 | 74.1577 | -0.092392 | 0.02014 | 0.086276 | 0.090922 | -0.0028049 | 0.018325 | 0.012646 | -0.0013123 | 0.024099 | 0.083719 | 0.091604 | -0.0014633 | 0.014896 | 0.0096279 | 0.045776 | 74.1669 | -47.9919 | 73.938 | -41.4917 | -45.0439 | -46.4767 | 0.15564 | 74.1119 | -47.9416 | 74.2081 | -41.9128 | -44.5221 | -45.8405 | -0.11902 | 0.009155 | -0.21973 | -0.22888 | 6.3812 | 6.5002 | 1.3962 | 2.829 | 1.5152 | 2.948 | -0.068664 | -0.045777 | 0.050354 | 0.096131 | 5.9601 | 6.0287 | 2.0325 | 3.3508 | 2.1011 | 3.4195
1776.1719 | 122.3282 | -48.0011 | 74.0524 | -0.16106 | 0.020568 | 0.085422 | 0.089854 | -0.0015231 | 0.01747 | 0.01051 | 0.032257 | 0.026235 | 0.085001 | 0.087331 | -0.00082252 | 0.017032 | 0.015823 | 0.050354 | 74.0662 | -48.0972 | 74.0067 | -41.5283 | -45.2408 | -46.5271 | 0.15564 | 74.0341 | -48.0286 | 74.0616 | -42.0776 | -44.5267 | -45.8908 | -0.09613 | 0.013733 | -0.045776 | -0.059509 | 6.4728 | 6.5689 | 1.474 | 2.7603 | 1.5701 | 2.8564 | -0.027465 | -0.01831 | 0.009155 | 0.027465 | 5.9235 | 5.9509 | 2.1103 | 3.4744 | 2.1378 | 3.5019
1777.1719 | 122.3831 | -48.0331 | 74.025 | -0.10765 | 0.020568 | 0.085849 | 0.088786 | 0.00018591 | 0.018325 | 0.010937 | 0.016998 | 0.025808 | 0.085428 | 0.09139 | -0.00060892 | 0.016605 | 0.010696 | 0.16022 | 74.0021 | -48.1842 | 73.9014 | -41.6565 | -45.2545 | -46.582 | 0.14648 | 74.0158 | -48.0515 | 74.0616 | -42.0593 | -44.6594 | -45.9732 | -0.15106 | -0.022888 | -0.1236 | -0.10071 | 6.3766 | 6.5277 | 1.4511 | 2.7786 | 1.6022 | 2.9297 | -0.018311 | -0.009155 | 0.036621 | 0.045776 | 5.9738 | 5.9921 | 2.0599 | 3.3737 | 2.0782 | 3.392
1778.1719 | 122.4014 | -48.0698 | 74.0158 | -0.13054 | 0.019927 | 0.086062 | 0.091563 | -0.0019503 | 0.015975 | 0.010297 | -0.0013123 | 0.026876 | 0.083719 | 0.10079 | -0.0027451 | 0.014469 | 0.0081327 | 0.027466 | 73.9792 | -48.1522 | 73.9334 | -41.6199 | -45.3323 | -46.5912 | 0.009155 | 73.9609 | -48.1018 | 74.0753 | -42.1234 | -44.7098 | -45.9869 | -0.082398 | -0.036621 | -0.082397 | -0.045776 | 6.4499 | 6.5323 | 1.4786 | 2.7374 | 1.561 | 2.8198 | -0.032044 | -0.054932 | 0.059509 | 0.11444 | 5.9464 | 5.9784 | 2.0828 | 3.36 | 2.1149 | 3.392
1779.1719 | 122.2778 | -48.1247 | 73.9563 | -0.10765 | 0.018218 | 0.084994 | 0.092418 | -0.0013095 | 0.017043 | 0.0096558 | -0.062347 | 0.029226 | 0.086496 | 0.093953 | -0.0012497 | 0.016818 | 0.011551 | -0.041199 | 73.9929 | -48.2895 | 73.7823 | -41.8579 | -45.5154 | -46.7789 | -0.009155 | 73.9655 | -48.1934 | 73.8968 | -42.215 | -44.8654 | -46.1334 | -0.16479 | 0.036621 | -0.17395 | -0.21057 | 6.2668 | 6.4316 | 1.3458 | 2.6093 | 1.5106 | 2.774 | -0.068664 | 0.009155 | -0.059509 | -0.068664 | 5.9097 | 5.9784 | 1.9913 | 3.2593 | 2.0599 | 3.3279
1780.1719 | 122.3602 | -48.1613 | 73.8556 | -0.10002 | 0.020568 | 0.08649 | 0.089213 | 0.0016813 | 0.018325 | 0.010297 | 0.1055 | 0.026449 | 0.086496 | 0.093953 | 0.0011001 | 0.015537 | 0.013901 | 0.07782 | 73.9105 | -48.262 | 73.8098 | -41.7343 | -45.4514 | -46.6461 | -0.036621 | 73.8373 | -48.2437 | 73.9334 | -42.16 | -44.8242 | -46.1243 | -0.10071 | 0.054931 | -0.045777 | -0.10071 | 6.427 | 6.5277 | 1.5152 | 2.71 | 1.6159 | 2.8107 | -0.082397 | -0.018311 | 0.07782 | 0.096131 | 6.0013 | 6.0837 | 2.037 | 3.3371 | 2.1194 | 3.4195
1781.1719 | 122.4655 | -48.2986 | 73.7778 | -0.15343 | 0.019927 | 0.086276 | 0.091349 | -2.7692e-05 | 0.021529 | 0.010083 | 0.010895 | 0.026449 | 0.086282 | 0.083059 | -0.00082252 | 0.013828 | 0.013473 | 0.022888 | 73.7366 | -48.4268 | 73.7045 | -41.8396 | -45.5383 | -46.8201 | 0.14191 | 73.7183 | -48.3398 | 73.819 | -42.2287 | -44.9203 | -46.376 | -0.12817 | -0.041199 | -0.073242 | -0.032043 | 6.459 | 6.5872 | 1.4786 | 2.7603 | 1.6068 | 2.8885 | -0.041199 | -0.059509 | 0.041199 | 0.10071 | 6.0699 | 6.1111 | 1.9226 | 3.3783 | 1.9638 | 3.4195
1782.1719 | 122.3831 | -48.3215 | 73.7091 | -0.092392 | 0.020781 | 0.085422 | 0.091563 | 0.0029629 | 0.015334 | 0.01051 | -0.016571 | 0.025594 | 0.085001 | 0.083486 | -0.00018152 | 0.015537 | 0.011978 | -0.013733 | 73.7137 | -48.4909 | 73.613 | -41.9724 | -45.6482 | -46.8933 | 0.07782 | 73.7183 | -48.3902 | 73.7686 | -42.3019 | -44.9936 | -46.257 | -0.16937 | 0.004578 | -0.09613 | -0.10071 | 6.3492 | 6.5186 | 1.4282 | 2.6733 | 1.5976 | 2.8427 | -0.068665 | 0.009156 | 0.05951 | 0.050354 | 6.0196 | 6.0883 | 2.0645 | 3.3279 | 2.1332 | 3.3966
1783.1719 | 122.4106 | -48.3124 | 73.7274 | -0.11528 | 0.0195 | 0.085422 | 0.091563 | -2.7692e-05 | 0.018538 | 0.008374 | 0.0047913 | 0.025808 | 0.085001 | 0.091817 | 0.00045928 | 0.014255 | 0.012192 | 0.004578 | 73.7366 | -48.4314 | 73.6679 | -41.89 | -45.6024 | -46.8521 | 0.11444 | 73.7228 | -48.3719 | 73.7869 | -42.2607 | -44.9524 | -46.2616 | -0.11902 | 0.009155 | -0.059509 | -0.068664 | 6.4224 | 6.5414 | 1.4603 | 2.71 | 1.5793 | 2.829 | -0.059509 | -0.004578 | 0.059509 | 0.064087 | 6.0516 | 6.1111 | 2.0508 | 3.36 | 2.1103 | 3.4195
1784.1719 | 122.2504 | -48.3078 | 73.7915 | -0.10002 | 0.019286 | 0.084353 | 0.091349 | -0.00088229 | 0.017043 | 0.0098694 | -0.059296 | 0.026662 | 0.082864 | 0.095022 | -0.0033859 | 0.017459 | 0.011551 | 0.032043 | 73.7869 | -48.4314 | 73.6633 | -41.9174 | -45.6161 | -46.8475 | 0.059509 | 73.7686 | -48.3398 | 73.732 | -42.2882 | -45.0256 | -46.312 | -0.1236 | -0.004578 | -0.12817 | -0.1236 | 6.3904 | 6.514 | 1.4603 | 2.6917 | 1.5839 | 2.8152 | -0.032044 | -0.022888 | -0.059509 | -0.036621 | 6.0196 | 6.0516 | 1.9958 | 3.2822 | 2.0279 | 3.3142
1785.1719 | 122.3831 | -48.0469 | 73.9883 | -0.061874 | 0.018431 | 0.085849 | 0.091349 | -0.0028049 | 0.018965 | 0.012646 | -0.025726 | 0.028371 | 0.08436 | 0.089254 | -0.0012497 | 0.013187 | 0.011123 | -0.022888 | 73.9746 | -48.23 | 73.8876 | -41.7206 | -45.3827 | -46.6141 | 0.041199 | 73.9929 | -48.111 | 74.0479 | -42.0181 | -44.7281 | -46.0098 | -0.1831 | -0.013733 | -0.10071 | -0.086975 | 6.3263 | 6.5094 | 1.4328 | 2.6642 | 1.6159 | 2.8473 | -0.064087 | 0.004578 | 0.05951 | 0.054932 | 6.0287 | 6.0928 | 2.037 | 3.3188 | 2.1011 | 3.3829
1786.1719 | 122.3511 | -48.1247 | 73.9288 | -0.16869 | 0.018218 | 0.087558 | 0.091136 | -0.0040867 | 0.019393 | 0.0077332 | -0.05014 | 0.024953 | 0.084787 | 0.097585 | 3.2078e-05 | 0.016178 | 0.0087735 | 0.064087 | 73.9151 | -48.2391 | 73.851 | -41.7526 | -45.4239 | -46.6782 | 0.11902 | 73.9243 | -48.175 | 73.9563 | -42.1097 | -44.838 | -46.0968 | -0.11444 | -0.013733 | -0.07782 | -0.064087 | 6.3721 | 6.4865 | 1.4465 | 2.7008 | 1.561 | 2.8152 | -0.050354 | -0.004578 | 0.027466 | 0.032044 | 6.015 | 6.0654 | 2.0279 | 3.2867 | 2.0782 | 3.3371
1787.1719 | 122.3557 | -48.1155 | 73.9563 | -0.11528 | 0.019927 | 0.087985 | 0.089854 | -0.00066869 | 0.017684 | 0.0085876 | -0.016571 | 0.02538 | 0.083932 | 0.089895 | -0.00060892 | 0.016818 | 0.011978 | -0.022888 | 73.938 | -48.2437 | 73.8602 | -41.7252 | -45.4102 | -46.637 | 0.068665 | 73.9197 | -48.1613 | 74.0204 | -42.041 | -44.7876 | -46.0464 | -0.12817 | -0.018311 | -0.096131 | -0.07782 | 6.3904 | 6.5186 | 1.4786 | 2.7054 | 1.6067 | 2.8336 | -0.045776 | -0.036621 | 0.064087 | 0.10071 | 6.0745 | 6.1203 | 2.0691 | 3.3279 | 2.1149 | 3.3737
1788.1719 | 122.3602 | -48.2666 | 73.8007 | -0.1458 | 0.019713 | 0.086062 | 0.092418 | 0.0018949 | 0.01747 | 0.012219 | -0.016571 | 0.025594 | 0.082864 | 0.09609 | -0.00060892 | 0.01511 | 0.014541 | -0.018311 | 73.7961 | -48.3856 | 73.7 | -41.8991 | -45.607 | -46.8109 | 0.045776 | 73.7457 | -48.2895 | 73.8556 | -42.1875 | -44.9478 | -46.2021 | -0.11902 | -0.004577 | -0.10071 | -0.096131 | 6.3675 | 6.4865 | 1.4557 | 2.6596 | 1.5747 | 2.7786 | -0.022888 | -0.054931 | 0.054932 | 0.10986 | 6.0791 | 6.102 | 2.0645 | 3.3188 | 2.0874 | 3.3417
1789.1719 | 122.3373 | -48.2941 | 73.7274 | -0.11528 | 0.020995 | 0.086917 | 0.089427 | -0.0021641 | 0.018325 | 0.01051 | -0.025726 | 0.027303 | 0.082437 | 0.084768 | -0.00018152 | 0.017459 | 0.010269 | 0.036621 | 73.7045 | -48.4543 | 73.6816 | -41.89 | -45.607 | -46.8201 | 0.14191 | 73.7274 | -48.3627 | 73.7228 | -42.247 | -44.9295 | -46.225 | -0.16022 | -0.022888 | -0.045776 | -0.022888 | 6.4041 | 6.5643 | 1.474 | 2.6871 | 1.6342 | 2.8473 | -0.068665 | 0 | -0.004578 | -0.004578 | 6.0471 | 6.1157 | 2.0691 | 3.3646 | 2.1378 | 3.4332
1790.1719 | 122.374 | -48.2391 | 73.819 | -0.13817 | 0.02014 | 0.08649 | 0.091349 | 0.0048855 | 0.017043 | 0.0079468 | -0.059296 | 0.026021 | 0.08201 | 0.090749 | -0.00039512 | 0.018314 | 0.012832 | -0.064087 | 73.8144 | -48.3994 | 73.6816 | -41.9312 | -45.6528 | -46.8201 | 0.009155 | 73.7869 | -48.317 | 73.8327 | -42.2516 | -44.9661 | -46.2204 | -0.16022 | -0.004578 | -0.13733 | -0.13275 | 6.308 | 6.4682 | 1.4191 | 2.5864 | 1.5793 | 2.7466 | -0.07782 | -0.032044 | 0.013733 | 0.045777 | 5.9875 | 6.0654 | 2.0187 | 3.273 | 2.0966 | 3.3508
1791.1719 | 122.2504 | -48.3353 | 73.8007 | -0.10765 | 0.019927 | 0.084994 | 0.090281 | -0.0032321 | 0.018325 | 0.011151 | -0.080658 | 0.027944 | 0.084787 | 0.097158 | 0.0025955 | 0.017032 | 0.010055 | -0.013733 | 73.7595 | -48.4589 | 73.5947 | -42.0136 | -45.7306 | -46.9299 | 0.018311 | 73.7961 | -48.3856 | 73.7137 | -42.3981 | -45.1126 | -46.3577 | -0.1236 | -0.041199 | -0.20599 | -0.16479 | 6.3217 | 6.4453 | 1.4053 | 2.6047 | 1.5289 | 2.7283 | -0.050354 | -0.004577 | -0.086975 | -0.082398 | 5.9372 | 5.9875 | 1.9775 | 3.2227 | 2.0279 | 3.273
1792.1719 | 122.3465 | -48.23 | 73.8098 | -0.1992 | 0.0195 | 0.08414 | 0.092631 | -0.0034457 | 0.017684 | 0.01051 | -0.03183 | 0.026235 | 0.082651 | 0.089254 | -0.0018905 | 0.016178 | 0.013473 | -0.073242 | 73.9059 | -48.2849 | 73.7686 | -41.8213 | -45.5475 | -46.7514 | -0.018311 | 73.9014 | -48.1979 | 73.9197 | -42.215 | -44.9707 | -46.257 | -0.054932 | 0.096131 | -0.041198 | -0.13733 | 6.4087 | 6.4636 | 1.4786 | 2.6825 | 1.5335 | 2.7374 | 0.032043 | 0.091553 | 0.10986 | 0.018311 | 6.015 | 5.983 | 1.973 | 3.2593 | 1.9409 | 3.2272
1793.1719 | 122.3099 | -48.2666 | 73.7686 | -0.12291 | 0.019927 | 0.086703 | 0.091563 | 0.0016813 | 0.017684 | 0.011792 | -0.065399 | 0.024953 | 0.083078 | 0.09139 | -0.00039512 | 0.015537 | 0.012619 | -0.082397 | 73.7961 | -48.3902 | 73.6725 | -41.9128 | -45.6665 | -46.8292 | -0.050354 | 73.7732 | -48.3307 | 73.8052 | -42.3294 | -45.0027 | -46.2433 | -0.1236 | 0.027466 | -0.096131 | -0.1236 | 6.3538 | 6.4774 | 1.4374 | 2.6001 | 1.561 | 2.7237 | -0.064086 | 0.004577 | 0.036621 | 0.032044 | 5.9372 | 6.0013 | 2.0233 | 3.2639 | 2.0874 | 3.3279
1794.1719 | 122.3465 | -48.3261 | 73.7091 | -0.15343 | 0.018645 | 0.085208 | 0.090922 | -0.00088229 | 0.017684 | 0.011151 | -0.022675 | 0.026449 | 0.084146 | 0.094594 | -0.0010361 | 0.016818 | 0.011551 | -0.027466 | 73.732 | -48.4589 | 73.6404 | -41.9403 | -45.6894 | -46.8704 | 0.027466 | 73.6908 | -48.3627 | 73.7961 | -42.3477 | -45.0165 | -46.2662 | -0.13275 | 0.022889 | -0.068664 | -0.091553 | 6.3858 | 6.5186 | 1.4557 | 2.6367 | 1.5884 | 2.7695 | -0.036621 | -0.01831 | 0.086976 | 0.10529 | 5.9784 | 6.015 | 2.0599 | 3.3096 | 2.0966 | 3.3463
1795.1719 | 122.3328 | -48.2986 | 73.7366 | -0.21446 | 0.022704 | 0.086276 | 0.093486 | 0.0016813 | 0.018325 | 0.0098694 | -0.086761 | 0.025594 | 0.082651 | 0.092244 | -0.00039512 | 0.015964 | 0.0098417 | -0.045776 | 73.7183 | -48.4268 | 73.6313 | -41.9128 | -45.7077 | -46.8658 | -0.022888 | 73.732 | -48.349 | 73.7778 | -42.421 | -45.0714 | -46.312 | -0.12817 | -0.01831 | -0.10528 | -0.086975 | 6.3858 | 6.514 | 1.4328 | 2.5909 | 1.561 | 2.7191 | -0.050354 | -0.004577 | 0.041199 | 0.045776 | 5.8777 | 5.928 | 1.9867 | 3.2272 | 2.037 | 3.2776
1796.1875 | 122.2687 | -48.4268 | 73.6771 | -0.084763 | 0.020995 | 0.087131 | 0.093486 | 0.0012539 | 0.018325 | 0.013074 | -0.12643 | 0.026662 | 0.082651 | 0.097585 | -0.0010361 | 0.016178 | 0.0085599 | -0.018311 | 73.6954 | -48.5779 | 73.4711 | -42.0959 | -45.8313 | -47.0489 | -0.022888 | 73.6771 | -48.4817 | 73.6267 | -42.5446 | -45.2179 | -46.4539 | -0.15106 | 0.018311 | -0.20599 | -0.22431 | 6.3309 | 6.4819 | 1.3779 | 2.5955 | 1.5289 | 2.7466 | -0.054931 | 0 | -0.050354 | -0.050354 | 5.8823 | 5.9372 | 1.973 | 3.2089 | 2.0279 | 3.2639
1797.1719 | 122.3236 | -48.3719 | 73.6679 | -0.10765 | 0.019713 | 0.086917 | 0.092631 | -0.00088229 | 0.018325 | 0.011578 | -0.028778 | 0.027517 | 0.083932 | 0.092244 | 0.0011001 | 0.014896 | 0.013687 | -0.07782 | 73.7045 | -48.4772 | 73.5855 | -41.9586 | -45.7718 | -46.9162 | -0.041199 | 73.6816 | -48.3994 | 73.7274 | -42.4622 | -45.0989 | -46.3394 | -0.10529 | 0.036621 | -0.082398 | -0.11902 | 6.4133 | 6.5186 | 1.4557 | 2.6001 | 1.561 | 2.7054 | -0.027466 | 0.013733 | 0.059509 | 0.045776 | 5.9097 | 5.9372 | 2.0325 | 3.273 | 2.0599 | 3.3005
1798.1719 | 122.3145 | -48.4497 | 73.6084 | -0.1992 | 0.019072 | 0.085422 | 0.092631 | 0.00039951 | 0.017897 | 0.011365 | -0.040985 | 0.02538 | 0.082651 | 0.093526 | -0.00082252 | 0.015537 | 0.016464 | -0.068665 | 73.5901 | -48.5687 | 73.494 | -42.0136 | -45.8542 | -46.9849 | -0.022888 | 73.5809 | -48.4589 | 73.6267 | -42.5491 | -45.1767 | -46.4172 | -0.11902 | -0.01831 | -0.11444 | -0.09613 | 6.4362 | 6.5552 | 1.4648 | 2.5955 | 1.5839 | 2.7145 | -0.009155 | -0.027465 | 0.018311 | 0.045776 | 5.9006 | 5.9097 | 2.0325 | 3.273 | 2.0416 | 3.2822
1799.1719 | 122.3694 | -48.5641 | 73.4711 | -0.10002 | 0.018859 | 0.089053 | 0.088359 | -0.0030185 | 0.01747 | 0.012006 | -0.040985 | 0.024526 | 0.085428 | 0.08904 | 0.00045928 | 0.015964 | 0.011978 | 0 | 73.4665 | -48.7015 | 73.407 | -42.1005 | -45.9 | -47.0764 | 0.10529 | 73.4573 | -48.6053 | 73.5672 | -42.5903 | -45.2454 | -46.5088 | -0.13733 | -0.004577 | -0.064087 | -0.05951 | 6.4636 | 6.601 | 1.4877 | 2.6642 | 1.6251 | 2.8015 | -0.041199 | -0.013733 | 0.096131 | 0.10986 | 5.9738 | 6.015 | 2.0554 | 3.3188 | 2.0966 | 3.36
1800.1719 | 122.3145 | -48.6237 | 73.4619 | -0.15343 | 0.020354 | 0.087985 | 0.090495 | -0.00024129 | 0.017043 | 0.012219 | -0.044037 | 0.026235 | 0.083719 | 0.09139 | -0.00082252 | 0.015964 | 0.013687 | -0.07782 | 73.4528 | -48.761 | 73.3383 | -42.2058 | -46.0007 | -47.1542 | -0.032043 | 73.439 | -48.6511 | 73.4756 | -42.7048 | -45.3369 | -46.582 | -0.13733 | -0.009155 | -0.1236 | -0.11444 | 6.4178 | 6.5552 | 1.4694 | 2.623 | 1.6068 | 2.7603 | -0.027466 | -0.022888 | 0.013733 | 0.036621 | 5.9189 | 5.9463 | 2.0416 | 3.2867 | 2.0691 | 3.3142
1801.1719 | 122.3236 | -48.6237 | 73.4024 | -0.10765 | 0.020568 | 0.087344 | 0.08964 | -0.0019503 | 0.017256 | 0.010724 | -0.059296 | 0.025594 | 0.083719 | 0.090749 | -0.00039512 | 0.013828 | 0.012192 | -0.059509 | 73.407 | -48.7473 | 73.3109 | -42.215 | -46.0144 | -47.1634 | -0.059509 | 73.4207 | -48.6694 | 73.4665 | -42.7277 | -45.3781 | -46.5912 | -0.1236 | 0.004577 | -0.091553 | -0.09613 | 6.4087 | 6.5323 | 1.4603 | 2.6093 | 1.5839 | 2.7328 | -0.045777 | 0.01831 | 0.064087 | 0.045777 | 5.896 | 5.9418 | 2.0325 | 3.2455 | 2.0782 | 3.2913
1802.1719 | 122.3007 | -48.88 | 73.114 | -0.18394 | 0.021422 | 0.087558 | 0.091136 | 0.00061311 | 0.018965 | 0.011792 | -0.089813 | 0.026021 | 0.083932 | 0.093313 | -0.00018152 | 0.017673 | 0.012832 | -0.022888 | 73.1186 | -49.0265 | 73.082 | -42.4255 | -46.2799 | -47.4243 | 0.036621 | 73.1369 | -48.9258 | 73.2422 | -42.9565 | -45.6161 | -46.8475 | -0.14648 | 0.004577 | -0.032044 | -0.036621 | 6.4545 | 6.601 | 1.4557 | 2.6001 | 1.6022 | 2.7466 | -0.045776 | 0.022888 | 0.12817 | 0.10528 | 5.9235 | 5.9692 | 2.0325 | 3.2639 | 2.0782 | 3.3096
1803.1719 | 122.3465 | -48.9807 | 73.0087 | -0.13054 | 0.019713 | 0.091189 | 0.091136 | -0.0040867 | 0.020034 | 0.010937 | -0.080658 | 0.026662 | 0.082864 | 0.101 | -0.0012497 | 0.016605 | 0.013687 | -0.26093 | 72.9675 | -49.1638 | 72.9675 | -42.5262 | -46.4035 | -47.5479 | 0.018311 | 72.9904 | -49.0768 | 73.1232 | -43.0756 | -45.7352 | -46.9711 | -0.1831 | -0.041199 | -0.041199 | 0 | 6.4545 | 6.6376 | 1.4328 | 2.5772 | 1.6159 | 2.7603 | -0.09613 | -0.018311 | 0.11444 | 0.13275 | 5.9052 | 6.0013 | 2.0096 | 3.2455 | 2.1057 | 3.3417
1804.1719 | 122.4518 | -48.8754 | 73.1735 | -0.13817 | 0.020568 | 0.086703 | 0.089854 | -0.0053683 | 0.017897 | 0.012433 | -0.034882 | 0.024953 | 0.080942 | 0.092244 | 3.2078e-05 | 0.013614 | 0.011123 | 0.027466 | 73.1781 | -48.9853 | 73.1415 | -42.366 | -46.1426 | -47.3236 | 0.13733 | 73.1598 | -48.912 | 73.2742 | -42.8833 | -45.5246 | -46.9528 | -0.10986 | 0.004578 | -0.032044 | -0.036622 | 6.5094 | 6.6193 | 1.5518 | 2.7328 | 1.6617 | 2.8427 | -0.036621 | -0.013733 | 0.10071 | 0.11444 | 5.9921 | 6.0287 | 1.9226 | 3.3508 | 1.9592 | 3.3875
1805.1719 | 122.3419 | -48.8708 | 73.114 | -0.13817 | 0.021849 | 0.087131 | 0.092845 | 0.0016813 | 0.01747 | 0.011792 | -0.077606 | 0.026662 | 0.080728 | 0.090322 | 0.00045928 | 0.017032 | 0.011764 | -0.050354 | 73.114 | -49.0494 | 73.0682 | -42.4942 | -46.2936 | -47.438 | 0.013733 | 73.114 | -48.9395 | 73.201 | -43.0115 | -45.6436 | -46.8796 | -0.17853 | 0 | -0.045777 | -0.045777 | 6.3766 | 6.5552 | 1.4328 | 2.5772 | 1.6113 | 2.7557 | -0.068664 | 0 | 0.086975 | 0.086975 | 5.8594 | 5.928 | 1.9913 | 3.2272 | 2.0599 | 3.2959
1806.1719 | 122.3557 | -48.9258 | 73.0911 | -0.16869 | 0.021422 | 0.085422 | 0.091136 | 0.0014675 | 0.017684 | 0.010724 | -0.08371 | 0.025808 | 0.082223 | 0.091817 | -0.0012497 | 0.018955 | 0.012405 | -0.041199 | 73.0728 | -49.086 | 73.0179 | -42.5079 | -46.3074 | -47.4609 | 0.009155 | 73.0865 | -48.9807 | 73.1827 | -43.0252 | -45.6619 | -46.8979 | -0.16022 | -0.01831 | -0.073242 | -0.054932 | 6.4178 | 6.5781 | 1.4648 | 2.6184 | 1.6251 | 2.7786 | -0.054932 | -0.004577 | 0.091553 | 0.09613 | 5.9006 | 5.9555 | 2.0279 | 3.2639 | 2.0828 | 3.3188
1807.1719 | 122.319 | -49.0173 | 72.9996 | -0.19157 | 0.020568 | 0.085849 | 0.092204 | 0.0012539 | 0.01747 | 0.0088012 | -0.071503 | 0.025594 | 0.085214 | 0.094594 | 0.00067288 | 0.01511 | 0.014541 | -0.068665 | 73.0453 | -49.1501 | 72.9492 | -42.5995 | -46.463 | -47.5662 | -0.050354 | 73.0179 | -49.0768 | 73.0682 | -43.1534 | -45.8038 | -47.0215 | -0.13275 | 0.045776 | -0.050354 | -0.09613 | 6.4178 | 6.5506 | 1.4511 | 2.5543 | 1.5839 | 2.6871 | -0.059509 | 0.01831 | 0.068664 | 0.050354 | 5.864 | 5.9235 | 1.9958 | 3.2135 | 2.0554 | 3.273
1808.1719 | 122.3099 | -49.1867 | 72.8439 | -0.1458 | 0.020568 | 0.087771 | 0.09199 | 0.00082671 | 0.018538 | 0.011578 | -0.044037 | 0.026662 | 0.080087 | 0.084127 | -0.00039512 | 0.013614 | 0.014541 | -0.10529 | 72.8577 | -49.3011 | 72.7295 | -42.8238 | -46.6003 | -47.731 | -0.041199 | 72.8394 | -49.2508 | 72.8989 | -43.309 | -45.9641 | -47.1863 | -0.11444 | 0.013733 | -0.11444 | -0.12817 | 6.3629 | 6.4774 | 1.4557 | 2.5864 | 1.5701 | 2.7008 | -0.064086 | -0.004578 | 0.054932 | 0.05951 | 5.8777 | 5.9418 | 2.0004 | 3.2227 | 2.0645 | 3.2867
1809.1719 | 122.3419 | -49.0906 | 72.9172 | -0.11528 | 0.022063 | 0.079867 | 0.090709 | -0.0025913 | 0.018538 | 0.010297 | -0.05014 | 0.026449 | 0.08201 | 0.095235 | 3.2078e-05 | 0.011264 | 0.014328 | 0.036621 | 72.9218 | -49.2691 | 72.8668 | -42.7368 | -46.4127 | -47.5937 | 0.14191 | 72.9263 | -49.1501 | 72.9996 | -43.1854 | -45.8313 | -47.081 | -0.17853 | 0.004578 | -0.050354 | -0.054932 | 6.3538 | 6.5323 | 1.4969 | 2.6779 | 1.6754 | 2.8564 | -0.059509 | 0.009156 | 0.082398 | 0.073242 | 5.9052 | 5.9647 | 2.0096 | 3.2593 | 2.0691 | 3.3188
1810.1719 | 122.3511 | -49.1638 | 72.8119 | -0.15343 | 0.022917 | 0.087771 | 0.090709 | 0.0016813 | 0.017256 | 0.010297 | -0.12338 | 0.026235 | 0.083078 | 0.091176 | 0.00067288 | 0.016178 | 0.011978 | -0.064087 | 72.8943 | -49.3057 | 72.8394 | -42.7185 | -46.5179 | -47.6624 | 0 | 72.8897 | -49.1776 | 73.0087 | -43.2312 | -45.8633 | -47.0764 | -0.14191 | 0.082397 | 0.027465 | -0.054932 | 6.4453 | 6.5872 | 1.5015 | 2.6459 | 1.6434 | 2.7878 | -0.013733 | 0.077819 | 0.19684 | 0.11902 | 5.9326 | 5.9464 | 2.0874 | 3.3005 | 2.1011 | 3.3142
1811.1719 | 122.3145 | -49.086 | 72.9538 | -0.19157 | 0.020354 | 0.086703 | 0.088786 | -0.00045509 | 0.017256 | 0.012646 | -0.068451 | 0.02538 | 0.083505 | 0.089895 | 0.00024568 | 0.016391 | 0.01561 | -0.050354 | 72.9034 | -49.2371 | 72.8531 | -42.7094 | -46.4172 | -47.5937 | 0.022888 | 72.9218 | -49.1089 | 73.0591 | -43.158 | -45.8221 | -47.0169 | -0.15106 | -0.050354 | -0.10071 | -0.050354 | 6.3766 | 6.5277 | 1.4923 | 2.6688 | 1.6434 | 2.8198 | -0.022888 | -0.032043 | 0.10529 | 0.13733 | 5.928 | 5.9509 | 2.0691 | 3.2639 | 2.092 | 3.2867
1812.1719 | 122.3602 | -49.0768 | 72.9401 | -0.13054 | 0.020781 | 0.085422 | 0.08964 | -0.00066869 | 0.016616 | 0.012646 | -0.062347 | 0.027517 | 0.083719 | 0.088186 | -0.00018152 | 0.017887 | 0.0098417 | -0.045776 | 73.0041 | -49.1638 | 72.9492 | -42.5629 | -46.4218 | -47.6395 | -0.14191 | 72.8943 | -49.173 | 73.0316 | -43.1076 | -45.8084 | -47.0352 | -0.086975 | 0.064087 | 0.009156 | -0.054931 | 6.514 | 6.601 | 1.4374 | 2.655 | 1.5244 | 2.742 | -0.096131 | -0.045776 | 0.091553 | 0.13733 | 5.9692 | 6.0654 | 2.0416 | 3.2684 | 2.1378 | 3.3646
1813.1719 | 122.4197 | -49.2416 | 72.7295 | -0.16869 | 0.021422 | 0.087985 | 0.091136 | -0.0017367 | 0.021102 | 0.0081604 | -0.086761 | 0.026662 | 0.083932 | 0.090749 | -0.0023177 | 0.015537 | 0.0094143 | -0.041199 | 72.7707 | -49.379 | 72.7432 | -42.7734 | -46.6049 | -47.7402 | -0.027466 | 72.8119 | -49.2874 | 72.9126 | -43.2953 | -45.9869 | -47.1771 | -0.13733 | 0.041199 | 0.013733 | -0.027466 | 6.4682 | 6.6055 | 1.5015 | 2.6367 | 1.6388 | 2.774 | -0.045777 | 0.082398 | 0.18311 | 0.10071 | 5.9464 | 5.9921 | 2.0645 | 3.2547 | 2.1103 | 3.3005
1814.1719 | 122.3557 | -49.3744 | 72.6791 | -0.26024 | 0.022063 | 0.085849 | 0.090495 | -0.0015231 | 0.018538 | 0.011578 | -0.025726 | 0.026235 | 0.083719 | 0.089681 | 0.00088648 | 0.016818 | 0.012405 | -0.050354 | 72.6242 | -49.5529 | 72.5327 | -42.9382 | -46.7606 | -47.8958 | 0 | 72.6608 | -49.4156 | 72.7203 | -43.4143 | -46.1243 | -47.3282 | -0.17853 | -0.054931 | -0.14648 | -0.091553 | 6.4362 | 6.6147 | 1.4786 | 2.6138 | 1.6571 | 2.7924 | -0.041198 | -0.01831 | 0.041199 | 0.059509 | 5.9601 | 6.0013 | 2.0462 | 3.2501 | 2.0874 | 3.2913
1815.1719 | 122.374 | -49.4568 | 72.5647 | -0.18394 | 0.022063 | 0.087344 | 0.092204 | 0.0048855 | 0.018752 | 0.011365 | -0.10202 | 0.025594 | 0.082651 | 0.086477 | -0.0025315 | 0.017459 | 0.011551 | -0.027466 | 72.5189 | -49.6124 | 72.4823 | -42.9977 | -46.8567 | -48.0057 | -0.10071 | 72.4869 | -49.5895 | 72.6242 | -43.4875 | -46.2158 | -47.4289 | -0.15564 | -0.045776 | -0.082397 | -0.036621 | 6.459 | 6.6147 | 1.4511 | 2.6001 | 1.6068 | 2.7557 | -0.13275 | -0.07782 | 0.05951 | 0.13733 | 5.9692 | 6.102 | 2.0279 | 3.241 | 2.1606 | 3.3737
1816.1719 | 122.3419 | -49.617 | 72.4319 | -0.22972 | 0.020354 | 0.085849 | 0.092845 | -0.0019503 | 0.018538 | 0.013501 | -0.13254 | 0.026876 | 0.081796 | 0.095449 | -0.0018905 | 0.015537 | 0.0092007 | -0.10986 | 72.4457 | -49.7589 | 72.3129 | -43.2037 | -47.0444 | -48.1934 | -0.086975 | 72.4365 | -49.6262 | 72.5006 | -43.6615 | -46.3806 | -47.58 | -0.14191 | 0.013733 | -0.11902 | -0.13275 | 6.4133 | 6.5552 | 1.4236 | 2.5726 | 1.5656 | 2.7145 | -0.009156 | 0.004577 | 0.068664 | 0.064087 | 5.9555 | 5.9647 | 2.037 | 3.2364 | 2.0462 | 3.2455
1817.1719 | 122.3099 | -49.617 | 72.4274 | -0.1458 | 0.021636 | 0.085635 | 0.093913 | -0.0021641 | 0.016829 | 0.0092286 | -0.13254 | 0.02773 | 0.084573 | 0.095662 | -0.0014633 | 0.017246 | 0.016037 | -0.10529 | 72.4457 | -49.7772 | 72.2534 | -43.2312 | -47.0947 | -48.2162 | -0.10986 | 72.4319 | -49.6857 | 72.4686 | -43.6935 | -46.431 | -47.6257 | -0.16022 | 0.018311 | -0.17395 | -0.19226 | 6.3858 | 6.546 | 1.4008 | 2.5223 | 1.561 | 2.6825 | -0.068665 | 0.004578 | 0.041199 | 0.036621 | 5.9235 | 5.9921 | 1.9913 | 3.186 | 2.0599 | 3.2547
1818.1719 | 122.3282 | -49.4385 | 72.5784 | -0.16869 | 0.020354 | 0.086062 | 0.093272 | -0.00024129 | 0.020034 | 0.0077332 | -0.11118 | 0.027944 | 0.082437 | 0.092244 | -0.0014633 | 0.017246 | 0.010696 | -0.11444 | 72.6059 | -49.6078 | 72.4686 | -43.0435 | -46.8796 | -48.0286 | -0.064087 | 72.6059 | -49.498 | 72.6837 | -43.5013 | -46.2204 | -47.3877 | -0.16937 | 0.027466 | -0.10986 | -0.13733 | 6.395 | 6.5643 | 1.4099 | 2.5589 | 1.5793 | 2.7283 | -0.059509 | 0.027466 | 0.10529 | 0.07782 | 5.9372 | 5.9967 | 2.0508 | 3.2181 | 2.1103 | 3.2776
1819.1719 | 122.3007 | -49.4339 | 72.5464 | -0.15343 | 0.019927 | 0.086062 | 0.090281 | -0.00045509 | 0.015975 | 0.007306 | 0.0047913 | 0.025594 | 0.083078 | 0.091176 | 0.0017411 | 0.014896 | 0.01326 | -0.082397 | 72.6791 | -49.5346 | 72.5235 | -42.9886 | -46.7926 | -47.9233 | -0.054932 | 72.6288 | -49.4522 | 72.6608 | -43.4692 | -46.1884 | -47.3785 | -0.10071 | 0.13275 | -0.022888 | -0.15564 | 6.4453 | 6.546 | 1.5106 | 2.6413 | 1.6113 | 2.742 | -0.01831 | 0.082397 | 0.11444 | 0.032044 | 5.9647 | 5.983 | 2.0554 | 3.2455 | 2.0737 | 3.2639
1820.1719 | 122.2641 | -49.4797 | 72.551 | -0.16869 | 0.0195 | 0.085208 | 0.094554 | 0.00082671 | 0.017043 | 0.010297 | -0.11118 | 0.026876 | 0.082437 | 0.094594 | 0.00024568 | 0.018527 | 0.01326 | -0.09613 | 72.6105 | -49.6399 | 72.4136 | -43.0801 | -46.9208 | -48.0743 | -0.11444 | 72.6013 | -49.5255 | 72.5876 | -43.602 | -46.3257 | -47.5021 | -0.16022 | 0.05951 | -0.13733 | -0.19684 | 6.3995 | 6.5598 | 1.4053 | 2.5589 | 1.5656 | 2.7191 | -0.045777 | 0.050354 | 0.036621 | -0.013733 | 5.8777 | 5.9235 | 1.9775 | 3.154 | 2.0233 | 3.1998
1821.1719 | 122.3099 | -49.5163 | 72.5281 | -0.1458 | 0.020995 | 0.08649 | 0.091563 | 0.00039951 | 0.019393 | 0.014142 | -0.10812 | 0.025808 | 0.082864 | 0.097585 | -0.0014633 | 0.016605 | 0.010696 | -0.10071 | 72.5555 | -49.6857 | 72.3724 | -43.1305 | -46.9711 | -48.0881 | -0.091553 | 72.5281 | -49.5804 | 72.5876 | -43.5745 | -46.3211 | -47.4976 | -0.16937 | 0.027466 | -0.15564 | -0.1831 | 6.3858 | 6.5552 | 1.4282 | 2.5452 | 1.5976 | 2.7145 | -0.064087 | 0 | 0.059509 | 0.059509 | 5.9418 | 6.0059 | 2.0187 | 3.1952 | 2.0828 | 3.2593
1822.1719 | 122.3053 | -49.5483 | 72.5098 | -0.11528 | 0.021636 | 0.085849 | 0.091777 | 0.00061311 | 0.018111 | 0.010724 | -0.12643 | 0.027089 | 0.081369 | 0.097799 | 0.0013139 | 0.015964 | 0.0094143 | -0.11902 | 72.496 | -49.704 | 72.3038 | -43.1854 | -47.0352 | -48.1613 | -0.11444 | 72.5006 | -49.6307 | 72.5098 | -43.6478 | -46.4172 | -47.5891 | -0.15564 | -0.013733 | -0.20599 | -0.19226 | 6.3629 | 6.5186 | 1.387 | 2.5131 | 1.5427 | 2.6688 | -0.082397 | -0.009156 | 0 | 0.009156 | 5.9006 | 5.983 | 1.9592 | 3.1311 | 2.0416 | 3.2135
1823.1719 | 122.2275 | -49.7314 | 72.3633 | -0.13054 | 0.022063 | 0.087771 | 0.093272 | -2.7692e-05 | 0.017897 | 0.0092286 | -0.13254 | 0.026662 | 0.085001 | 0.090322 | 0.0011001 | 0.017673 | 0.013687 | -0.091553 | 72.3358 | -49.8734 | 72.139 | -43.4326 | -47.2275 | -48.3719 | -0.16937 | 72.3495 | -49.7772 | 72.3038 | -43.8812 | -46.6003 | -47.7814 | -0.14191 | -0.027466 | -0.2243 | -0.19684 | 6.2988 | 6.4407 | 1.3596 | 2.504 | 1.5015 | 2.6459 | -0.045777 | -0.013733 | -0.059509 | -0.045776 | 5.8502 | 5.896 | 1.9501 | 3.1311 | 1.9959 | 3.1769
1824.1719 | 122.3511 | -49.8093 | 72.1985 | -0.13817 | 0.023131 | 0.088199 | 0.093058 | -2.7692e-05 | 0.018752 | 0.0079468 | 0.0078429 | 0.026235 | 0.085214 | 0.090108 | 0.00067288 | 0.016391 | 0.014541 | -0.059509 | 72.1939 | -49.9649 | 72.2305 | -43.2861 | -47.2916 | -48.3353 | 0.018311 | 72.1756 | -49.8642 | 72.2992 | -43.8171 | -46.5637 | -47.7676 | -0.15564 | -0.004577 | 0.032044 | 0.036621 | 6.5231 | 6.6788 | 1.474 | 2.5177 | 1.6296 | 2.6733 | -0.054932 | -0.022888 | 0.10071 | 0.1236 | 5.9921 | 6.0471 | 2.0416 | 3.2455 | 2.0966 | 3.3005
1825.1719 | 122.2366 | -49.8962 | 72.1939 | -0.15343 | 0.021636 | 0.085208 | 0.091349 | 0.00061311 | 0.020034 | 0.0098694 | -0.11423 | 0.027517 | 0.083932 | 0.096517 | 0.00067288 | 0.017246 | 0.015182 | -0.082397 | 72.2031 | -50.0565 | 71.9604 | -43.5242 | -47.3831 | -48.5184 | -0.13733 | 72.1985 | -49.9283 | 72.1619 | -44.014 | -46.7651 | -47.937 | -0.16022 | 0.009155 | -0.23346 | -0.24262 | 6.3721 | 6.5323 | 1.3779 | 2.5131 | 1.5381 | 2.6733 | -0.032044 | 0.004577 | -0.032044 | -0.036621 | 5.8823 | 5.9143 | 1.9592 | 3.1311 | 1.9913 | 3.1631
1826.1719 | 122.2458 | -49.9512 | 72.1664 | -0.17632 | 0.021422 | 0.085422 | 0.092204 | 0.00082671 | 0.014266 | 0.0096558 | -0.10507 | 0.027517 | 0.086069 | 0.089895 | 0.0011001 | 0.015537 | 0.010696 | -0.073242 | 72.1436 | -50.1068 | 71.9284 | -43.5791 | -47.4197 | -48.5733 | -0.13733 | 72.1344 | -49.9924 | 72.0795 | -44.0826 | -46.8155 | -48.0057 | -0.15564 | -0.022888 | -0.23804 | -0.21515 | 6.3721 | 6.5277 | 1.3779 | 2.5314 | 1.5335 | 2.6871 | -0.041199 | -0.032044 | -0.086975 | -0.054931 | 5.8685 | 5.9097 | 1.9455 | 3.1357 | 1.9867 | 3.1769
1827.1719 | 122.2092 | -49.9649 | 72.1024 | -0.12291 | 0.02014 | 0.087771 | 0.091777 | 0.0012539 | 0.016188 | 0.0079468 | -0.089813 | 0.026449 | 0.084573 | 0.089681 | 0.0013139 | 0.016605 | 0.011978 | -0.045776 | 72.1527 | -50.0977 | 71.9238 | -43.5287 | -47.3877 | -48.5596 | -0.091553 | 72.0886 | -50.0244 | 72.084 | -44.0552 | -46.8384 | -48.0148 | -0.13275 | 0.050354 | -0.17853 | -0.22888 | 6.4362 | 6.5689 | 1.4053 | 2.5772 | 1.5381 | 2.71 | -0.059509 | -0.013733 | -0.018311 | -0.004578 | 5.9097 | 5.9692 | 1.9501 | 3.1265 | 2.0096 | 3.186
1828.1719 | 122.2183 | -49.9649 | 72.0978 | -0.16106 | 0.019927 | 0.083713 | 0.093272 | 0.00082671 | 0.017897 | 0.0096558 | -0.08371 | 0.027517 | 0.083292 | 0.08669 | -0.00018152 | 0.016818 | 0.015396 | -0.086975 | 72.1344 | -50.1114 | 71.9147 | -43.5425 | -47.4106 | -48.587 | -0.13275 | 72.1161 | -50.0107 | 72.0612 | -44.0735 | -46.8292 | -48.0011 | -0.14648 | 0.036621 | -0.18311 | -0.21973 | 6.4224 | 6.5689 | 1.3779 | 2.5543 | 1.5244 | 2.7008 | -0.045776 | 0.018311 | -0.036621 | -0.054932 | 5.8914 | 5.9372 | 1.9638 | 3.1357 | 2.0096 | 3.1815
1829.1719 | 122.3785 | -50.2579 | 71.7453 | -0.18394 | 0.020781 | 0.087344 | 0.090068 | 0.0042447 | 0.019606 | 0.0092286 | -0.016571 | 0.024526 | 0.08436 | 0.094381 | -0.0012497 | 0.017032 | 0.013901 | -0.032043 | 71.7865 | -50.3494 | 71.6949 | -43.7439 | -47.6028 | -48.7793 | -0.004578 | 71.7545 | -50.2945 | 71.8506 | -44.133 | -47.0627 | -48.2391 | -0.091552 | 0.041199 | -0.050354 | -0.091553 | 6.514 | 6.6055 | 1.4786 | 2.655 | 1.5701 | 2.7466 | -0.036621 | 0.009156 | 0.10529 | 0.09613 | 6.1249 | 6.1615 | 2.0187 | 3.1952 | 2.0554 | 3.2318
1830.1719 | 122.287 | -50.3998 | 71.6721 | -0.13817 | 0.023131 | 0.087131 | 0.09199 | -0.0021641 | 0.020034 | 0.0088012 | -0.080658 | 0.029012 | 0.085641 | 0.092672 | 0.00024568 | 0.017887 | 0.015823 | -0.027466 | 71.6721 | -50.5508 | 71.5118 | -43.9316 | -47.7905 | -48.9716 | 0.18768 | 71.7224 | -50.3403 | 71.7041 | -44.4443 | -47.1909 | -48.3948 | -0.15106 | 0 | -0.16022 | -0.16022 | 6.4682 | 6.6193 | 1.4282 | 2.6093 | 1.5793 | 2.7603 | 0.059509 | 0.050354 | 0.032044 | -0.01831 | 5.9555 | 5.896 | 2.005 | 3.2089 | 1.9455 | 3.1494
1831.1719 | 122.2824 | -50.4547 | 71.6537 | -0.061874 | 0.020568 | 0.087131 | 0.092845 | -0.0021641 | 0.017256 | 0.011365 | -0.044037 | 0.027089 | 0.085641 | 0.096517 | -0.00039512 | 0.016605 | 0.012619 | 0 | 71.6537 | -50.5829 | 71.4706 | -43.9911 | -47.8271 | -49.0311 | 0.045776 | 71.6125 | -50.4776 | 71.6171 | -44.5038 | -47.2183 | -48.4314 | -0.12817 | 0 | -0.18311 | -0.18311 | 6.4636 | 6.5918 | 1.4236 | 2.6276 | 1.5518 | 2.7557 | -0.022888 | -0.041199 | -0.036622 | 0.004577 | 5.9509 | 5.9738 | 2.0233 | 3.2364 | 2.0462 | 3.2593
1832.1719 | 122.2778 | -50.5371 | 71.5622 | -0.092392 | 0.020995 | 0.085422 | 0.092418 | 0.0018949 | 0.017043 | 0.011365 | -0.071503 | 0.027089 | 0.084787 | 0.085836 | -0.0021041 | 0.01511 | 0.01326 | -0.013733 | 71.5393 | -50.6699 | 71.402 | -43.9774 | -47.8546 | -49.0173 | 0.013733 | 71.553 | -50.5463 | 71.5302 | -44.6136 | -47.3557 | -48.5596 | -0.13275 | -0.022888 | -0.16022 | -0.13733 | 6.5598 | 6.6925 | 1.5198 | 2.6825 | 1.6525 | 2.8152 | -0.009156 | -0.009155 | -0.032044 | -0.022889 | 5.9235 | 5.9326 | 1.9775 | 3.1815 | 1.9867 | 3.1906
1833.1719 | 122.3007 | -50.6012 | 71.4523 | -0.13817 | 0.020568 | 0.087771 | 0.092845 | -0.0032321 | 0.019179 | 0.01051 | -0.059296 | 0.026021 | 0.089059 | 0.088613 | -0.00060892 | 0.017246 | 0.0186 | -0.013733 | 71.402 | -50.8209 | 71.2509 | -44.1788 | -48.0469 | -49.2508 | 0.032043 | 71.3928 | -50.7019 | 71.64 | -44.6686 | -47.3465 | -48.6099 | -0.21973 | -0.050354 | -0.20142 | -0.15106 | 6.4224 | 6.6422 | 1.3504 | 2.5543 | 1.5701 | 2.774 | -0.10071 | -0.05951 | 0.18768 | 0.24719 | 5.9326 | 6.0333 | 1.9913 | 3.2547 | 2.092 | 3.3554
1834.1719 | 122.2916 | -50.7156 | 71.3699 | -0.046616 | 0.020781 | 0.087131 | 0.090709 | -0.0055819 | 0.016829 | 0.013715 | -0.034882 | 0.029012 | 0.08735 | 0.092244 | -0.0010361 | 0.017032 | 0.011337 | -0.12817 | 71.3928 | -50.8575 | 71.1823 | -44.2703 | -48.1476 | -49.3195 | -0.11444 | 71.3745 | -50.7523 | 71.3654 | -44.7968 | -47.5296 | -48.7106 | -0.14191 | 0.022888 | -0.18768 | -0.21057 | 6.4453 | 6.5872 | 1.3962 | 2.5681 | 1.5381 | 2.71 | -0.036621 | 0.004578 | -0.004578 | -0.009156 | 5.9189 | 5.9555 | 2.005 | 3.186 | 2.0416 | 3.2227
1835.1719 | 122.3007 | -50.6195 | 71.4157 | -0.16106 | 0.018859 | 0.085208 | 0.092631 | -0.0034457 | 0.016829 | 0.0094422 | -0.05014 | 0.025167 | 0.083932 | 0.092244 | 0.00024568 | 0.018955 | 0.014755 | -0.018311 | 71.4523 | -50.7889 | 71.3242 | -44.0643 | -47.9599 | -49.1318 | 0.009155 | 71.4706 | -50.6561 | 71.4294 | -44.7189 | -47.4609 | -48.6465 | -0.16937 | 0.036622 | -0.091552 | -0.12817 | 6.5552 | 6.7245 | 1.4877 | 2.6596 | 1.6571 | 2.829 | -0.036621 | 0.054932 | 0.013733 | -0.041199 | 5.9006 | 5.9372 | 1.973 | 3.1586 | 2.0096 | 3.1952
1836.1719 | 122.2916 | -50.5829 | 71.4798 | -0.12291 | 0.020568 | 0.084353 | 0.091563 | 0.0018949 | 0.018538 | 0.0088012 | -0.068451 | 0.026021 | 0.085428 | 0.088827 | 3.2078e-05 | 0.020664 | 0.010055 | -0.009155 | 71.4661 | -50.7614 | 71.2875 | -44.165 | -48.0011 | -49.2096 | 0.027466 | 71.4523 | -50.6516 | 71.4523 | -44.6503 | -47.4014 | -48.6237 | -0.17853 | -0.013733 | -0.19226 | -0.17853 | 6.4178 | 6.5964 | 1.3733 | 2.5818 | 1.5518 | 2.7603 | -0.068664 | -0.027465 | -0.027465 | 0 | 5.9326 | 6.0013 | 1.9592 | 3.1815 | 2.0279 | 3.2501
1837.1719 | 122.2961 | -50.7111 | 71.3745 | -0.046616 | 0.020995 | 0.085849 | 0.091777 | 0.0014675 | 0.017043 | 0.0092286 | -0.037933 | 0.025594 | 0.083505 | 0.092244 | 0.00024568 | 0.014255 | 0.013046 | -0.041199 | 71.402 | -50.8392 | 71.1914 | -44.3161 | -48.0835 | -49.2783 | 0.004578 | 71.3791 | -50.7523 | 71.3608 | -44.751 | -47.5204 | -48.7106 | -0.12817 | 0.027466 | -0.18311 | -0.21057 | 6.395 | 6.5231 | 1.4328 | 2.6276 | 1.561 | 2.7557 | -0.041198 | 0.004577 | -0.013733 | -0.01831 | 5.9601 | 6.0013 | 2.0004 | 3.1906 | 2.0416 | 3.2318
1838.1719 | 122.2641 | -50.6699 | 71.4386 | -0.10002 | 0.018645 | 0.083499 | 0.088572 | -0.0015231 | 0.017684 | 0.011578 | -0.016571 | 0.026662 | 0.086069 | 0.090535 | 0.0011001 | 0.016391 | 0.012192 | -0.032043 | 71.4478 | -50.7889 | 71.2097 | -44.3024 | -48.0743 | -49.2783 | -0.013733 | 71.4111 | -50.7065 | 71.3974 | -44.7327 | -47.4747 | -48.6877 | -0.11902 | 0.009155 | -0.22888 | -0.23804 | 6.3675 | 6.4865 | 1.3916 | 2.5955 | 1.5106 | 2.7145 | -0.036621 | -0.027466 | -0.041199 | -0.013733 | 5.9372 | 5.9738 | 1.9821 | 3.1952 | 2.0187 | 3.2318
1839.1719 | 122.2321 | -50.6653 | 71.4386 | -0.16869 | 0.023131 | 0.086703 | 0.090922 | 0.0010403 | 0.017256 | 0.0064514 | -0.03183 | 0.026235 | 0.086496 | 0.090749 | 0.0015275 | 0.017246 | 0.014114 | -0.091553 | 71.4249 | -50.7935 | 71.2509 | -44.0964 | -48.0331 | -49.3607 | -0.10071 | 71.4157 | -50.6882 | 71.3882 | -44.7693 | -47.5388 | -48.6969 | -0.12817 | -0.013733 | -0.18768 | -0.17395 | 6.5689 | 6.6971 | 1.3046 | 2.6321 | 1.4328 | 2.7603 | -0.022888 | -0.022889 | -0.050354 | -0.027465 | 5.896 | 5.9189 | 1.9684 | 3.1265 | 1.9913 | 3.1494
1840.1719 | 122.3007 | -50.7889 | 71.2921 | -0.1458 | 0.020995 | 0.086703 | 0.090709 | -0.0019503 | 0.020247 | 0.010297 | -0.059296 | 0.026876 | 0.085641 | 0.087118 | -0.0021041 | 0.013828 | 0.011123 | -0.009155 | 71.2967 | -50.9125 | 71.0999 | -44.4489 | -48.1659 | -49.3698 | 0.041199 | 71.2738 | -50.8209 | 71.2738 | -44.8425 | -47.58 | -48.7885 | -0.1236 | 0.004578 | -0.19226 | -0.19684 | 6.34 | 6.4636 | 1.4191 | 2.623 | 1.5427 | 2.7466 | -0.032044 | -0.01831 | -0.01831 | 0 | 5.9464 | 5.9784 | 2.0004 | 3.2089 | 2.0325 | 3.241
1841.1719 | 122.2687 | -50.7706 | 71.2601 | -0.084763 | 0.020781 | 0.085849 | 0.091563 | 0.00061311 | 0.020674 | 0.010083 | -0.03183 | 0.026449 | 0.086282 | 0.09908 | -0.00082252 | 0.015537 | 0.012832 | 0.013733 | 71.3013 | -50.9583 | 71.109 | -44.4305 | -48.1796 | -49.3835 | 0.041199 | 71.2601 | -50.8392 | 71.2601 | -44.8517 | -47.5845 | -48.7839 | -0.18768 | 0.041199 | -0.15106 | -0.19226 | 6.34 | 6.5277 | 1.387 | 2.5909 | 1.5747 | 2.7786 | -0.068664 | 0 | 0 | 0 | 5.9189 | 5.9875 | 1.9867 | 3.186 | 2.0554 | 3.2547
1842.1719 | 122.3282 | -50.8347 | 71.2097 | -0.11528 | 0.022063 | 0.085635 | 0.091777 | -0.0025913 | 0.015761 | 0.0098694 | -0.080658 | 0.02538 | 0.085001 | 0.089254 | 3.2078e-05 | 0.017887 | 0.010269 | 0.013733 | 71.2418 | -50.9674 | 71.0724 | -44.4534 | -48.2071 | -49.4156 | 0.045776 | 71.2143 | -50.8621 | 71.2326 | -44.8975 | -47.6532 | -48.8434 | -0.13275 | 0.032043 | -0.13733 | -0.16937 | 6.3812 | 6.514 | 1.4191 | 2.6276 | 1.5518 | 2.7603 | -0.027466 | 0.004577 | 0.022888 | 0.018311 | 5.9372 | 5.9647 | 1.9913 | 3.1815 | 2.0187 | 3.2089
1843.1719 | 122.2916 | -50.798 | 71.2692 | -0.092392 | 0.020354 | 0.085635 | 0.092418 | -0.0010959 | 0.017256 | 0.011151 | -0.05014 | 0.026449 | 0.085214 | 0.086904 | -0.0016769 | 0.016605 | 0.012192 | -0.004578 | 71.283 | -50.9628 | 71.1273 | -44.4443 | -48.1796 | -49.3744 | 0.041199 | 71.2875 | -50.8118 | 71.2738 | -44.8929 | -47.6349 | -48.7885 | -0.16479 | 0.013733 | -0.14191 | -0.15564 | 6.3538 | 6.5186 | 1.4236 | 2.6184 | 1.5884 | 2.7832 | -0.013733 | 0.018311 | 0.004578 | -0.013733 | 5.9052 | 5.9189 | 2.0096 | 3.1631 | 2.0233 | 3.1769
1844.1719 | 122.3236 | -50.853 | 71.1868 | -0.092392 | 0.019713 | 0.087131 | 0.088786 | -2.7692e-05 | 0.018538 | 0.0079468 | -0.037933 | 0.024526 | 0.084573 | 0.086904 | -0.0014633 | 0.017246 | 0.014969 | -0.10986 | 71.2189 | -51.0086 | 71.0312 | -44.5633 | -48.2849 | -49.4659 | -0.064087 | 71.1411 | -50.8987 | 71.2326 | -44.9249 | -47.6761 | -48.8388 | -0.15564 | 0.032043 | -0.15564 | -0.18768 | 6.2897 | 6.4453 | 1.387 | 2.5681 | 1.5427 | 2.7237 | -0.045777 | -0.045777 | 0.045776 | 0.091553 | 5.928 | 5.9738 | 2.0142 | 3.1769 | 2.0599 | 3.2227
1845.1719 | 122.2458 | -50.8575 | 71.2509 | -0.10765 | 0.020781 | 0.084781 | 0.088572 | 0.0036039 | 0.018111 | 0.0088012 | -0.010467 | 0.026235 | 0.082223 | 0.095449 | -0.0023177 | 0.01575 | 0.014969 | -0.027466 | 71.3058 | -50.9491 | 71.0495 | -44.5679 | -48.2117 | -49.4247 | -0.013733 | 71.2509 | -50.8804 | 71.2006 | -44.9432 | -47.6761 | -48.8754 | -0.091553 | 0.054931 | -0.20142 | -0.25635 | 6.2897 | 6.3812 | 1.4328 | 2.6459 | 1.5244 | 2.7374 | -0.022888 | 0 | -0.050354 | -0.050354 | 5.9143 | 5.9372 | 1.9821 | 3.1815 | 2.005 | 3.2043
1846.1719 | 122.3007 | -50.9079 | 71.1594 | -0.12291 | 0.022277 | 0.084781 | 0.09199 | 0.0025357 | 0.016829 | 0.01051 | -0.059296 | 0.027089 | 0.086282 | 0.093953 | 0.00024568 | 0.017887 | 0.012405 | 0 | 71.1227 | -51.0956 | 70.9671 | -44.6365 | -48.3261 | -49.5438 | 0.036621 | 71.1411 | -50.9491 | 71.1365 | -45.0348 | -47.7951 | -48.967 | -0.18768 | -0.036621 | -0.19226 | -0.15564 | 6.2714 | 6.459 | 1.3641 | 2.5818 | 1.5518 | 2.7695 | -0.041199 | -0.018311 | -0.022888 | -0.004577 | 5.8731 | 5.9143 | 1.9409 | 3.1128 | 1.9821 | 3.154
1847.1719 | 122.2778 | -50.8484 | 71.2051 | -0.10765 | 0.022704 | 0.085849 | 0.092204 | -0.0019503 | 0.018752 | 0.01051 | -0.037933 | 0.026662 | 0.08201 | 0.09139 | -0.0014633 | 0.015537 | 0.0074917 | -0.009155 | 71.2326 | -51.0178 | 71.0266 | -44.5908 | -48.2437 | -49.4705 | 0.045776 | 71.2189 | -50.9171 | 71.1914 | -44.989 | -47.6715 | -48.8617 | -0.16937 | 0.027466 | -0.17853 | -0.20599 | 6.2576 | 6.427 | 1.3779 | 2.6047 | 1.5472 | 2.774 | -0.068664 | 0.013733 | -0.013733 | -0.027466 | 5.8594 | 5.928 | 1.9867 | 3.1769 | 2.0554 | 3.2455
1848.1719 | 122.2687 | -50.9399 | 71.1548 | -0.11528 | 0.020781 | 0.085208 | 0.093486 | 0.0048855 | 0.019179 | 0.010083 | -0.059296 | 0.026662 | 0.084787 | 0.087118 | -0.00018152 | 0.015537 | 0.013687 | 0.004578 | 71.0495 | -51.1551 | 70.8847 | -44.7006 | -48.3902 | -49.617 | 0.041199 | 71.0495 | -51.0681 | 71.0312 | -45.1492 | -47.8867 | -48.8754 | -0.21515 | -0.10529 | -0.27008 | -0.16479 | 6.2393 | 6.4545 | 1.3229 | 2.5497 | 1.5381 | 2.7649 | -0.12817 | -0.10529 | -0.1236 | -0.018311 | 5.7907 | 5.9189 | 2.0645 | 3.0533 | 2.1927 | 3.1815
1849.1719 | 122.3602 | -51.004 | 71.0129 | -0.13817 | 0.025695 | 0.086276 | 0.090922 | -0.00066869 | 0.019393 | 0.013287 | -0.044037 | 0.026449 | 0.083505 | 0.084981 | -0.00018152 | 0.016605 | 0.012192 | -0.091553 | 71.0312 | -51.1642 | 70.8755 | -44.7372 | -48.4268 | -49.6216 | -0.045776 | 70.99 | -51.1185 | 71.0724 | -45.1263 | -47.85 | -48.9944 | -0.16022 | 0.018311 | -0.13733 | -0.15564 | 6.2668 | 6.427 | 1.3824 | 2.5772 | 1.5427 | 2.7374 | -0.11444 | -0.022888 | 0.05951 | 0.082398 | 5.8777 | 5.9921 | 2.0096 | 3.154 | 2.124 | 3.2684
1850.1719 | 122.2641 | -51.0315 | 71.0037 | -0.084763 | 0.023558 | 0.087131 | 0.090709 | -0.0023777 | 0.016616 | 0.0098694 | -0.025726 | 0.025594 | 0.08436 | 0.092244 | 3.2078e-05 | 0.015964 | 0.012832 | 0.018311 | 71.0037 | -51.1963 | 70.8572 | -44.7556 | -48.436 | -49.6399 | 0.045776 | 70.9808 | -51.1047 | 70.9946 | -45.1996 | -47.9004 | -49.0585 | -0.16479 | 0 | -0.14648 | -0.14648 | 6.2759 | 6.4407 | 1.3916 | 2.5955 | 1.5564 | 2.7603 | -0.073242 | -0.022888 | -0.009155 | 0.013733 | 5.8319 | 5.9052 | 1.973 | 3.1311 | 2.0462 | 3.2043
1851.1719 | 122.2595 | -51.1139 | 70.9671 | -0.077133 | 0.018218 | 0.086062 | 0.090922 | -0.0017367 | 0.017256 | 0.0090148 | -0.034882 | 0.026662 | 0.086496 | 0.086263 | -0.0014633 | 0.013614 | 0.013473 | -0.022888 | 70.9991 | -51.2375 | 70.7748 | -44.8471 | -48.4909 | -49.7086 | -0.045776 | 70.9717 | -51.1597 | 70.9351 | -45.2591 | -47.937 | -49.1272 | -0.1236 | 0.032044 | -0.19226 | -0.22431 | 6.2668 | 6.3904 | 1.4053 | 2.623 | 1.5289 | 2.7466 | -0.045776 | 0.004578 | -0.032043 | -0.036621 | 5.8548 | 5.9006 | 1.9867 | 3.1769 | 2.0325 | 3.2227
1852.1719 | 122.2778 | -51.1551 | 70.9259 | -0.13817 | 0.022704 | 0.087558 | 0.089427 | 0.0018949 | 0.020461 | 0.011365 | -0.062347 | 0.026449 | 0.087137 | 0.087972 | 0.0017411 | 0.015323 | 0.013046 | -0.10071 | 70.8939 | -51.3336 | 70.6696 | -44.6823 | -48.5229 | -49.8871 | -0.10071 | 70.903 | -51.2192 | 70.871 | -45.3827 | -48.0698 | -49.2188 | -0.17853 | -0.032043 | -0.25635 | -0.2243 | 6.4728 | 6.6513 | 1.268 | 2.6321 | 1.4465 | 2.8107 | -0.064087 | -0.022888 | -0.054931 | -0.032043 | 5.7724 | 5.8365 | 1.9363 | 3.0853 | 2.0004 | 3.1494
1853.1719 | 122.2961 | -51.3062 | 70.7748 | -0.10002 | 0.021209 | 0.083499 | 0.091136 | -0.00066869 | 0.019393 | 0.0070922 | -0.044037 | 0.027944 | 0.083505 | 0.091817 | -0.00060892 | 0.017887 | 0.014541 | -0.009155 | 70.7977 | -51.4343 | 70.5963 | -45.0302 | -48.6877 | -49.8825 | 0.050354 | 70.7657 | -51.3474 | 70.7474 | -45.4468 | -48.1522 | -49.3149 | -0.12817 | 0.022888 | -0.17853 | -0.20142 | 6.2759 | 6.4041 | 1.4236 | 2.6184 | 1.5518 | 2.7466 | -0.041199 | -0.009155 | -0.027466 | -0.018311 | 5.8594 | 5.9006 | 1.9913 | 3.154 | 2.0325 | 3.1952
1854.1719 | 122.2961 | -51.4481 | 70.6055 | -0.18394 | 0.020995 | 0.086276 | 0.093058 | -0.0025913 | 0.018325 | 0.0094422 | -0.044037 | 0.025808 | 0.084146 | 0.088399 | 0.00067288 | 0.016391 | 0.011337 | -0.13733 | 70.6558 | -51.6037 | 70.4041 | -45.195 | -48.912 | -50.0931 | -0.11444 | 70.5688 | -51.5305 | 70.5734 | -45.6894 | -48.3582 | -49.5026 | -0.15564 | 0.050354 | -0.20142 | -0.25177 | 6.2531 | 6.4087 | 1.355 | 2.536 | 1.5106 | 2.6917 | -0.082398 | -0.036621 | -0.032044 | 0.004577 | 5.7587 | 5.8411 | 1.9455 | 3.0899 | 2.0279 | 3.1723
1855.1719 | 122.2778 | -51.5121 | 70.5643 | -0.10765 | 0.021422 | 0.08649 | 0.089854 | 0.0031765 | 0.017684 | 0.0079468 | -0.016571 | 0.025167 | 0.085001 | 0.090749 | -0.00060892 | 0.015323 | 0.01326 | -0.022888 | 70.5734 | -51.6724 | 70.3537 | -45.2591 | -48.9395 | -50.1251 | 0 | 70.5414 | -51.5625 | 70.5093 | -45.7123 | -48.3719 | -49.5438 | -0.16022 | 0.009155 | -0.21057 | -0.21973 | 6.2531 | 6.4133 | 1.387 | 2.5726 | 1.5472 | 2.7328 | -0.050354 | -0.022888 | -0.054932 | -0.032044 | 5.7999 | 5.8502 | 1.9684 | 3.1403 | 2.0187 | 3.1906
1856.1719 | 122.2687 | -51.5442 | 70.5139 | -0.077133 | 0.022917 | 0.084994 | 0.091349 | 0.0025357 | 0.018538 | 0.0085876 | -0.034882 | 0.025167 | 0.083505 | 0.088399 | 0.00067288 | 0.013187 | 0.016037 | 0.013733 | 70.5231 | -51.6861 | 70.3537 | -45.2133 | -48.9166 | -50.1389 | 0.068665 | 70.5368 | -51.59 | 70.5139 | -45.726 | -48.4085 | -49.5346 | -0.14191 | 0.009155 | -0.16022 | -0.16937 | 6.3309 | 6.4728 | 1.4053 | 2.6276 | 1.5472 | 2.7695 | -0.045777 | 0.022888 | 0 | -0.022888 | 5.8182 | 5.864 | 2.0096 | 3.1357 | 2.0554 | 3.1815
1857.1875 | 122.3282 | -51.5121 | 70.5139 | -0.13817 | 0.021422 | 0.083713 | 0.088572 | -0.0028049 | 0.018752 | 0.010083 | 0.068878 | 0.024953 | 0.084146 | 0.082204 | 0.00024568 | 0.012973 | 0.012192 | -0.10986 | 70.5002 | -51.6541 | 70.3674 | -45.2087 | -48.912 | -50.0931 | -0.027466 | 70.4819 | -51.5671 | 70.5093 | -45.7123 | -48.3536 | -49.5255 | -0.14191 | -0.013733 | -0.14648 | -0.13275 | 6.3034 | 6.4453 | 1.4191 | 2.6001 | 1.561 | 2.742 | -0.054932 | -0.032043 | -0.004578 | 0.027465 | 5.7999 | 5.8548 | 1.9867 | 3.1586 | 2.0416 | 3.2135
1858.1719 | 122.2321 | -51.6449 | 70.4269 | -0.13054 | 0.018859 | 0.083713 | 0.089854 | -0.0013095 | 0.020461 | 0.008374 | -0.025726 | 0.026235 | 0.085641 | 0.086477 | -0.0012497 | 0.014896 | 0.015182 | 0.19684 | 70.4361 | -51.7914 | 70.2347 | -45.369 | -49.0631 | -50.3174 | -0.036621 | 70.3766 | -51.7364 | 70.3812 | -45.8588 | -48.5229 | -49.7086 | -0.14648 | 0.009155 | -0.19226 | -0.20142 | 6.2759 | 6.4224 | 1.3275 | 2.5818 | 1.474 | 2.7283 | -0.091553 | -0.050354 | -0.045776 | 0.004578 | 5.7861 | 5.8777 | 1.9363 | 3.1219 | 2.0279 | 3.2135
1859.1719 | 122.2549 | -51.7914 | 70.3262 | -0.092392 | 0.019072 | 0.084781 | 0.091349 | -0.00088229 | 0.016188 | 0.011792 | 0.010895 | 0.025594 | 0.085001 | 0.088827 | -0.00018152 | 0.014682 | 0.012832 | -0.032043 | 70.3491 | -51.915 | 70.3125 | -45.4651 | -49.3652 | -50.3769 | -0.041199 | 70.3171 | -51.8143 | 70.2347 | -46.0098 | -48.6282 | -49.7818 | -0.1236 | 0.022888 | -0.013733 | -0.036621 | 6.3263 | 6.4499 | 1.4145 | 2.4261 | 1.5381 | 2.5497 | -0.022888 | -0.009155 | -0.091553 | -0.082398 | 5.7816 | 5.8044 | 2.0096 | 3.1631 | 2.0325 | 3.186
1860.1719 | 122.2641 | -51.7868 | 70.3033 | -0.084763 | 0.021209 | 0.083713 | 0.090068 | -0.0013095 | 0.015975 | 0.010724 | -0.056244 | 0.027517 | 0.085428 | 0.087972 | 0.00088648 | 0.014041 | 0.010483 | -0.032043 | 70.3079 | -51.9333 | 70.0974 | -45.5109 | -49.1638 | -50.386 | -0.004578 | 70.285 | -51.8417 | 70.253 | -45.9778 | -48.6557 | -49.8001 | -0.14649 | 0.004577 | -0.20599 | -0.21057 | 6.2759 | 6.4224 | 1.4008 | 2.623 | 1.5472 | 2.7695 | -0.054932 | -0.018311 | -0.050354 | -0.032043 | 5.809 | 5.864 | 1.9867 | 3.1311 | 2.0416 | 3.186
1861.1719 | 122.2733 | -51.4847 | 70.61 | -0.10002 | 0.021209 | 0.084781 | 0.09199 | -0.00066869 | 0.016616 | 0.0075196 | -0.025726 | 0.025594 | 0.083719 | 0.093526 | -0.0012497 | 0.014041 | 0.014969 | -0.022888 | 70.6238 | -51.6403 | 70.3857 | -45.2499 | -48.8846 | -50.1022 | -0.009155 | 70.5826 | -51.5396 | 70.5551 | -45.7443 | -48.3582 | -49.53 | -0.15564 | 0.013733 | -0.2243 | -0.23804 | 6.2347 | 6.3904 | 1.3824 | 2.6001 | 1.5381 | 2.7557 | -0.054932 | -0.027465 | -0.054931 | -0.027466 | 5.7404 | 5.7953 | 1.9547 | 3.1265 | 2.0096 | 3.1815
1862.1875 | 122.3465 | -51.6998 | 70.3125 | -0.1992 | 0.022917 | 0.087344 | 0.091349 | 0.00039951 | 0.018965 | 0.0094422 | -0.034882 | 0.024099 | 0.081583 | 0.086904 | -0.0018905 | 0.014896 | 0.011337 | -0.07782 | 70.3171 | -51.8555 | 70.2347 | -45.3735 | -49.0997 | -50.2853 | 0.009155 | 70.285 | -51.7593 | 70.3629 | -45.932 | -48.555 | -49.6811 | -0.15564 | 0.004578 | -0.07782 | -0.082398 | 6.3263 | 6.4819 | 1.4145 | 2.6001 | 1.5701 | 2.7557 | -0.059509 | -0.027466 | 0.050354 | 0.07782 | 5.7678 | 5.8273 | 2.0187 | 3.1448 | 2.0782 | 3.2043
1863.1719 | 122.3099 | -51.7502 | 70.3217 | -0.084763 | 0.020568 | 0.085635 | 0.091563 | -0.0021641 | 0.017897 | 0.0075196 | -0.047089 | 0.026449 | 0.08436 | 0.079 | -0.00060892 | 0.01511 | 0.010483 | -0.013733 | 70.3445 | -51.9012 | 70.1614 | -45.488 | -49.1364 | -50.3448 | 0.041199 | 70.3354 | -51.8051 | 70.3171 | -45.9961 | -48.6099 | -49.7452 | -0.15106 | 0.022888 | -0.16022 | -0.1831 | 6.2622 | 6.4133 | 1.4053 | 2.6138 | 1.5564 | 2.7649 | -0.054932 | 0.013733 | -0.004577 | -0.01831 | 5.7541 | 5.809 | 2.005 | 3.1403 | 2.0599 | 3.1952
1864.1719 | 122.2733 | -51.8784 | 70.2209 | -0.12291 | 0.023345 | 0.084781 | 0.090922 | -0.0013095 | 0.017043 | 0.0085876 | -0.025726 | 0.024953 | 0.082223 | 0.097158 | -0.00018152 | 0.01575 | 0.014755 | -0.018311 | 70.2347 | -52.0294 | 70.0287 | -45.7123 | -49.2462 | -50.473 | 0.022888 | 70.1797 | -51.9012 | 70.1797 | -46.1151 | -48.7289 | -49.8779 | -0.15106 | 0.013733 | -0.19226 | -0.20599 | 6.1661 | 6.3171 | 1.4053 | 2.6321 | 1.5564 | 2.7832 | -0.022888 | -0.041198 | -0.041198 | 0 | 5.7632 | 5.7861 | 2.0004 | 3.1494 | 2.0233 | 3.1723
1865.1719 | 122.2504 | -51.6449 | 70.4681 | -0.13054 | 0.021209 | 0.083713 | 0.091136 | -0.00066869 | 0.016829 | 0.01051 | -0.11728 | 0.026021 | 0.082651 | 0.089681 | 0.0011001 | 0.01575 | 0.0072781 | -0.022888 | 70.4773 | -51.8097 | 70.2209 | -45.5612 | -48.6969 | -50.0839 | -0.013733 | 70.4636 | -51.6632 | 70.3583 | -45.9412 | -48.5229 | -49.704 | -0.16479 | 0.009155 | -0.24719 | -0.25635 | 6.0837 | 6.2485 | 1.561 | 2.948 | 1.7258 | 3.1128 | -0.018311 | -0.004578 | -0.10986 | -0.10529 | 5.7037 | 5.722 | 1.9409 | 3.1219 | 1.9592 | 3.1403
1866.1719 | 122.2137 | -51.7548 | 70.3491 | -0.13817 | 0.021209 | 0.082858 | 0.08964 | -0.0034457 | 0.018752 | 0.01051 | -0.022675 | 0.028158 | 0.083719 | 0.09844 | 0.00045928 | 0.014896 | 0.011978 | -0.036621 | 70.34 | -51.8784 | 70.134 | -45.6161 | -49.1455 | -50.3586 | 0.004578 | 70.3354 | -51.7639 | 70.2805 | -46.0556 | -48.6877 | -49.8093 | -0.1236 | -0.009155 | -0.21515 | -0.20599 | 6.1386 | 6.2622 | 1.3962 | 2.6093 | 1.5198 | 2.7328 | -0.009155 | -0.013733 | -0.068664 | -0.054931 | 5.6992 | 5.7083 | 1.9455 | 3.067 | 1.9547 | 3.0762
1867.1719 | 122.2778 | -51.7822 | 70.3171 | -0.077133 | 0.021636 | 0.084781 | 0.091777 | 0.0016813 | 0.019606 | 0.0081604 | -0.028778 | 0.026021 | 0.087564 | 0.08669 | 0.00067288 | 0.013614 | 0.0096279 | -0.027466 | 70.34 | -51.9333 | 70.0836 | -45.6573 | -49.1913 | -50.3906 | -0.022888 | 70.2988 | -51.8188 | 70.2393 | -46.0968 | -48.674 | -49.8276 | -0.15106 | 0.022888 | -0.23346 | -0.25635 | 6.1249 | 6.2759 | 1.3916 | 2.5909 | 1.5427 | 2.742 | -0.036621 | -0.018311 | -0.07782 | -0.059509 | 5.6854 | 5.722 | 1.9547 | 3.1082 | 1.9913 | 3.1448
1868.1719 | 122.2504 | -51.8967 | 70.2026 | -0.24498 | 0.020354 | 0.083499 | 0.09434 | -0.00045509 | 0.020247 | 0.0094422 | -0.028778 | 0.025594 | 0.083719 | 0.094381 | -0.0010361 | 0.013828 | 0.016037 | 0.022888 | 70.2438 | -52.0111 | 69.9829 | -45.7581 | -49.3011 | -50.4959 | -0.018311 | 70.1843 | -51.9287 | 70.1111 | -46.2158 | -48.7885 | -49.9374 | -0.11444 | 0.041198 | -0.21973 | -0.26093 | 6.1386 | 6.2531 | 1.4008 | 2.5955 | 1.5152 | 2.71 | -0.032044 | -0.018311 | -0.091553 | -0.073242 | 5.6808 | 5.7129 | 1.9592 | 3.1082 | 1.9913 | 3.1403
1869.1719 | 122.2549 | -51.8921 | 70.2118 | -0.15343 | 0.020995 | 0.08649 | 0.092631 | -2.7692e-05 | 0.018111 | 0.010297 | -0.074554 | 0.026662 | 0.082437 | 0.09374 | -0.0014633 | 0.01511 | 0.011123 | -0.10071 | 70.2026 | -52.0477 | 69.9326 | -45.8176 | -49.3607 | -50.528 | -0.14648 | 70.1706 | -51.9516 | 70.1477 | -46.3806 | -48.5733 | -49.8276 | -0.15564 | -0.009155 | -0.27924 | -0.27008 | 6.0745 | 6.2302 | 1.3641 | 2.5314 | 1.5198 | 2.6871 | -0.059509 | -0.041199 | -0.064087 | -0.022888 | 5.5115 | 5.571 | 2.0645 | 3.3188 | 2.124 | 3.3783
1870.1719 | 122.2824 | -51.6037 | 70.4636 | -0.069504 | 0.021849 | 0.084567 | 0.090922 | -0.0055819 | 0.017897 | 0.0062378 | -0.03183 | 0.027089 | 0.08436 | 0.095235 | -0.00039512 | 0.015323 | 0.013473 | -0.018311 | 70.5002 | -51.7502 | 70.253 | -45.5246 | -49.0036 | -50.2304 | 0 | 70.459 | -51.6815 | 70.4315 | -45.932 | -48.4909 | -49.6262 | -0.14648 | 0.036621 | -0.21057 | -0.24719 | 6.0791 | 6.2256 | 1.3733 | 2.6001 | 1.5198 | 2.7466 | -0.07782 | -0.004578 | -0.032043 | -0.027465 | 5.6717 | 5.7495 | 1.9775 | 3.1128 | 2.0554 | 3.1906
1871.1719 | 122.2641 | -51.7502 | 70.3445 | -0.069504 | 0.022704 | 0.085849 | 0.090281 | -0.0025913 | 0.019393 | 0.0094422 | -0.047089 | 0.026021 | 0.084573 | 0.089254 | 3.2078e-05 | 0.01511 | 0.014114 | -0.027466 | 70.3537 | -51.8646 | 70.1431 | -45.6528 | -49.1409 | -50.3632 | -0.004578 | 70.3308 | -51.7868 | 70.2988 | -46.0464 | -48.6145 | -49.7772 | -0.11444 | 0.009156 | -0.20142 | -0.21057 | 6.0974 | 6.2119 | 1.387 | 2.6093 | 1.5015 | 2.7237 | -0.036621 | -0.013732 | -0.045776 | -0.032044 | 5.7037 | 5.7404 | 1.973 | 3.1357 | 2.0096 | 3.1723
1872.1719 | 122.2687 | -51.6495 | 70.4681 | -0.054245 | 0.019927 | 0.085208 | 0.08964 | 0.0012539 | 0.016188 | 0.011365 | -0.034882 | 0.026021 | 0.083505 | 0.086904 | 3.2078e-05 | 0.016391 | 0.012619 | -0.032043 | 70.4636 | -51.7868 | 70.2255 | -45.5841 | -49.0448 | -50.2625 | -0.018311 | 70.4086 | -51.709 | 70.3674 | -45.9549 | -48.5275 | -49.6902 | -0.13733 | -0.004578 | -0.24262 | -0.23804 | 6.0654 | 6.2027 | 1.387 | 2.6047 | 1.5244 | 2.742 | -0.059509 | -0.05951 | -0.10071 | -0.041198 | 5.6946 | 5.7541 | 1.9592 | 3.1219 | 2.0187 | 3.1815
1873.1719 | 122.2321 | -51.6678 | 70.459 | -0.1458 | 0.019927 | 0.084567 | 0.090922 | -0.00066869 | 0.017897 | 0.010724 | -0.013519 | 0.024099 | 0.082651 | 0.082204 | 0.00045928 | 0.014041 | 0.014969 | -0.018311 | 70.4727 | -51.8051 | 70.2209 | -45.5795 | -49.0494 | -50.2625 | -0.018311 | 70.4178 | -51.709 | 70.3537 | -45.9824 | -48.5321 | -49.704 | -0.13733 | 0.013733 | -0.23804 | -0.25177 | 6.0883 | 6.2256 | 1.4053 | 2.6184 | 1.5427 | 2.7557 | -0.041198 | -0.041198 | -0.10528 | -0.064087 | 5.6854 | 5.7266 | 1.9638 | 3.1357 | 2.005 | 3.1769
1874.1719 | 122.2458 | -51.5259 | 70.5505 | -0.13817 | 0.021209 | 0.08414 | 0.091349 | -0.0028049 | 0.015548 | 0.011365 | -0.080658 | 0.025808 | 0.084787 | 0.093313 | -0.0010361 | 0.016178 | 0.012192 | -0.022888 | 70.546 | -51.7044 | 70.34 | -45.4651 | -48.9441 | -50.148 | 0.027466 | 70.5276 | -51.5717 | 70.5231 | -45.8633 | -48.4222 | -49.5529 | -0.17853 | -0.004578 | -0.21057 | -0.20599 | 6.0608 | 6.2393 | 1.3779 | 2.5818 | 1.5564 | 2.7603 | -0.045776 | -0.022888 | -0.027466 | -0.004578 | 5.6625 | 5.7083 | 1.973 | 3.1036 | 2.0187 | 3.1494
1875.1719 | 122.2275 | -51.7044 | 70.3766 | -0.061874 | 0.019927 | 0.083713 | 0.088145 | -0.00066869 | 0.013839 | 0.0070922 | 0.02005 | 0.025594 | 0.08436 | 0.091176 | -0.00039512 | 0.012546 | 0.01625 | -0.013733 | 70.3812 | -51.8372 | 70.372 | -45.5475 | -49.3515 | -50.3998 | -0.027466 | 70.3995 | -51.7502 | 70.3079 | -46.1563 | -48.5916 | -49.7635 | -0.13275 | 0.004578 | -0.004578 | -0.009156 | 6.1569 | 6.2897 | 1.3046 | 2.3529 | 1.4374 | 2.4857 | -0.045776 | 0.022888 | -0.068665 | -0.091553 | 5.5481 | 5.5939 | 1.9409 | 3.1128 | 1.9867 | 3.1586
1876.1719 | 122.2412 | -51.4893 | 70.6284 | -0.092392 | 0.022704 | 0.084781 | 0.090495 | -0.0032321 | 0.016402 | 0.0088012 | 0.10245 | 0.025167 | 0.084573 | 0.087331 | 0.0023819 | 0.015323 | 0.012832 | -0.032043 | 70.6192 | -51.6083 | 70.3857 | -45.4697 | -48.8983 | -50.0885 | -0.013733 | 70.6604 | -51.4938 | 70.5368 | -45.9412 | -48.381 | -49.5392 | -0.11902 | -0.009155 | -0.24262 | -0.23346 | 6.0196 | 6.1386 | 1.4008 | 2.5909 | 1.5198 | 2.71 | -0.004577 | 0.032043 | -0.091553 | -0.1236 | 5.5481 | 5.5527 | 1.9501 | 3.1082 | 1.9546 | 3.1128
1877.1719 | 122.2549 | -51.3748 | 70.7108 | -0.077133 | 0.021849 | 0.084781 | 0.091777 | 0.0029629 | 0.016402 | 0.0085876 | -0.056244 | 0.024312 | 0.083719 | 0.089254 | -0.0016769 | 0.017246 | 0.0092007 | -0.041199 | 70.7108 | -51.5121 | 70.5231 | -45.3644 | -48.7747 | -49.9878 | 0.004578 | 70.7108 | -51.4297 | 70.6604 | -45.7993 | -48.2483 | -49.4156 | -0.13733 | 0 | -0.18768 | -0.18768 | 6.0104 | 6.1478 | 1.387 | 2.6001 | 1.5244 | 2.7374 | -0.054932 | 0 | -0.050354 | -0.050354 | 5.5756 | 5.6305 | 1.9592 | 3.1265 | 2.0142 | 3.1815
1878.1719 | 122.287 | -51.4755 | 70.6055 | -0.13054 | 0.021849 | 0.085422 | 0.090922 | -0.0019503 | 0.018965 | 0.011365 | -0.03183 | 0.027089 | 0.085001 | 0.093099 | 0.00067288 | 0.016391 | 0.012832 | -0.032043 | 70.6421 | -51.5945 | 70.3995 | -45.4559 | -48.8617 | -50.0748 | 0.004578 | 70.6329 | -51.4847 | 70.5643 | -45.9686 | -48.3627 | -49.498 | -0.11902 | 0.036621 | -0.20599 | -0.24262 | 6.0196 | 6.1386 | 1.4008 | 2.6138 | 1.5198 | 2.7328 | -0.009155 | 0.027466 | -0.041199 | -0.068665 | 5.5069 | 5.5161 | 1.9775 | 3.1128 | 1.9867 | 3.1219
1879.1719 | 122.2549 | -51.2741 | 70.7977 | -0.13054 | 0.022063 | 0.085208 | 0.090281 | -0.0025913 | 0.017897 | 0.011578 | -0.034882 | 0.024526 | 0.084146 | 0.092885 | -0.0023177 | 0.015537 | 0.013046 | -0.009155 | 70.784 | -51.4389 | 70.6284 | -45.2866 | -48.6557 | -49.8917 | 0.032043 | 70.7703 | -51.3245 | 70.7611 | -45.7123 | -48.143 | -49.2874 | -0.16479 | -0.013732 | -0.16937 | -0.15564 | 5.9875 | 6.1523 | 1.3824 | 2.6184 | 1.5472 | 2.7832 | -0.050354 | -0.027465 | -0.036621 | -0.009156 | 5.5618 | 5.6122 | 1.9867 | 3.1311 | 2.037 | 3.1815
1880.1719 | 122.2458 | -51.3794 | 70.7611 | -0.10002 | 0.020781 | 0.085849 | 0.090281 | -2.7692e-05 | 0.016829 | 0.0092286 | 0.013947 | 0.025167 | 0.086496 | 0.089681 | 0.0015275 | 0.014469 | 0.0098417 | 0.004578 | 70.7291 | -51.4801 | 70.5368 | -45.3598 | -48.7473 | -49.9512 | 0.018311 | 70.7062 | -51.4114 | 70.6421 | -45.8633 | -48.262 | -49.4064 | -0.10071 | -0.032043 | -0.2243 | -0.19226 | 6.0196 | 6.1203 | 1.4282 | 2.6321 | 1.5289 | 2.7328 | -0.032043 | -0.054931 | -0.11902 | -0.064087 | 5.5161 | 5.5481 | 1.973 | 3.1174 | 2.005 | 3.1494
1881.1719 | 122.2275 | -51.4114 | 70.6696 | -0.069504 | 0.021209 | 0.086276 | 0.092845 | 0.00061311 | 0.019606 | 0.011578 | 0.016998 | 0.025808 | 0.084573 | 0.093099 | 0.00024568 | 0.014469 | 0.012405 | -0.050354 | 70.6924 | -51.5625 | 70.4361 | -45.4285 | -48.8388 | -50.0381 | 0.09613 | 70.7199 | -51.4297 | 70.5826 | -46.0052 | -48.3398 | -49.4888 | -0.15106 | 0.022888 | -0.23346 | -0.25635 | 5.983 | 6.134 | 1.3733 | 2.5726 | 1.5244 | 2.7237 | -0.018311 | 0.050354 | -0.086975 | -0.13733 | 5.4062 | 5.4245 | 1.9226 | 3.0716 | 1.9409 | 3.0899
1882.1719 | 122.2687 | -51.4893 | 70.5963 | -0.084763 | 0.0195 | 0.083499 | 0.091136 | -0.0032321 | 0.018325 | 0.0090148 | -0.022675 | 0.024099 | 0.083292 | 0.090535 | -0.0018905 | 0.011692 | 0.015182 | -0.022888 | 70.5917 | -51.6357 | 70.3125 | -45.5612 | -48.9532 | -50.148 | -0.022888 | 70.5368 | -51.5671 | 70.5505 | -46.2021 | -48.4039 | -49.5346 | -0.14648 | -0.004577 | -0.28381 | -0.27924 | 5.928 | 6.0745 | 1.3412 | 2.536 | 1.4877 | 2.6825 | -0.07782 | -0.059509 | -0.045776 | 0.013733 | 5.2872 | 5.365 | 1.9547 | 3.0853 | 2.0325 | 3.1631
1883.1719 | 122.1588 | -51.4755 | 70.6512 | -0.038986 | 0.020568 | 0.087344 | 0.089213 | -2.7692e-05 | 0.015761 | 0.011365 | -0.016571 | 0.02538 | 0.085428 | 0.082845 | -0.0012497 | 0.012119 | 0.01625 | -0.032043 | 70.5917 | -51.6357 | 70.3903 | -45.4742 | -48.9029 | -50.1022 | 0.018311 | 70.5872 | -51.5488 | 70.5048 | -46.0648 | -48.3994 | -49.5804 | -0.16022 | -0.059509 | -0.26093 | -0.20142 | 6.0013 | 6.1615 | 1.3733 | 2.5726 | 1.5335 | 2.7328 | -0.073242 | -0.064087 | -0.14648 | -0.082397 | 5.4108 | 5.484 | 1.8951 | 3.0762 | 1.9684 | 3.1494
1884.1719 | 122.2366 | -51.4893 | 70.61 | -0.16869 | 0.023986 | 0.084994 | 0.093058 | 0.0023221 | 0.017256 | 0.011578 | -0.11118 | 0.026021 | 0.083078 | 0.090322 | -0.0010361 | 0.018741 | 0.01091 | -0.10986 | 70.5826 | -51.6861 | 70.3262 | -45.5658 | -48.9807 | -50.1617 | -0.11902 | 70.5826 | -51.5625 | 70.4224 | -45.8588 | -48.3856 | -49.5667 | -0.19684 | -0.027465 | -0.28381 | -0.25635 | 5.9235 | 6.1203 | 1.3275 | 2.5085 | 1.5244 | 2.7054 | -0.073242 | -0.027465 | -0.18768 | -0.16022 | 5.6305 | 5.7037 | 1.9226 | 3.1036 | 1.9958 | 3.1769
1885.1719 | 122.2595 | -51.5305 | 70.5643 | -0.069504 | 0.022277 | 0.084353 | 0.088786 | 0.0014675 | 0.016402 | 0.010937 | -0.13864 | 0.025808 | 0.084146 | 0.090963 | 3.2078e-05 | 0.014896 | 0.012619 | -0.036621 | 70.5826 | -51.6724 | 70.34 | -45.52 | -48.5916 | -50.0839 | 0.018311 | 70.5688 | -51.5533 | 70.4819 | -46.0831 | -48.4314 | -49.585 | -0.14191 | 0.018311 | -0.2243 | -0.24262 | 6.0104 | 6.1523 | 1.4465 | 2.9388 | 1.5884 | 3.0807 | -0.022888 | 0.004578 | -0.082397 | -0.086975 | 5.4474 | 5.4703 | 1.9455 | 3.0991 | 1.9684 | 3.1219
1886.1719 | 122.2549 | -51.5762 | 70.5139 | -0.077133 | 0.021849 | 0.082858 | 0.08964 | -2.7692e-05 | 0.017684 | 0.0088012 | -0.0043641 | 0.025167 | 0.082223 | 0.092672 | 0.00067288 | 0.012973 | 0.014328 | 0.013733 | 70.491 | -51.6861 | 70.3537 | -45.5475 | -48.9212 | -50.1343 | 0.054932 | 70.4773 | -51.5991 | 70.4498 | -46.1517 | -48.4543 | -49.6124 | -0.10986 | -0.022888 | -0.16022 | -0.13733 | 6.0287 | 6.1386 | 1.442 | 2.655 | 1.5518 | 2.7649 | -0.022888 | -0.036621 | -0.064087 | -0.027466 | 5.4245 | 5.4474 | 1.9638 | 3.1219 | 1.9867 | 3.1448
1887.1719 | 122.2412 | -51.5991 | 70.5139 | -0.10765 | 0.022277 | 0.084994 | 0.091349 | 0.00018591 | 0.018325 | 0.011578 | -0.028778 | 0.02538 | 0.086282 | 0.08904 | 0.00024568 | 0.016178 | 0.014755 | -0.027466 | 70.5048 | -51.7044 | 70.3079 | -45.6299 | -48.9853 | -50.1892 | 0.009155 | 70.491 | -51.6174 | 70.4361 | -46.1838 | -48.4726 | -49.6536 | -0.10529 | -0.009155 | -0.20599 | -0.19684 | 5.9692 | 6.0745 | 1.4099 | 2.6138 | 1.5152 | 2.7191 | -0.018311 | -0.022888 | -0.07782 | -0.054932 | 5.4153 | 5.4337 | 1.9455 | 3.1265 | 1.9638 | 3.1448
1888.1719 | 122.2229 | -51.2466 | 70.8481 | -0.1458 | 0.02249 | 0.083713 | 0.089 | -0.00088229 | 0.018752 | 0.0079468 | 0.0078429 | 0.026662 | 0.082651 | 0.089895 | 0.0019547 | 0.013614 | 0.012405 | -0.018311 | 70.8023 | -51.416 | 70.6604 | -45.3415 | -48.7015 | -49.8642 | 0.013733 | 70.8252 | -51.2833 | 70.7291 | -45.8862 | -48.1476 | -49.3469 | -0.16937 | -0.045776 | -0.18768 | -0.14191 | 5.9052 | 6.0745 | 1.3824 | 2.5452 | 1.5518 | 2.7145 | -0.036621 | -0.022888 | -0.11902 | -0.09613 | 5.3604 | 5.397 | 1.8997 | 3.0991 | 1.9363 | 3.1357
1889.1719 | 122.2641 | -51.2558 | 70.8298 | -0.10002 | 0.02249 | 0.087771 | 0.09199 | 0.0044583 | 0.018752 | 0.012646 | -0.056244 | 0.027517 | 0.083505 | 0.087118 | -0.0014633 | 0.014469 | 0.012405 | 0 | 70.7932 | -51.3886 | 70.6741 | -45.1538 | -48.6053 | -49.7589 | 0.018311 | 70.8435 | -51.2558 | 70.8115 | -45.8588 | -48.1567 | -49.292 | -0.13275 | -0.036621 | -0.15564 | -0.11902 | 6.102 | 6.2347 | 1.4969 | 2.6505 | 1.6296 | 2.7832 | 0 | 0.013733 | -0.018311 | -0.032044 | 5.397 | 5.397 | 1.9638 | 3.0991 | 1.9638 | 3.0991
1890.1719 | 122.2183 | -51.3107 | 70.7932 | -0.13054 | 0.02014 | 0.084781 | 0.089427 | 0.0014675 | 0.017256 | 0.0090148 | 0.0047913 | 0.025594 | 0.082437 | 0.088399 | -0.00082252 | 0.016818 | 0.012619 | -0.027466 | 70.8115 | -51.4297 | 70.5688 | -45.4376 | -48.7198 | -49.91 | 0 | 70.7886 | -51.3382 | 70.7016 | -45.9229 | -48.2071 | -49.3835 | -0.11902 | 0.01831 | -0.2243 | -0.24261 | 5.8731 | 5.9921 | 1.4008 | 2.5909 | 1.5198 | 2.71 | -0.027466 | -0.004578 | -0.091553 | -0.086975 | 5.3879 | 5.4153 | 1.9272 | 3.1036 | 1.9547 | 3.1311
1891.1719 | 122.2366 | -51.3748 | 70.7062 | -0.13817 | 0.021636 | 0.087558 | 0.090068 | -0.0077183 | 0.016829 | 0.010724 | 0.0017395 | 0.025167 | 0.085214 | 0.083059 | 0.00045928 | 0.013187 | 0.013687 | 0.004578 | 70.7108 | -51.4801 | 70.5139 | -45.488 | -48.7473 | -49.9466 | 0.013733 | 70.7016 | -51.3886 | 70.6421 | -45.9778 | -48.2666 | -49.4476 | -0.10529 | 0.004577 | -0.19226 | -0.19684 | 5.8868 | 5.9921 | 1.4282 | 2.6276 | 1.5335 | 2.7328 | -0.013733 | -0.004578 | -0.064087 | -0.059509 | 5.397 | 5.4108 | 1.9272 | 3.1082 | 1.9409 | 3.1219
1892.1719 | 122.2458 | -51.5213 | 70.5872 | -0.12291 | 0.020995 | 0.086062 | 0.091136 | -0.00088229 | 0.018965 | 0.010724 | -0.013519 | 0.022817 | 0.082437 | 0.092885 | -0.00082252 | 0.016178 | 0.012832 | 0.036621 | 70.5597 | -51.5945 | 70.4361 | -45.5658 | -48.8388 | -50.061 | 0.064087 | 70.578 | -51.5396 | 70.5139 | -46.138 | -48.3719 | -49.5621 | -0.073242 | -0.027466 | -0.15106 | -0.1236 | 5.9555 | 6.0287 | 1.4603 | 2.6825 | 1.5335 | 2.7557 | -0.018311 | -0.009155 | -0.073242 | -0.064087 | 5.3833 | 5.4016 | 1.9592 | 3.1494 | 1.9775 | 3.1677
1893.1719 | 122.3602 | -51.3931 | 70.6055 | -0.13817 | 0.019286 | 0.083713 | 0.090922 | -0.00045509 | 0.01747 | 0.0098694 | -0.12949 | 0.024312 | 0.083292 | 0.090322 | -0.00039512 | 0.016178 | 0.010055 | 0.036621 | 70.5826 | -51.5396 | 70.5505 | -45.4376 | -48.4909 | -49.8505 | 0.11902 | 70.5597 | -51.4664 | 70.6696 | -45.9961 | -48.2437 | -49.4431 | -0.14648 | -0.022888 | -0.054932 | -0.032044 | 5.9555 | 6.102 | 1.5427 | 2.9022 | 1.6891 | 3.0487 | -0.073243 | -0.045777 | 0.064087 | 0.10986 | 5.397 | 5.4703 | 1.9501 | 3.1494 | 2.0233 | 3.2227
1894.1719 | 122.3053 | -51.6083 | 70.459 | -0.031357 | 0.023345 | 0.085422 | 0.093699 | 0.00082671 | 0.016829 | 0.014783 | -0.074554 | 0.02474 | 0.082651 | 0.098867 | 0.0011001 | 0.015537 | 0.012619 | 0.027466 | 70.4773 | -51.7593 | 70.3125 | -45.6894 | -48.9716 | -50.1938 | 0.07782 | 70.459 | -51.6449 | 70.4544 | -46.2479 | -48.5138 | -49.6765 | -0.15106 | 0.018311 | -0.14648 | -0.16479 | 5.9189 | 6.0699 | 1.4145 | 2.6367 | 1.5656 | 2.7878 | -0.036621 | 0 | -0.004577 | -0.004577 | 5.3604 | 5.397 | 1.9318 | 3.0945 | 1.9684 | 3.1311
1895.1719 | 122.3236 | -51.4709 | 70.5368 | -0.11528 | 0.020354 | 0.084994 | 0.08964 | 0.0016813 | 0.016616 | 0.0085876 | -0.010467 | 0.024526 | 0.082864 | 0.089467 | -0.00082252 | 0.012119 | 0.0092007 | 0.036621 | 70.546 | -51.622 | 70.4361 | -45.7809 | -48.8754 | -49.9374 | 0.10071 | 70.5093 | -51.5488 | 70.5963 | -46.0876 | -48.3398 | -49.5026 | -0.15106 | 0.009155 | -0.10071 | -0.10986 | 5.69 | 5.8411 | 1.5335 | 2.5955 | 1.6846 | 2.7466 | -0.07782 | -0.027466 | 0.059509 | 0.086975 | 5.3833 | 5.4611 | 1.9684 | 3.1311 | 2.0462 | 3.2089
1896.1719 | 122.3648 | -51.384 | 70.6604 | -0.16106 | 0.020354 | 0.085422 | 0.08964 | -0.00066869 | 0.017043 | 0.010297 | 0.02005 | 0.022176 | 0.080728 | 0.090322 | -0.0018905 | 0.016605 | 0.011551 | 0.050354 | 70.6375 | -51.4893 | 70.6146 | -45.3873 | -48.6786 | -49.8734 | 0.14648 | 70.6421 | -51.3977 | 70.7336 | -46.0052 | -48.4268 | -49.4843 | -0.10529 | -0.022888 | -0.045776 | -0.022888 | 5.9967 | 6.102 | 1.5106 | 2.7054 | 1.6159 | 2.8107 | -0.013733 | -0.01831 | 0.073243 | 0.091553 | 5.3787 | 5.3925 | 1.8997 | 2.9572 | 1.9135 | 2.9709
1897.1719 | 122.3877 | -51.4481 | 70.5139 | -0.15343 | 0.018859 | 0.084353 | 0.091777 | -0.0021641 | 0.019179 | 0.0077332 | 0.023102 | 0.023671 | 0.080301 | 0.092672 | -0.00018152 | 0.013187 | 0.013046 | 0.027466 | 70.5276 | -51.6037 | 70.4956 | -45.5246 | -48.8068 | -49.9924 | 0.16022 | 70.5414 | -51.5167 | 70.6192 | -46.0785 | -48.2986 | -49.5071 | -0.15564 | 0.013733 | -0.018311 | -0.032044 | 5.9235 | 6.0791 | 1.4557 | 2.6413 | 1.6113 | 2.7969 | -0.068665 | 0.027466 | 0.10529 | 0.07782 | 5.3696 | 5.4382 | 1.9409 | 3.1494 | 2.0096 | 3.2181
1898.1719 | 122.3373 | -51.4847 | 70.5093 | -0.13054 | 0.02249 | 0.084567 | 0.088145 | 0.0012539 | 0.017684 | 0.0088012 | 0.0017395 | 0.025167 | 0.083932 | 0.087118 | 0.00088648 | 0.0181 | 0.013687 | 0.027466 | 70.5231 | -51.6357 | 70.4636 | -45.5429 | -48.8068 | -50.0244 | 0.15564 | 70.4956 | -51.5533 | 70.5597 | -46.0648 | -48.3994 | -49.5575 | -0.15106 | 0.013733 | -0.045776 | -0.059509 | 5.9418 | 6.0928 | 1.4603 | 2.6779 | 1.6113 | 2.829 | -0.068665 | -0.013733 | 0.050354 | 0.064087 | 5.4199 | 5.4886 | 1.9272 | 3.0853 | 1.9958 | 3.154
1899.1719 | 122.2366 | -51.4023 | 70.7108 | -0.12291 | 0.02014 | 0.083926 | 0.093913 | 0.0016813 | 0.016616 | 0.012646 | -0.068451 | 0.027303 | 0.082864 | 0.092672 | -0.0023177 | 0.01511 | 0.011764 | 0 | 70.6879 | -51.5442 | 70.491 | -45.5933 | -48.8159 | -50.0244 | -0.009155 | 70.6787 | -51.4709 | 70.6329 | -46.1014 | -48.3307 | -49.498 | -0.14191 | -0.022888 | -0.21973 | -0.19684 | 5.809 | 5.9509 | 1.3779 | 2.5864 | 1.5198 | 2.7283 | -0.068664 | -0.032043 | -0.077819 | -0.045776 | 5.3009 | 5.3696 | 1.9043 | 3.0716 | 1.973 | 3.1403
1900.1719 | 122.319 | -51.4252 | 70.6146 | -0.13817 | 0.020568 | 0.08649 | 0.090495 | -0.0013095 | 0.017684 | 0.0096558 | 0.0017395 | 0.023244 | 0.082651 | 0.091176 | -0.0010361 | 0.015323 | 0.0094143 | -0.041199 | 70.5826 | -51.5854 | 70.491 | -45.5429 | -48.8251 | -50.0107 | 0.032043 | 70.5505 | -51.4893 | 70.6421 | -46.0968 | -48.3124 | -49.4614 | -0.16022 | -0.032043 | -0.1236 | -0.091553 | 5.8823 | 6.0425 | 1.4145 | 2.6001 | 1.5747 | 2.7603 | -0.064087 | -0.064087 | 0.027466 | 0.091553 | 5.3284 | 5.3925 | 1.9638 | 3.1128 | 2.0279 | 3.1769
1901.1719 | 122.374 | -51.2833 | 70.7291 | -0.13817 | 0.020781 | 0.085635 | 0.088145 | -0.0062229 | 0.01747 | 0.010724 | -0.0013123 | 0.024953 | 0.082651 | 0.092672 | -0.00060892 | 0.011692 | 0.011123 | -0.009155 | 70.7382 | -51.4435 | 70.6375 | -45.4239 | -48.6374 | -49.8276 | 0.073242 | 70.7108 | -51.3611 | 70.7977 | -45.9274 | -48.1384 | -49.3195 | -0.16022 | 0.009155 | -0.091553 | -0.10071 | 5.8594 | 6.0196 | 1.4557 | 2.6459 | 1.6159 | 2.8061 | -0.07782 | -0.018311 | 0.068664 | 0.086975 | 5.3558 | 5.4337 | 1.9638 | 3.1448 | 2.0416 | 3.2227
1902.1719 | 122.3602 | -51.5854 | 70.4315 | -0.069504 | 0.018218 | 0.084781 | 0.090068 | -0.0021641 | 0.015761 | 0.0090148 | 0.0078429 | 0.023458 | 0.084146 | 0.087972 | -0.00060892 | 0.014682 | 0.012832 | 0.009155 | 70.4498 | -51.7273 | 70.3857 | -45.6711 | -48.9029 | -50.116 | 0.15106 | 70.4269 | -51.6449 | 70.5322 | -46.2021 | -48.3856 | -49.6033 | -0.14191 | 0.01831 | -0.045777 | -0.064087 | 5.9143 | 6.0562 | 1.4694 | 2.6825 | 1.6113 | 2.8244 | -0.059509 | -0.004578 | 0.10071 | 0.10529 | 5.3833 | 5.4428 | 1.9821 | 3.1998 | 2.0416 | 3.2593
1903.1719 | 122.3648 | -51.5579 | 70.4315 | -0.16869 | 0.022063 | 0.083926 | 0.090922 | 0.0033901 | 0.016829 | 0.012006 | -0.016571 | 0.022603 | 0.081369 | 0.0837 | 0.00088648 | 0.014896 | 0.0094143 | -0.009155 | 70.4361 | -51.7227 | 70.3766 | -45.7031 | -48.9441 | -50.1205 | 0.10071 | 70.4178 | -51.6403 | 70.5231 | -46.3028 | -48.4268 | -49.6262 | -0.1648 | 0.004577 | -0.054932 | -0.059509 | 5.8548 | 6.0196 | 1.4374 | 2.6138 | 1.6022 | 2.7786 | -0.082398 | -0.013733 | 0.091552 | 0.10528 | 5.2551 | 5.3375 | 1.9318 | 3.1311 | 2.0142 | 3.2135
1904.1719 | 122.3373 | -51.2833 | 70.7382 | -0.16106 | 0.0195 | 0.084781 | 0.090709 | -0.00024129 | 0.014907 | 0.0088012 | -0.010467 | 0.022603 | 0.082864 | 0.087331 | -0.0025315 | 0.01575 | 0.0074917 | -0.036621 | 70.7245 | -51.4297 | 70.6512 | -45.4376 | -48.6649 | -49.8367 | 0.054932 | 70.7199 | -51.329 | 70.7932 | -46.0785 | -48.143 | -49.3378 | -0.14648 | -0.013733 | -0.086975 | -0.073242 | 5.8456 | 5.9921 | 1.4465 | 2.6184 | 1.593 | 2.7649 | -0.045777 | -0.01831 | 0.054932 | 0.073242 | 5.2048 | 5.2505 | 1.9455 | 3.1403 | 1.9913 | 3.186
1905.1719 | 122.3373 | -51.9287 | 70.1111 | -0.1458 | 0.023772 | 0.086276 | 0.090922 | 0.00018591 | 0.017043 | 0.0098694 | -0.016571 | 0.021963 | 0.083078 | 0.09139 | -0.00018152 | 0.013828 | 0.011764 | -0.036621 | 70.0882 | -52.0752 | 70.0104 | -46.0693 | -49.3057 | -50.4822 | 0.022888 | 70.0836 | -51.9699 | 70.1523 | -46.8979 | -48.7885 | -49.9649 | -0.14648 | -0.022888 | -0.10071 | -0.07782 | 5.8594 | 6.0059 | 1.4465 | 2.623 | 1.593 | 2.7695 | -0.041199 | -0.027466 | 0.041199 | 0.068665 | 5.0308 | 5.072 | 1.9638 | 3.1403 | 2.005 | 3.1815
1906.1719 | 122.3053 | -51.8692 | 70.1477 | -0.084763 | 0.019286 | 0.085208 | 0.09199 | -0.00066869 | 0.019606 | 0.0092286 | -0.025726 | 0.024312 | 0.082864 | 0.092458 | 0.00088648 | 0.014255 | 0.0098417 | -0.091553 | 70.1569 | -52.0157 | 70.0378 | -46.0419 | -49.2691 | -50.4318 | -0.032043 | 70.1248 | -51.9058 | 70.2072 | -46.9025 | -48.7839 | -49.942 | -0.14648 | 0.009155 | -0.10986 | -0.11902 | 5.8273 | 5.9738 | 1.4374 | 2.6001 | 1.5839 | 2.7466 | -0.036621 | -0.022888 | 0.059509 | 0.082397 | 4.9667 | 5.0034 | 1.9272 | 3.0853 | 1.9638 | 3.1219
1907.1719 | 122.3145 | -51.622 | 70.4178 | -0.18394 | 0.020568 | 0.085422 | 0.08964 | 0.00061311 | 0.017897 | 0.011792 | -0.053192 | 0.024099 | 0.082651 | 0.093526 | -0.00060892 | 0.014469 | 0.010269 | -0.068665 | 70.4132 | -51.7685 | 70.3079 | -45.7809 | -49.0128 | -50.1892 | -0.022888 | 70.4086 | -51.6769 | 70.4315 | -46.6461 | -48.5321 | -49.6857 | -0.14648 | -0.004578 | -0.10986 | -0.10529 | 5.8411 | 5.9875 | 1.4328 | 2.6093 | 1.5793 | 2.7557 | -0.054932 | -0.009156 | 0.013733 | 0.022889 | 4.9759 | 5.0308 | 1.9363 | 3.0899 | 1.9913 | 3.1448
1908.1719 | 122.3007 | -51.6449 | 70.34 | -0.18394 | 0.021209 | 0.085208 | 0.090068 | -0.0010959 | 0.018965 | 0.010297 | -0.03183 | 0.023244 | 0.082223 | 0.085836 | -0.0021041 | 0.012119 | 0.0081327 | -0.073242 | 70.3949 | -51.7776 | 70.2942 | -45.8176 | -49.0311 | -50.2075 | -0.036621 | 70.3629 | -51.6769 | 70.4269 | -46.6599 | -48.5458 | -49.7269 | -0.13275 | 0.054931 | -0.045777 | -0.10071 | 5.8273 | 5.9601 | 1.4374 | 2.6138 | 1.5701 | 2.7466 | -0.032044 | 0.022888 | 0.086975 | 0.064087 | 4.985 | 5.0171 | 1.918 | 3.0991 | 1.9501 | 3.1311
1909.1875 | 122.2687 | -51.7181 | 70.3766 | -0.092392 | 0.021422 | 0.086062 | 0.092204 | 0.00039951 | 0.019179 | 0.01051 | -0.14169 | 0.02474 | 0.081155 | 0.093099 | -0.0012497 | 0.017673 | 0.010696 | -0.027466 | 70.3903 | -51.8692 | 70.2347 | -45.8908 | -49.054 | -50.2945 | -0.004578 | 70.3812 | -51.7273 | 70.3171 | -46.7422 | -48.6465 | -49.8093 | -0.15106 | 0.013733 | -0.14191 | -0.15564 | 5.8273 | 5.9784 | 1.4236 | 2.6642 | 1.5747 | 2.8152 | -0.009155 | 0.004578 | -0.059509 | -0.064087 | 4.9759 | 4.985 | 1.9089 | 3.0716 | 1.918 | 3.0807
1910.1719 | 122.3236 | -51.6953 | 70.3033 | -0.13054 | 0.020568 | 0.085208 | 0.091563 | -0.0019503 | 0.017256 | 0.011578 | -0.056244 | 0.022603 | 0.082223 | 0.079854 | -0.00082252 | 0.014255 | 0.010269 | -0.064087 | 70.3217 | -51.86 | 70.2118 | -45.9 | -49.1135 | -50.2716 | -0.064087 | 70.3079 | -51.7868 | 70.3583 | -46.7194 | -48.6237 | -49.8047 | -0.16479 | 0.01831 | -0.091553 | -0.10986 | 5.7953 | 5.9601 | 1.4236 | 2.5818 | 1.5884 | 2.7466 | -0.091553 | 0.004577 | 0.054931 | 0.050354 | 4.9759 | 5.0674 | 1.8906 | 3.0716 | 1.9821 | 3.1631
1911.1719 | 122.3282 | -51.7044 | 70.3217 | -0.11528 | 0.022277 | 0.084567 | 0.090495 | -0.0038729 | 0.019393 | 0.014142 | -0.040985 | 0.022817 | 0.082864 | 0.083486 | 3.2078e-05 | 0.014682 | 0.014328 | -0.073242 | 70.3857 | -51.7548 | 70.285 | -45.8267 | -49.0173 | -50.1846 | -0.045776 | 70.3812 | -51.6907 | 70.4407 | -46.6415 | -48.674 | -49.7955 | -0.050354 | 0.064087 | -0.036621 | -0.10071 | 5.8777 | 5.928 | 1.5198 | 2.6871 | 1.5701 | 2.7374 | 0.013733 | 0.05951 | 0.11902 | 0.059509 | 5.0629 | 5.0491 | 1.9089 | 3.0304 | 1.8951 | 3.0167
1912.1719 | 122.3236 | -51.5671 | 70.4498 | -0.10765 | 0.0195 | 0.085635 | 0.090709 | 0.0027493 | 0.015761 | 0.010724 | -0.05014 | 0.026662 | 0.080728 | 0.095022 | -0.0033859 | 0.014041 | 0.011764 | -0.068665 | 70.4636 | -51.6907 | 70.3537 | -45.7581 | -48.967 | -50.1389 | -0.045776 | 70.4681 | -51.6037 | 70.491 | -46.5958 | -48.4909 | -49.6399 | -0.1236 | 0.013733 | -0.09613 | -0.10986 | 5.809 | 5.9326 | 1.4282 | 2.6001 | 1.5518 | 2.7237 | -0.036621 | 0.018311 | 0.041199 | 0.022888 | 4.9713 | 5.0079 | 1.9272 | 3.0762 | 1.9638 | 3.1128
1913.1719 | 122.2412 | -51.4572 | 70.6146 | -0.13817 | 0.022277 | 0.084781 | 0.091777 | -2.7692e-05 | 0.017684 | 0.01286 | -0.14474 | 0.025808 | 0.08201 | 0.096517 | -0.0025315 | 0.015323 | 0.010055 | -0.009155 | 70.6375 | -51.622 | 70.4224 | -45.7169 | -48.8846 | -50.0885 | 0.027466 | 70.61 | -51.503 | 70.5872 | -46.5134 | -48.4085 | -49.5621 | -0.16479 | 0.022888 | -0.19226 | -0.21515 | 5.7404 | 5.9052 | 1.3687 | 2.5726 | 1.5335 | 2.7374 | -0.045777 | -0.004578 | -0.027466 | -0.022888 | 4.9438 | 4.9896 | 1.8951 | 3.0487 | 1.9409 | 3.0945
1914.1719 | 122.2916 | -51.7273 | 70.2805 | -0.061874 | 0.022917 | 0.084994 | 0.08964 | -0.0030185 | 0.018111 | 0.01051 | -0.016571 | 0.023885 | 0.083719 | 0.09139 | -0.0014633 | 0.014041 | 0.010483 | 0.022888 | 70.2713 | -51.8738 | 70.2209 | -45.9183 | -49.0723 | -50.2579 | 0.091553 | 70.2988 | -51.7868 | 70.3125 | -46.756 | -48.642 | -49.8276 | -0.14648 | -0.009156 | -0.05951 | -0.050354 | 5.809 | 5.9555 | 1.4694 | 2.655 | 1.6159 | 2.8015 | -0.059509 | 0.01831 | 0.032043 | 0.013733 | 4.9713 | 5.0308 | 1.8997 | 3.0853 | 1.9592 | 3.1448
1915.1719 | 122.2778 | -51.7731 | 70.2255 | -0.15343 | 0.020568 | 0.086917 | 0.090922 | -0.00066869 | 0.018965 | 0.0081604 | -0.05014 | 0.024526 | 0.081583 | 0.089254 | 0.0011001 | 0.014682 | 0.011978 | -0.054932 | 70.2209 | -51.915 | 70.1523 | -45.9778 | -49.1776 | -50.3403 | -0.013733 | 70.2438 | -51.8326 | 70.2942 | -46.8018 | -48.6832 | -49.8642 | -0.14191 | -0.004578 | -0.073242 | -0.068664 | 5.7953 | 5.9372 | 1.4328 | 2.5955 | 1.5747 | 2.7374 | -0.05951 | 0.01831 | 0.068664 | 0.050354 | 4.9713 | 5.0308 | 1.9089 | 3.0899 | 1.9684 | 3.1494
1916.1719 | 122.3373 | -51.7227 | 70.285 | -0.1458 | 0.021849 | 0.08414 | 0.089854 | 0.0010403 | 0.017256 | 0.013287 | -0.037933 | 0.022817 | 0.081583 | 0.086904 | 3.2078e-05 | 0.014255 | 0.0098417 | -0.064087 | 70.2621 | -51.8784 | 70.2164 | -45.9183 | -49.118 | -50.2991 | 0.009155 | 70.2896 | -51.7914 | 70.3537 | -46.7377 | -48.6328 | -49.7864 | -0.15564 | -0.022888 | -0.068664 | -0.045776 | 5.8044 | 5.9601 | 1.4236 | 2.6047 | 1.5793 | 2.7603 | -0.068665 | 0.004578 | 0.068665 | 0.064087 | 4.985 | 5.0537 | 1.9363 | 3.0899 | 2.005 | 3.1586
1917.1719 | 122.319 | -51.8097 | 70.1889 | -0.18394 | 0.019713 | 0.084781 | 0.091136 | 0.00061311 | 0.018538 | 0.0098694 | -0.068451 | 0.024099 | 0.08201 | 0.088399 | -0.0010361 | 0.014682 | 0.0066373 | -0.036621 | 70.1614 | -51.9608 | 70.1294 | -46.0052 | -49.205 | -50.3906 | 0.013733 | 70.1614 | -51.8555 | 70.2805 | -46.8063 | -48.7335 | -49.8779 | -0.15106 | -0.027466 | -0.059509 | -0.032043 | 5.8044 | 5.9555 | 1.4191 | 2.6047 | 1.5701 | 2.7557 | -0.045777 | -0.027466 | 0.091553 | 0.11902 | 5.0034 | 5.0491 | 1.9318 | 3.0762 | 1.9775 | 3.1219
1918.1875 | 122.2961 | -51.915 | 70.1843 | -0.13817 | 0.021636 | 0.083285 | 0.090922 | 0.0021085 | 0.017684 | 0.010937 | -0.12643 | 0.024526 | 0.08201 | 0.08434 | -0.00018152 | 0.016178 | 0.0096279 | -0.027466 | 70.2026 | -52.0203 | 70.015 | -46.1792 | -49.2874 | -50.5051 | -0.032043 | 70.2072 | -51.915 | 70.1706 | -46.9025 | -48.8297 | -49.9786 | -0.10529 | 0.018311 | -0.16937 | -0.18768 | 5.7358 | 5.8411 | 1.4099 | 2.6276 | 1.5152 | 2.7328 | 0 | 0.022888 | -0.013733 | -0.036621 | 5.0125 | 5.0125 | 1.9363 | 3.0853 | 1.9363 | 3.0853
1919.1719 | 122.3785 | -51.8463 | 70.1523 | -0.1458 | 0.021422 | 0.083926 | 0.087504 | -0.0055819 | 0.018965 | 0.0094422 | 0.0047913 | 0.022603 | 0.082651 | 0.089681 | 0.00067288 | 0.01276 | 0.011123 | -0.027466 | 70.1614 | -52.0157 | 70.0836 | -46.0464 | -49.2188 | -50.386 | 0.059509 | 70.1431 | -51.9012 | 70.2438 | -46.8201 | -48.7289 | -49.8962 | -0.16937 | 0.009155 | -0.068665 | -0.07782 | 5.7999 | 5.9692 | 1.4603 | 2.6276 | 1.6296 | 2.7969 | -0.054932 | -0.009156 | 0.091552 | 0.10071 | 5.0262 | 5.0812 | 1.9501 | 3.1174 | 2.005 | 3.1723
1920.1719 | 122.3969 | -51.1414 | 70.871 | -0.16869 | 0.019072 | 0.087771 | 0.092418 | 0.0010403 | 0.018752 | 0.010724 | -0.077606 | 0.023031 | 0.079233 | 0.088186 | -0.0014633 | 0.014255 | 0.0098417 | -0.045776 | 70.9076 | -51.2466 | 70.8664 | -45.3552 | -48.4772 | -49.649 | -0.004578 | 70.8801 | -51.1414 | 71.0266 | -46.2479 | -48.2346 | -49.2371 | -0.10529 | 0.036621 | -0.004578 | -0.041199 | 5.7861 | 5.8914 | 1.4923 | 2.6642 | 1.5976 | 2.7695 | 0 | 0.009155 | 0.15564 | 0.14648 | 4.8935 | 4.8935 | 1.9043 | 2.9068 | 1.9043 | 2.9068
1921.1719 | 122.3236 | -51.297 | 70.6924 | -0.15343 | 0.019072 | 0.086276 | 0.092418 | -0.00024129 | 0.015761 | 0.010083 | -0.10812 | 0.024526 | 0.082223 | 0.092458 | -0.00060892 | 0.013614 | 0.011978 | -0.054932 | 70.7108 | -51.4801 | 70.6284 | -45.5795 | -48.6786 | -49.8917 | -0.004578 | 70.6924 | -51.3519 | 70.7932 | -46.3165 | -48.2208 | -49.3423 | -0.18311 | 0.01831 | -0.064087 | -0.082397 | 5.7175 | 5.9006 | 1.4053 | 2.6184 | 1.5884 | 2.8015 | -0.054932 | 0 | 0.10071 | 0.10071 | 4.9805 | 5.0354 | 1.9547 | 3.0762 | 2.0096 | 3.1311
1922.1719 | 122.319 | -51.2604 | 70.7108 | -0.19157 | 0.02014 | 0.08649 | 0.090495 | 0.0016813 | 0.01982 | 0.013287 | -0.20273 | 0.026876 | 0.079874 | 0.09139 | -0.0014633 | 0.016178 | 0.0098417 | -0.045776 | 70.7108 | -51.4252 | 70.665 | -45.5521 | -48.674 | -49.8459 | 0.045776 | 70.6604 | -51.3702 | 70.8206 | -46.2891 | -48.1979 | -49.3378 | -0.16479 | 0 | -0.045776 | -0.045776 | 5.7083 | 5.8731 | 1.4145 | 2.5864 | 1.5793 | 2.7512 | -0.10986 | -0.050354 | 0.10986 | 0.16022 | 4.9713 | 5.0812 | 1.9226 | 3.0624 | 2.0325 | 3.1723
1923.1875 | 122.2549 | -51.3153 | 70.7474 | -0.13817 | 0.02014 | 0.08414 | 0.090281 | 0.00061311 | 0.017043 | 0.0070922 | -0.03183 | 0.023671 | 0.081155 | 0.094167 | 3.2078e-05 | 0.015323 | 0.0096279 | 0.041199 | 70.7474 | -51.4297 | 70.6421 | -45.5795 | -48.6649 | -49.855 | 0.091553 | 70.7153 | -51.3474 | 70.7245 | -46.3303 | -48.262 | -49.4202 | -0.11444 | 0 | -0.10529 | -0.10529 | 5.7358 | 5.8502 | 1.4603 | 2.6505 | 1.5747 | 2.7649 | -0.032043 | -0.032043 | -0.022888 | 0.009155 | 4.985 | 5.0171 | 1.8951 | 3.0533 | 1.9272 | 3.0853
1924.1719 | 122.3511 | -51.4572 | 70.5597 | -0.077133 | 0.020568 | 0.083713 | 0.090495 | -0.00024129 | 0.016829 | 0.0090148 | -0.019623 | 0.022603 | 0.084146 | 0.086477 | -0.00082252 | 0.015537 | 0.0085599 | -0.050354 | 70.5322 | -51.622 | 70.4269 | -45.8267 | -48.88 | -50.0473 | 0.009155 | 70.5368 | -51.5396 | 70.6009 | -46.4539 | -48.3673 | -49.5438 | -0.16479 | -0.027465 | -0.13275 | -0.10529 | 5.6305 | 5.7953 | 1.4099 | 2.5772 | 1.5747 | 2.742 | -0.082398 | -0.022888 | 0.041199 | 0.064087 | 5.0034 | 5.0858 | 1.9135 | 3.0899 | 1.9958 | 3.1723
1925.1719 | 122.3694 | -51.5442 | 70.4361 | -0.1992 | 0.02249 | 0.085849 | 0.09199 | -2.7692e-05 | 0.014693 | 0.0096558 | -0.16306 | 0.025167 | 0.080728 | 0.092031 | -0.00082252 | 0.014469 | 0.011123 | -0.054932 | 70.4178 | -51.7044 | 70.3903 | -45.9045 | -48.9395 | -50.116 | -0.027466 | 70.3949 | -51.6312 | 70.5551 | -46.5637 | -48.4406 | -49.617 | -0.16022 | -0.01831 | -0.045776 | -0.027466 | 5.6396 | 5.7999 | 1.4282 | 2.6047 | 1.5884 | 2.7649 | -0.086976 | -0.041199 | 0.11902 | 0.16022 | 4.9805 | 5.0674 | 1.9272 | 3.1036 | 2.0142 | 3.1906
1926.1719 | 122.4106 | -51.4389 | 70.5734 | -0.17632 | 0.020995 | 0.086276 | 0.093699 | 0.0023221 | 0.018538 | 0.011151 | -0.068451 | 0.023885 | 0.079874 | 0.08904 | -0.0029587 | 0.014041 | 0.011123 | -0.041199 | 70.578 | -51.5808 | 70.5597 | -45.7855 | -48.7976 | -49.9649 | -0.004578 | 70.5826 | -51.4435 | 70.7199 | -46.4676 | -48.5962 | -49.5941 | -0.14191 | 0.004578 | -0.013733 | -0.018311 | 5.6534 | 5.7953 | 1.474 | 2.6413 | 1.6159 | 2.7832 | -0.004577 | 0.009156 | 0.14648 | 0.13733 | 4.9713 | 4.9759 | 1.8448 | 2.8427 | 1.8494 | 2.8473
1927.1719 | 122.3099 | -51.1322 | 70.9122 | -0.16869 | 0.020568 | 0.084353 | 0.092204 | 0.00018591 | 0.019179 | 0.010937 | -0.13864 | 0.02538 | 0.083505 | 0.090535 | -0.0016769 | 0.018314 | 0.012832 | -0.086975 | 70.9259 | -51.2512 | 70.7016 | -45.6619 | -48.6328 | -49.8001 | -0.10529 | 70.8618 | -51.2192 | 70.9259 | -46.1929 | -48.1201 | -49.2325 | -0.11902 | 0.013733 | -0.21057 | -0.2243 | 5.4703 | 5.5893 | 1.3321 | 2.4994 | 1.4511 | 2.6184 | -0.086975 | -0.050354 | 0.013733 | 0.064087 | 4.9393 | 5.0262 | 1.8997 | 3.0121 | 1.9867 | 3.0991
1928.1719 | 122.2504 | -51.6037 | 70.4498 | -0.092392 | 0.0195 | 0.083926 | 0.094554 | -0.0025913 | 0.018325 | 0.011365 | -0.13254 | 0.025167 | 0.083932 | 0.09139 | -0.00082252 | 0.014682 | 0.0081327 | -0.13733 | 70.4727 | -51.741 | 70.2805 | -46.1014 | -49.0585 | -50.2258 | -0.1236 | 70.4544 | -51.6586 | 70.459 | -46.6187 | -48.5916 | -49.7177 | -0.13733 | 0.022888 | -0.16937 | -0.19226 | 5.5023 | 5.6396 | 1.3779 | 2.5452 | 1.5152 | 2.6825 | -0.054931 | 0.004578 | 0.009155 | 0.004577 | 4.985 | 5.04 | 1.886 | 3.0121 | 1.9409 | 3.067
1929.1719 | 122.287 | -51.7548 | 70.2484 | -0.10765 | 0.023772 | 0.08649 | 0.090922 | -0.0036593 | 0.017897 | 0.0081604 | -0.0043641 | 0.022603 | 0.08436 | 0.086477 | 0.00088648 | 0.014682 | 0.0092007 | -0.045776 | 70.253 | -51.8829 | 70.1477 | -46.2021 | -49.173 | -50.322 | -0.050354 | 70.2393 | -51.828 | 70.3033 | -46.7514 | -48.674 | -49.8505 | -0.12817 | 0.004578 | -0.10071 | -0.10529 | 5.5527 | 5.6808 | 1.4328 | 2.5818 | 1.561 | 2.71 | -0.073242 | -0.009155 | 0.054932 | 0.064087 | 5.0034 | 5.0766 | 1.9043 | 3.0808 | 1.9775 | 3.154
1930.1719 | 122.2641 | -51.5442 | 70.4865 | -0.084763 | 0.020781 | 0.084781 | 0.090922 | 0.00082671 | 0.017256 | 0.011792 | -0.098968 | 0.023031 | 0.082864 | 0.093313 | 0.00024568 | 0.016818 | 0.011123 | -0.09613 | 70.5185 | -51.6953 | 70.3445 | -46.0419 | -49.0036 | -50.1617 | -0.10529 | 70.5185 | -51.5945 | 70.5231 | -46.5591 | -48.5275 | -49.6582 | -0.15106 | 0.032044 | -0.14191 | -0.17395 | 5.5023 | 5.6534 | 1.3824 | 2.5406 | 1.5335 | 2.6916 | -0.050354 | 0.032044 | 0.036621 | 0.004577 | 4.985 | 5.0354 | 1.886 | 3.0167 | 1.9363 | 3.067
1931.1719 | 122.2733 | -51.8326 | 70.1752 | -0.1458 | 0.018859 | 0.084994 | 0.092418 | -2.7692e-05 | 0.017043 | 0.012006 | -0.11728 | 0.023885 | 0.082223 | 0.085195 | -0.0010361 | 0.016178 | 0.014114 | -0.12817 | 70.1935 | -52.0248 | 70.015 | -46.3715 | -49.3332 | -50.5051 | -0.11444 | 70.166 | -51.8967 | 70.1981 | -46.8842 | -48.8525 | -49.9741 | -0.19226 | 0.01831 | -0.16022 | -0.17853 | 5.4611 | 5.6534 | 1.3275 | 2.4994 | 1.5198 | 2.6917 | -0.064086 | -0.009155 | 0.022888 | 0.032043 | 4.9484 | 5.0125 | 1.8585 | 2.98 | 1.9226 | 3.0441
1932.1719 | 122.3328 | -51.3428 | 70.6467 | -0.17632 | 0.020781 | 0.087344 | 0.091136 | 0.0018949 | 0.020034 | 0.010297 | -0.065399 | 0.024312 | 0.083719 | 0.086263 | 3.2078e-05 | 0.014041 | 0.0096279 | -0.050354 | 70.7062 | -51.4709 | 70.6512 | -45.7581 | -48.6969 | -49.8596 | -0.009155 | 70.6879 | -51.3519 | 70.8069 | -46.2616 | -48.4817 | -49.5346 | -0.12817 | 0.05951 | 0.004578 | -0.054932 | 5.5847 | 5.7129 | 1.4832 | 2.6459 | 1.6113 | 2.774 | -0.009156 | 0.041199 | 0.16022 | 0.11902 | 5.0812 | 5.0903 | 1.8082 | 2.861 | 1.8173 | 2.8702
1933.1719 | 122.2595 | -51.5991 | 70.4407 | -0.16106 | 0.019072 | 0.084781 | 0.095622 | 0.00082671 | 0.016188 | 0.01286 | -0.098968 | 0.025167 | 0.082864 | 0.088399 | 0.00067288 | 0.015323 | 0.01091 | -0.10071 | 70.4727 | -51.7685 | 70.2576 | -46.1426 | -49.0723 | -50.2304 | -0.10071 | 70.4224 | -51.6403 | 70.4773 | -46.6141 | -48.5733 | -49.704 | -0.16937 | 0.032043 | -0.18311 | -0.21515 | 5.4565 | 5.6259 | 1.3687 | 2.5269 | 1.5381 | 2.6962 | -0.041199 | -0.018311 | 0.036621 | 0.054932 | 4.985 | 5.0262 | 1.8951 | 3.0258 | 1.9363 | 3.067
1934.1719 | 122.168 | -51.5259 | 70.5505 | -0.16106 | 0.02014 | 0.084781 | 0.094767 | -0.0021641 | 0.021102 | 0.011151 | -0.12033 | 0.024953 | 0.080301 | 0.085836 | 0.0023819 | 0.018527 | 0.012405 | -0.09613 | 70.5276 | -51.709 | 70.3033 | -46.1655 | -49.0311 | -50.2121 | -0.1236 | 70.5368 | -51.59 | 70.3949 | -46.3165 | -48.468 | -49.6582 | -0.1831 | -0.022888 | -0.24719 | -0.2243 | 5.3604 | 5.5435 | 1.3138 | 2.4948 | 1.4969 | 2.6779 | -0.064087 | -0.013733 | -0.15564 | -0.14191 | 5.2094 | 5.2734 | 1.8677 | 3.0579 | 1.9318 | 3.1219
1935.1719 | 122.2183 | -51.4481 | 70.6192 | -0.12291 | 0.024199 | 0.086917 | 0.093486 | 0.0021085 | 0.018752 | 0.011365 | -0.11118 | 0.024312 | 0.081155 | 0.092244 | 0.00067288 | 0.014041 | 0.010269 | -0.073242 | 70.6375 | -51.5991 | 70.372 | -46.0968 | -48.9532 | -50.1297 | -0.15564 | 70.6284 | -51.4755 | 70.5872 | -46.4767 | -48.4955 | -49.6033 | -0.15106 | 0.01831 | -0.24719 | -0.2655 | 5.3513 | 5.5023 | 1.3184 | 2.4948 | 1.4694 | 2.6459 | -0.027466 | 0.009155 | -0.032044 | -0.041199 | 4.9713 | 4.9988 | 1.8448 | 2.9526 | 1.8723 | 2.98
1936.1719 | 122.3007 | -53.4348 | 68.5181 | -0.1458 | 0.021422 | 0.084994 | 0.093058 | -2.7692e-05 | 0.01512 | 0.01051 | 0.013947 | 0.026449 | 0.082864 | 0.093526 | 3.2078e-05 | 0.017032 | 0.012405 | -0.27924 | 69.5526 | -52.6978 | 68.8385 | -47.6349 | -50.6744 | -51.5305 | -0.59052 | 68.9667 | -53.1555 | 68.6279 | -48.4909 | -50.2579 | -51.3931 | 0.737 | 1.0345 | 0.32043 | -0.71411 | 5.7999 | 5.0629 | 1.9043 | 2.7603 | 1.1673 | 2.0233 | 0.27923 | 0.44861 | 0.10986 | -0.33874 | 4.9438 | 4.6646 | 2.0416 | 3.1769 | 1.7624 | 2.8976
1937.1719 | 122.1634 | -51.796 | 70.2896 | -0.19157 | 0.021209 | 0.086276 | 0.092845 | 0.0010403 | 0.018325 | 0.0096558 | -0.092865 | 0.025167 | 0.083505 | 0.097799 | 0.0011001 | 0.01575 | 0.011978 | -0.10529 | 70.2576 | -51.9699 | 70.0241 | -46.4035 | -49.2874 | -50.4639 | -0.14648 | 70.253 | -51.8509 | 70.1477 | -46.6507 | -48.6282 | -49.8596 | -0.17395 | -0.032044 | -0.2655 | -0.23346 | 5.3925 | 5.5664 | 1.3321 | 2.5085 | 1.506 | 2.6825 | -0.054932 | -0.036621 | -0.14191 | -0.10529 | 5.1453 | 5.2002 | 1.9363 | 3.1677 | 1.9913 | 3.2227
1938.1719 | 122.2321 | -51.5945 | 70.4681 | -0.10002 | 0.021209 | 0.083713 | 0.092845 | 0.0012539 | 0.018325 | 0.0098694 | -0.10812 | 0.024526 | 0.085214 | 0.087118 | 0.00045928 | 0.016605 | 0.014114 | -0.11444 | 70.491 | -51.7227 | 70.2576 | -46.2067 | -49.0723 | -50.2579 | -0.16022 | 70.5002 | -51.6357 | 70.4224 | -46.6324 | -48.6511 | -49.7223 | -0.12817 | 0.022888 | -0.21057 | -0.23346 | 5.3879 | 5.5161 | 1.3367 | 2.5223 | 1.4648 | 2.6505 | -0.041199 | 0.032043 | -0.045777 | -0.07782 | 4.9622 | 5.0034 | 1.8723 | 2.9434 | 1.9135 | 2.9846
1939.1719 | 122.287 | -51.6357 | 70.3308 | -0.12291 | 0.02014 | 0.085635 | 0.087931 | 0.0010403 | 0.018538 | 0.010083 | 0.026154 | 0.021535 | 0.08436 | 0.084768 | 0.0013139 | 0.011905 | 0.011123 | -0.064087 | 70.4681 | -51.6724 | 70.3125 | -46.1197 | -49.054 | -50.2304 | -0.1236 | 70.34 | -51.7136 | 70.3766 | -46.6278 | -48.6008 | -49.7635 | -0.036621 | 0.13733 | -0.018311 | -0.15564 | 5.5161 | 5.5527 | 1.4053 | 2.5818 | 1.442 | 2.6184 | -0.07782 | 0.009155 | 0.045776 | 0.036621 | 5.0079 | 5.0858 | 1.8723 | 3.035 | 1.9501 | 3.1128
1940.1719 | 122.2733 | -51.7364 | 70.3445 | -0.13054 | 0.019713 | 0.084567 | 0.090922 | 0.00039951 | 0.018752 | 0.01051 | -0.11728 | 0.024526 | 0.08201 | 0.09908 | 0.00024568 | 0.015323 | 0.012619 | -0.10986 | 70.3308 | -51.9104 | 70.079 | -46.4035 | -49.2462 | -50.4135 | -0.14191 | 70.3033 | -51.8051 | 70.3354 | -46.7102 | -48.8937 | -49.9374 | -0.17395 | -0.013732 | -0.2655 | -0.25177 | 5.3329 | 5.5069 | 1.3229 | 2.4902 | 1.4969 | 2.6642 | -0.068665 | -0.041198 | -0.009155 | 0.032043 | 5.0262 | 5.0949 | 1.799 | 2.8427 | 1.8677 | 2.9114
1941.1719 | 122.2092 | -51.6129 | 70.4544 | -0.12291 | 0.020354 | 0.084567 | 0.091777 | 0.0010403 | 0.015975 | 0.0096558 | -0.095917 | 0.025167 | 0.081369 | 0.090749 | -0.00018152 | 0.014469 | 0.015182 | -0.054932 | 70.4544 | -51.7593 | 70.2713 | -46.2296 | -49.0723 | -50.2716 | -0.073242 | 70.4498 | -51.6724 | 70.3949 | -46.6644 | -48.6465 | -49.7772 | -0.14648 | 0 | -0.18311 | -0.18311 | 5.3833 | 5.5298 | 1.3412 | 2.5406 | 1.4877 | 2.6871 | -0.059509 | -0.004578 | -0.05951 | -0.054932 | 4.9484 | 5.0079 | 1.8356 | 2.9663 | 1.8951 | 3.0258
1942.1719 | 122.2366 | -51.6953 | 70.3674 | -0.17632 | 0.021209 | 0.085422 | 0.092631 | 0.0027493 | 0.018752 | 0.0098694 | -0.08371 | 0.023885 | 0.081796 | 0.092672 | -0.00082252 | 0.012973 | 0.012832 | -0.091553 | 70.3217 | -51.8692 | 70.2072 | -46.2845 | -49.1409 | -50.3311 | -0.054932 | 70.3674 | -51.7502 | 70.3445 | -46.7011 | -48.7335 | -49.8596 | -0.17395 | -0.045777 | -0.16022 | -0.11444 | 5.4108 | 5.5847 | 1.3641 | 2.5543 | 1.5381 | 2.7283 | -0.054932 | 0 | -0.022889 | -0.022889 | 4.9942 | 5.0491 | 1.8356 | 2.9617 | 1.8906 | 3.0167
1943.1719 | 122.2046 | -51.5671 | 70.5368 | -0.17632 | 0.021849 | 0.085635 | 0.09199 | 0.00018591 | 0.015975 | 0.010724 | -0.056244 | 0.026235 | 0.081796 | 0.09139 | -0.00018152 | 0.016178 | 0.01561 | -0.045776 | 70.5505 | -51.6678 | 70.3171 | -46.2158 | -49.0402 | -50.1984 | -0.15564 | 70.5368 | -51.59 | 70.4361 | -46.6415 | -48.5916 | -49.6902 | -0.10071 | 0.013733 | -0.21973 | -0.23346 | 5.3513 | 5.452 | 1.3687 | 2.5269 | 1.4694 | 2.6276 | -0.022888 | 0 | -0.10071 | -0.10071 | 4.9255 | 4.9484 | 1.8768 | 2.9755 | 1.8997 | 2.9984
1944.1719 | 122.3373 | -51.7044 | 70.253 | -0.15343 | 0.023558 | 0.088412 | 0.090281 | 0.00082671 | 0.015975 | 0.010724 | -0.11118 | 0.022817 | 0.081369 | 0.081777 | -0.0010361 | 0.013401 | 0.010483 | 0.009155 | 70.2988 | -51.8005 | 70.2118 | -46.2662 | -49.1409 | -50.3494 | 0.09613 | 70.2255 | -51.7868 | 70.3491 | -46.6461 | -48.6694 | -49.8047 | -0.09613 | 0.045776 | -0.041199 | -0.086975 | 5.4382 | 5.5344 | 1.355 | 2.5635 | 1.4511 | 2.6596 | -0.082397 | -0.027466 | 0.09613 | 0.1236 | 5.0583 | 5.1407 | 1.8997 | 3.035 | 1.9821 | 3.1174
1945.1719 | 122.2458 | -51.6037 | 70.4681 | -0.11528 | 0.019286 | 0.084994 | 0.093913 | 0.0016813 | 0.017043 | 0.011151 | -0.095917 | 0.02239 | 0.083932 | 0.092244 | -0.00018152 | 0.017032 | 0.01091 | -0.036621 | 70.4544 | -51.741 | 70.2805 | -46.2616 | -49.0997 | -50.2762 | -0.07782 | 70.4132 | -51.6678 | 70.4132 | -46.637 | -48.642 | -49.7498 | -0.13733 | -0.013733 | -0.18768 | -0.17395 | 5.3421 | 5.4794 | 1.3275 | 2.504 | 1.4648 | 2.6413 | -0.064087 | -0.054932 | -0.054932 | 0 | 4.9667 | 5.0308 | 1.8539 | 2.9617 | 1.918 | 3.0258
1946.1719 | 122.2092 | -51.6861 | 70.3629 | -0.12291 | 0.022917 | 0.087344 | 0.092631 | -0.0028049 | 0.019393 | 0.012006 | -0.080658 | 0.023244 | 0.083078 | 0.092672 | 0.00045928 | 0.015323 | 0.014969 | -0.09613 | 70.34 | -51.8829 | 70.1385 | -46.4081 | -49.2371 | -50.4089 | -0.15564 | 70.3308 | -51.7868 | 70.2667 | -46.7743 | -48.7839 | -49.8734 | -0.19684 | -0.022888 | -0.2243 | -0.20142 | 5.278 | 5.4749 | 1.2772 | 2.449 | 1.474 | 2.6459 | -0.10071 | -0.032043 | -0.09613 | -0.064087 | 4.9118 | 5.0125 | 1.8127 | 2.9022 | 1.9135 | 3.0029
1947.1719 | 121.9345 | -51.5167 | 70.6284 | -0.15343 | 0.020781 | 0.086062 | 0.091349 | -0.0010959 | 0.018325 | 0.0098694 | -0.071503 | 0.024312 | 0.083078 | 0.090322 | 0.0013139 | 0.014896 | 0.013046 | -0.10986 | 70.5872 | -51.6541 | 70.34 | -46.2158 | -48.9899 | -50.1434 | -0.12817 | 70.61 | -51.5121 | 70.5688 | -46.5088 | -48.5321 | -49.6124 | -0.13733 | -0.041199 | -0.28839 | -0.24719 | 5.3009 | 5.4382 | 1.3733 | 2.5269 | 1.5106 | 2.6642 | 0.004578 | -0.018311 | -0.059509 | -0.041198 | 5.0079 | 5.0034 | 1.9043 | 2.9846 | 1.8997 | 2.98
1948.1719 | 122.1954 | -51.7319 | 70.3812 | -0.092392 | 0.021849 | 0.087558 | 0.091777 | 0.00082671 | 0.01747 | 0.0094422 | -0.074554 | 0.025808 | 0.081796 | 0.089467 | 0.00024568 | 0.015964 | 0.010696 | -0.022888 | 70.4178 | -51.796 | 70.1569 | -46.3852 | -49.1913 | -50.4684 | -1.3046 | 69.4702 | -52.7481 | 70.3125 | -46.7331 | -48.7885 | -49.8367 | -0.064086 | 0.036621 | -0.22431 | -0.26093 | 5.3467 | 5.4108 | 1.2634 | 2.5406 | 1.3275 | 2.6047 | -1.0162 | -0.91095 | -0.068665 | 0.84229 | 4.9988 | 6.015 | 1.8951 | 2.9434 | 2.9114 | 3.9597
1949.1719 | 122.1909 | -51.8692 | 70.2209 | -0.1458 | 0.022063 | 0.087771 | 0.09199 | -0.0028049 | 0.017256 | 0.0094422 | -0.065399 | 0.023671 | 0.083719 | 0.094167 | 0.0011001 | 0.014682 | 0.013687 | -0.082397 | 70.1843 | -52.034 | 70.0012 | -46.5591 | -49.3378 | -50.5142 | -0.064087 | 70.2026 | -51.9241 | 70.1523 | -46.875 | -48.9212 | -50.0244 | -0.16479 | -0.036621 | -0.21973 | -0.1831 | 5.3101 | 5.4749 | 1.355 | 2.5314 | 1.5198 | 2.6962 | -0.054931 | -0.01831 | -0.068664 | -0.050354 | 4.9942 | 5.0491 | 1.8448 | 2.948 | 1.8997 | 3.0029
1950.1719 | 122.2046 | -51.7822 | 70.285 | -0.12291 | 0.0195 | 0.089053 | 0.092631 | 0.0025357 | 0.016188 | 0.011578 | -0.062347 | 0.027517 | 0.081369 | 0.092458 | -0.0010361 | 0.015323 | 0.013901 | -0.027466 | 70.3125 | -51.8921 | 70.1019 | -46.4722 | -49.2233 | -50.4227 | -0.059509 | 70.2667 | -51.8188 | 70.2255 | -46.7972 | -48.8159 | -49.9283 | -0.10986 | 0.027466 | -0.1831 | -0.21057 | 5.3101 | 5.4199 | 1.3596 | 2.5589 | 1.4694 | 2.6688 | -0.036621 | -0.01831 | -0.059509 | -0.041199 | 4.985 | 5.0217 | 1.8539 | 2.9663 | 1.8906 | 3.0029
1951.1719 | 122.2137 | -51.8372 | 70.253 | -0.092392 | 0.0195 | 0.087771 | 0.092204 | 0.0023221 | 0.016829 | 0.0085876 | -0.077606 | 0.025594 | 0.081369 | 0.094381 | 0.0011001 | 0.014682 | 0.010269 | -0.07782 | 70.285 | -51.915 | 70.0745 | -46.5134 | -49.2371 | -50.4272 | -0.068665 | 70.2759 | -51.828 | 70.1752 | -46.8201 | -48.848 | -49.9603 | -0.07782 | 0.032043 | -0.17853 | -0.21057 | 5.3238 | 5.4016 | 1.4099 | 2.6001 | 1.4877 | 2.6779 | 0.009155 | 0.022888 | -0.07782 | -0.10071 | 5.0171 | 5.0079 | 1.8768 | 2.9892 | 1.8677 | 2.98
1952.1719 | 122.2 | -51.7181 | 70.3903 | -0.13054 | 0.020568 | 0.085208 | 0.092631 | -2.7692e-05 | 0.015975 | 0.01051 | -0.059296 | 0.023885 | 0.082864 | 0.088399 | 0.0017411 | 0.015323 | 0.012832 | -0.013733 | 70.4041 | -51.7731 | 70.1706 | -46.4172 | -49.1318 | -50.3357 | -0.022888 | 70.3583 | -51.7456 | 70.3125 | -46.7331 | -48.7564 | -49.8688 | -0.054931 | 0.013733 | -0.21973 | -0.23346 | 5.3009 | 5.3558 | 1.3824 | 2.5864 | 1.4374 | 2.6413 | -0.027465 | -0.032044 | -0.07782 | -0.045776 | 4.985 | 5.0125 | 1.8494 | 2.9617 | 1.8768 | 2.9892
1953.1719 | 122.2183 | -51.6495 | 70.4727 | -0.077133 | 0.021209 | 0.085849 | 0.091777 | 0.00039951 | 0.018752 | 0.01051 | -0.053192 | 0.025167 | 0.082223 | 0.088613 | 0.00088648 | 0.015323 | 0.011764 | -0.073242 | 70.4865 | -51.7227 | 70.2209 | -46.4401 | -49.0952 | -50.2487 | -0.11902 | 70.4865 | -51.6678 | 70.3812 | -46.6278 | -48.6832 | -49.791 | -0.073242 | 0.013733 | -0.25177 | -0.2655 | 5.2094 | 5.2826 | 1.4008 | 2.5543 | 1.474 | 2.6276 | -0.018311 | 0.013733 | -0.091552 | -0.10529 | 5.0217 | 5.04 | 1.8585 | 2.9663 | 1.8768 | 2.9846
1954.1719 | 122.3053 | -51.7868 | 70.1843 | -0.16869 | 0.020995 | 0.086917 | 0.08964 | 0.00039951 | 0.018752 | 0.0090148 | 0.029205 | 0.023671 | 0.081796 | 0.081563 | 0.00067288 | 0.01511 | 0.010055 | -0.07782 | 70.1706 | -51.9745 | 70.0378 | -46.5546 | -49.2416 | -50.3906 | 0.009155 | 70.134 | -51.8921 | 70.2164 | -46.7697 | -48.8022 | -49.9146 | -0.18768 | -0.013733 | -0.14648 | -0.13275 | 5.2322 | 5.4199 | 1.3962 | 2.5452 | 1.5839 | 2.7328 | -0.10529 | -0.050354 | 0.032044 | 0.082398 | 5.0171 | 5.1224 | 1.8723 | 2.9846 | 1.9775 | 3.0899
1955.1719 | 122.2595 | -51.7181 | 70.34 | -0.16869 | 0.024413 | 0.085422 | 0.092845 | -0.0019503 | 0.017256 | 0.012219 | -0.062347 | 0.024312 | 0.080515 | 0.091604 | 3.2078e-05 | 0.014896 | 0.013687 | -0.013733 | 70.4178 | -51.8188 | 70.134 | -46.5225 | -49.205 | -50.3815 | -0.004578 | 70.3125 | -51.796 | 70.3217 | -46.6965 | -48.7198 | -49.8276 | -0.10071 | 0.07782 | -0.20599 | -0.28381 | 5.1956 | 5.2963 | 1.3367 | 2.5131 | 1.4374 | 2.6138 | -0.077819 | -0.027466 | -0.018311 | 0.009155 | 5.0217 | 5.0995 | 1.8906 | 2.9984 | 1.9684 | 3.0762
1956.1719 | 122.2504 | -51.8967 | 70.1752 | -0.10765 | 0.019713 | 0.086276 | 0.091777 | -2.7692e-05 | 0.017256 | 0.010937 | -0.071503 | 0.023458 | 0.083719 | 0.087118 | -0.00060892 | 0.013828 | 0.0077053 | -0.009155 | 70.1843 | -52.0157 | 70.0745 | -46.5866 | -49.3057 | -50.5051 | 0.027466 | 70.1797 | -51.8967 | 70.1797 | -46.8246 | -48.8892 | -50.0107 | -0.11902 | 0.009155 | -0.10071 | -0.10986 | 5.3101 | 5.4291 | 1.3916 | 2.5909 | 1.5106 | 2.71 | 0 | 0.004578 | 0.004578 | 0 | 5.072 | 5.072 | 1.886 | 3.0075 | 1.886 | 3.0075
1957.1719 | 122.2412 | -52.0889 | 69.9783 | -0.13817 | 0.020568 | 0.085208 | 0.090709 | -0.0028049 | 0.016616 | 0.012006 | -0.040985 | 0.025594 | 0.085214 | 0.088399 | 3.2078e-05 | 0.015537 | 0.012832 | -0.032043 | 69.96 | -52.2491 | 69.7861 | -46.8567 | -49.5712 | -50.7523 | -0.073242 | 69.9554 | -52.1576 | 69.928 | -47.0673 | -49.1592 | -50.2167 | -0.16022 | -0.018311 | -0.19226 | -0.17395 | 5.2322 | 5.3925 | 1.3367 | 2.5177 | 1.4969 | 2.6779 | -0.068665 | -0.022889 | -0.050354 | -0.027465 | 5.0217 | 5.0903 | 1.8723 | 2.9297 | 1.9409 | 2.9984
1958.1719 | 122.2229 | -52.2903 | 69.7678 | -0.1458 | 0.020354 | 0.084567 | 0.091777 | 0.00082671 | 0.016616 | 0.011151 | -0.03183 | 0.023458 | 0.08201 | 0.092244 | -0.0010361 | 0.014255 | 0.0094143 | -0.041199 | 69.7678 | -52.4368 | 69.5801 | -47.0718 | -49.7818 | -50.9491 | -0.059509 | 69.7495 | -52.3544 | 69.722 | -47.2778 | -49.3607 | -50.4501 | -0.14649 | 0 | -0.18768 | -0.18768 | 5.2185 | 5.365 | 1.3412 | 2.5085 | 1.4877 | 2.655 | -0.064087 | -0.01831 | -0.045776 | -0.027466 | 5.0125 | 5.0766 | 1.8402 | 2.9297 | 1.9043 | 2.9938
1959.1719 | 122.3648 | -52.2125 | 69.7403 | -0.1992 | 0.019927 | 0.084353 | 0.089213 | 0.0025357 | 0.016616 | 0.0098694 | 0.0047913 | 0.02538 | 0.085428 | 0.086049 | -0.0010361 | 0.014469 | 0.010483 | -0.045776 | 69.7495 | -52.3911 | 69.6991 | -46.9666 | -49.6674 | -50.8347 | 0.009155 | 69.7632 | -52.2903 | 69.8364 | -47.1359 | -49.2142 | -50.2945 | -0.17853 | 0.009156 | -0.041198 | -0.050354 | 5.246 | 5.4245 | 1.3779 | 2.5452 | 1.5564 | 2.7237 | -0.07782 | 0.022889 | 0.096131 | 0.073242 | 5.0766 | 5.1544 | 1.918 | 2.9984 | 1.9958 | 3.0762
1960.1719 | 122.2183 | -52.1805 | 69.8914 | -0.11528 | 0.019286 | 0.086703 | 0.092204 | 0.0027493 | 0.015975 | 0.012006 | -0.062347 | 0.023244 | 0.081155 | 0.08904 | 3.2078e-05 | 0.01511 | 0.012405 | -0.10986 | 69.8639 | -52.3636 | 69.6533 | -47.0352 | -49.7223 | -50.8896 | -0.14648 | 69.8685 | -52.2537 | 69.7815 | -47.2137 | -49.3057 | -50.3448 | -0.1831 | -0.027465 | -0.23804 | -0.21057 | 5.1453 | 5.3284 | 1.2909 | 2.4582 | 1.474 | 2.6413 | -0.073242 | -0.022888 | -0.10986 | -0.086975 | 4.9667 | 5.04 | 1.8356 | 2.8748 | 1.9089 | 2.948
1961.1719 | 122.2229 | -52.2995 | 69.7861 | -0.10765 | 0.023131 | 0.085635 | 0.091563 | -0.0013095 | 0.018752 | 0.013074 | -0.080658 | 0.022817 | 0.083719 | 0.097799 | -0.0012497 | 0.012332 | 0.014114 | -0.073242 | 69.7632 | -52.4506 | 69.5572 | -47.113 | -49.8093 | -50.9811 | -0.15106 | 69.7586 | -52.3407 | 69.6625 | -47.3099 | -49.4247 | -50.473 | -0.15106 | -0.022888 | -0.22888 | -0.20599 | 5.1865 | 5.3375 | 1.3184 | 2.4902 | 1.4694 | 2.6413 | -0.041198 | -0.027466 | -0.1236 | -0.09613 | 4.9896 | 5.0308 | 1.8265 | 2.8748 | 1.8677 | 2.916
1962.1719 | 122.2366 | -52.1667 | 69.8914 | -0.10765 | 0.020995 | 0.085849 | 0.092204 | 0.00039951 | 0.015761 | 0.010083 | -0.053192 | 0.024312 | 0.081583 | 0.087972 | 0.00024568 | 0.017032 | 0.012405 | -0.041199 | 69.8959 | -52.2995 | 69.6945 | -46.9711 | -49.6674 | -50.8301 | -0.036621 | 69.8776 | -52.2308 | 69.841 | -47.1268 | -49.2508 | -50.3128 | -0.13275 | 0.004578 | -0.19684 | -0.20142 | 5.1956 | 5.3284 | 1.3367 | 2.4994 | 1.4694 | 2.6321 | -0.064087 | -0.013732 | -0.050354 | -0.036622 | 5.04 | 5.1041 | 1.8539 | 2.916 | 1.918 | 2.98
1963.1719 | 122.2733 | -52.3911 | 69.6487 | -0.10002 | 0.019072 | 0.085208 | 0.091349 | 0.0025357 | 0.017256 | 0.008374 | -0.047089 | 0.024953 | 0.081796 | 0.093526 | 3.2078e-05 | 0.014041 | 0.01326 | -0.032043 | 69.6716 | -52.5284 | 69.429 | -47.2458 | -49.9054 | -51.0681 | -0.018311 | 69.6304 | -52.446 | 69.6579 | -47.3007 | -49.4522 | -50.4959 | -0.13733 | 0.022888 | -0.21973 | -0.24262 | 5.1453 | 5.2826 | 1.3229 | 2.4857 | 1.4603 | 2.623 | -0.054932 | -0.018311 | 0.009155 | 0.027466 | 5.0903 | 5.1453 | 1.8951 | 2.9388 | 1.9501 | 2.9938
1964.1719 | 122.2229 | -52.2858 | 69.8318 | -0.10002 | 0.02014 | 0.083285 | 0.092631 | 0.00039951 | 0.018325 | 0.007306 | -0.074554 | 0.023031 | 0.083719 | 0.090963 | 0.00024568 | 0.013187 | 0.011551 | -0.091553 | 69.7998 | -52.4231 | 69.5572 | -47.1588 | -49.8138 | -51.1505 | -0.009155 | 69.8135 | -52.2858 | 69.7449 | -47.2 | -49.3469 | -50.4089 | -0.13733 | -0.032043 | -0.27466 | -0.24262 | 5.127 | 5.2643 | 1.1353 | 2.4719 | 1.2726 | 2.6093 | 0 | -0.01831 | -0.086975 | -0.068665 | 5.0858 | 5.0858 | 1.8768 | 2.9388 | 1.8768 | 2.9388
1965.1719 | 122.2 | -52.066 | 69.9966 | -0.1458 | 0.020354 | 0.085849 | 0.091563 | 0.00082671 | 0.021529 | 0.0081604 | -0.074554 | 0.024953 | 0.082651 | 0.088186 | -0.00039512 | 0.016605 | 0.012619 | -0.054932 | 70.0378 | -52.2034 | 69.7723 | -46.9528 | -49.5895 | -50.7797 | -0.17853 | 69.9829 | -52.1484 | 69.9463 | -47.0306 | -49.173 | -50.2304 | -0.13733 | 0.041199 | -0.2243 | -0.2655 | 5.1132 | 5.2505 | 1.2863 | 2.4765 | 1.4236 | 2.6138 | -0.082397 | -0.013733 | -0.050354 | -0.036621 | 5.0354 | 5.1178 | 1.8356 | 2.8931 | 1.918 | 2.9755
1966.1719 | 122.2595 | -51.9516 | 70.1157 | -0.16869 | 0.020568 | 0.085635 | 0.091349 | -0.00045509 | 0.01512 | 0.0092286 | -0.040985 | 0.02474 | 0.079874 | 0.084127 | 0.0011001 | 0.012546 | 0.012619 | -0.068665 | 70.1385 | -52.0706 | 69.928 | -46.843 | -49.4614 | -50.6104 | -0.23346 | 70.0974 | -51.9745 | 70.0516 | -46.8338 | -49.0585 | -50.1068 | -0.11902 | 0.022888 | -0.18768 | -0.21057 | 5.1086 | 5.2277 | 1.3412 | 2.4902 | 1.4603 | 2.6093 | -0.022888 | -0.018311 | -0.064087 | -0.045776 | 5.1178 | 5.1407 | 1.8448 | 2.8931 | 1.8677 | 2.916
1967.1719 | 122.2229 | -52.2583 | 69.8364 | -0.16869 | 0.021636 | 0.085208 | 0.090709 | -0.0015231 | 0.018111 | 0.011365 | -0.071503 | 0.027089 | 0.082223 | 0.084981 | -0.00039512 | 0.015537 | 0.014328 | -0.07782 | 69.8181 | -52.3911 | 69.6259 | -47.1359 | -49.7498 | -50.9399 | -0.3479 | 69.8364 | -52.2491 | 69.7952 | -47.168 | -49.3286 | -50.3769 | -0.13275 | -0.018311 | -0.21057 | -0.19226 | 5.1224 | 5.2551 | 1.3184 | 2.5085 | 1.4511 | 2.6413 | 0.009155 | 0 | -0.041199 | -0.041199 | 5.0903 | 5.0812 | 1.8814 | 2.9297 | 1.8723 | 2.9205
1968.1719 | 122.2412 | -52.2583 | 69.8181 | -0.1458 | 0.021636 | 0.086703 | 0.089854 | 0.0014675 | 0.019179 | 0.010083 | -0.068451 | 0.024099 | 0.085001 | 0.087331 | -0.00039512 | 0.01511 | 0.0186 | -0.091553 | 69.8135 | -52.3956 | 69.6121 | -47.1725 | -49.791 | -50.9399 | -0.10529 | 69.7998 | -52.2903 | 69.722 | -47.1268 | -49.3469 | -50.3906 | -0.13733 | -0.004577 | -0.20599 | -0.20142 | 5.0858 | 5.2231 | 1.3184 | 2.4673 | 1.4557 | 2.6047 | -0.032043 | -0.01831 | -0.09613 | -0.07782 | 5.1315 | 5.1636 | 1.8677 | 2.9114 | 1.8997 | 2.9434
1969.1875 | 122.2183 | -52.0844 | 69.9829 | -0.13054 | 0.021209 | 0.083499 | 0.087931 | -0.0060093 | 0.018325 | 0.0094422 | 0.050568 | 0.022817 | 0.081796 | 0.081777 | -0.0010361 | 0.012332 | 0.014328 | 0.036621 | 69.9921 | -52.1622 | 69.8181 | -46.9162 | -49.4843 | -50.6653 | 0.050354 | 69.9646 | -52.0935 | 69.928 | -46.9254 | -49.1364 | -50.2487 | -0.077819 | 0.009155 | -0.16479 | -0.17395 | 5.1682 | 5.246 | 1.4191 | 2.6001 | 1.4969 | 2.6779 | -0.009155 | -0.01831 | -0.054931 | -0.036621 | 5.159 | 5.1682 | 1.8356 | 2.948 | 1.8448 | 2.9572
1970.1719 | 122.2137 | -52.3224 | 69.809 | -0.10002 | 0.020354 | 0.087558 | 0.092204 | 0.0021085 | 0.017897 | 0.0068786 | -0.077606 | 0.02538 | 0.084146 | 0.089895 | -0.00082252 | 0.016818 | 0.0062101 | -0.07782 | 69.7586 | -52.4826 | 69.548 | -47.2504 | -49.8367 | -50.9903 | -0.11444 | 69.754 | -52.359 | 69.7174 | -47.1909 | -49.4247 | -50.4227 | -0.16022 | -0.050354 | -0.26093 | -0.21057 | 5.072 | 5.2322 | 1.3321 | 2.4857 | 1.4923 | 2.6459 | -0.036621 | -0.054932 | -0.091553 | -0.036621 | 5.1315 | 5.1682 | 1.8997 | 2.8976 | 1.9363 | 2.9343
1971.1719 | 122.2137 | -52.359 | 69.6854 | -0.13054 | 0.021422 | 0.084994 | 0.091136 | 0.0023221 | 0.016188 | 0.0075196 | -0.089813 | 0.023244 | 0.08436 | 0.0837 | -0.0012497 | 0.013614 | 0.012192 | -0.1236 | 69.7311 | -52.446 | 69.5343 | -47.2687 | -49.8322 | -51.0132 | -0.16479 | 69.7449 | -52.3636 | 69.6762 | -47.2595 | -49.411 | -50.4868 | -0.086975 | 0.045776 | -0.15106 | -0.19684 | 5.0903 | 5.1773 | 1.3458 | 2.5269 | 1.4328 | 2.6138 | -0.004577 | 0.059509 | -0.009156 | -0.068665 | 5.0995 | 5.1041 | 1.8723 | 2.948 | 1.8768 | 2.9526
1972.1719 | 122.2687 | -52.1622 | 69.9188 | -0.17632 | 0.021636 | 0.087131 | 0.091563 | 0.00039951 | 0.015761 | 0.011792 | -0.065399 | 0.025167 | 0.082864 | 0.086263 | -0.00039512 | 0.012332 | 0.012405 | -0.068665 | 69.9554 | -52.2675 | 69.722 | -47.0764 | -49.6445 | -50.7889 | -0.14191 | 69.96 | -52.1851 | 69.9188 | -47.0993 | -49.2828 | -50.267 | -0.10529 | 0.036621 | -0.19684 | -0.23346 | 5.0858 | 5.191 | 1.3733 | 2.5177 | 1.4786 | 2.623 | -0.022889 | 0.041199 | 0 | -0.041199 | 5.0629 | 5.0858 | 1.8951 | 2.8793 | 1.918 | 2.9022
1973.1719 | 122.2321 | -52.2263 | 69.8181 | -0.12291 | 0.0195 | 0.084994 | 0.090281 | 0.0021085 | 0.018325 | 0.0090148 | -0.059296 | 0.023671 | 0.082864 | 0.087972 | 0.0019547 | 0.01575 | 0.015396 | -0.09613 | 69.7998 | -52.4231 | 69.6075 | -47.2137 | -49.7772 | -50.9125 | -0.10071 | 69.8135 | -52.2949 | 69.8135 | -47.1039 | -49.3286 | -50.3677 | -0.19684 | -0.01831 | -0.21057 | -0.19226 | 5.0125 | 5.2094 | 1.3138 | 2.449 | 1.5106 | 2.6459 | -0.068665 | -0.004577 | -0.004577 | 0 | 5.1224 | 5.191 | 1.8585 | 2.8976 | 1.9272 | 2.9663
1974.1719 | 122.3145 | -114.4089 | 7.6172 | -0.10765 | 0.013091 | 0.095889 | 0.088786 | 0.0031765 | 0.022597 | 0.010083 | -0.092865 | 0.080068 | 0.088846 | 0.08669 | -0.0014633 | 0.017459 | 0.013901 | -0.16479 | 7.6584 | -114.6149 | 7.3792 | -105.4962 | -111.232 | -112.6694 | -0.16479 | 7.6492 | -114.5416 | 7.4387 | -106.4667 | -110.6644 | -111.7355 | -0.20599 | 0.041199 | -0.23804 | -0.27924 | 8.9127 | 9.1187 | 1.7395 | 3.1769 | 1.9455 | 3.3829 | -0.13275 | 0.032044 | -0.17853 | -0.21057 | 7.9422 | 8.075 | 2.6733 | 3.7445 | 2.8061 | 3.8773
1975.1719 | 122.3282 | -116.3956 | 5.658 | 0.029678 | 0.016509 | 0.096316 | 0.086863 | 0.0012539 | 0.020034 | 0.011792 | -0.065399 | 0.080495 | 0.08671 | 0.084554 | 0.00024568 | 0.018314 | 0.013901 | -0.19226 | 5.7129 | -116.5787 | 5.4016 | -107.3364 | -113.1592 | -114.5782 | -0.15564 | 5.6992 | -116.4917 | 5.4611 | -108.2932 | -112.6099 | -113.6627 | -0.18311 | 0.054932 | -0.25635 | -0.31128 | 9.0591 | 9.2422 | 1.8173 | 3.2364 | 2.0004 | 3.4195 | -0.09613 | 0.041199 | -0.19684 | -0.23804 | 8.1024 | 8.1985 | 2.7328 | 3.7857 | 2.829 | 3.8818
1976.1719 | 122.2595 | -116.3956 | 5.7449 | -0.084763 | 0.013945 | 0.094394 | 0.085582 | 0.0010403 | 0.021315 | 0.014783 | -0.10812 | 0.080068 | 0.087991 | 0.084554 | -0.0023177 | 0.018527 | 0.017532 | -0.11444 | 5.7312 | -116.5466 | 5.4337 | -107.2952 | -113.15 | -114.5737 | -0.19226 | 5.7495 | -116.423 | 5.5298 | -108.2565 | -112.5961 | -113.6307 | -0.15106 | -0.013733 | -0.31128 | -0.29755 | 9.1003 | 9.2514 | 1.8219 | 3.2455 | 1.973 | 3.3966 | -0.027466 | 0.004578 | -0.21515 | -0.21973 | 8.139 | 8.1665 | 2.7649 | 3.7994 | 2.7924 | 3.8269
1977.1719 | 122.2458 | -115.416 | 6.7245 | -0.11528 | 0.015441 | 0.092257 | 0.086863 | -0.00045509 | 0.020247 | 0.013287 | -0.10812 | 0.079427 | 0.087778 | 0.082631 | 0.0013139 | 0.018955 | 0.016037 | -0.082397 | 6.7108 | -115.5624 | 6.4316 | -106.4072 | -112.1567 | -113.5986 | -0.15564 | 6.7612 | -115.4068 | 6.5048 | -107.2861 | -111.6486 | -112.6556 | -0.14648 | -0.013733 | -0.29297 | -0.27923 | 9.0088 | 9.1553 | 1.8173 | 3.2593 | 1.9638 | 3.4058 | 0.009156 | 0.036621 | -0.21973 | -0.25635 | 8.1299 | 8.1207 | 2.7603 | 3.7674 | 2.7512 | 3.7582
1978.1719 | 122.2549 | -116.5741 | 5.5481 | -0.15343 | 0.014373 | 0.093112 | 0.088359 | -0.00045509 | 0.023452 | 0.013287 | -0.15695 | 0.080068 | 0.08671 | 0.082631 | -0.00018152 | 0.019168 | 0.016891 | -0.10529 | 5.5664 | -116.7023 | 5.2277 | -107.5882 | -113.324 | -114.8117 | -0.036621 | 5.5847 | -116.5878 | 5.3238 | -108.3618 | -112.7975 | -113.8046 | -0.12817 | 0.01831 | -0.32043 | -0.33874 | 8.9859 | 9.1141 | 1.7624 | 3.2501 | 1.8906 | 3.3783 | -0.013733 | 0.036621 | -0.2243 | -0.26093 | 8.2123 | 8.226 | 2.7695 | 3.7766 | 2.7832 | 3.7903
1979.1875 | 122.2321 | -116.5009 | 5.6396 | -0.10002 | 0.015654 | 0.093753 | 0.0843 | -0.0023777 | 0.020034 | 0.010724 | 0.0047913 | 0.081991 | 0.086496 | 0.0837 | -0.00018152 | 0.020236 | 0.013687 | 0.004578 | 5.603 | -116.6245 | 5.365 | -107.3547 | -113.1454 | -114.6057 | 0.027466 | 5.6213 | -116.5009 | 5.3925 | -108.252 | -112.6602 | -113.736 | -0.1236 | -0.036621 | -0.27466 | -0.23804 | 9.1461 | 9.2697 | 1.8951 | 3.3554 | 2.0187 | 3.479 | 0 | -0.01831 | -0.24719 | -0.22888 | 8.2489 | 8.2489 | 2.7649 | 3.8406 | 2.7649 | 3.8406
1980.1719 | 122.2916 | -116.597 | 5.4977 | -0.10002 | 0.013732 | 0.093753 | 0.086436 | 0.00018591 | 0.021956 | 0.013074 | -0.092865 | 0.081777 | 0.089273 | 0.084768 | 0.00045928 | 0.019382 | 0.013473 | -0.13275 | 5.5344 | -116.7572 | 5.2185 | -107.5745 | -113.3377 | -114.7705 | -0.10986 | 5.5298 | -116.6107 | 5.3421 | -108.3115 | -112.7563 | -113.8138 | -0.16022 | 0.036621 | -0.27924 | -0.31586 | 9.0225 | 9.1827 | 1.8265 | 3.2593 | 1.9867 | 3.4195 | -0.013733 | 0.032043 | -0.15564 | -0.18768 | 8.2855 | 8.2993 | 2.7832 | 3.8406 | 2.7969 | 3.8544
1981.1719 | 122.2641 | -116.7618 | 5.3787 | -0.17632 | 0.015441 | 0.094821 | 0.085795 | -0.0015231 | 0.021315 | 0.012433 | -0.077606 | 0.080495 | 0.090555 | 0.084554 | 3.2078e-05 | 0.021091 | 0.017319 | -0.10071 | 5.3925 | -116.89 | 5.0812 | -107.6752 | -113.475 | -114.917 | -0.14648 | 5.3879 | -116.7389 | 5.191 | -108.4076 | -112.9211 | -113.9694 | -0.12817 | 0.013733 | -0.29755 | -0.31128 | 9.0866 | 9.2148 | 1.8448 | 3.2867 | 1.973 | 3.4149 | 0.022888 | 0.009155 | -0.18768 | -0.19684 | 8.3542 | 8.3313 | 2.7924 | 3.8406 | 2.7695 | 3.8177
1982.1719 | 122.2687 | -116.7297 | 5.397 | -0.016098 | 0.014373 | 0.092257 | 0.087291 | -0.0019503 | 0.020034 | 0.010724 | -0.092865 | 0.080709 | 0.085855 | 0.082204 | -0.00039512 | 0.019809 | 0.017105 | -0.18768 | 5.4016 | -116.8945 | 5.1224 | -107.7026 | -113.4705 | -114.9124 | -0.1236 | 5.3925 | -116.7389 | 5.1773 | -108.3755 | -112.8983 | -113.9328 | -0.16479 | 0.004577 | -0.27466 | -0.27924 | 9.0271 | 9.1919 | 1.8173 | 3.2593 | 1.9821 | 3.4241 | -0.009156 | -0.004578 | -0.21973 | -0.21515 | 8.3542 | 8.3633 | 2.7969 | 3.8315 | 2.8061 | 3.8406
1983.1719 | 122.2183 | -115.7227 | 6.4041 | -0.10765 | 0.015654 | 0.095035 | 0.087077 | 0.0029629 | 0.019179 | 0.013501 | -0.095917 | 0.080068 | 0.088205 | 0.081563 | 0.00045928 | 0.019382 | 0.016037 | -0.10529 | 6.4041 | -115.8554 | 6.102 | -106.7871 | -112.4863 | -113.942 | -0.14648 | 6.4133 | -115.7318 | 6.1844 | -107.3959 | -111.9186 | -112.9532 | -0.13275 | 0 | -0.30212 | -0.30212 | 8.9355 | 9.0683 | 1.7807 | 3.2364 | 1.9135 | 3.3691 | -0.009156 | 0.009155 | -0.21973 | -0.22888 | 8.3267 | 8.3359 | 2.7695 | 3.804 | 2.7786 | 3.8132
1984.1875 | 122.3785 | -115.6723 | 6.3446 | -0.046616 | 0.015654 | 0.092471 | 0.083659 | -0.0023777 | 0.023238 | 0.0096558 | 0.041412 | 0.077718 | 0.087564 | 0.074941 | 3.2078e-05 | 0.017887 | 0.013687 | -0.018311 | 6.3629 | -115.8005 | 6.2073 | -106.572 | -112.3215 | -113.7772 | 0.11444 | 6.3492 | -115.7364 | 6.2759 | -107.2128 | -111.763 | -112.8387 | -0.12817 | 0.018311 | -0.13733 | -0.15564 | 9.1003 | 9.2285 | 1.8951 | 3.3508 | 2.0233 | 3.479 | -0.064087 | 0.004578 | -0.068664 | -0.073242 | 8.4595 | 8.5236 | 2.8336 | 3.9093 | 2.8976 | 3.9734
1985.1719 | 122.2092 | -115.8554 | 6.2714 | -0.1458 | 0.01715 | 0.091617 | 0.085368 | 0.00018591 | 0.021529 | 0.013715 | -0.11423 | 0.079641 | 0.089914 | 0.086477 | -0.0021041 | 0.016178 | 0.011978 | -0.14648 | 6.2759 | -115.9332 | 6.0104 | -106.842 | -112.6099 | -114.0427 | -0.15564 | 6.2714 | -115.8417 | 6.0608 | -107.4417 | -112.056 | -113.0722 | -0.07782 | 0.004578 | -0.26093 | -0.2655 | 9.0134 | 9.0912 | 1.8127 | 3.2455 | 1.8906 | 3.3234 | 0.013733 | 0 | -0.21057 | -0.21057 | 8.4137 | 8.4 | 2.7832 | 3.7994 | 2.7695 | 3.7857
1986.1719 | 122.2549 | -115.9241 | 6.221 | -0.11528 | 0.012877 | 0.092685 | 0.086222 | -0.00045509 | 0.018965 | 0.011578 | -0.095917 | 0.08135 | 0.088632 | 0.086477 | 0.0011001 | 0.017032 | 0.012832 | -0.10986 | 6.1981 | -116.066 | 5.8777 | -106.9473 | -112.5732 | -114.1022 | -0.1236 | 6.2164 | -115.9195 | 5.9784 | -107.4829 | -112.1109 | -113.1363 | -0.14191 | -0.022888 | -0.34332 | -0.32043 | 8.9767 | 9.1187 | 1.8219 | 3.3508 | 1.9638 | 3.4927 | 0.004577 | -0.004577 | -0.24261 | -0.23804 | 8.4412 | 8.4366 | 2.7878 | 3.8132 | 2.7832 | 3.8086
1987.1719 | 122.2961 | -115.8966 | 6.1478 | -0.1458 | 0.015227 | 0.092685 | 0.087077 | -0.00066869 | 0.022597 | 0.013074 | -0.095917 | 0.080709 | 0.089273 | 0.084554 | -0.0012497 | 0.019809 | 0.013046 | -0.11444 | 6.1615 | -116.0751 | 5.896 | -106.9199 | -112.6648 | -114.1296 | -0.10071 | 6.1707 | -115.947 | 6.0196 | -107.4142 | -112.1017 | -113.118 | -0.17853 | 0.013733 | -0.25177 | -0.2655 | 8.9767 | 9.1553 | 1.767 | 3.2318 | 1.9455 | 3.4103 | -0.050354 | 0.022888 | -0.12817 | -0.15106 | 8.4824 | 8.5327 | 2.7786 | 3.7949 | 2.829 | 3.8452
1988.1719 | 122.2916 | -116.8304 | 5.2551 | -0.15343 | 0.013945 | 0.09418 | 0.086863 | 0.0025357 | 0.021956 | 0.011151 | -0.08371 | 0.082418 | 0.088632 | 0.083059 | -0.00039512 | 0.020664 | 0.013901 | -0.10986 | 5.2505 | -116.9724 | 4.9896 | -107.7667 | -113.5803 | -115.0223 | -0.10529 | 5.2872 | -116.8442 | 5.1086 | -108.2428 | -112.9807 | -113.974 | -0.14191 | -0.004578 | -0.2655 | -0.26093 | 9.0637 | 9.2056 | 1.8082 | 3.2501 | 1.9501 | 3.392 | -0.013733 | 0.032043 | -0.14648 | -0.17853 | 8.5876 | 8.6014 | 2.8564 | 3.8498 | 2.8702 | 3.8635
1989.1719 | 122.319 | -116.6794 | 5.4016 | -0.12291 | 0.015868 | 0.093112 | 0.083873 | -0.0028049 | 0.021315 | 0.012006 | 0.0047913 | 0.080709 | 0.089059 | 0.077291 | 0.0011001 | 0.017887 | 0.016464 | 0.09613 | 5.3604 | -116.8488 | 5.2002 | -107.5287 | -113.2599 | -114.7842 | 0.12817 | 5.3787 | -116.6748 | 5.2551 | -107.9865 | -112.7609 | -113.8092 | -0.16937 | -0.041198 | -0.20142 | -0.16022 | 9.1507 | 9.3201 | 1.8951 | 3.4195 | 2.0645 | 3.5889 | 0.004577 | -0.022888 | -0.14648 | -0.1236 | 8.6929 | 8.6884 | 2.8702 | 3.9185 | 2.8656 | 3.9139
1990.1719 | 122.2733 | -116.3727 | 5.7358 | -0.12291 | 0.016082 | 0.093326 | 0.087077 | 0.0029629 | 0.023024 | 0.01286 | -0.10507 | 0.080922 | 0.089914 | 0.077718 | -0.0014633 | 0.019596 | 0.01625 | -0.13275 | 5.7037 | -116.5466 | 5.4428 | -107.3364 | -113.1271 | -114.5874 | -0.11902 | 5.7083 | -116.3956 | 5.5435 | -107.7667 | -112.5275 | -113.5666 | -0.17395 | -0.032044 | -0.29297 | -0.26093 | 9.0363 | 9.2102 | 1.7853 | 3.2455 | 1.9592 | 3.4195 | -0.022888 | -0.027466 | -0.19226 | -0.1648 | 8.606 | 8.6288 | 2.8061 | 3.8452 | 2.829 | 3.8681
1991.1719 | 122.168 | -116.7252 | 5.4016 | -0.16869 | 0.01715 | 0.093966 | 0.08665 | -0.00088229 | 0.021529 | 0.011578 | -0.11423 | 0.081563 | 0.087991 | 0.087545 | 0.0023819 | 0.021091 | 0.01561 | -0.10986 | 5.3787 | -116.89 | 5.1132 | -107.6935 | -113.475 | -114.9353 | -0.15564 | 5.4337 | -116.7343 | 5.072 | -107.8629 | -112.8525 | -113.9191 | -0.16479 | -0.022888 | -0.28839 | -0.2655 | 9.0317 | 9.1965 | 1.7899 | 3.2501 | 1.9547 | 3.4149 | -0.009155 | 0.032044 | -0.32959 | -0.36163 | 8.8623 | 8.8715 | 2.8061 | 3.8727 | 2.8152 | 3.8818
1992.1719 | 122.3373 | -116.5421 | 5.5161 | -0.10765 | 0.015013 | 0.09418 | 0.086436 | 0.0014675 | 0.018111 | 0.0098694 | -0.10202 | 0.081777 | 0.090128 | 0.088186 | 0.00088648 | 0.018314 | 0.015396 | -0.10986 | 5.5069 | -116.7297 | 5.2872 | -107.4463 | -113.2965 | -114.7202 | -0.068665 | 5.5298 | -116.5833 | 5.4153 | -107.7621 | -112.6236 | -113.6673 | -0.18768 | -0.009155 | -0.22888 | -0.21973 | 9.0958 | 9.2834 | 1.8219 | 3.2455 | 2.0096 | 3.4332 | -0.041199 | 0.013733 | -0.10071 | -0.11444 | 8.7799 | 8.8211 | 2.8748 | 3.9185 | 2.916 | 3.9597
1993.1719 | 122.2229 | -116.6153 | 5.4977 | -0.24498 | 0.012023 | 0.09183 | 0.089854 | -0.0017367 | 0.021743 | 0.013287 | -0.10812 | 0.084127 | 0.088419 | 0.081777 | 0.00067288 | 0.019809 | 0.012619 | -0.004578 | 5.452 | -116.7755 | 5.2185 | -107.5058 | -113.4201 | -114.8483 | -0.11444 | 5.5069 | -116.6382 | 5.3055 | -107.8812 | -112.7655 | -113.7955 | -0.16022 | -0.045777 | -0.27924 | -0.23346 | 9.1095 | 9.2697 | 1.767 | 3.1952 | 1.9272 | 3.3554 | -0.022889 | 0.009155 | -0.19226 | -0.20142 | 8.7341 | 8.757 | 2.8198 | 3.8498 | 2.8427 | 3.8727
1994.1719 | 122.2595 | -116.9266 | 5.1498 | -0.13817 | 0.013732 | 0.09418 | 0.086222 | -0.0034457 | 0.018965 | 0.014996 | -0.022675 | 0.080068 | 0.090128 | 0.081563 | -0.0018905 | 0.015964 | 0.014114 | -0.064087 | 5.0491 | -117.128 | 4.8981 | -107.8033 | -113.6398 | -115.1367 | 0.004578 | 5.1086 | -116.9632 | 4.9393 | -108.1375 | -113.0493 | -114.0381 | -0.20142 | -0.10071 | -0.25177 | -0.15106 | 9.1232 | 9.3246 | 1.7899 | 3.2867 | 1.9913 | 3.4882 | -0.036621 | -0.041198 | -0.21057 | -0.16937 | 8.7891 | 8.8257 | 2.8885 | 3.8773 | 2.9251 | 3.9139
1995.1719 | 122.2458 | -117.4622 | 4.6143 | -0.11528 | 0.013518 | 0.092471 | 0.085582 | -0.0021641 | 0.021956 | 0.011792 | -0.11728 | 0.082204 | 0.089487 | 0.079427 | -0.00060892 | 0.021091 | 0.014541 | -0.10529 | 4.5731 | -117.6682 | 4.3213 | -108.3344 | -114.2303 | -115.6998 | -0.10529 | 4.5731 | -117.5537 | 4.454 | -108.5907 | -113.5757 | -114.6194 | -0.20599 | -0.041199 | -0.29297 | -0.25177 | 9.1278 | 9.3338 | 1.7624 | 3.2318 | 1.9684 | 3.4378 | -0.091553 | -0.041199 | -0.16022 | -0.11902 | 8.8715 | 8.963 | 2.8427 | 3.8864 | 2.9343 | 3.978
1996.1719 | 122.2641 | -116.9861 | 5.1041 | -0.092392 | 0.013518 | 0.091617 | 0.086436 | 0.0018949 | 0.020461 | 0.012646 | -0.11728 | 0.084127 | 0.088419 | 0.0837 | -0.00018152 | 0.019809 | 0.017319 | -0.11444 | 5.0812 | -117.16 | 4.8294 | -107.84 | -113.7314 | -115.1871 | -0.12817 | 5.1086 | -117.0547 | 4.9072 | -108.1238 | -113.1363 | -114.1525 | -0.17395 | -0.022888 | -0.27466 | -0.25177 | 9.1461 | 9.3201 | 1.799 | 3.2547 | 1.973 | 3.4286 | -0.068665 | 0.004578 | -0.19684 | -0.20142 | 8.8623 | 8.931 | 2.8336 | 3.8498 | 2.9022 | 3.9185
1997.1719 | 122.2916 | -116.9678 | 5.1544 | -0.12291 | 0.015441 | 0.094394 | 0.087504 | 0.0031765 | 0.020461 | 0.011792 | -0.13559 | 0.083913 | 0.090982 | 0.08669 | 3.2078e-05 | 0.019382 | 0.012832 | -0.10071 | 5.1132 | -117.1463 | 4.8019 | -107.9224 | -113.768 | -115.2191 | -0.14191 | 5.1498 | -116.9998 | 5.0034 | -108.1604 | -113.2233 | -114.1846 | -0.17853 | -0.041199 | -0.35248 | -0.31128 | 9.0454 | 9.2239 | 1.7487 | 3.1998 | 1.9272 | 3.3783 | -0.032044 | -0.004578 | -0.15106 | -0.14648 | 8.8074 | 8.8394 | 2.7832 | 3.7445 | 2.8152 | 3.7765
1998.1719 | 122.2046 | -116.5009 | 5.6396 | -0.12291 | 0.015654 | 0.092044 | 0.085582 | 0.0031765 | 0.023452 | 0.013928 | -0.13559 | 0.082418 | 0.089487 | 0.085409 | 0.0017411 | 0.018527 | 0.01561 | -0.082397 | 5.6168 | -116.6519 | 5.3146 | -107.4326 | -113.2553 | -114.711 | -0.15106 | 5.658 | -116.5192 | 5.4016 | -107.6157 | -112.6648 | -113.6536 | -0.15106 | -0.022888 | -0.32501 | -0.30212 | 9.0683 | 9.2194 | 1.7899 | 3.2455 | 1.9409 | 3.3966 | -0.018311 | 0.018311 | -0.23804 | -0.25635 | 8.8852 | 8.9035 | 2.8473 | 3.8361 | 2.8656 | 3.8544
1999.1719 | 122.2549 | -116.6382 | 5.4108 | -0.15343 | 0.015227 | 0.092257 | 0.082591 | 0.00039951 | 0.02217 | 0.011365 | 0.0047913 | 0.080495 | 0.087137 | 0.081777 | -0.0031723 | 0.017459 | 0.016678 | -0.086975 | 5.4565 | -116.7435 | 5.2505 | -107.4646 | -113.2965 | -114.7247 | -0.059509 | 5.4611 | -116.6199 | 5.3055 | -107.6111 | -112.6968 | -113.7222 | -0.10529 | 0.045776 | -0.16022 | -0.20599 | 9.1736 | 9.2789 | 1.9135 | 3.3417 | 2.0187 | 3.447 | 0.018311 | 0.050354 | -0.10529 | -0.15564 | 9.0271 | 9.0088 | 2.916 | 3.9413 | 2.8976 | 3.923
2000.1719 | 122.2229 | -116.6611 | 5.4565 | -0.13817 | 0.014373 | 0.092898 | 0.090068 | -0.00045509 | 0.020461 | 0.012219 | -0.13254 | 0.081991 | 0.090982 | 0.085195 | -0.0010361 | 0.02045 | 0.016464 | -0.10071 | 5.4382 | -116.8121 | 5.1315 | -107.6157 | -113.0722 | -114.7705 | -0.13275 | 5.4932 | -116.6565 | 5.2414 | -107.6706 | -112.8387 | -113.8229 | -0.15106 | -0.018311 | -0.32501 | -0.3067 | 9.0454 | 9.1965 | 1.8906 | 3.5889 | 2.0416 | 3.7399 | 0.004578 | 0.036621 | -0.21515 | -0.25177 | 8.9905 | 8.9859 | 2.8381 | 3.8223 | 2.8336 | 3.8177
2001.1719 | 122.2733 | -116.7938 | 5.2917 | -0.13054 | 0.016082 | 0.091403 | 0.085154 | -0.0019503 | 0.022597 | 0.011151 | -0.10202 | 0.08135 | 0.08735 | 0.08669 | 0.00088648 | 0.0181 | 0.016464 | -0.10529 | 5.278 | -116.9632 | 5.0354 | -107.666 | -113.5437 | -115.0223 | -0.11444 | 5.3146 | -116.8213 | 5.1086 | -107.7576 | -112.9211 | -113.9465 | -0.16937 | -0.013733 | -0.25635 | -0.24261 | 9.1278 | 9.2972 | 1.7715 | 3.2501 | 1.9409 | 3.4195 | -0.027466 | 0.022888 | -0.18311 | -0.20599 | 9.0363 | 9.0637 | 2.8473 | 3.8727 | 2.8748 | 3.9001
2002.1719 | 122.2595 | -116.8808 | 5.2231 | -0.13817 | 0.015868 | 0.092898 | 0.087504 | -0.0034457 | 0.022383 | 0.011578 | -0.12033 | 0.082631 | 0.090341 | 0.083486 | 0.00045928 | 0.019382 | 0.015396 | -0.1236 | 5.2094 | -117.0319 | 4.9393 | -107.7301 | -113.6169 | -115.0772 | -0.1236 | 5.2139 | -116.9083 | 5.0217 | -107.7942 | -113.031 | -114.0244 | -0.15106 | -0.013732 | -0.28381 | -0.27008 | 9.1507 | 9.3018 | 1.8036 | 3.2639 | 1.9547 | 3.4149 | -0.027466 | -0.009155 | -0.20142 | -0.19226 | 9.0866 | 9.1141 | 2.8564 | 3.8498 | 2.8839 | 3.8773
2003.1719 | 122.2321 | -116.89 | 5.2094 | -0.11528 | 0.01245 | 0.093539 | 0.087931 | 0.0010403 | 0.021743 | 0.01051 | -0.12338 | 0.081777 | 0.088419 | 0.081777 | 0.0013139 | 0.017887 | 0.016464 | -0.1236 | 5.2231 | -117.0685 | 4.9164 | -107.7209 | -113.6307 | -115.0864 | -0.1236 | 5.1956 | -116.9312 | 5.0217 | -107.7759 | -113.0447 | -114.0381 | -0.17853 | 0.013732 | -0.29297 | -0.3067 | 9.169 | 9.3475 | 1.8036 | 3.2593 | 1.9821 | 3.4378 | -0.041198 | -0.013733 | -0.18768 | -0.17395 | 9.1141 | 9.1553 | 2.8519 | 3.8452 | 2.8931 | 3.8864
2004.1719 | 122.287 | -117.0914 | 4.9667 | -0.1458 | 0.013518 | 0.095889 | 0.086436 | -0.0010959 | 0.024306 | 0.012006 | -0.044037 | 0.08135 | 0.088419 | 0.081991 | 0.00024568 | 0.017246 | 0.018387 | -0.041199 | 5.0217 | -117.1509 | 4.8981 | -107.7576 | -113.6719 | -115.1505 | -0.018311 | 5.0308 | -117.0593 | 4.953 | -107.785 | -113.2507 | -114.2715 | -0.059509 | 0.054931 | -0.068665 | -0.1236 | 9.3338 | 9.3933 | 1.9409 | 3.4195 | 2.0004 | 3.479 | 0.032044 | 0.064087 | -0.013733 | -0.07782 | 9.3063 | 9.2743 | 2.8198 | 3.8406 | 2.7878 | 3.8086
2005.1719 | 122.2137 | -116.9495 | 5.1315 | -0.10002 | 0.013945 | 0.09418 | 0.087077 | -0.0010959 | 0.017043 | 0.012646 | -0.11423 | 0.082204 | 0.090341 | 0.089895 | 0.00067288 | 0.017887 | 0.013901 | -0.10986 | 5.1224 | -117.1463 | 4.9164 | -107.8354 | -113.6856 | -115.123 | -0.14191 | 5.1773 | -117.0044 | 4.9118 | -107.7667 | -113.118 | -114.1113 | -0.19684 | -0.009156 | -0.21515 | -0.20599 | 9.1141 | 9.3109 | 1.8265 | 3.2639 | 2.0233 | 3.4607 | -0.054932 | 0.045776 | -0.21973 | -0.2655 | 9.1827 | 9.2377 | 2.8381 | 3.8315 | 2.8931 | 3.8864
2006.1719 | 122.2321 | -117.0502 | 5.0308 | -0.13054 | 0.014586 | 0.093539 | 0.08665 | 0.0010403 | 0.020034 | 0.011365 | 0.0017395 | 0.082204 | 0.090341 | 0.085195 | -0.0014633 | 0.018955 | 0.017532 | -0.11444 | 5.0217 | -117.2379 | 4.7195 | -107.8125 | -114.0747 | -115.242 | -0.18768 | 5.0491 | -117.0959 | 4.8157 | -107.8537 | -113.2507 | -114.2166 | -0.18768 | -0.009156 | -0.31128 | -0.30212 | 9.2377 | 9.4254 | 1.8082 | 2.9755 | 1.9959 | 3.1631 | -0.045776 | 0.01831 | -0.21515 | -0.23346 | 9.1965 | 9.2422 | 2.8336 | 3.7994 | 2.8793 | 3.8452
2007.1719 | 122.2183 | -117.0364 | 5.0629 | -0.11528 | 0.014373 | 0.094821 | 0.087931 | 0.0016813 | 0.023238 | 0.014142 | -0.11423 | 0.081777 | 0.089487 | 0.080709 | -0.00082252 | 0.020023 | 0.015182 | -0.10071 | 5.0583 | -117.2104 | 4.7607 | -107.8629 | -113.7909 | -115.2466 | -0.29297 | 5.0812 | -116.9952 | 4.8523 | -107.8033 | -113.2095 | -114.1891 | -0.17395 | -0.004577 | -0.30212 | -0.29755 | 9.1736 | 9.3475 | 1.7899 | 3.2455 | 1.9638 | 3.4195 | 0.041199 | 0.018311 | -0.21057 | -0.22888 | 9.2331 | 9.1919 | 2.8473 | 3.8269 | 2.8061 | 3.7857
2008.1719 | 122.2733 | -117.1051 | 5.0079 | -0.16106 | 0.012877 | 0.092685 | 0.087077 | 0.0018949 | 0.020461 | 0.01286 | -0.08371 | 0.081991 | 0.087991 | 0.08135 | 0.00045928 | 0.018314 | 0.020309 | -0.14648 | 4.985 | -117.2424 | 4.7104 | -107.8537 | -113.8229 | -115.3061 | -0.12817 | 5.0262 | -117.1097 | 4.8111 | -107.7896 | -113.2553 | -114.1571 | -0.13733 | -0.022889 | -0.29755 | -0.27466 | 9.2514 | 9.3887 | 1.799 | 3.2822 | 1.9363 | 3.4195 | -0.004577 | 0.01831 | -0.19684 | -0.21515 | 9.3155 | 9.3201 | 2.948 | 3.8498 | 2.9526 | 3.8544
2009.1719 | 122.2916 | -117.0593 | 4.9713 | -0.17632 | 0.014373 | 0.093326 | 0.082591 | -0.0010959 | 0.02217 | 0.011792 | -0.025726 | 0.080495 | 0.086496 | 0.077504 | -0.00018152 | 0.019382 | 0.016678 | -0.050354 | 4.9347 | -117.2058 | 4.8203 | -107.7301 | -113.7314 | -115.1916 | 0 | 4.9667 | -117.0868 | 4.8798 | -107.6935 | -113.1683 | -114.1388 | -0.14648 | -0.036621 | -0.15106 | -0.11444 | 9.3292 | 9.4757 | 1.8677 | 3.3279 | 2.0142 | 3.4744 | -0.027466 | -0.004577 | -0.091552 | -0.086975 | 9.3658 | 9.3933 | 2.9205 | 3.891 | 2.948 | 3.9185
2010.1719 | 122.2229 | -117.2424 | 4.9026 | -0.10765 | 0.015013 | 0.092044 | 0.086863 | 0.0033901 | 0.020034 | 0.01051 | -0.12643 | 0.083059 | 0.090341 | 0.081563 | -0.0029587 | 0.019596 | 0.013046 | -0.064087 | 4.866 | -117.3615 | 4.6188 | -107.9361 | -113.9328 | -115.416 | -0.064087 | 4.8981 | -117.2333 | 4.6692 | -107.8583 | -113.3835 | -114.3585 | -0.11902 | -0.036621 | -0.28381 | -0.24719 | 9.3063 | 9.4254 | 1.8265 | 3.3096 | 1.9455 | 3.4287 | 0.009156 | -0.004578 | -0.23346 | -0.22888 | 9.3842 | 9.375 | 2.8839 | 3.8589 | 2.8748 | 3.8498
2011.1719 | 122.1817 | -117.2836 | 4.8477 | -0.12291 | 0.018218 | 0.093753 | 0.085795 | 0.00039951 | 0.020461 | 0.012006 | -0.10812 | 0.083913 | 0.090768 | 0.0837 | 0.00088648 | 0.0181 | 0.019027 | -0.10529 | 4.7974 | -117.4393 | 4.509 | -108.0505 | -114.0427 | -115.4892 | -0.16022 | 4.8294 | -117.2928 | 4.5959 | -107.9086 | -113.4567 | -114.4318 | -0.15564 | -0.050354 | -0.33875 | -0.28839 | 9.2331 | 9.3887 | 1.7944 | 3.241 | 1.9501 | 3.3966 | -0.009156 | -0.01831 | -0.25177 | -0.23346 | 9.375 | 9.3842 | 2.8519 | 3.8269 | 2.861 | 3.8361
2012.1719 | 122.2275 | -117.3248 | 4.7699 | -0.13817 | 0.012664 | 0.094607 | 0.085582 | 0.0068081 | 0.026442 | 0.011365 | -0.092865 | 0.084554 | 0.089914 | 0.079641 | -0.00018152 | 0.018955 | 0.016891 | -0.11902 | 4.7562 | -117.4759 | 4.5364 | -107.8445 | -114.0701 | -115.567 | -0.13733 | 4.7974 | -117.3615 | 4.5639 | -107.8857 | -113.5025 | -114.4547 | -0.15106 | -0.013732 | -0.23346 | -0.21973 | 9.4803 | 9.6313 | 1.7578 | 3.2547 | 1.9089 | 3.4058 | -0.036621 | 0.027466 | -0.20599 | -0.23346 | 9.4391 | 9.4757 | 2.8702 | 3.8223 | 2.9068 | 3.8589
2013.1719 | 122.2183 | -117.2745 | 4.8248 | -0.092392 | 0.018004 | 0.094607 | 0.085795 | -0.00045509 | 0.020461 | 0.01286 | -0.1478 | 0.08135 | 0.090768 | 0.079427 | 0.0017411 | 0.021091 | 0.018814 | -0.11444 | 4.7974 | -117.453 | 4.5273 | -108.0185 | -114.0472 | -115.5899 | -0.032043 | 4.8477 | -117.3019 | 4.5959 | -107.7988 | -113.4659 | -114.4135 | -0.17853 | -0.027466 | -0.29755 | -0.27008 | 9.256 | 9.4345 | 1.6846 | 3.2272 | 1.8631 | 3.4058 | -0.027466 | 0.022888 | -0.22888 | -0.25177 | 9.4757 | 9.5032 | 2.861 | 3.8086 | 2.8885 | 3.8361
2014.1719 | 122.2733 | -117.2562 | 4.7974 | -0.17632 | 0.015868 | 0.093753 | 0.082591 | -0.0034457 | 0.02217 | 0.0096558 | -0.037933 | 0.082204 | 0.091196 | 0.077077 | 0.0017411 | 0.019168 | 0.011978 | -0.068665 | 4.7562 | -117.4164 | 4.6326 | -107.84 | -113.9511 | -115.3748 | 0.022888 | 4.8019 | -117.2562 | 4.6829 | -107.6797 | -113.3331 | -114.3311 | -0.16022 | -0.041198 | -0.16479 | -0.1236 | 9.4162 | 9.5764 | 1.8814 | 3.3051 | 2.0416 | 3.4653 | 0 | 0.004578 | -0.11444 | -0.11902 | 9.5764 | 9.5764 | 2.9251 | 3.923 | 2.9251 | 3.923
2015.1719 | 122.3236 | -117.2836 | 4.7104 | -0.1458 | 0.016295 | 0.092471 | 0.085582 | 0.00082671 | 0.022383 | 0.011792 | -0.059296 | 0.083913 | 0.090982 | 0.08669 | -0.0018905 | 0.018955 | 0.017532 | -0.10986 | 4.7241 | -117.4805 | 4.5319 | -107.9498 | -114.0289 | -115.4892 | -0.036621 | 4.747 | -117.3248 | 4.6417 | -107.6935 | -113.3926 | -114.3585 | -0.19684 | 0.013733 | -0.17853 | -0.19226 | 9.3338 | 9.5306 | 1.7944 | 3.2547 | 1.9913 | 3.4515 | -0.041199 | 0.036621 | -0.068664 | -0.10529 | 9.5901 | 9.6313 | 2.9251 | 3.891 | 2.9663 | 3.9322
2016.1719 | 122.2412 | -117.3477 | 4.7562 | -0.17632 | 0.014159 | 0.090762 | 0.085368 | -0.00024129 | 0.020461 | 0.013715 | -0.10812 | 0.08434 | 0.093332 | 0.087331 | 0.00045928 | 0.0181 | 0.015396 | -0.10529 | 4.7699 | -117.4896 | 4.4678 | -108.0139 | -114.093 | -115.5441 | -0.1236 | 4.7607 | -117.3706 | 4.5502 | -107.7713 | -113.5208 | -114.4638 | -0.14191 | 0.013732 | -0.28839 | -0.30212 | 9.3338 | 9.4757 | 1.8036 | 3.2547 | 1.9455 | 3.3966 | -0.022888 | 0.004577 | -0.20599 | -0.21057 | 9.5764 | 9.5993 | 2.8839 | 3.8269 | 2.9068 | 3.8498
2017.1719 | 122.2916 | -117.3889 | 4.6509 | -0.16106 | 0.013945 | 0.093112 | 0.084941 | -0.0028049 | 0.023665 | 0.014996 | -0.059296 | 0.082845 | 0.088419 | 0.084127 | 0.0015275 | 0.017459 | 0.016464 | -0.11444 | 4.6051 | -117.5858 | 4.4128 | -108.0185 | -114.1525 | -115.6082 | -0.086975 | 4.66 | -117.4347 | 4.5227 | -107.7484 | -113.5208 | -114.4455 | -0.19684 | -0.045776 | -0.23804 | -0.19226 | 9.3704 | 9.5673 | 1.7807 | 3.2364 | 1.9775 | 3.4332 | -0.045776 | 0.009155 | -0.12817 | -0.13733 | 9.6405 | 9.6863 | 2.9434 | 3.8681 | 2.9892 | 3.9139
2018.1719 | 122.168 | -117.485 | 4.6143 | -0.1458 | 0.018218 | 0.095248 | 0.086009 | -0.00066869 | 0.020461 | 0.012646 | -0.12033 | 0.085195 | 0.088846 | 0.087972 | 0.0015275 | 0.019168 | 0.011764 | -0.09613 | 4.5731 | -117.6773 | 4.2984 | -108.1833 | -114.2532 | -115.7181 | -0.15106 | 4.6188 | -117.54 | 4.3762 | -107.872 | -113.7039 | -114.6423 | -0.19226 | -0.041199 | -0.31586 | -0.27466 | 9.3018 | 9.494 | 1.767 | 3.2318 | 1.9592 | 3.4241 | -0.054932 | 0.004577 | -0.23804 | -0.24261 | 9.613 | 9.668 | 2.8427 | 3.7811 | 2.8976 | 3.8361
2019.1719 | 122.2961 | -117.5308 | 4.5593 | -0.11528 | 0.013732 | 0.093326 | 0.08665 | 0.0021085 | 0.022597 | 0.0090148 | -0.10507 | 0.082418 | 0.091196 | 0.080495 | -0.0010361 | 0.018314 | 0.013473 | -0.11444 | 4.5593 | -117.6819 | 4.3121 | -108.1604 | -114.2624 | -115.7318 | -0.10071 | 4.5822 | -117.5491 | 4.3991 | -107.7942 | -113.6398 | -114.5874 | -0.15106 | 0 | -0.24719 | -0.24719 | 9.3704 | 9.5215 | 1.799 | 3.2684 | 1.9501 | 3.4195 | -0.01831 | 0.022888 | -0.16022 | -0.1831 | 9.7366 | 9.7549 | 2.9434 | 3.891 | 2.9617 | 3.9093
2020.1719 | 122.2 | -117.5674 | 4.5456 | -0.13054 | 0.013945 | 0.092898 | 0.087291 | 0.0053129 | 0.019179 | 0.011365 | -0.013519 | 0.083913 | 0.089487 | 0.087118 | 0.00024568 | 0.017673 | 0.013473 | -0.12817 | 4.5273 | -117.7139 | 4.3121 | -108.3115 | -114.4135 | -115.5624 | -0.2243 | 4.5822 | -117.5903 | 4.3167 | -107.8537 | -113.7497 | -114.679 | -0.14648 | -0.01831 | -0.23346 | -0.21515 | 9.256 | 9.4025 | 2.005 | 3.154 | 2.1515 | 3.3005 | -0.022888 | 0.036621 | -0.22888 | -0.2655 | 9.7137 | 9.7366 | 2.8885 | 3.8177 | 2.9114 | 3.8406
2021.1719 | 122.2641 | -117.2379 | 4.8157 | -0.13054 | 0.014159 | 0.091617 | 0.085795 | 0.00018591 | 0.020888 | 0.013287 | -0.10202 | 0.082204 | 0.088419 | 0.079 | 0.00045928 | 0.019168 | 0.014969 | -0.1236 | 4.8019 | -117.4393 | 4.5502 | -107.8812 | -114.0289 | -115.4663 | -0.10071 | 4.8294 | -117.2882 | 4.6509 | -107.5104 | -113.3972 | -114.3356 | -0.20142 | -0.013733 | -0.2655 | -0.25177 | 9.3567 | 9.5581 | 1.7715 | 3.2089 | 1.973 | 3.4103 | -0.050354 | 0.013733 | -0.16479 | -0.17853 | 9.7275 | 9.7778 | 2.9022 | 3.8406 | 2.9526 | 3.891
2022.1719 | 122.2366 | -117.247 | 4.8157 | -0.10765 | 0.014373 | 0.093539 | 0.085154 | 0.0036039 | 0.021529 | 0.013287 | -0.077606 | 0.083486 | 0.089059 | 0.085622 | 0.0013139 | 0.018955 | 0.013687 | -0.11444 | 4.8248 | -117.4347 | 4.5364 | -107.8812 | -114.0106 | -115.4709 | -0.11444 | 4.8477 | -117.247 | 4.6371 | -107.5058 | -113.4109 | -114.3494 | -0.18768 | 0.009155 | -0.27924 | -0.28839 | 9.3658 | 9.5535 | 1.7761 | 3.2364 | 1.9638 | 3.4241 | 0 | 0.032043 | -0.17853 | -0.21057 | 9.7412 | 9.7412 | 2.8976 | 3.8361 | 2.8976 | 3.8361
2023.1719 | 122.2275 | -117.2562 | 4.8111 | -0.13817 | 0.014373 | 0.09418 | 0.087077 | -0.0013095 | 0.021743 | 0.012219 | -0.12033 | 0.081991 | 0.090768 | 0.085409 | -0.00039512 | 0.020664 | 0.018814 | -0.10986 | 4.7012 | -117.54 | 4.4174 | -107.9727 | -114.1663 | -115.5991 | -0.1236 | 4.7241 | -117.4072 | 4.4952 | -107.515 | -113.4384 | -114.3448 | -0.28381 | -0.10986 | -0.39368 | -0.28381 | 9.2834 | 9.5673 | 1.6571 | 3.0899 | 1.9409 | 3.3737 | -0.15106 | -0.086975 | -0.31586 | -0.22888 | 9.7412 | 9.8923 | 2.9114 | 3.8177 | 3.0624 | 3.9688
2024.1719 | 122.3145 | -117.366 | 4.6829 | -0.19157 | 0.015227 | 0.093539 | 0.085582 | -2.7692e-05 | 0.018965 | 0.01051 | 0.0078429 | 0.082204 | 0.090555 | 0.077932 | -0.00060892 | 0.018527 | 0.015396 | -0.073242 | 4.66 | -117.5262 | 4.454 | -107.9132 | -114.1251 | -115.5441 | -0.073242 | 4.6646 | -117.4301 | 4.5319 | -107.5058 | -113.4979 | -114.4272 | -0.16022 | -0.022888 | -0.22888 | -0.20599 | 9.4528 | 9.613 | 1.8219 | 3.241 | 1.9821 | 3.4012 | -0.064087 | -0.01831 | -0.15106 | -0.13275 | 9.8602 | 9.9243 | 2.9388 | 3.8681 | 3.0029 | 3.9322
2025.1719 | 122.1863 | -117.5217 | 4.5685 | -0.12291 | 0.014586 | 0.094607 | 0.087504 | -0.0038729 | 0.019393 | 0.014142 | -0.12033 | 0.084768 | 0.089059 | 0.084981 | -0.0018905 | 0.018955 | 0.014969 | -0.091553 | 4.5456 | -117.7094 | 4.2618 | -108.1512 | -114.3585 | -115.7913 | -0.1236 | 4.5685 | -117.5583 | 4.335 | -107.7576 | -113.736 | -114.6515 | -0.18768 | -0.022888 | -0.3067 | -0.28381 | 9.3704 | 9.5581 | 1.7303 | 3.1631 | 1.918 | 3.3508 | -0.036622 | 0 | -0.23346 | -0.23346 | 9.7641 | 9.8007 | 2.8702 | 3.7857 | 2.9068 | 3.8223
2026.1719 | 122.2229 | -117.691 | 4.4128 | -0.1458 | 0.015013 | 0.092685 | 0.087291 | -0.0047275 | 0.02217 | 0.013928 | -0.086761 | 0.084981 | 0.089273 | 0.083272 | 3.2078e-05 | 0.020023 | 0.014969 | -0.10986 | 4.39 | -117.8329 | 4.0741 | -108.2748 | -114.3951 | -115.9241 | -0.17395 | 4.4128 | -117.7277 | 4.1885 | -107.8262 | -113.9603 | -114.798 | -0.14191 | -0.022888 | -0.33875 | -0.31586 | 9.4162 | 9.5581 | 1.767 | 3.2959 | 1.9089 | 3.4378 | -0.036621 | 0 | -0.2243 | -0.2243 | 9.8648 | 9.9014 | 2.8931 | 3.7308 | 2.9297 | 3.7674
2027.1719 | 122.2046 | -117.6865 | 4.4037 | -0.13054 | 0.016722 | 0.094394 | 0.088359 | 0.00082671 | 0.021956 | 0.01286 | -0.028778 | 0.083272 | 0.089487 | 0.084127 | 0.00045928 | 0.019382 | 0.015823 | -0.10071 | 4.3716 | -117.8513 | 4.2068 | -108.316 | -114.5966 | -115.7043 | -0.21973 | 4.422 | -117.7094 | 4.1794 | -107.8491 | -113.9008 | -114.798 | -0.16479 | -0.032044 | -0.19684 | -0.16479 | 9.3704 | 9.5352 | 1.9821 | 3.0899 | 2.1469 | 3.2547 | -0.022889 | 0.01831 | -0.2243 | -0.24261 | 9.8373 | 9.8602 | 2.8885 | 3.7857 | 2.9114 | 3.8086
2028.1719 | 122.2504 | -117.7643 | 4.3167 | -0.15343 | 0.015013 | 0.091403 | 0.087077 | 0.0027493 | 0.021956 | 0.013287 | -0.077606 | 0.082845 | 0.088419 | 0.085195 | -0.00082252 | 0.019809 | 0.01326 | -0.1236 | 4.2984 | -117.9474 | 4.0237 | -108.316 | -114.5782 | -115.9882 | -0.12817 | 4.3167 | -117.8284 | 4.1428 | -107.8445 | -113.9694 | -114.8483 | -0.18311 | -0.01831 | -0.29297 | -0.27466 | 9.4482 | 9.6313 | 1.7761 | 3.186 | 1.9592 | 3.3691 | -0.064087 | 0 | -0.17395 | -0.17395 | 9.9197 | 9.9838 | 2.916 | 3.7949 | 2.98 | 3.8589
2029.1875 | 122.2137 | -117.8055 | 4.3259 | -0.069504 | 0.013732 | 0.090976 | 0.084086 | -0.0010959 | 0.021102 | 0.01286 | -0.092865 | 0.08434 | 0.090128 | 0.078786 | 0.0017411 | 0.017032 | 0.016678 | -0.050354 | 4.3121 | -117.9337 | 4.0283 | -108.3206 | -114.5508 | -116.011 | -0.036621 | 4.3304 | -117.8101 | 4.0604 | -107.8491 | -113.9832 | -114.8895 | -0.12817 | -0.013733 | -0.29755 | -0.28381 | 9.4849 | 9.613 | 1.7944 | 3.2547 | 1.9226 | 3.3829 | -0.004578 | 0.004577 | -0.2655 | -0.27008 | 9.9564 | 9.9609 | 2.916 | 3.8223 | 2.9205 | 3.8269
2030.1719 | 122.2183 | -117.7322 | 4.3671 | -0.069504 | 0.0148 | 0.091617 | 0.087077 | -2.7692e-05 | 0.021529 | 0.011578 | -0.10202 | 0.0837 | 0.0897 | 0.085409 | 0.0015275 | 0.016818 | 0.020736 | -0.086975 | 4.3533 | -117.8787 | 4.0604 | -108.2565 | -114.5279 | -115.9424 | -0.14191 | 4.3579 | -117.7597 | 4.1611 | -107.785 | -113.9786 | -114.8209 | -0.14648 | -0.013732 | -0.3067 | -0.29297 | 9.4757 | 9.6222 | 1.7899 | 3.2043 | 1.9363 | 3.3508 | -0.027466 | -0.009155 | -0.20599 | -0.19684 | 9.9472 | 9.9747 | 2.9114 | 3.7537 | 2.9388 | 3.7811
2031.1719 | 122.2046 | -117.7139 | 4.3808 | -0.16869 | 0.013732 | 0.092044 | 0.087931 | -2.7692e-05 | 0.020247 | 0.014355 | -0.12949 | 0.083059 | 0.088846 | 0.080495 | 0.00088648 | 0.019382 | 0.013687 | -0.086975 | 4.3259 | -117.9245 | 4.0512 | -108.2657 | -114.5508 | -115.9744 | -0.15564 | 4.3808 | -117.7734 | 4.1519 | -107.7805 | -113.9786 | -114.8529 | -0.21057 | -0.054931 | -0.32959 | -0.27466 | 9.4482 | 9.6588 | 1.7395 | 3.1631 | 1.9501 | 3.3737 | -0.059509 | 0 | -0.22888 | -0.22888 | 9.9335 | 9.993 | 2.861 | 3.7354 | 2.9205 | 3.7949
2032.1719 | 122.1863 | -117.8055 | 4.3213 | -0.13817 | 0.017363 | 0.093753 | 0.085795 | 0.0033901 | 0.020247 | 0.013715 | -0.12338 | 0.083272 | 0.092264 | 0.088613 | 0.00045928 | 0.020664 | 0.015396 | -0.091553 | 4.2618 | -117.984 | 3.9642 | -108.316 | -114.6103 | -116.0156 | -0.13733 | 4.3167 | -117.8421 | 4.0695 | -107.8125 | -114.0427 | -114.9216 | -0.17853 | -0.059509 | -0.35706 | -0.29755 | 9.4894 | 9.668 | 1.7899 | 3.1952 | 1.9684 | 3.3737 | -0.036621 | -0.004578 | -0.25177 | -0.24719 | 9.993 | 10.0296 | 2.8839 | 3.7628 | 2.9205 | 3.7994
2033.1719 | 122.2504 | -117.7917 | 4.2755 | -0.16869 | 0.015013 | 0.093326 | 0.086222 | 0.0023221 | 0.021315 | 0.0090148 | -0.10507 | 0.084127 | 0.088205 | 0.084554 | 0.0017411 | 0.018314 | 0.012619 | -0.1236 | 4.2801 | -117.9474 | 4.0146 | -108.284 | -114.5966 | -116.0065 | -0.11902 | 4.3167 | -117.7917 | 4.1061 | -107.7393 | -113.9832 | -114.8895 | -0.15564 | 0.004577 | -0.26093 | -0.2655 | 9.5078 | 9.6634 | 1.7853 | 3.1952 | 1.9409 | 3.3508 | 0 | 0.041198 | -0.16937 | -0.21057 | 10.0525 | 10.0525 | 2.9022 | 3.8086 | 2.9022 | 3.8086
2034.1719 | 122.3099 | -117.7277 | 4.3259 | -0.17632 | 0.015654 | 0.091403 | 0.082377 | -0.0019503 | 0.020888 | 0.013715 | -0.013519 | 0.08135 | 0.0897 | 0.082845 | 0.0015275 | 0.016391 | 0.011764 | -0.050354 | 4.2755 | -117.8696 | 4.1428 | -108.1146 | -114.4318 | -115.8417 | 0.004578 | 4.3167 | -117.7231 | 4.2068 | -107.6248 | -113.8504 | -114.7751 | -0.14191 | -0.050354 | -0.18311 | -0.13275 | 9.613 | 9.7549 | 1.886 | 3.2959 | 2.0279 | 3.4378 | 0.004578 | -0.009156 | -0.11902 | -0.10986 | 10.1028 | 10.0983 | 2.9526 | 3.8773 | 2.948 | 3.8727
2035.1719 | 122.2458 | -117.865 | 4.2206 | -0.15343 | 0.017791 | 0.095675 | 0.086009 | 0.0025357 | 0.021956 | 0.012646 | -0.092865 | 0.082631 | 0.090128 | 0.078786 | -0.0016769 | 0.020236 | 0.016464 | -0.22888 | 4.248 | -118.0023 | 3.9734 | -108.3481 | -114.6469 | -116.0431 | -0.13733 | 4.2618 | -117.8833 | 4.0054 | -107.7896 | -114.0884 | -114.9399 | -0.13733 | 0.027466 | -0.24719 | -0.27466 | 9.5169 | 9.6542 | 1.8219 | 3.2181 | 1.9592 | 3.3554 | -0.018311 | 0.041199 | -0.21515 | -0.25635 | 10.0754 | 10.0937 | 2.9251 | 3.7766 | 2.9434 | 3.7949
2036.1719 | 122.2 | -117.7734 | 4.3213 | -0.13817 | 0.015868 | 0.093539 | 0.085368 | 0.0010403 | 0.021956 | 0.014569 | -0.086761 | 0.08135 | 0.089487 | 0.088613 | 0.00024568 | 0.017887 | 0.014755 | -0.10529 | 4.2709 | -117.9474 | 3.9871 | -108.0551 | -114.505 | -116.0339 | -0.11902 | 4.335 | -117.8009 | 4.1107 | -107.7118 | -114.0106 | -114.8758 | -0.17395 | -0.050354 | -0.33417 | -0.28381 | 9.7183 | 9.8923 | 1.7395 | 3.2684 | 1.9135 | 3.4424 | -0.027466 | 0.013733 | -0.21057 | -0.2243 | 10.0616 | 10.0891 | 2.8976 | 3.7628 | 2.9251 | 3.7903
2037.1719 | 122.2412 | -117.9886 | 4.0878 | -0.15343 | 0.014586 | 0.094394 | 0.084513 | -0.00088229 | 0.020888 | 0.010937 | -0.095917 | 0.082204 | 0.091196 | 0.084127 | 0.0015275 | 0.0181 | 0.01326 | -0.11902 | 4.0924 | -118.158 | 3.8132 | -108.4076 | -114.8163 | -116.185 | -0.11902 | 4.097 | -117.9977 | 3.9139 | -107.8537 | -114.1663 | -115.0497 | -0.16937 | 0.004577 | -0.27466 | -0.27924 | 9.581 | 9.7504 | 1.8036 | 3.1723 | 1.973 | 3.3417 | -0.009156 | 0.009155 | -0.17395 | -0.18311 | 10.1349 | 10.144 | 2.9388 | 3.8223 | 2.948 | 3.8315
2038.1719 | 122.1725 | -118.0206 | 4.1016 | -0.12291 | 0.017791 | 0.092044 | 0.085582 | -0.0047275 | 0.018752 | 0.01521 | -0.11423 | 0.082845 | 0.090555 | 0.080709 | -0.00018152 | 0.017246 | 0.019882 | -0.082397 | 4.0466 | -118.1763 | 3.7857 | -108.4579 | -114.8575 | -116.2308 | -0.10529 | 4.1016 | -118.0298 | 3.8544 | -107.8949 | -114.299 | -115.1184 | -0.15564 | -0.054931 | -0.31586 | -0.26092 | 9.5627 | 9.7183 | 1.7899 | 3.1631 | 1.9455 | 3.3188 | -0.009155 | 0 | -0.24719 | -0.24719 | 10.1257 | 10.1349 | 2.9022 | 3.7216 | 2.9114 | 3.7308
2039.1719 | 122.2916 | -117.952 | 4.0924 | -0.18394 | 0.016936 | 0.092471 | 0.083873 | -0.0028049 | 0.022383 | 0.013287 | 0.0017395 | 0.081991 | 0.087778 | 0.077504 | 0.0017411 | 0.01575 | 0.016037 | -0.068665 | 4.0466 | -118.0664 | 3.9139 | -108.2611 | -114.6561 | -116.0477 | -0.009155 | 4.0924 | -117.9703 | 3.9871 | -107.7347 | -114.0839 | -114.9719 | -0.11444 | -0.045776 | -0.17853 | -0.13275 | 9.6909 | 9.8053 | 1.9043 | 3.2959 | 2.0187 | 3.4103 | -0.018311 | 0 | -0.10528 | -0.10528 | 10.2173 | 10.2356 | 2.98 | 3.8681 | 2.9984 | 3.8864
2040.1719 | 122.287 | -117.9382 | 4.1382 | -0.16106 | 0.016295 | 0.092685 | 0.084513 | -0.0025913 | 0.019606 | 0.0094422 | -0.0043641 | 0.082845 | 0.0897 | 0.076864 | 3.2078e-05 | 0.018314 | 0.017746 | -0.082397 | 4.0787 | -118.0801 | 3.923 | -108.2657 | -114.6973 | -116.0751 | -0.045776 | 4.1245 | -117.9611 | 3.9551 | -107.7072 | -114.093 | -114.9857 | -0.14191 | -0.05951 | -0.21515 | -0.15564 | 9.6725 | 9.8145 | 1.8631 | 3.241 | 2.005 | 3.3829 | -0.022889 | -0.013733 | -0.18311 | -0.16937 | 10.231 | 10.2539 | 2.9526 | 3.8452 | 2.9755 | 3.8681
2041.1719 | 122.1909 | -117.8284 | 4.2618 | -0.069504 | 0.016509 | 0.093753 | 0.085582 | 0.0021085 | 0.021529 | 0.01286 | -0.11118 | 0.083913 | 0.088205 | 0.088827 | -0.00039512 | 0.017887 | 0.01625 | -0.086975 | 4.2435 | -118.0115 | 3.9642 | -108.2748 | -114.6698 | -116.0797 | -0.14191 | 4.303 | -117.8787 | 4.0192 | -107.7118 | -114.1022 | -114.9445 | -0.18311 | -0.018311 | -0.29755 | -0.27924 | 9.5535 | 9.7366 | 1.7487 | 3.1586 | 1.9318 | 3.3417 | -0.050354 | 0.041199 | -0.24261 | -0.28381 | 10.1166 | 10.1669 | 2.8839 | 3.7262 | 2.9343 | 3.7766
2042.1719 | 122.1954 | -117.9153 | 4.184 | -0.10002 | 0.015441 | 0.093326 | 0.087077 | 0.0023221 | 0.020461 | 0.011151 | -0.23325 | 0.085195 | 0.0897 | 0.080922 | -0.0029587 | 0.018314 | 0.013046 | -0.11444 | 4.1473 | -118.1122 | 3.8681 | -108.3344 | -114.7934 | -116.1621 | -0.032043 | 4.2068 | -117.9199 | 3.9368 | -107.7347 | -113.5849 | -114.9948 | -0.19684 | -0.036621 | -0.31586 | -0.27924 | 9.581 | 9.7778 | 1.7532 | 3.1219 | 1.9501 | 3.3188 | -0.004578 | 0.022888 | -0.24719 | -0.27008 | 10.1807 | 10.1852 | 2.9205 | 4.3304 | 2.9251 | 4.335
2043.1719 | 122.2137 | -117.7185 | 4.4312 | -0.084763 | 0.014159 | 0.093326 | 0.086009 | 0.00039951 | 0.021102 | 0.011578 | -0.10507 | 0.082418 | 0.088205 | 0.090749 | -0.00082252 | 0.015323 | 0.013473 | -0.12817 | 4.3854 | -117.8787 | 4.0833 | -108.165 | -114.5828 | -115.9195 | -0.14648 | 4.4037 | -117.7368 | 4.2206 | -107.547 | -113.9694 | -114.798 | -0.16022 | -0.045776 | -0.3479 | -0.30212 | 9.5535 | 9.7137 | 1.799 | 3.1357 | 1.9592 | 3.2959 | -0.01831 | -0.027465 | -0.21057 | -0.18311 | 10.1715 | 10.1898 | 2.9205 | 3.7491 | 2.9388 | 3.7674
2044.1719 | 122.287 | -117.6956 | 4.3304 | -0.16869 | 0.016936 | 0.093326 | 0.084727 | -0.00045509 | 0.021743 | 0.011151 | -0.022675 | 0.081777 | 0.086923 | 0.074727 | 0.00088648 | 0.017246 | 0.017959 | -0.059509 | 4.303 | -117.8787 | 4.1565 | -107.9865 | -114.4867 | -115.8279 | 0 | 4.3167 | -117.7414 | 4.2343 | -107.3959 | -113.8916 | -114.7522 | -0.1831 | -0.027465 | -0.17395 | -0.14648 | 9.7092 | 9.8923 | 1.8677 | 3.2089 | 2.0508 | 3.392 | -0.045776 | -0.013733 | -0.09613 | -0.082397 | 10.2997 | 10.3455 | 2.9434 | 3.804 | 2.9892 | 3.8498
2045.1875 | 122.3328 | -117.6773 | 4.3625 | -0.11528 | 0.015868 | 0.093112 | 0.084513 | -0.00066869 | 0.020888 | 0.010083 | 0.016998 | 0.082418 | 0.089914 | 0.070455 | 3.2078e-05 | 0.017673 | 0.01326 | -0.054932 | 4.3808 | -117.8467 | 4.1565 | -107.9544 | -114.418 | -115.7822 | 0.009155 | 4.3625 | -117.7231 | 4.2389 | -107.3639 | -113.8138 | -114.7156 | -0.16937 | 0.01831 | -0.20599 | -0.2243 | 9.7229 | 9.8923 | 1.8951 | 3.2593 | 2.0645 | 3.4287 | -0.045776 | 0 | -0.1236 | -0.1236 | 10.3134 | 10.3592 | 2.9617 | 3.8635 | 3.0075 | 3.9093
2046.1719 | 122.2 | -117.7139 | 4.3625 | -0.15343 | 0.014373 | 0.09183 | 0.08665 | 0.0038175 | 0.018538 | 0.012646 | -0.14169 | 0.081563 | 0.089914 | 0.08135 | -0.0023177 | 0.018741 | 0.012405 | -0.064087 | 4.3488 | -117.8558 | 4.1245 | -108.0368 | -114.5416 | -115.9058 | -0.036621 | 4.3762 | -117.7322 | 4.1611 | -107.4463 | -113.974 | -114.8209 | -0.14191 | -0.013733 | -0.23804 | -0.2243 | 9.6771 | 9.819 | 1.8082 | 3.1723 | 1.9501 | 3.3142 | -0.018311 | 0.013733 | -0.20142 | -0.21515 | 10.2676 | 10.286 | 2.8931 | 3.7399 | 2.9114 | 3.7582
2047.1719 | 122.2 | -117.7094 | 4.4037 | -0.16869 | 0.014373 | 0.092685 | 0.086009 | -0.0017367 | 0.020461 | 0.011151 | -0.11423 | 0.0837 | 0.091837 | 0.087545 | -0.0012497 | 0.01575 | 0.013901 | -0.091553 | 4.335 | -117.8787 | 4.0878 | -108.1055 | -114.5554 | -115.947 | -0.11902 | 4.3671 | -117.7277 | 4.1336 | -107.46 | -114.1113 | -114.8209 | -0.16937 | -0.068665 | -0.31586 | -0.24719 | 9.6039 | 9.7733 | 1.7624 | 3.154 | 1.9318 | 3.3234 | -0.01831 | -0.036622 | -0.27008 | -0.23346 | 10.2493 | 10.2676 | 2.8885 | 3.598 | 2.9068 | 3.6163
2048.1719 | 122.168 | -117.6956 | 4.4128 | -0.13054 | 0.015654 | 0.092898 | 0.085582 | 0.00018591 | 0.019179 | 0.01286 | -0.11118 | 0.082631 | 0.088632 | 0.085195 | 3.2078e-05 | 0.016391 | 0.014969 | -0.059509 | 4.3671 | -117.8696 | 4.1061 | -108.046 | -114.537 | -115.9103 | -0.11444 | 4.4037 | -117.7277 | 4.1748 | -107.4051 | -113.9877 | -114.7888 | -0.17395 | -0.045777 | -0.3067 | -0.26093 | 9.6497 | 9.8236 | 1.7853 | 3.1586 | 1.9592 | 3.3325 | -0.032043 | -0.009155 | -0.23804 | -0.22888 | 10.2905 | 10.3226 | 2.9068 | 3.7079 | 2.9388 | 3.7399
2049.1719 | 122.2824 | -117.6544 | 4.3762 | -0.16869 | 0.013304 | 0.092685 | 0.082377 | -0.0021641 | 0.019179 | 0.012433 | -0.040985 | 0.081563 | 0.089914 | 0.079427 | 0.00067288 | 0.016818 | 0.015823 | -0.050354 | 4.3076 | -117.8146 | 4.216 | -107.8903 | -114.4455 | -115.7822 | 0.004578 | 4.3671 | -117.6727 | 4.2801 | -107.2495 | -113.855 | -114.7018 | -0.16022 | -0.068665 | -0.16022 | -0.091553 | 9.7641 | 9.9243 | 1.8723 | 3.2089 | 2.0325 | 3.3691 | -0.01831 | -0.009156 | -0.096131 | -0.086975 | 10.405 | 10.4233 | 2.9526 | 3.7994 | 2.9709 | 3.8177
2050.1875 | 122.2549 | -117.6453 | 4.3991 | -0.1458 | 0.014586 | 0.095462 | 0.083018 | 0.0012539 | 0.020674 | 0.01051 | 0.0047913 | 0.081991 | 0.089059 | 0.082204 | 0.00067288 | 0.017459 | 0.016678 | -0.091553 | 4.3945 | -117.7917 | 4.2068 | -107.9269 | -114.4318 | -115.773 | -0.054932 | 4.4037 | -117.6682 | 4.2526 | -107.254 | -113.855 | -114.679 | -0.14648 | -0.004578 | -0.19226 | -0.18768 | 9.7183 | 9.8648 | 1.8723 | 3.2135 | 2.0187 | 3.36 | -0.022888 | 0.004578 | -0.14648 | -0.15106 | 10.3912 | 10.4141 | 2.9663 | 3.7903 | 2.9892 | 3.8132
2051.1719 | 122.1771 | -117.7231 | 4.4037 | -0.061874 | 0.015868 | 0.092044 | 0.086222 | 0.0014675 | 0.021743 | 0.013501 | -0.13864 | 0.081991 | 0.091196 | 0.087972 | -0.00018152 | 0.017459 | 0.011978 | -0.073242 | 4.3442 | -117.9199 | 4.0924 | -108.0917 | -114.6011 | -115.9332 | -0.1236 | 4.3625 | -117.746 | 4.129 | -107.3868 | -113.9969 | -114.8163 | -0.19684 | -0.05951 | -0.31128 | -0.25177 | 9.6313 | 9.8282 | 1.7899 | 3.1219 | 1.9867 | 3.3188 | -0.022889 | -0.041199 | -0.27466 | -0.23346 | 10.3363 | 10.3592 | 2.9068 | 3.7262 | 2.9297 | 3.7491
2052.1719 | 122.1771 | -117.7277 | 4.3579 | -0.092392 | 0.016509 | 0.092471 | 0.087718 | -2.7692e-05 | 0.020247 | 0.011151 | -0.11423 | 0.084127 | 0.089273 | 0.085622 | 3.2078e-05 | 0.018955 | 0.014969 | -0.054932 | 4.3304 | -117.9016 | 4.0695 | -108.0322 | -114.6057 | -115.9424 | -0.10071 | 4.3488 | -117.7826 | 4.1245 | -107.4005 | -114.0106 | -114.83 | -0.17395 | -0.027466 | -0.28839 | -0.26093 | 9.6954 | 9.8694 | 1.7853 | 3.1219 | 1.9592 | 3.2959 | -0.054932 | -0.009155 | -0.23346 | -0.2243 | 10.3271 | 10.3821 | 2.8976 | 3.717 | 2.9526 | 3.772
2053.1719 | 122.2366 | -117.7643 | 4.335 | -0.13817 | 0.015227 | 0.092898 | 0.085368 | 0.0012539 | 0.021315 | 0.013074 | -0.11423 | 0.08434 | 0.089487 | 0.085836 | 0.0015275 | 0.018741 | 0.012619 | -0.022888 | 4.2892 | -117.9337 | 4.0833 | -107.785 | -114.4455 | -115.8279 | -0.004578 | 4.3991 | -117.7643 | 4.1428 | -107.373 | -114.0106 | -114.8346 | -0.16937 | -0.045776 | -0.25177 | -0.20599 | 9.9792 | 10.1486 | 1.9363 | 3.3188 | 2.1057 | 3.4882 | 0 | 0.064087 | -0.19226 | -0.25635 | 10.3912 | 10.3912 | 2.9297 | 3.7537 | 2.9297 | 3.7537
2054.1719 | 122.2778 | -117.778 | 4.303 | -0.12291 | 0.014159 | 0.090976 | 0.086863 | 0.00018591 | 0.021315 | 0.0096558 | -0.065399 | 0.0837 | 0.090341 | 0.084127 | -0.00018152 | 0.018955 | 0.017746 | -0.11444 | 4.2892 | -117.9657 | 4.0146 | -108.078 | -114.6698 | -115.9744 | -0.09613 | 4.2801 | -117.8101 | 4.1428 | -107.3547 | -114.0106 | -114.8438 | -0.18768 | -0.013733 | -0.28839 | -0.27466 | 9.7 | 9.8877 | 1.8036 | 3.1082 | 1.9913 | 3.2959 | -0.032044 | -0.022889 | -0.16022 | -0.13733 | 10.4233 | 10.4553 | 2.9343 | 3.7674 | 2.9663 | 3.7994
2055.1719 | 122.2092 | -117.7963 | 4.2892 | -0.092392 | 0.0148 | 0.092898 | 0.086222 | -0.00066869 | 0.018752 | 0.013287 | -0.10202 | 0.08135 | 0.088419 | 0.090108 | 0.00024568 | 0.016818 | 0.015823 | -0.022888 | 4.3167 | -117.952 | 4.0466 | -108.0688 | -114.6332 | -115.9561 | -0.022888 | 4.303 | -117.8375 | 4.1107 | -107.3822 | -114.0564 | -114.8758 | -0.15564 | 0.027465 | -0.24262 | -0.27008 | 9.7275 | 9.8831 | 1.8402 | 3.1631 | 1.9958 | 3.3188 | -0.041198 | 0.013733 | -0.17853 | -0.19226 | 10.4141 | 10.4553 | 2.9205 | 3.7399 | 2.9617 | 3.7811
2056.1719 | 122.2092 | -117.8329 | 4.2526 | -0.18394 | 0.015868 | 0.093112 | 0.088572 | -0.00088229 | 0.020034 | 0.0092286 | -0.10812 | 0.084127 | 0.090128 | 0.079213 | -0.00018152 | 0.016391 | 0.014969 | -0.054932 | 4.2297 | -118.0023 | 3.9642 | -108.1055 | -114.7018 | -116.0156 | -0.086975 | 4.2435 | -117.865 | 4.01 | -107.4326 | -114.2395 | -114.9399 | -0.16937 | -0.022889 | -0.28839 | -0.2655 | 9.7275 | 9.8969 | 1.8173 | 3.1311 | 1.9867 | 3.3005 | -0.032043 | -0.009156 | -0.24261 | -0.23346 | 10.4004 | 10.4324 | 2.8931 | 3.5934 | 2.9251 | 3.6255
2057.1719 | 122.1863 | -117.8513 | 4.2755 | -0.069504 | 0.015868 | 0.092898 | 0.085582 | 0.0038175 | 0.020461 | 0.012433 | -0.11118 | 0.081777 | 0.092264 | 0.08135 | -0.0010361 | 0.018527 | 0.019241 | -0.050354 | 4.1885 | -117.9977 | 4.0054 | -108.0917 | -114.6927 | -116.0294 | -0.045776 | 4.2618 | -117.8513 | 4.0283 | -107.4188 | -114.1113 | -114.9399 | -0.14648 | -0.086975 | -0.27008 | -0.18311 | 9.7595 | 9.906 | 1.8219 | 3.1586 | 1.9684 | 3.3051 | 0 | -0.013733 | -0.24719 | -0.23346 | 10.4324 | 10.4324 | 2.9114 | 3.7399 | 2.9114 | 3.7399
2058.1719 | 122.2229 | -117.8467 | 4.2389 | -0.12291 | 0.016936 | 0.090762 | 0.084727 | -2.7692e-05 | 0.020888 | 0.012219 | -0.092865 | 0.081991 | 0.089914 | 0.079427 | -0.0018905 | 0.019809 | 0.015396 | -0.050354 | 4.2572 | -117.9794 | 3.9917 | -108.0826 | -114.6881 | -116.0156 | -0.013733 | 4.2343 | -117.8787 | 4.0375 | -107.4051 | -114.1159 | -114.9536 | -0.13275 | 0.01831 | -0.24719 | -0.2655 | 9.7641 | 9.8969 | 1.8311 | 3.1586 | 1.9638 | 3.2913 | -0.032043 | -0.004578 | -0.20142 | -0.19684 | 10.4416 | 10.4736 | 2.8931 | 3.7308 | 2.9251 | 3.7628
2059.1719 | 122.2733 | -117.7917 | 4.2114 | -0.13054 | 0.015227 | 0.091403 | 0.086222 | 0.0014675 | 0.022811 | 0.0092286 | -0.059296 | 0.080709 | 0.092905 | 0.083272 | -0.0012497 | 0.020236 | 0.017319 | -0.11444 | 4.2343 | -118.0161 | 3.9825 | -108.0872 | -114.6881 | -115.9973 | -0.059509 | 4.2343 | -117.8513 | 4.0878 | -107.2906 | -114.0289 | -114.8483 | -0.2243 | 0.022888 | -0.22888 | -0.25177 | 9.7046 | 9.9289 | 1.7944 | 3.1036 | 2.0187 | 3.3279 | -0.059509 | 0.022888 | -0.1236 | -0.14648 | 10.5011 | 10.5606 | 2.9434 | 3.7628 | 3.0029 | 3.8223
2060.1719 | 122.2366 | -117.8467 | 4.2526 | -0.15343 | 0.015013 | 0.092898 | 0.082804 | -0.0019503 | 0.021315 | 0.010297 | -0.053192 | 0.084127 | 0.091837 | 0.080068 | -0.00039512 | 0.018314 | 0.012405 | -0.022888 | 4.2252 | -118.0344 | 3.9825 | -108.0643 | -114.6652 | -116.0019 | 0.018311 | 4.2389 | -117.8558 | 4.0512 | -107.3273 | -114.0656 | -114.9033 | -0.18768 | -0.027466 | -0.27008 | -0.24261 | 9.7824 | 9.9701 | 1.8448 | 3.1815 | 2.0325 | 3.3691 | -0.009155 | -0.013733 | -0.20142 | -0.18768 | 10.5194 | 10.5286 | 2.9434 | 3.7811 | 2.9526 | 3.7903
2061.1719 | 122.2504 | -117.8009 | 4.2664 | -0.092392 | 0.016082 | 0.091617 | 0.08665 | 0.00061311 | 0.019393 | 0.012219 | -0.053192 | 0.083272 | 0.090341 | 0.084768 | -0.0016769 | 0.019168 | 0.014328 | 0.013733 | 4.2343 | -117.9657 | 4.01 | -108.0414 | -114.6423 | -115.9882 | 0.018311 | 4.2618 | -117.8467 | 4.0787 | -107.2586 | -114.0564 | -114.8941 | -0.16479 | -0.032043 | -0.25635 | -0.2243 | 9.7595 | 9.9243 | 1.8127 | 3.1586 | 1.9775 | 3.3234 | -0.045777 | -0.004577 | -0.18768 | -0.18311 | 10.5423 | 10.5881 | 2.9068 | 3.7445 | 2.9526 | 3.7903
2062.1719 | 122.2641 | -117.8238 | 4.2709 | -0.092392 | 0.015227 | 0.092471 | 0.085795 | -0.0010959 | 0.020674 | 0.012006 | -0.025726 | 0.082631 | 0.088846 | 0.083913 | 0.0011001 | 0.017246 | 0.016891 | 0.004578 | 4.2389 | -118.0298 | 3.9871 | -108.046 | -114.6744 | -115.9927 | 0.050354 | 4.2389 | -117.8787 | 4.0558 | -107.2723 | -114.0839 | -114.9078 | -0.20599 | -0.032043 | -0.28381 | -0.25177 | 9.7778 | 9.9838 | 1.8311 | 3.1494 | 2.037 | 3.3554 | -0.054931 | -0.032043 | -0.21515 | -0.18311 | 10.5515 | 10.6064 | 2.916 | 3.7399 | 2.9709 | 3.7949
2063.1719 | 122.1909 | -117.746 | 4.3121 | -0.077133 | 0.015441 | 0.094394 | 0.084727 | 0.00039951 | 0.020674 | 0.013928 | -0.065399 | 0.081563 | 0.089059 | 0.082845 | -0.0016769 | 0.017887 | 0.017105 | 0.009155 | 4.1931 | -118.0161 | 3.9917 | -108.0505 | -114.6881 | -115.9927 | 0.036621 | 4.2526 | -117.8375 | 4.0604 | -107.2495 | -113.8687 | -114.7888 | -0.27008 | -0.11902 | -0.32044 | -0.20142 | 9.6954 | 9.9655 | 1.7532 | 3.0579 | 2.0233 | 3.3279 | -0.091552 | -0.059509 | -0.25177 | -0.19226 | 10.4965 | 10.5881 | 2.9572 | 3.8773 | 3.0487 | 3.9688
2064.1719 | 122.1771 | -117.8421 | 4.2709 | -0.16869 | 0.016509 | 0.092898 | 0.087077 | 0.0018949 | 0.020461 | 0.0088012 | -0.080658 | 0.082845 | 0.091623 | 0.084981 | 0.00088648 | 0.017032 | 0.014541 | -0.09613 | 4.248 | -117.9977 | 3.9276 | -108.1467 | -114.7842 | -116.0431 | -0.12817 | 4.303 | -117.8741 | 4.01 | -107.3044 | -114.1754 | -114.949 | -0.15564 | -0.022888 | -0.34332 | -0.32043 | 9.6954 | 9.8511 | 1.799 | 3.0579 | 1.9547 | 3.2135 | -0.032044 | 0.032044 | -0.26092 | -0.29297 | 10.5377 | 10.5698 | 2.8931 | 3.6667 | 2.9251 | 3.6987
2065.1719 | 122.2275 | -117.8375 | 4.248 | -0.092392 | 0.015868 | 0.092685 | 0.085582 | 0.00082671 | 0.018111 | 0.012006 | -0.065399 | 0.082204 | 0.093546 | 0.079854 | 3.2078e-05 | 0.0181 | 0.016891 | -0.022888 | 4.2252 | -117.9977 | 3.9963 | -108.0093 | -114.6744 | -116.0019 | 0.036621 | 4.2664 | -117.865 | 4.0558 | -107.2037 | -114.0839 | -114.9124 | -0.16022 | -0.022888 | -0.25177 | -0.22888 | 9.8282 | 9.9884 | 1.8356 | 3.1631 | 1.9959 | 3.3234 | -0.027466 | 0.01831 | -0.19226 | -0.21057 | 10.6338 | 10.6613 | 2.9251 | 3.7537 | 2.9526 | 3.7811
2066.1719 | 122.1863 | -117.8192 | 4.2664 | -0.13054 | 0.013732 | 0.091403 | 0.084941 | 0.0048855 | 0.022597 | 0.011792 | -0.034882 | 0.081991 | 0.092264 | 0.079213 | 0.00024568 | 0.017887 | 0.016678 | -0.032043 | 4.2572 | -117.952 | 3.9963 | -108.046 | -114.6744 | -116.0065 | -0.041199 | 4.3121 | -117.8558 | 4.0329 | -107.2128 | -114.0976 | -114.9307 | -0.13275 | -0.009155 | -0.27008 | -0.26093 | 9.7733 | 9.906 | 1.8127 | 3.1448 | 1.9455 | 3.2776 | -0.036621 | 0.045777 | -0.23346 | -0.27924 | 10.6064 | 10.643 | 2.8885 | 3.7216 | 2.9251 | 3.7582
2067.1719 | 122.1817 | -117.8421 | 4.2755 | -0.061874 | 0.015013 | 0.090549 | 0.084086 | -0.0019503 | 0.018752 | 0.012219 | -0.016571 | 0.080068 | 0.091409 | 0.086904 | 0.0019547 | 0.017673 | 0.013046 | -0.045776 | 4.2206 | -118.0252 | 3.9413 | -108.1558 | -114.7156 | -116.0156 | -0.004578 | 4.2526 | -117.8101 | 3.9825 | -107.2495 | -114.1754 | -114.9628 | -0.18311 | -0.054932 | -0.33417 | -0.27924 | 9.6863 | 9.8694 | 1.8265 | 3.1265 | 2.0096 | 3.3096 | 0.032043 | -0.022888 | -0.29297 | -0.27008 | 10.5927 | 10.5606 | 2.8793 | 3.6667 | 2.8473 | 3.6346
2068.1719 | 122.2137 | -117.897 | 4.2206 | -0.17632 | 0.013518 | 0.090976 | 0.087077 | -0.0010959 | 0.020674 | 0.013501 | -0.034882 | 0.082845 | 0.090128 | 0.078359 | -0.00060892 | 0.017887 | 0.014755 | 0 | 4.2435 | -118.0115 | 3.9276 | -108.1604 | -114.7476 | -116.0477 | -0.013733 | 4.2252 | -117.897 | 3.9734 | -107.2723 | -114.1754 | -114.9857 | -0.11444 | 0.022888 | -0.29297 | -0.31586 | 9.7366 | 9.8511 | 1.8494 | 3.1494 | 1.9638 | 3.2639 | 0 | 0.004578 | -0.24719 | -0.25177 | 10.6247 | 10.6247 | 2.9114 | 3.7216 | 2.9114 | 3.7216
2069.1719 | 122.2092 | -117.8787 | 4.2297 | -0.10002 | 0.017363 | 0.095035 | 0.087718 | 0.00061311 | 0.019606 | 0.011792 | -0.092865 | 0.084768 | 0.091837 | 0.081563 | -0.00039512 | 0.017459 | 0.014755 | -0.054932 | 4.1977 | -118.0527 | 3.9413 | -108.0872 | -114.7614 | -116.0706 | -0.07782 | 4.2206 | -117.8879 | 3.9642 | -107.2586 | -114.2029 | -114.9948 | -0.17395 | -0.032043 | -0.28839 | -0.25635 | 9.7916 | 9.9655 | 1.8082 | 3.1174 | 1.9821 | 3.2913 | -0.009155 | -0.009155 | -0.2655 | -0.25635 | 10.6201 | 10.6293 | 2.8839 | 3.6758 | 2.8931 | 3.685
2070.1719 | 122.2504 | -117.9108 | 4.2206 | -0.12291 | 0.014586 | 0.095462 | 0.084086 | 0.0021085 | 0.020034 | 0.01286 | -0.037933 | 0.082418 | 0.091837 | 0.078573 | -0.00039512 | 0.0181 | 0.01561 | -0.041199 | 4.2252 | -118.0481 | 3.9139 | -108.11 | -114.743 | -116.0568 | -0.009155 | 4.2343 | -117.9016 | 3.9688 | -107.2357 | -114.18 | -114.9857 | -0.13733 | 0.004578 | -0.3067 | -0.31128 | 9.8007 | 9.9381 | 1.8539 | 3.1677 | 1.9913 | 3.3051 | 0.009156 | 0.013733 | -0.25177 | -0.2655 | 10.675 | 10.6659 | 2.9251 | 3.7308 | 2.916 | 3.7216
2071.1719 | 122.2137 | -117.9108 | 4.1748 | -0.084763 | 0.014586 | 0.094607 | 0.085154 | -0.0021641 | 0.020674 | 0.013715 | -0.053192 | 0.082418 | 0.090768 | 0.081136 | -0.00082252 | 0.018314 | 0.012192 | -0.086975 | 4.2114 | -118.0161 | 3.9001 | -108.1558 | -114.7705 | -116.0614 | -0.032043 | 4.2114 | -117.9153 | 3.9368 | -107.2586 | -114.212 | -115.0085 | -0.10528 | 0.036621 | -0.27466 | -0.31128 | 9.7549 | 9.8602 | 1.8494 | 3.1403 | 1.9547 | 3.2455 | -0.004577 | 0.036621 | -0.23804 | -0.27466 | 10.6522 | 10.6567 | 2.9022 | 3.6987 | 2.9068 | 3.7033
2072.1719 | 122.2229 | -117.8513 | 4.2343 | -0.077133 | 0.014159 | 0.091403 | 0.086436 | 0.0027493 | 0.018538 | 0.012433 | -0.019623 | 0.083059 | 0.092264 | 0.083059 | 0.00088648 | 0.017032 | 0.013901 | -0.036621 | 4.2252 | -118.0298 | 4.1336 | -107.8629 | -114.8163 | -116.0843 | -0.018311 | 4.2892 | -117.8696 | 3.9825 | -107.1899 | -114.2075 | -114.9719 | -0.17853 | -0.009155 | -0.10071 | -0.091553 | 9.9884 | 10.1669 | 1.767 | 3.035 | 1.9455 | 3.2135 | -0.018311 | 0.054932 | -0.25177 | -0.3067 | 10.6613 | 10.6796 | 2.8793 | 3.6438 | 2.8976 | 3.6621
2073.1719 | 122.1954 | -117.8421 | 4.2343 | -0.1458 | 0.014373 | 0.092685 | 0.084941 | 0.00061311 | 0.021102 | 0.010724 | 0.0047913 | 0.080068 | 0.090768 | 0.077077 | 0.00045928 | 0.016391 | 0.014114 | 0.022888 | 4.2023 | -117.9703 | 4.0054 | -108.0688 | -114.6652 | -115.9607 | 0.064087 | 4.2343 | -117.8604 | 4.0009 | -107.1396 | -114.1571 | -114.9582 | -0.12817 | -0.032043 | -0.22888 | -0.19684 | 9.7733 | 9.9014 | 1.8814 | 3.1769 | 2.0096 | 3.3051 | -0.018311 | 0 | -0.23346 | -0.23346 | 10.7025 | 10.7208 | 2.8839 | 3.685 | 2.9022 | 3.7033
2074.1719 | 122.2321 | -117.8284 | 4.2938 | -0.11528 | 0.013091 | 0.093326 | 0.087291 | 0.0038175 | 0.019393 | 0.010724 | -0.080658 | 0.081777 | 0.091196 | 0.080495 | -0.0016769 | 0.0181 | 0.012619 | 0.018311 | 4.1565 | -118.0573 | 3.9642 | -108.11 | -114.7568 | -116.0248 | 0.041199 | 4.184 | -117.9016 | 4.2801 | -107.0572 | -114.0747 | -114.9124 | -0.22888 | -0.13733 | -0.32959 | -0.19226 | 9.7183 | 9.9472 | 1.8036 | 3.0716 | 2.0325 | 3.3005 | -0.073242 | -0.10986 | -0.013733 | 0.09613 | 10.7712 | 10.8444 | 2.916 | 3.7537 | 2.9892 | 3.8269
2075.1719 | 122.2229 | -117.7917 | 4.3442 | -0.069504 | 0.013304 | 0.092471 | 0.085154 | -0.0010959 | 0.019393 | 0.014355 | -0.013519 | 0.081777 | 0.091837 | 0.079213 | 0.00067288 | 0.01511 | 0.014328 | 0.009155 | 4.2709 | -117.9062 | 4.0421 | -108.0414 | -114.6423 | -115.9241 | 0.009155 | 4.3076 | -117.7917 | 4.0558 | -107.048 | -114.1022 | -114.8987 | -0.11444 | -0.073242 | -0.30212 | -0.22888 | 9.7504 | 9.8648 | 1.8677 | 3.1494 | 1.9821 | 3.2639 | 0 | -0.036621 | -0.28839 | -0.25177 | 10.7437 | 10.7437 | 2.8931 | 3.6896 | 2.8931 | 3.6896
2076.1719 | 122.2275 | -117.7872 | 4.3488 | -0.092392 | 0.016295 | 0.092685 | 0.084086 | -0.00045509 | 0.020247 | 0.012433 | 0.0017395 | 0.082204 | 0.088632 | 0.078573 | -0.00060892 | 0.016818 | 0.015182 | 0.027466 | 4.2847 | -117.9199 | 4.1061 | -107.9498 | -114.5874 | -115.892 | 0.054932 | 4.3167 | -117.7734 | 4.1199 | -106.9931 | -114.0701 | -114.8804 | -0.13275 | -0.064087 | -0.24261 | -0.17853 | 9.8373 | 9.9701 | 1.8951 | 3.1998 | 2.0279 | 3.3325 | 0.013733 | -0.032044 | -0.22888 | -0.19684 | 10.7941 | 10.7803 | 2.9068 | 3.717 | 2.8931 | 3.7033
2077.1719 | 122.1909 | -117.7277 | 4.39 | -0.16106 | 0.013732 | 0.092471 | 0.084086 | 0.0027493 | 0.01512 | 0.0096558 | -0.019623 | 0.081563 | 0.090128 | 0.081563 | -0.0010361 | 0.014896 | 0.012832 | -0.022888 | 4.3579 | -117.865 | 4.1016 | -107.9819 | -114.6103 | -115.8508 | -0.004578 | 4.3991 | -117.7414 | 4.1473 | -106.9565 | -114.0381 | -114.8529 | -0.13733 | -0.032044 | -0.28839 | -0.25635 | 9.7458 | 9.8831 | 1.8768 | 3.1174 | 2.0142 | 3.2547 | -0.013733 | 0.009155 | -0.24262 | -0.25177 | 10.7712 | 10.7849 | 2.8748 | 3.6896 | 2.8885 | 3.7033
2078.1719 | 122.2275 | -117.7139 | 4.3854 | -0.11528 | 0.014373 | 0.092471 | 0.0843 | -0.00088229 | 0.019393 | 0.010297 | -0.013519 | 0.082631 | 0.090128 | 0.084981 | -0.00082252 | 0.01511 | 0.016891 | 0.004578 | 4.335 | -117.8925 | 4.1199 | -107.9681 | -114.624 | -115.8691 | 0.018311 | 4.3762 | -117.7414 | 4.1199 | -106.9565 | -114.0244 | -114.8392 | -0.17853 | -0.050354 | -0.2655 | -0.21515 | 9.7458 | 9.9243 | 1.8448 | 3.0899 | 2.0233 | 3.2684 | -0.027466 | -0.009155 | -0.2655 | -0.25635 | 10.7574 | 10.7849 | 2.8748 | 3.6896 | 2.9022 | 3.717
2079.1719 | 122.2412 | -117.6086 | 4.4678 | -0.13054 | 0.015227 | 0.093326 | 0.085368 | 0.00061311 | 0.019606 | 0.016278 | -0.077606 | 0.082631 | 0.090982 | 0.086477 | -0.0010361 | 0.018527 | 0.016891 | -0.009155 | 4.4586 | -117.7826 | 4.2389 | -107.8491 | -114.5096 | -115.7822 | 0.032043 | 4.4724 | -117.6315 | 4.2801 | -106.8375 | -113.9236 | -114.711 | -0.17395 | -0.009155 | -0.22888 | -0.21973 | 9.7595 | 9.9335 | 1.8265 | 3.0991 | 2.0004 | 3.273 | -0.022888 | 0.004578 | -0.18768 | -0.19226 | 10.7712 | 10.7941 | 2.8976 | 3.685 | 2.9205 | 3.7079
2080.1719 | 122.3328 | -117.5079 | 4.5547 | -0.13054 | 0.015654 | 0.091617 | 0.083873 | -0.00024129 | 0.018538 | 0.010083 | 0.032257 | 0.080495 | 0.0897 | 0.080495 | 0.00024568 | 0.01511 | 0.01625 | 0.009155 | 4.5273 | -117.6453 | 4.4403 | -107.7667 | -114.5096 | -115.5441 | 0.13275 | 4.5044 | -117.5217 | 4.4312 | -106.6864 | -113.736 | -114.5462 | -0.13733 | -0.027466 | -0.11444 | -0.086975 | 9.7412 | 9.8785 | 1.9638 | 2.9984 | 2.1011 | 3.1357 | -0.013732 | -0.050354 | -0.1236 | -0.073243 | 10.8215 | 10.8353 | 2.9617 | 3.772 | 2.9755 | 3.7857
2081.1719 | 122.3282 | -117.6407 | 4.4037 | -0.13054 | 0.015441 | 0.093326 | 0.0843 | -0.0038729 | 0.019606 | 0.008374 | 0.029205 | 0.078786 | 0.089487 | 0.079427 | 0.0023819 | 0.017246 | 0.014328 | 0.041199 | 4.4037 | -117.7643 | 4.2709 | -107.7209 | -114.4318 | -115.7089 | 0.12817 | 4.4037 | -117.6819 | 4.2847 | -106.7688 | -113.9328 | -114.7064 | -0.1236 | 0 | -0.13275 | -0.13275 | 9.9197 | 10.0433 | 1.9318 | 3.2089 | 2.0554 | 3.3325 | -0.041199 | 0 | -0.11902 | -0.11902 | 10.8719 | 10.9131 | 2.9343 | 3.7079 | 2.9755 | 3.7491
2082.1719 | 122.3053 | -117.6819 | 4.335 | -0.092392 | 0.013732 | 0.093326 | 0.083445 | -0.0019503 | 0.018752 | 0.013074 | -0.016571 | 0.080922 | 0.089273 | 0.082631 | -0.0021041 | 0.0181 | 0.013901 | 0.027466 | 4.2847 | -117.8604 | 4.2114 | -107.8812 | -114.5187 | -115.8005 | 0.11444 | 4.3396 | -117.7094 | 4.2206 | -106.8604 | -113.9786 | -114.7705 | -0.17853 | -0.050354 | -0.1236 | -0.073242 | 9.8007 | 9.9792 | 1.8814 | 3.1631 | 2.0599 | 3.3417 | -0.027466 | 0.004578 | -0.11444 | -0.11902 | 10.8215 | 10.849 | 2.9114 | 3.7033 | 2.9388 | 3.7308
2083.1719 | 122.3007 | -117.7322 | 4.3213 | -0.13054 | 0.014586 | 0.091403 | 0.085582 | -0.0040867 | 0.021102 | 0.0092286 | -0.047089 | 0.080495 | 0.088419 | 0.087118 | -0.0023177 | 0.016605 | 0.012405 | 0.10986 | 4.2389 | -117.9337 | 4.1428 | -107.8445 | -114.537 | -115.8371 | 0.13733 | 4.2938 | -117.7414 | 4.1702 | -107.0068 | -114.0198 | -114.8209 | -0.20142 | -0.082397 | -0.17853 | -0.096131 | 9.8877 | 10.0891 | 1.8951 | 3.1952 | 2.0966 | 3.3966 | -0.009155 | -0.027466 | -0.15106 | -0.1236 | 10.7254 | 10.7346 | 2.9114 | 3.7125 | 2.9205 | 3.7216
2084.1719 | 122.1909 | -117.691 | 4.335 | -0.13817 | 0.015227 | 0.093112 | 0.088359 | 0.0033901 | 0.021529 | 0.013928 | -0.10812 | 0.082204 | 0.090555 | 0.086263 | -0.00082252 | 0.016818 | 0.013473 | 0.027466 | 4.2343 | -118.0115 | 4.0466 | -107.9773 | -114.7064 | -115.9882 | 0.045776 | 4.2801 | -117.8421 | 4.097 | -107.1671 | -114.1434 | -114.8163 | -0.32044 | -0.10071 | -0.28839 | -0.18768 | 9.7137 | 10.0342 | 1.7029 | 2.9846 | 2.0233 | 3.3051 | -0.15106 | -0.054932 | -0.23804 | -0.18311 | 10.524 | 10.675 | 2.8748 | 3.5477 | 3.0258 | 3.6987
2085.1719 | 122.2595 | -117.8146 | 4.2618 | -0.13817 | 0.015441 | 0.090976 | 0.086222 | 0.0010403 | 0.017684 | 0.011151 | -0.013519 | 0.080495 | 0.089487 | 0.078786 | 0.00045928 | 0.017673 | 0.014328 | 0.041199 | 4.2023 | -117.9611 | 4.0695 | -107.959 | -114.6698 | -115.9332 | 0.11444 | 4.2389 | -117.8421 | 4.0924 | -107.1533 | -114.1296 | -114.9445 | -0.14649 | -0.059509 | -0.19226 | -0.13275 | 9.8557 | 10.0021 | 1.8814 | 3.1448 | 2.0279 | 3.2913 | -0.027466 | -0.022888 | -0.16937 | -0.14649 | 10.6613 | 10.6888 | 2.8702 | 3.685 | 2.8976 | 3.7125
2086.1719 | 122.2778 | -117.8192 | 4.248 | -0.10002 | 0.014159 | 0.093326 | 0.083445 | -0.00066869 | 0.020674 | 0.013501 | 0.010895 | 0.079854 | 0.087137 | 0.077077 | -0.00060892 | 0.015323 | 0.0098417 | 0.045776 | 4.2023 | -117.9794 | 4.0878 | -107.9315 | -114.6469 | -115.9332 | 0.11902 | 4.2252 | -117.8329 | 4.1061 | -107.1396 | -114.1251 | -114.9216 | -0.16022 | -0.045776 | -0.16022 | -0.11444 | 9.8877 | 10.0479 | 1.886 | 3.1723 | 2.0462 | 3.3325 | -0.013733 | -0.022888 | -0.14191 | -0.11902 | 10.6796 | 10.6934 | 2.8976 | 3.6942 | 2.9114 | 3.7079
2087.1719 | 122.4014 | -117.9153 | 4.1656 | -0.077133 | 0.014586 | 0.090976 | 0.08195 | -0.00066869 | 0.018965 | 0.011151 | 0.010895 | 0.080495 | 0.086282 | 0.076436 | 0.0011001 | 0.015537 | 0.010055 | 0.013733 | 4.1107 | -118.0206 | 4.0466 | -107.9636 | -114.6927 | -115.947 | 0.14191 | 4.1428 | -117.8787 | 4.0558 | -107.1762 | -114.1434 | -115.1825 | -0.10529 | -0.054931 | -0.11902 | -0.064087 | 9.9518 | 10.0571 | 1.9684 | 3.2227 | 2.0737 | 3.3279 | 0.036621 | -0.022888 | -0.10986 | -0.086975 | 10.7391 | 10.7025 | 2.7328 | 3.772 | 2.6962 | 3.7354
2088.1719 | 122.3282 | -117.16 | 4.8431 | -0.12291 | 0.016295 | 0.090762 | 0.086009 | 0.0016813 | 0.022597 | 0.012433 | -0.0043641 | 0.079854 | 0.089273 | 0.080922 | -0.0014633 | 0.01575 | 0.011978 | 0.013733 | 4.8111 | -117.3981 | 4.7058 | -107.3273 | -114.0427 | -115.2832 | 0.14648 | 4.8431 | -117.2012 | 4.7333 | -106.5994 | -113.4888 | -114.2807 | -0.23804 | -0.032044 | -0.13733 | -0.10529 | 9.8328 | 10.0708 | 1.8768 | 3.1174 | 2.1149 | 3.3554 | -0.041199 | 0 | -0.10986 | -0.10986 | 10.5606 | 10.6018 | 2.8793 | 3.6713 | 2.9205 | 3.7125
2089.1719 | 122.2916 | -117.5949 | 4.454 | -0.12291 | 0.014159 | 0.093112 | 0.084941 | 0.0027493 | 0.017256 | 0.013501 | -0.028778 | 0.078359 | 0.08735 | 0.081136 | -0.00060892 | 0.015964 | 0.01561 | -0.027466 | 4.4037 | -117.7689 | 4.2709 | -107.7759 | -114.4913 | -115.7135 | 0.045776 | 4.4312 | -117.6086 | 4.3167 | -107.0114 | -113.9008 | -114.6881 | -0.17395 | -0.050354 | -0.18311 | -0.13275 | 9.819 | 9.993 | 1.8814 | 3.1036 | 2.0554 | 3.2776 | -0.013733 | -0.022889 | -0.13733 | -0.11444 | 10.5835 | 10.5972 | 2.9068 | 3.6942 | 2.9205 | 3.7079
2090.1719 | 122.2412 | -117.8558 | 4.2297 | -0.10002 | 0.015441 | 0.092898 | 0.086009 | -0.0023777 | 0.019606 | 0.014142 | -0.047089 | 0.079641 | 0.088205 | 0.079213 | -0.0014633 | 0.01511 | 0.014114 | 0.027466 | 4.1702 | -118.0115 | 4.0466 | -108.0048 | -114.7247 | -115.979 | 0.07782 | 4.2206 | -117.8833 | 4.0146 | -107.3273 | -114.1983 | -114.9857 | -0.15564 | -0.059509 | -0.18311 | -0.1236 | 9.8511 | 10.0067 | 1.8768 | 3.1311 | 2.0325 | 3.2867 | -0.027466 | -0.009155 | -0.21515 | -0.20599 | 10.5286 | 10.556 | 2.8702 | 3.6575 | 2.8976 | 3.685
2091.1719 | 122.2275 | -117.6498 | 4.4128 | -0.092392 | 0.013732 | 0.092685 | 0.083445 | 0.0016813 | 0.018752 | 0.012433 | -0.034882 | 0.080282 | 0.089273 | 0.082631 | 3.2078e-05 | 0.017032 | 0.013046 | 0.022888 | 4.3488 | -117.8192 | 4.2206 | -107.8217 | -114.5508 | -115.7822 | 0.10071 | 4.3945 | -117.659 | 4.2252 | -107.1442 | -114.0152 | -114.7934 | -0.16937 | -0.064087 | -0.19226 | -0.12817 | 9.8282 | 9.9976 | 1.8677 | 3.0991 | 2.037 | 3.2684 | -0.009156 | -0.018311 | -0.18768 | -0.16937 | 10.5057 | 10.5148 | 2.8564 | 3.6346 | 2.8656 | 3.6438
2092.1719 | 122.2458 | -117.5629 | 4.4998 | -0.13817 | 0.016082 | 0.092471 | 0.0843 | -0.0021641 | 0.018538 | 0.011578 | -0.05014 | 0.078359 | 0.090768 | 0.079 | 0.0013139 | 0.016391 | 0.013901 | 0.032043 | 4.4403 | -117.7231 | 4.3442 | -107.7438 | -114.4226 | -115.686 | 0.10529 | 4.4861 | -117.5812 | 4.3304 | -107.0938 | -113.9099 | -114.6927 | -0.16022 | -0.059509 | -0.15564 | -0.096131 | 9.819 | 9.9792 | 1.8768 | 3.1403 | 2.037 | 3.3005 | -0.018311 | -0.013733 | -0.16937 | -0.15564 | 10.4691 | 10.4874 | 2.8702 | 3.653 | 2.8885 | 3.6713
2093.1719 | 122.2641 | -117.3706 | 4.6829 | -0.061874 | 0.014373 | 0.091189 | 0.085368 | -0.00045509 | 0.016188 | 0.015851 | -0.053192 | 0.080495 | 0.087991 | 0.085195 | -0.0025315 | 0.012332 | 0.014114 | 0.032043 | 4.628 | -117.5171 | 4.5181 | -107.5653 | -114.2532 | -115.5029 | 0.10986 | 4.6783 | -117.3981 | 4.5181 | -106.9061 | -113.7085 | -114.5142 | -0.14648 | -0.054931 | -0.16479 | -0.10986 | 9.8053 | 9.9518 | 1.8677 | 3.1174 | 2.0142 | 3.2639 | -0.027466 | -0.004577 | -0.16479 | -0.16022 | 10.4645 | 10.4919 | 2.8564 | 3.6621 | 2.8839 | 3.6896
2094.1719 | 122.2549 | -117.4438 | 4.6417 | -0.10765 | 0.015654 | 0.092471 | 0.087718 | 0.00082671 | 0.021743 | 0.012433 | -0.15085 | 0.080495 | 0.088846 | 0.087972 | -0.0016769 | 0.017887 | 0.014969 | -0.004578 | 4.5731 | -117.6636 | 4.3854 | -107.6019 | -114.3494 | -115.6082 | 0.064087 | 4.6509 | -117.4622 | 4.4678 | -106.9977 | -113.8412 | -114.5874 | -0.21973 | -0.068665 | -0.25635 | -0.18768 | 9.8419 | 10.0616 | 1.8356 | 3.0945 | 2.0554 | 3.3142 | -0.01831 | 0.009155 | -0.17395 | -0.18311 | 10.4462 | 10.4645 | 2.8564 | 3.6026 | 2.8748 | 3.6209
2095.1719 | 122.2229 | -117.4438 | 4.5731 | -0.12291 | 0.0148 | 0.092044 | 0.083873 | -0.00045509 | 0.019179 | 0.010937 | -0.013519 | 0.079854 | 0.088846 | 0.080495 | 3.2078e-05 | 0.014041 | 0.014755 | 0.027466 | 4.5776 | -117.6086 | 4.4769 | -107.5424 | -114.2624 | -115.5075 | 0.13275 | 4.5685 | -117.4622 | 4.2755 | -106.8329 | -113.7634 | -114.5645 | -0.16479 | 0.004578 | -0.09613 | -0.10071 | 9.9014 | 10.0662 | 1.9363 | 3.1815 | 2.1011 | 3.3463 | -0.01831 | -0.004578 | -0.29755 | -0.29297 | 10.611 | 10.6293 | 2.8793 | 3.6804 | 2.8976 | 3.6987
2096.1719 | 122.3053 | -117.4026 | 4.6509 | -0.10765 | 0.015868 | 0.092685 | 0.085795 | 0.00082671 | 0.020034 | 0.014569 | -0.025726 | 0.078145 | 0.090555 | 0.085195 | 0.0011001 | 0.016391 | 0.01561 | 0.022888 | 4.5685 | -117.5629 | 4.5181 | -107.5333 | -114.2395 | -115.5121 | 0.11444 | 4.6555 | -117.4164 | 4.5364 | -106.9519 | -113.7451 | -114.505 | -0.16022 | -0.082398 | -0.13275 | -0.050354 | 9.8694 | 10.0296 | 1.8906 | 3.1631 | 2.0508 | 3.3234 | -0.013733 | 0.004578 | -0.11444 | -0.11902 | 10.4507 | 10.4645 | 2.8976 | 3.6575 | 2.9114 | 3.6713
2097.1719 | 122.2092 | -117.2882 | 4.8157 | -0.092392 | 0.013304 | 0.092898 | 0.087291 | 0.0018949 | 0.017684 | 0.011151 | -0.062347 | 0.079641 | 0.088419 | 0.077077 | -0.0016769 | 0.016818 | 0.013901 | 0.032043 | 4.7653 | -117.421 | 4.6234 | -107.4463 | -114.1617 | -115.4022 | 0.068665 | 4.8065 | -117.2882 | 4.5959 | -106.897 | -113.6948 | -114.4363 | -0.13275 | -0.050354 | -0.19226 | -0.14191 | 9.8419 | 9.9747 | 1.886 | 3.1265 | 2.0187 | 3.2593 | 0 | -0.009155 | -0.21973 | -0.21057 | 10.3912 | 10.3912 | 2.8519 | 3.5934 | 2.8519 | 3.5934
2098.1719 | 122.168 | -117.3935 | 4.747 | -0.061874 | 0.015227 | 0.092044 | 0.084727 | -0.0032321 | 0.017684 | 0.011365 | -0.089813 | 0.079854 | 0.091196 | 0.080709 | 0.00088648 | 0.017887 | 0.012619 | -0.013733 | 4.6692 | -117.5354 | 4.4403 | -107.3959 | -114.3356 | -115.5396 | 0.013733 | 4.7287 | -117.3935 | 4.4952 | -107.0526 | -113.8275 | -114.5599 | -0.14191 | -0.07782 | -0.3067 | -0.22888 | 9.9976 | 10.1395 | 1.8539 | 3.0579 | 1.9958 | 3.1998 | 0 | -0.01831 | -0.25177 | -0.23346 | 10.3409 | 10.3409 | 2.8336 | 3.566 | 2.8336 | 3.566
2099.1719 | 122.2046 | -117.3935 | 4.6829 | 0.060196 | 0.0148 | 0.093753 | 0.088145 | 0.0010403 | 0.019606 | 0.011578 | -0.1478 | 0.079427 | 0.088419 | 0.079641 | -0.00082252 | 0.017887 | 0.014755 | -0.14648 | 4.6875 | -117.572 | 4.4128 | -107.6935 | -114.418 | -115.6219 | -0.09613 | 4.6555 | -117.4118 | 4.4907 | -107.0618 | -113.8458 | -114.5645 | -0.17853 | 0.004578 | -0.27008 | -0.27466 | 9.7 | 9.8785 | 1.7715 | 2.9755 | 1.9501 | 3.154 | -0.01831 | -0.027465 | -0.19226 | -0.16479 | 10.3317 | 10.35 | 2.829 | 3.5477 | 2.8473 | 3.566
2100.1719 | 122.2412 | -117.2333 | 4.834 | -0.10002 | 0.016082 | 0.093753 | 0.082804 | -0.0040867 | 0.019179 | 0.010724 | -0.044037 | 0.077291 | 0.08735 | 0.077504 | -0.00060892 | 0.015537 | 0.018814 | 0.036621 | 4.7012 | -117.4301 | 4.6234 | -107.4051 | -114.1479 | -115.3793 | 0.10071 | 4.7699 | -117.2745 | 4.6646 | -106.8604 | -113.6215 | -114.386 | -0.19684 | -0.13275 | -0.21057 | -0.07782 | 9.8282 | 10.025 | 1.8539 | 3.0853 | 2.0508 | 3.2822 | -0.041199 | -0.064087 | -0.16937 | -0.10529 | 10.3729 | 10.4141 | 2.8473 | 3.6118 | 2.8885 | 3.653
2101.1875 | 122.1954 | -117.1646 | 4.9118 | -0.084763 | 0.015013 | 0.093966 | 0.088359 | 0.0018949 | 0.020247 | 0.013715 | -0.10812 | 0.080282 | 0.090341 | 0.082204 | -0.0016769 | 0.016818 | 0.016037 | -0.10071 | 4.9393 | -117.3248 | 4.6783 | -107.4326 | -114.1663 | -115.3702 | -0.11444 | 4.9438 | -117.1692 | 4.7516 | -106.8787 | -113.6398 | -114.3265 | -0.16022 | 0.027466 | -0.23346 | -0.26092 | 9.7321 | 9.8923 | 1.7944 | 2.9984 | 1.9547 | 3.1586 | -0.004577 | 0.032044 | -0.16022 | -0.19226 | 10.2859 | 10.2905 | 2.8381 | 3.5248 | 2.8427 | 3.5294
2102.1719 | 122.2733 | -117.3065 | 4.747 | -0.12291 | 0.012877 | 0.092257 | 0.083659 | -0.0013095 | 0.020674 | 0.011578 | -0.034882 | 0.079 | 0.086496 | 0.082204 | -0.0010361 | 0.014896 | 0.013901 | 0.041199 | 4.6921 | -117.453 | 4.5914 | -107.4005 | -114.1846 | -115.3885 | 0.10071 | 4.7607 | -117.3294 | 4.6051 | -106.9565 | -113.6993 | -114.4272 | -0.14648 | -0.054931 | -0.15564 | -0.10071 | 9.906 | 10.0525 | 1.918 | 3.1219 | 2.0645 | 3.2684 | -0.022888 | 0.013733 | -0.14191 | -0.15564 | 10.35 | 10.3729 | 2.8793 | 3.6072 | 2.9022 | 3.6301
2103.1719 | 122.1954 | -117.485 | 4.6005 | -0.061874 | 0.014373 | 0.090549 | 0.085154 | -0.0010959 | 0.020461 | 0.012433 | -0.059296 | 0.080068 | 0.088846 | 0.081991 | -0.0010361 | 0.014041 | 0.01326 | 0.013733 | 4.5776 | -117.6224 | 4.3488 | -107.634 | -114.4272 | -115.6448 | 0.036621 | 4.5914 | -117.485 | 4.3671 | -107.1991 | -113.9374 | -114.711 | -0.13733 | -0.022888 | -0.25177 | -0.22888 | 9.8511 | 9.9884 | 1.8402 | 3.0579 | 1.9775 | 3.1952 | 0 | -0.009155 | -0.23346 | -0.22431 | 10.2859 | 10.2859 | 2.774 | 3.5477 | 2.774 | 3.5477
2104.1719 | 122.2137 | -117.3523 | 4.7516 | -0.077133 | 0.014159 | 0.089908 | 0.085154 | 0.0016813 | 0.018325 | 0.013715 | -0.089813 | 0.080068 | 0.087778 | 0.0837 | 3.2078e-05 | 0.017032 | 0.011978 | -0.032043 | 4.7195 | -117.5262 | 4.4769 | -107.5745 | -114.3219 | -115.5396 | -0.013733 | 4.747 | -117.3569 | 4.5319 | -107.103 | -113.8138 | -114.5187 | -0.17395 | -0.032044 | -0.27466 | -0.24261 | 9.7778 | 9.9518 | 1.8127 | 3.0304 | 1.9867 | 3.2043 | -0.004578 | -0.004578 | -0.21973 | -0.21515 | 10.2493 | 10.2539 | 2.8336 | 3.5385 | 2.8381 | 3.5431
2105.1719 | 122.1817 | -117.6086 | 4.5273 | -0.084763 | 0.012877 | 0.093112 | 0.084086 | 0.00039951 | 0.017684 | 0.0096558 | -0.044037 | 0.078786 | 0.088632 | 0.078359 | 0.00067288 | 0.016391 | 0.014328 | 0.018311 | 4.454 | -117.7322 | 4.2801 | -107.7484 | -114.5187 | -115.7272 | 0.050354 | 4.4861 | -117.627 | 4.2572 | -107.3181 | -114.0564 | -114.7705 | -0.1236 | -0.073242 | -0.24719 | -0.17395 | 9.8602 | 9.9838 | 1.8814 | 3.0899 | 2.005 | 3.2135 | -0.01831 | -0.041199 | -0.27008 | -0.22888 | 10.2905 | 10.3088 | 2.8381 | 3.5522 | 2.8564 | 3.5706
2106.1719 | 122.2366 | -117.5903 | 4.454 | -0.1458 | 0.013518 | 0.091617 | 0.08665 | 0.0036039 | 0.019393 | 0.012433 | -0.12949 | 0.079641 | 0.090341 | 0.079427 | -0.00018152 | 0.018527 | 0.012619 | -0.10986 | 4.4266 | -117.8284 | 4.1977 | -107.872 | -114.6515 | -115.8325 | -0.09613 | 4.454 | -117.6407 | 4.2892 | -107.373 | -114.1068 | -114.7705 | -0.23804 | -0.027466 | -0.25635 | -0.22888 | 9.7183 | 9.9564 | 1.7578 | 2.9388 | 1.9958 | 3.1769 | -0.050354 | 0 | -0.16479 | -0.16479 | 10.2173 | 10.2676 | 2.8198 | 3.4836 | 2.8702 | 3.5339
2107.1719 | 122.2092 | -117.8787 | 4.2206 | -0.11528 | 0.014159 | 0.089694 | 0.082804 | -0.00088229 | 0.018538 | 0.011365 | -0.03183 | 0.080495 | 0.090768 | 0.078145 | 0.00067288 | 0.015537 | 0.012619 | 0.032043 | 4.1794 | -118.0023 | 4.0192 | -107.9865 | -114.7797 | -115.9836 | 0.054932 | 4.2068 | -117.8787 | 3.9871 | -107.579 | -114.3402 | -115.0497 | -0.1236 | -0.041199 | -0.20142 | -0.16022 | 9.8923 | 10.0159 | 1.8951 | 3.0991 | 2.0187 | 3.2227 | 0 | -0.013733 | -0.23346 | -0.21973 | 10.2997 | 10.2997 | 2.829 | 3.5385 | 2.829 | 3.5385
2108.1719 | 122.3053 | -117.7185 | 4.3167 | -0.023727 | 0.01715 | 0.095248 | 0.084513 | 0.00018591 | 0.020034 | 0.013287 | -0.019623 | 0.078359 | 0.088846 | 0.079427 | -0.0010361 | 0.01575 | 0.012832 | -0.068665 | 4.2435 | -117.897 | 4.1794 | -107.8537 | -114.6332 | -115.8279 | 0.13275 | 4.2892 | -117.7551 | 4.184 | -107.4051 | -114.1251 | -114.8529 | -0.17853 | -0.073242 | -0.13733 | -0.064087 | 9.8648 | 10.0433 | 1.8906 | 3.0853 | 2.0691 | 3.2639 | -0.036621 | -0.027465 | -0.13275 | -0.10529 | 10.3134 | 10.35 | 2.8656 | 3.5934 | 2.9022 | 3.6301
2109.1719 | 122.2641 | -117.6315 | 4.4449 | -0.11528 | 0.013945 | 0.090549 | 0.087504 | 0.00018591 | 0.018538 | 0.012646 | -0.10812 | 0.081136 | 0.088632 | 0.081136 | -0.00018152 | 0.016605 | 0.016891 | -0.004578 | 4.4083 | -117.8513 | 4.2068 | -107.8262 | -114.6011 | -115.8142 | 0.032043 | 4.4357 | -117.659 | 4.2801 | -107.3547 | -114.093 | -114.7522 | -0.21973 | -0.036621 | -0.23804 | -0.20142 | 9.8053 | 10.025 | 1.8173 | 3.0304 | 2.037 | 3.2501 | -0.027466 | -0.009155 | -0.16479 | -0.15564 | 10.2768 | 10.3043 | 2.8793 | 3.5385 | 2.9068 | 3.566
2110.1719 | 122.2137 | -117.334 | 4.7379 | -0.077133 | 0.016082 | 0.093539 | 0.085582 | -0.0047275 | 0.01747 | 0.012646 | -0.034882 | 0.079427 | 0.088205 | 0.0743 | 0.0013139 | 0.01575 | 0.014328 | 0.041199 | 4.6967 | -117.4988 | 4.5456 | -107.5287 | -114.2624 | -115.4526 | 0.10071 | 4.7195 | -117.3386 | 4.5593 | -107.0709 | -113.7634 | -114.473 | -0.16479 | -0.041199 | -0.19226 | -0.15106 | 9.8053 | 9.9701 | 1.8814 | 3.0716 | 2.0462 | 3.2364 | -0.004578 | -0.018311 | -0.17853 | -0.16022 | 10.2631 | 10.2676 | 2.861 | 3.5706 | 2.8656 | 3.5751
2111.1875 | 122.3007 | -117.366 | 4.6555 | -0.1992 | 0.015013 | 0.09418 | 0.086863 | -0.00024129 | 0.020034 | 0.014355 | -0.12643 | 0.078145 | 0.088632 | 0.079641 | -0.0012497 | 0.017887 | 0.01326 | -0.10529 | 4.6188 | -117.5812 | 4.4724 | -107.6202 | -114.3494 | -115.5533 | -0.022888 | 4.6326 | -117.4026 | 4.5639 | -107.103 | -113.8184 | -114.4821 | -0.21515 | -0.036622 | -0.18311 | -0.14648 | 9.7458 | 9.9609 | 1.8127 | 3.0167 | 2.0279 | 3.2318 | -0.036621 | -0.022889 | -0.091553 | -0.068664 | 10.2631 | 10.2997 | 2.8839 | 3.5477 | 2.9205 | 3.5843
2112.1719 | 122.1817 | -117.4484 | 4.6463 | -0.038986 | 0.012023 | 0.090976 | 0.085582 | -0.00088229 | 0.016829 | 0.011365 | -0.089813 | 0.080709 | 0.088205 | 0.083272 | -0.0014633 | 0.014682 | 0.011764 | -0.036621 | 4.6371 | -117.5995 | 4.4403 | -107.7484 | -114.4821 | -115.5991 | -0.041199 | 4.6921 | -117.4255 | 4.4174 | -107.2403 | -113.9694 | -114.6194 | -0.15106 | -0.009155 | -0.20599 | -0.19684 | 9.7 | 9.8511 | 1.8494 | 2.9663 | 2.0004 | 3.1174 | 0.022888 | 0.045777 | -0.22888 | -0.27466 | 10.2081 | 10.1852 | 2.829 | 3.479 | 2.8061 | 3.4561
2113.1719 | 122.1771 | -117.3431 | 4.7607 | -0.077133 | 0.015868 | 0.092471 | 0.082164 | -0.0030185 | 0.019606 | 0.014569 | -0.028778 | 0.078359 | 0.090768 | 0.080282 | 0.0013139 | 0.015537 | 0.01561 | 0 | 4.7058 | -117.4896 | 4.4998 | -107.5928 | -114.3127 | -115.4984 | 0.032043 | 4.747 | -117.3477 | 4.5181 | -107.135 | -113.855 | -114.5554 | -0.14648 | -0.054931 | -0.26092 | -0.20599 | 9.7504 | 9.8969 | 1.8448 | 3.0304 | 1.9913 | 3.1769 | -0.004577 | -0.013733 | -0.24261 | -0.22888 | 10.2081 | 10.2127 | 2.7878 | 3.4882 | 2.7924 | 3.4927
2114.1719 | 122.2092 | -117.2928 | 4.7836 | -0.13817 | 0.012023 | 0.092685 | 0.087504 | -0.0023777 | 0.017684 | 0.016492 | -0.12949 | 0.078359 | 0.089487 | 0.086263 | 0.00024568 | 0.018527 | 0.011764 | -0.009155 | 4.7379 | -117.4896 | 4.7699 | -107.5012 | -114.1434 | -115.4663 | 0.022888 | 4.8569 | -117.3065 | 4.6234 | -107.0709 | -113.8 | -114.4547 | -0.19684 | -0.045776 | -0.013733 | 0.032043 | 9.7916 | 9.9884 | 1.8265 | 3.1494 | 2.0233 | 3.3463 | -0.013733 | 0.073243 | -0.16022 | -0.23346 | 10.2219 | 10.2356 | 2.8381 | 3.4927 | 2.8519 | 3.5065
2115.1719 | 122.2595 | -117.2699 | 4.7836 | -0.22972 | 0.016722 | 0.093966 | 0.083018 | 0.0023221 | 0.021956 | 0.010083 | -0.022675 | 0.077077 | 0.090128 | 0.078786 | -0.0018905 | 0.015964 | 0.015396 | 0.054932 | 4.7241 | -117.4118 | 4.6463 | -107.4738 | -114.1754 | -115.3564 | 0.13733 | 4.7836 | -117.2516 | 4.6646 | -106.9565 | -113.6948 | -114.3814 | -0.14191 | -0.059509 | -0.13733 | -0.07782 | 9.7961 | 9.938 | 1.9135 | 3.0945 | 2.0554 | 3.2364 | 0.01831 | 0 | -0.11902 | -0.11902 | 10.3134 | 10.2951 | 2.8885 | 3.5751 | 2.8702 | 3.5568
2116.1719 | 122.2778 | -117.1371 | 4.8843 | -0.1458 | 0.016295 | 0.092471 | 0.086436 | 0.00039951 | 0.01982 | 0.011365 | -0.010467 | 0.078786 | 0.088205 | 0.078786 | -0.00082252 | 0.015323 | 0.013473 | -0.082397 | 4.7791 | -117.3981 | 4.7241 | -107.4783 | -114.1296 | -115.2924 | 0.13733 | 4.8386 | -117.215 | 4.7791 | -106.8604 | -113.6078 | -114.2807 | -0.26093 | -0.10528 | -0.16022 | -0.054932 | 9.6588 | 9.9197 | 1.8448 | 3.0075 | 2.1057 | 3.2684 | -0.07782 | -0.045776 | -0.10528 | -0.059509 | 10.2768 | 10.3546 | 2.8564 | 3.5294 | 2.9343 | 3.6072
2117.1719 | 122.2275 | -117.2516 | 4.8203 | -0.11528 | 0.013091 | 0.092471 | 0.082804 | 0.0021085 | 0.017897 | 0.011365 | -0.010467 | 0.078573 | 0.089273 | 0.082418 | 0.00045928 | 0.014682 | 0.016464 | 0.045776 | 4.8157 | -117.3615 | 4.5959 | -107.5653 | -114.212 | -115.3976 | 0.10986 | 4.7974 | -117.247 | 4.6692 | -106.9473 | -113.7177 | -114.3402 | -0.10986 | -0.004577 | -0.2243 | -0.21973 | 9.6863 | 9.7961 | 1.8539 | 3.0396 | 1.9638 | 3.1494 | 0.004578 | -0.022888 | -0.15106 | -0.12817 | 10.3043 | 10.2997 | 2.9114 | 3.5339 | 2.9068 | 3.5294
2118.1719 | 122.1909 | -117.7917 | 4.3121 | -0.077133 | 0.014586 | 0.088412 | 0.083659 | -0.0032321 | 0.020888 | 0.0088012 | -0.056244 | 0.079427 | 0.089487 | 0.077932 | -0.0010361 | 0.017246 | 0.013901 | -0.018311 | 4.2618 | -117.9474 | 4.0237 | -108.0963 | -114.8071 | -115.9561 | 0.022888 | 4.3167 | -117.8146 | 4.0649 | -107.5104 | -114.2944 | -114.9673 | -0.15564 | -0.050354 | -0.28839 | -0.23804 | 9.6954 | 9.8511 | 1.8356 | 2.9846 | 1.9913 | 3.1403 | -0.022888 | 0.004577 | -0.24719 | -0.25177 | 10.2814 | 10.3043 | 2.8244 | 3.4973 | 2.8473 | 3.5202
2119.1719 | 122.1817 | -117.2516 | 4.8477 | -0.12291 | 0.013945 | 0.092898 | 0.087077 | 0.00018591 | 0.016829 | 0.013074 | -0.13864 | 0.079427 | 0.088846 | 0.084768 | -0.0012497 | 0.016391 | 0.015823 | -0.086975 | 4.8111 | -117.4301 | 4.5959 | -107.5745 | -114.2853 | -115.2786 | -0.091553 | 4.8614 | -117.2607 | 4.6326 | -107.0343 | -113.8321 | -114.4455 | -0.17853 | -0.036621 | -0.25177 | -0.21515 | 9.6771 | 9.8557 | 1.973 | 2.9663 | 2.1515 | 3.1448 | -0.009155 | 0.013733 | -0.21515 | -0.22888 | 10.2173 | 10.2264 | 2.8061 | 3.4195 | 2.8152 | 3.4286
2120.1719 | 122.1909 | -117.4118 | 4.6875 | -0.10002 | 0.013518 | 0.090335 | 0.083018 | 0.00082671 | 0.018965 | 0.012646 | -0.03183 | 0.077291 | 0.088419 | 0.083486 | 0.0015275 | 0.014041 | 0.015182 | 0.013733 | 4.6509 | -117.5583 | 4.4403 | -107.7164 | -114.4135 | -115.5762 | 0.027466 | 4.6967 | -117.4164 | 4.4449 | -107.1487 | -113.9328 | -114.6057 | -0.14648 | -0.036621 | -0.24719 | -0.21057 | 9.6954 | 9.8419 | 1.8356 | 2.9984 | 1.9821 | 3.1448 | -0.004578 | 0.009155 | -0.24261 | -0.25177 | 10.2631 | 10.2676 | 2.8061 | 3.479 | 2.8107 | 3.4836
2121.1719 | 122.2092 | -117.3065 | 4.7836 | -0.13817 | 0.015013 | 0.092471 | 0.082377 | -0.0015231 | 0.016188 | 0.01286 | -0.013519 | 0.078573 | 0.089914 | 0.079641 | 3.2078e-05 | 0.016178 | 0.01625 | 0.022888 | 4.7562 | -117.4667 | 4.5547 | -107.5745 | -114.2807 | -115.4434 | 0.054932 | 4.7745 | -117.3386 | 4.5593 | -107.0023 | -113.8275 | -114.505 | -0.16022 | -0.027465 | -0.22888 | -0.20142 | 9.7321 | 9.8923 | 1.8631 | 3.0258 | 2.0233 | 3.186 | -0.032043 | -0.009155 | -0.2243 | -0.21515 | 10.3043 | 10.3363 | 2.8015 | 3.479 | 2.8336 | 3.511
2122.1719 | 122.2824 | -117.2424 | 4.7928 | -0.16869 | 0.015227 | 0.093539 | 0.085368 | 0.0014675 | 0.018965 | 0.012219 | -0.016571 | 0.078359 | 0.087778 | 0.083486 | -0.0014633 | 0.013614 | 0.013046 | 0.032043 | 4.747 | -117.4118 | 4.6646 | -107.46 | -114.2166 | -115.361 | 0.11444 | 4.7928 | -117.2699 | 4.6738 | -106.9382 | -113.7039 | -114.3997 | -0.16937 | -0.045777 | -0.12817 | -0.082397 | 9.7824 | 9.9518 | 1.8814 | 3.0258 | 2.0508 | 3.1952 | -0.027465 | 0 | -0.11902 | -0.11902 | 10.3043 | 10.3317 | 2.8427 | 3.5385 | 2.8702 | 3.566
2123.1719 | 122.2275 | -117.5446 | 4.5456 | -0.13817 | 0.012877 | 0.09183 | 0.0843 | -0.0021641 | 0.020674 | 0.011151 | -0.037933 | 0.077504 | 0.090341 | 0.086263 | 0.0013139 | 0.015323 | 0.01091 | 0.009155 | 4.4998 | -117.6956 | 4.3533 | -107.6935 | -114.505 | -115.6723 | 0.059509 | 4.5502 | -117.5171 | 4.3396 | -107.2449 | -114.0381 | -114.7293 | -0.15106 | -0.045776 | -0.19226 | -0.14648 | 9.8511 | 10.0021 | 1.8723 | 3.0396 | 2.0233 | 3.1906 | 0.027466 | 0.004578 | -0.20599 | -0.21057 | 10.2997 | 10.2722 | 2.8152 | 3.5065 | 2.7878 | 3.479
2124.1719 | 122.2366 | -117.3019 | 4.7882 | -0.092392 | 0.014586 | 0.090121 | 0.082804 | -0.0030185 | 0.020461 | 0.013715 | -0.059296 | 0.079427 | 0.090555 | 0.085195 | -0.0012497 | 0.01575 | 0.015182 | -0.041199 | 4.6921 | -117.5446 | 4.4449 | -107.7026 | -114.4226 | -115.5762 | 0.004578 | 4.715 | -117.4026 | 4.4724 | -107.135 | -113.7863 | -114.3814 | -0.24262 | -0.09613 | -0.34332 | -0.24719 | 9.5993 | 9.8419 | 1.7258 | 2.8793 | 1.9684 | 3.1219 | -0.10071 | -0.073242 | -0.31586 | -0.24262 | 10.1669 | 10.2676 | 2.9205 | 3.5156 | 3.0212 | 3.6163
2125.1719 | 122.1725 | -117.3752 | 4.747 | -0.16869 | 0.013945 | 0.091403 | 0.084513 | -0.0025913 | 0.017684 | 0.010724 | -0.065399 | 0.079427 | 0.089914 | 0.086049 | -0.0031723 | 0.01575 | 0.0087735 | -0.027466 | 4.7241 | -117.5308 | 4.4495 | -107.6431 | -114.4089 | -115.5624 | 0.004578 | 4.7516 | -117.3752 | 4.4861 | -107.1213 | -113.9511 | -114.5737 | -0.15564 | -0.022888 | -0.29755 | -0.27466 | 9.7321 | 9.8877 | 1.8127 | 2.9663 | 1.9684 | 3.1219 | 0 | 0.004578 | -0.26093 | -0.2655 | 10.2539 | 10.2539 | 2.8015 | 3.4241 | 2.8015 | 3.4241
2126.1719 | 122.1909 | -117.453 | 4.6509 | -0.084763 | 0.013945 | 0.091189 | 0.084086 | 0.00082671 | 0.017897 | 0.0098694 | -0.040985 | 0.078573 | 0.090768 | 0.082845 | -0.0010361 | 0.015323 | 0.012619 | -0.004578 | 4.5868 | -117.5995 | 4.3671 | -107.6797 | -114.4821 | -115.6357 | 0.022888 | 4.6509 | -117.485 | 4.4174 | -107.1716 | -114.006 | -114.6652 | -0.14648 | -0.064087 | -0.28381 | -0.21973 | 9.7733 | 9.9197 | 1.8173 | 2.9709 | 1.9638 | 3.1174 | -0.032043 | 0 | -0.23346 | -0.23346 | 10.2814 | 10.3134 | 2.7878 | 3.447 | 2.8198 | 3.479
2127.1719 | 122.2092 | -117.4118 | 4.715 | -0.084763 | 0.013091 | 0.090549 | 0.083445 | 0.0038175 | 0.018325 | 0.0098694 | -0.077606 | 0.077077 | 0.090768 | 0.079854 | 0.00045928 | 0.015537 | 0.011978 | -0.018311 | 4.6692 | -117.54 | 4.4403 | -107.5882 | -114.4501 | -115.5945 | -0.013733 | 4.7058 | -117.4301 | 4.4632 | -107.1075 | -113.9786 | -114.6103 | -0.12817 | -0.045777 | -0.27466 | -0.22888 | 9.8236 | 9.9518 | 1.8173 | 2.9617 | 1.9455 | 3.0899 | -0.018311 | -0.009155 | -0.25177 | -0.24261 | 10.3043 | 10.3226 | 2.8015 | 3.4332 | 2.8198 | 3.4515
2128.1719 | 122.1909 | -117.6224 | 4.4998 | -0.084763 | 0.014159 | 0.092257 | 0.084727 | 0.0010403 | 0.019606 | 0.011792 | -0.05014 | 0.078359 | 0.088419 | 0.081563 | 0.00045928 | 0.01511 | 0.014328 | 0 | 4.4678 | -117.7872 | 4.216 | -107.8079 | -114.6561 | -115.7776 | 0.022888 | 4.4998 | -117.6224 | 4.2435 | -107.2769 | -114.1388 | -114.83 | -0.16479 | -0.032044 | -0.28381 | -0.25177 | 9.8145 | 9.9792 | 1.8448 | 2.9663 | 2.0096 | 3.1311 | 0 | 0 | -0.25635 | -0.25635 | 10.3455 | 10.3455 | 2.7924 | 3.4836 | 2.7924 | 3.4836
2129.1719 | 122.2549 | -117.5262 | 4.5547 | -0.084763 | 0.015441 | 0.091617 | 0.085795 | 0.0016813 | 0.018752 | 0.013715 | -0.13864 | 0.079641 | 0.090768 | 0.08135 | 3.2078e-05 | 0.01511 | 0.015823 | 0.022888 | 4.4907 | -117.7231 | 4.3213 | -107.7164 | -114.5691 | -115.7135 | 0.045776 | 4.5547 | -117.5537 | 4.3945 | -107.1854 | -114.0839 | -114.6835 | -0.19684 | -0.064087 | -0.23346 | -0.16937 | 9.8099 | 10.0067 | 1.8127 | 2.9572 | 2.0096 | 3.154 | -0.027466 | 0 | -0.16022 | -0.16022 | 10.3409 | 10.3683 | 2.8427 | 3.4424 | 2.8702 | 3.4698
2130.1719 | 122.2092 | -117.5674 | 4.5502 | -0.077133 | 0.0148 | 0.09183 | 0.084727 | -0.00045509 | 0.017256 | 0.010937 | -0.034882 | 0.077504 | 0.0897 | 0.080709 | -0.00018152 | 0.014041 | 0.013046 | 0.013733 | 4.4998 | -117.691 | 4.2801 | -107.7576 | -114.6149 | -115.7455 | 0.018311 | 4.5319 | -117.5629 | 4.303 | -107.2495 | -114.1388 | -114.7568 | -0.1236 | -0.050354 | -0.27008 | -0.21973 | 9.8099 | 9.9335 | 1.8219 | 2.9526 | 1.9455 | 3.0762 | 0.004578 | -0.018311 | -0.24719 | -0.22888 | 10.318 | 10.3134 | 2.8107 | 3.4286 | 2.8061 | 3.4241
2131.1719 | 122.1909 | -117.3523 | 4.7699 | -0.038986 | 0.015654 | 0.091617 | 0.084513 | -2.7692e-05 | 0.017897 | 0.012006 | -0.044037 | 0.078786 | 0.090128 | 0.084768 | 3.2078e-05 | 0.01511 | 0.011123 | 0.013733 | 4.7379 | -117.4667 | 4.509 | -107.46 | -114.3585 | -115.48 | 0.027466 | 4.7745 | -117.3386 | 4.5227 | -107.0251 | -113.8596 | -114.5462 | -0.11444 | -0.032043 | -0.26093 | -0.22888 | 9.8923 | 10.0067 | 1.8723 | 2.9938 | 1.9867 | 3.1082 | 0.013733 | 0.004578 | -0.24719 | -0.25177 | 10.3271 | 10.3134 | 2.8061 | 3.4927 | 2.7924 | 3.479
2132.1719 | 122.1909 | -117.4759 | 4.6326 | -0.046616 | 0.014159 | 0.090549 | 0.084086 | 0.00039951 | 0.017043 | 0.0088012 | -0.056244 | 0.077932 | 0.089487 | 0.083059 | 0.0013139 | 0.015964 | 0.0098417 | -0.073242 | 4.6051 | -117.6407 | 4.5364 | -107.4738 | -114.6973 | -115.7684 | -0.10071 | 4.6738 | -117.4805 | 4.3945 | -107.1625 | -114.2075 | -114.679 | -0.16479 | -0.027465 | -0.09613 | -0.068665 | 10.0021 | 10.1669 | 1.7075 | 2.7786 | 1.8723 | 2.9434 | -0.004578 | 0.041199 | -0.23804 | -0.27924 | 10.3134 | 10.318 | 2.7969 | 3.2684 | 2.8015 | 3.273
2133.1719 | 122.1909 | -117.3798 | 4.7241 | -0.12291 | 0.013945 | 0.092044 | 0.085795 | 0.0012539 | 0.017684 | 0.013287 | -0.15695 | 0.078573 | 0.087778 | 0.088399 | 0.0011001 | 0.015964 | 0.012405 | -0.036621 | 4.6921 | -117.5766 | 4.422 | -107.2861 | -114.0976 | -115.5853 | 0.013733 | 4.7607 | -117.366 | 4.4632 | -107.0709 | -113.9694 | -114.5782 | -0.19684 | -0.032043 | -0.30212 | -0.27008 | 10.0937 | 10.2905 | 1.7944 | 3.2822 | 1.9913 | 3.479 | 0.013733 | 0.036621 | -0.26093 | -0.29755 | 10.3088 | 10.2951 | 2.8015 | 3.4103 | 2.7878 | 3.3966
2134.1719 | 122.1863 | -117.5125 | 4.5593 | -0.084763 | 0.015441 | 0.092898 | 0.08665 | -2.7692e-05 | 0.018965 | 0.011151 | -0.12033 | 0.079854 | 0.09205 | 0.0837 | 0.00045928 | 0.014896 | 0.012832 | -0.050354 | 4.5364 | -117.6682 | 4.3304 | -107.4142 | -114.6011 | -115.7318 | -0.059509 | 4.5822 | -117.5354 | 4.3579 | -107.2128 | -114.1159 | -114.7156 | -0.15564 | -0.022888 | -0.22888 | -0.20599 | 10.0983 | 10.2539 | 1.7807 | 2.9114 | 1.9363 | 3.067 | -0.022888 | 0.022888 | -0.20142 | -0.2243 | 10.2997 | 10.3226 | 2.7969 | 3.3966 | 2.8198 | 3.4195
2135.1719 | 122.1863 | -117.453 | 4.6875 | -0.13054 | 0.014159 | 0.092044 | 0.082591 | -0.00024129 | 0.016188 | 0.012646 | -0.022675 | 0.077718 | 0.090768 | 0.081136 | -0.0014633 | 0.014041 | 0.01561 | 0.032043 | 4.6509 | -117.5766 | 4.4632 | -107.3318 | -114.4318 | -115.5807 | 0.027466 | 4.6875 | -117.421 | 4.4449 | -107.0663 | -114.0152 | -114.6378 | -0.1236 | -0.036621 | -0.2243 | -0.18768 | 10.1212 | 10.2448 | 1.8723 | 3.0212 | 1.9958 | 3.1448 | 0.032044 | 0 | -0.24261 | -0.24261 | 10.3867 | 10.3546 | 2.8152 | 3.4378 | 2.7832 | 3.4058
2136.1719 | 122.2 | -117.5217 | 4.5868 | -0.084763 | 0.013304 | 0.090976 | 0.085795 | -2.7692e-05 | 0.018965 | 0.013928 | -0.028778 | 0.077504 | 0.090768 | 0.080709 | -0.00060892 | 0.013187 | 0.013473 | 0.013733 | 4.509 | -117.6819 | 4.3488 | -107.5562 | -114.5416 | -115.6677 | 0.032043 | 4.5959 | -117.5262 | 4.3533 | -107.1625 | -114.1525 | -114.711 | -0.16022 | -0.07782 | -0.23804 | -0.16022 | 9.9655 | 10.1257 | 1.8539 | 2.98 | 2.0142 | 3.1403 | -0.004578 | 0.009155 | -0.23346 | -0.24261 | 10.3592 | 10.3638 | 2.8107 | 3.3691 | 2.8152 | 3.3737
2137.1719 | 122.1405 | -117.6453 | 4.4769 | -0.084763 | 0.012877 | 0.092898 | 0.084513 | -0.00045509 | 0.019179 | 0.014569 | -0.05014 | 0.077718 | 0.087778 | 0.076009 | 0.0023819 | 0.014255 | 0.01561 | -0.018311 | 4.422 | -117.8284 | 4.1656 | -107.785 | -114.743 | -115.8188 | -0.004578 | 4.4357 | -117.6682 | 4.2023 | -107.3227 | -114.1846 | -114.8209 | -0.1831 | -0.054932 | -0.31128 | -0.25635 | 9.8602 | 10.0433 | 1.8265 | 2.9022 | 2.0096 | 3.0853 | -0.022888 | -0.041199 | -0.27466 | -0.23346 | 10.3226 | 10.3455 | 2.8244 | 3.4607 | 2.8473 | 3.4836
2138.1719 | 122.2183 | -117.5812 | 4.5593 | -0.092392 | 0.013304 | 0.089694 | 0.084086 | -0.0017367 | 0.018325 | 0.011365 | -0.077606 | 0.079 | 0.089914 | 0.079854 | -0.0021041 | 0.014255 | 0.017105 | -0.036621 | 4.4678 | -117.778 | 4.216 | -107.6889 | -114.6744 | -115.7684 | 0.004578 | 4.5136 | -117.5995 | 4.3945 | -107.1487 | -114.3219 | -114.8392 | -0.19684 | -0.091553 | -0.34332 | -0.25177 | 9.8923 | 10.0891 | 1.8127 | 2.9068 | 2.0096 | 3.1036 | -0.01831 | -0.045776 | -0.1648 | -0.11902 | 10.4324 | 10.4507 | 2.742 | 3.2593 | 2.7603 | 3.2776
2139.1719 | 122.2549 | -117.4576 | 4.6143 | -0.061874 | 0.0148 | 0.089053 | 0.086222 | 0.0016813 | 0.019179 | 0.011578 | -0.098968 | 0.078573 | 0.089914 | 0.087972 | -0.0023177 | 0.014896 | 0.015823 | 0 | 4.5776 | -117.6727 | 4.3533 | -107.5928 | -114.5416 | -115.6723 | 0.036621 | 4.6005 | -117.5034 | 4.4357 | -107.1671 | -114.0793 | -114.6698 | -0.21515 | -0.036621 | -0.26093 | -0.2243 | 9.8648 | 10.08 | 1.7853 | 2.916 | 2.0004 | 3.1311 | -0.045776 | -0.013733 | -0.17853 | -0.16479 | 10.2905 | 10.3363 | 2.7878 | 3.3783 | 2.8336 | 3.4241
2140.1719 | 122.2183 | -117.6544 | 4.4632 | -0.10002 | 0.013945 | 0.090549 | 0.083018 | -0.00024129 | 0.018111 | 0.01051 | -0.040985 | 0.077504 | 0.091196 | 0.081136 | -0.0018905 | 0.014896 | 0.012832 | -0.041199 | 4.4586 | -117.8009 | 4.1748 | -107.7713 | -114.7293 | -115.8554 | -0.009155 | 4.454 | -117.6453 | 4.216 | -107.3547 | -114.2349 | -114.8209 | -0.14648 | -0.004578 | -0.28839 | -0.28381 | 9.8831 | 10.0296 | 1.799 | 2.9251 | 1.9455 | 3.0716 | 0.009155 | -0.009155 | -0.24719 | -0.23804 | 10.2997 | 10.2905 | 2.8336 | 3.4195 | 2.8244 | 3.4103
2141.1719 | 122.2046 | -117.421 | 4.7058 | -0.1458 | 0.013518 | 0.092685 | 0.086222 | 0.00061311 | 0.016402 | 0.010297 | -0.05014 | 0.077932 | 0.090768 | 0.076223 | -0.00060892 | 0.013187 | 0.015182 | -0.036621 | 4.6967 | -117.5629 | 4.422 | -107.547 | -114.5187 | -115.6174 | -0.018311 | 4.7012 | -117.4072 | 4.4678 | -107.1167 | -114.0381 | -114.6011 | -0.14191 | -0.009156 | -0.28381 | -0.27466 | 9.874 | 10.0159 | 1.8036 | 2.9022 | 1.9455 | 3.0441 | 0.013732 | -0.004578 | -0.23804 | -0.23346 | 10.3043 | 10.2905 | 2.8198 | 3.3829 | 2.8061 | 3.3691
2142.1719 | 122.2229 | -117.3523 | 4.715 | -0.046616 | 0.011809 | 0.092685 | 0.083873 | 0.0018949 | 0.017043 | 0.012006 | -0.047089 | 0.077077 | 0.090555 | 0.082631 | 0.00024568 | 0.013828 | 0.01561 | -0.018311 | 4.6417 | -117.6086 | 4.39 | -107.5287 | -114.5233 | -115.6219 | 0.013733 | 4.6738 | -117.4347 | 4.4266 | -107.1167 | -114.0701 | -114.6332 | -0.25635 | -0.073242 | -0.32501 | -0.25177 | 9.8236 | 10.08 | 1.7303 | 2.829 | 1.9867 | 3.0853 | -0.082397 | -0.041199 | -0.28839 | -0.24719 | 10.2356 | 10.318 | 2.7191 | 3.2822 | 2.8015 | 3.3646
2143.1719 | 122.1771 | -117.3615 | 4.747 | -0.077133 | 0.012877 | 0.08884 | 0.083445 | -0.0021641 | 0.018965 | 0.0096558 | -0.016571 | 0.078145 | 0.091409 | 0.078359 | 0.00067288 | 0.014682 | 0.013046 | -0.041199 | 4.6875 | -117.5125 | 4.454 | -107.2083 | -114.4959 | -115.5899 | -0.004578 | 4.7287 | -117.3889 | 4.4769 | -107.0892 | -114.0152 | -114.6103 | -0.15106 | -0.059509 | -0.29297 | -0.23346 | 10.1532 | 10.3043 | 1.7715 | 2.8656 | 1.9226 | 3.0167 | -0.027466 | -0.01831 | -0.27008 | -0.25177 | 10.2722 | 10.2997 | 2.7512 | 3.3463 | 2.7786 | 3.3737
2144.1719 | 122.2046 | -117.5308 | 4.5639 | -0.1458 | 0.014586 | 0.092471 | 0.087077 | -0.00045509 | 0.018752 | 0.011578 | -0.12033 | 0.079427 | 0.0897 | 0.083059 | 0.00045928 | 0.014469 | 0.014541 | -0.022888 | 4.5364 | -117.7139 | 4.335 | -107.3776 | -114.6378 | -115.773 | -0.009155 | 4.5547 | -117.5446 | 4.3579 | -107.254 | -114.1525 | -114.7202 | -0.1831 | -0.027466 | -0.22888 | -0.20142 | 10.1532 | 10.3363 | 1.7578 | 2.8931 | 1.9409 | 3.0762 | -0.013733 | -0.009155 | -0.20599 | -0.19684 | 10.2768 | 10.2905 | 2.8107 | 3.3783 | 2.8244 | 3.392
2145.1719 | 122.1863 | -117.659 | 4.4449 | -0.038986 | 0.015013 | 0.088199 | 0.0843 | 0.0021085 | 0.017897 | 0.010297 | -0.037933 | 0.077932 | 0.090128 | 0.079427 | 0.00045928 | 0.013401 | 0.012832 | 0.19684 | 4.4403 | -117.746 | 4.1611 | -107.6111 | -114.7522 | -115.8737 | -0.022888 | 4.454 | -117.659 | 4.2068 | -107.3547 | -114.3127 | -114.8758 | -0.086975 | -0.004577 | -0.28381 | -0.27924 | 10.0479 | 10.1349 | 1.7853 | 2.9068 | 1.8723 | 2.9938 | 0 | 0.009156 | -0.23804 | -0.24719 | 10.3043 | 10.3043 | 2.7832 | 3.3463 | 2.7832 | 3.3463
2146.1719 | 122.2366 | -117.5079 | 4.6051 | -0.054245 | 0.015013 | 0.09418 | 0.083873 | -2.7692e-05 | 0.014907 | 0.010724 | -0.068451 | 0.080068 | 0.090982 | 0.085195 | -0.00060892 | 0.016818 | 0.016464 | -0.027466 | 4.5593 | -117.6865 | 4.3121 | -107.46 | -114.6011 | -115.7318 | 0 | 4.5822 | -117.5034 | 4.39 | -107.1625 | -114.1617 | -114.679 | -0.17853 | -0.045777 | -0.29297 | -0.24719 | 10.0479 | 10.2264 | 1.7761 | 2.9068 | 1.9547 | 3.0853 | 0.004578 | -0.022889 | -0.21515 | -0.19226 | 10.3455 | 10.3409 | 2.829 | 3.3463 | 2.8244 | 3.3417
2147.1719 | 122.2137 | -117.4072 | 4.6692 | -0.13054 | 0.0148 | 0.092471 | 0.083232 | -0.0045139 | 0.017897 | 0.010937 | -0.056244 | 0.078573 | 0.089273 | 0.080495 | 0.00024568 | 0.014896 | 0.012192 | -0.027466 | 4.6875 | -117.5858 | 4.4037 | -107.3227 | -114.505 | -115.6403 | 0.004578 | 4.6738 | -117.4164 | 4.4861 | -107.0572 | -114.0198 | -114.5966 | -0.17853 | 0.018311 | -0.2655 | -0.28381 | 10.0845 | 10.2631 | 1.767 | 2.9022 | 1.9455 | 3.0807 | -0.009155 | 0.004578 | -0.18311 | -0.18768 | 10.35 | 10.3592 | 2.8107 | 3.3875 | 2.8198 | 3.3966
2148.1719 | 122.2549 | -117.2379 | 4.8386 | -0.092392 | 0.012877 | 0.09183 | 0.08665 | 0.0010403 | 0.017256 | 0.01286 | -0.03183 | 0.078145 | 0.090128 | 0.083913 | -0.0014633 | 0.013828 | 0.013046 | -0.022888 | 4.7241 | -117.5217 | 4.4998 | -107.309 | -114.4272 | -115.567 | 0.009155 | 4.7653 | -117.3203 | 4.6371 | -106.8878 | -113.9236 | -114.4318 | -0.28381 | -0.11444 | -0.33874 | -0.2243 | 9.9289 | 10.2127 | 1.6708 | 2.8107 | 1.9546 | 3.0945 | -0.082397 | -0.073242 | -0.20142 | -0.12817 | 10.35 | 10.4324 | 2.8061 | 3.3142 | 2.8885 | 3.3966
2149.1719 | 122.1771 | -117.9428 | 4.1611 | -0.10002 | 0.013091 | 0.09653 | 0.085582 | 0.0038175 | 0.017043 | 0.013287 | -0.11423 | 0.079213 | 0.091409 | 0.081136 | -0.00082252 | 0.016178 | 0.014755 | -0.045776 | 4.1107 | -118.1442 | 3.8956 | -107.8583 | -115.0772 | -116.2033 | -0.036621 | 4.1656 | -117.9565 | 3.9413 | -107.5836 | -114.6011 | -115.1505 | -0.20142 | -0.050354 | -0.2655 | -0.21515 | 10.0845 | 10.2859 | 1.7395 | 2.8656 | 1.9409 | 3.067 | -0.013733 | 0.004577 | -0.21973 | -0.2243 | 10.3592 | 10.3729 | 2.7924 | 3.3417 | 2.8061 | 3.3554
2150.1719 | 122.1771 | -117.8009 | 4.3396 | -0.092392 | 0.013518 | 0.093539 | 0.084513 | -0.00066869 | 0.018325 | 0.012433 | -0.034882 | 0.079 | 0.089487 | 0.082631 | 0.0023819 | 0.014896 | 0.013046 | -0.050354 | 4.2755 | -117.984 | 4.0009 | -107.7026 | -114.9353 | -116.0385 | -0.009155 | 4.2801 | -117.8192 | 4.0421 | -107.4371 | -114.4043 | -114.9765 | -0.18311 | -0.064087 | -0.33875 | -0.27466 | 10.0983 | 10.2814 | 1.7624 | 2.8656 | 1.9455 | 3.0487 | -0.018311 | -0.05951 | -0.29755 | -0.23804 | 10.3638 | 10.3821 | 2.8244 | 3.3966 | 2.8427 | 3.4149
2151.1719 | 122.2137 | -117.485 | 4.6097 | -0.11528 | 0.013091 | 0.093326 | 0.083445 | 0.00018591 | 0.01747 | 0.008374 | -0.05014 | 0.077932 | 0.090982 | 0.08135 | -0.0014633 | 0.013187 | 0.015182 | -0.032043 | 4.6051 | -117.6407 | 4.3579 | -107.2495 | -114.6149 | -115.7227 | 0 | 4.6234 | -117.4667 | 4.3854 | -107.0618 | -114.1296 | -114.6652 | -0.15564 | -0.004577 | -0.25177 | -0.24719 | 10.2356 | 10.3912 | 1.7624 | 2.8702 | 1.918 | 3.0258 | 0.01831 | 0.013733 | -0.2243 | -0.23804 | 10.4233 | 10.405 | 2.8198 | 3.3554 | 2.8015 | 3.3371
2152.1719 | 122.2183 | -117.5491 | 4.5731 | -0.084763 | 0.013091 | 0.091403 | 0.083873 | 0.00082671 | 0.016188 | 0.010297 | -0.03183 | 0.080282 | 0.089059 | 0.076223 | -0.0012497 | 0.013828 | 0.012192 | -0.036621 | 4.5364 | -117.7368 | 4.2664 | -107.515 | -114.6973 | -115.7913 | -0.018311 | 4.5364 | -117.5446 | 4.3167 | -107.1487 | -114.2303 | -114.7476 | -0.18768 | -0.036621 | -0.3067 | -0.27008 | 10.0342 | 10.2219 | 1.7578 | 2.8519 | 1.9455 | 3.0396 | 0.004577 | -0.036621 | -0.25635 | -0.21973 | 10.4004 | 10.3958 | 2.8015 | 3.3188 | 2.7969 | 3.3142
2153.1719 | 122.2275 | -117.3935 | 4.6463 | -0.11528 | 0.012236 | 0.092471 | 0.084941 | 0.00061311 | 0.016829 | 0.0092286 | -0.074554 | 0.077077 | 0.090341 | 0.085622 | -0.0029587 | 0.014896 | 0.016464 | 0.009155 | 4.6051 | -117.6178 | 4.4083 | -107.254 | -114.5599 | -115.654 | 0.059509 | 4.6738 | -117.421 | 4.4815 | -107.0526 | -114.0747 | -114.624 | -0.2243 | -0.041198 | -0.23804 | -0.19684 | 10.1395 | 10.3638 | 1.7395 | 2.8336 | 1.9638 | 3.0579 | -0.027465 | 0.027466 | -0.16479 | -0.19226 | 10.3409 | 10.3683 | 2.7695 | 3.3188 | 2.7969 | 3.3463
2154.1719 | 122.1725 | -117.6132 | 4.4769 | -0.1458 | 0.01245 | 0.092257 | 0.083232 | -0.0021641 | 0.018538 | 0.012433 | -0.092865 | 0.077291 | 0.089273 | 0.081991 | 0.0011001 | 0.017246 | 0.012405 | -0.082397 | 4.4128 | -117.8558 | 4.1428 | -107.666 | -114.8071 | -115.9287 | -0.11444 | 4.4724 | -117.6453 | 4.2206 | -107.3135 | -114.3219 | -114.8712 | -0.24262 | -0.064087 | -0.33417 | -0.27008 | 9.9472 | 10.1898 | 1.6846 | 2.8061 | 1.9272 | 3.0487 | -0.032044 | -0.004578 | -0.25635 | -0.25177 | 10.2997 | 10.3317 | 2.742 | 3.2913 | 2.774 | 3.3234
2155.1719 | 122.2458 | -117.572 | 4.4952 | -0.031357 | 0.014373 | 0.093966 | 0.084513 | -0.0025913 | 0.01747 | 0.010724 | -0.044037 | 0.079 | 0.087778 | 0.077291 | 0.00024568 | 0.016605 | 0.014328 | -0.004578 | 4.4724 | -117.7597 | 4.248 | -107.4097 | -114.7156 | -115.8279 | 0.059509 | 4.509 | -117.5858 | 4.3076 | -107.1854 | -114.2441 | -114.7751 | -0.18768 | -0.022888 | -0.24719 | -0.2243 | 10.1624 | 10.35 | 1.7441 | 2.8564 | 1.9318 | 3.0441 | -0.013733 | 0.013733 | -0.18768 | -0.20142 | 10.3867 | 10.4004 | 2.7969 | 3.3279 | 2.8107 | 3.3417
2156.1719 | 122.2137 | -117.4347 | 4.6738 | -0.12291 | 0.012023 | 0.090976 | 0.084513 | -0.0049411 | 0.018538 | 0.012219 | -0.056244 | 0.079213 | 0.089914 | 0.079 | -0.0014633 | 0.01511 | 0.014755 | -0.050354 | 4.6463 | -117.6132 | 4.3762 | -107.3822 | -114.5966 | -115.7043 | -0.013733 | 4.6417 | -117.4484 | 4.422 | -107.048 | -114.1434 | -114.6286 | -0.17853 | -0.027466 | -0.29755 | -0.27008 | 10.0525 | 10.231 | 1.7303 | 2.8381 | 1.9089 | 3.0167 | -0.013733 | -0.032043 | -0.25177 | -0.21973 | 10.3867 | 10.4004 | 2.8061 | 3.2913 | 2.8198 | 3.3051
2157.1719 | 122.2412 | -117.4484 | 4.6371 | -0.11528 | 0.014586 | 0.092685 | 0.083232 | 0.0014675 | 0.019179 | 0.011578 | -0.037933 | 0.078145 | 0.088846 | 0.083272 | 0.0011001 | 0.013828 | 0.014328 | -0.027466 | 4.6005 | -117.6178 | 4.3533 | -107.3914 | -114.6286 | -115.7364 | -0.022888 | 4.6234 | -117.4438 | 4.4128 | -107.0435 | -114.2303 | -114.6698 | -0.16937 | -0.036621 | -0.28381 | -0.24719 | 10.0571 | 10.2264 | 1.712 | 2.8198 | 1.8814 | 2.9892 | 0.004577 | -0.013733 | -0.2243 | -0.21057 | 10.405 | 10.4004 | 2.7786 | 3.2181 | 2.774 | 3.2135
2158.1719 | 122.2092 | -117.6361 | 4.4586 | -0.12291 | 0.014373 | 0.09183 | 0.085795 | -0.00066869 | 0.017256 | 0.01051 | -0.065399 | 0.079213 | 0.090341 | 0.081991 | -0.0018905 | 0.01575 | 0.014755 | -0.013733 | 4.422 | -117.8146 | 4.2023 | -107.46 | -114.8026 | -115.892 | 0.068665 | 4.4449 | -117.6453 | 4.2526 | -107.1899 | -114.2258 | -114.798 | -0.17853 | -0.036621 | -0.25635 | -0.21973 | 10.1761 | 10.3546 | 1.7441 | 2.8336 | 1.9226 | 3.0121 | -0.009156 | -0.013733 | -0.20599 | -0.19226 | 10.4462 | 10.4553 | 2.8381 | 3.4103 | 2.8473 | 3.4195
2159.1719 | 122.1725 | -117.7551 | 4.3442 | -0.1458 | 0.014159 | 0.094821 | 0.084513 | 0.0068081 | 0.022383 | 0.011792 | -0.092865 | 0.077291 | 0.090128 | 0.08434 | -0.0021041 | 0.016391 | 0.011337 | -0.073242 | 4.2892 | -117.952 | 4.0329 | -107.753 | -114.9994 | -116.0477 | -0.12817 | 4.3579 | -117.7917 | 4.0741 | -107.4142 | -114.4318 | -114.9857 | -0.19684 | -0.054931 | -0.31128 | -0.25635 | 10.0021 | 10.199 | 1.7075 | 2.7557 | 1.9043 | 2.9526 | -0.036621 | 0.013733 | -0.27008 | -0.28381 | 10.3409 | 10.3775 | 2.7695 | 3.3234 | 2.8061 | 3.36
2160.1719 | 121.875 | -51.8463 | 70.2438 | -0.12291 | 0.018218 | 0.082004 | 0.089427 | 0.00018591 | 0.015334 | 0.0092286 | -0.047089 | 0.020254 | 0.080728 | 0.081563 | -0.0018905 | 0.014469 | 0.0098417 | 0.027466 | 70.1431 | -51.9745 | 70.0607 | -46.3211 | -49.7314 | -50.5692 | 0.068665 | 70.1752 | -51.8188 | 70.1843 | -45.639 | -49.2462 | -49.9878 | -0.12817 | -0.10071 | -0.1831 | -0.082397 | 5.5252 | 5.6534 | 1.2772 | 2.1149 | 1.4053 | 2.243 | 0.027465 | -0.068664 | -0.059509 | 0.009155 | 6.2073 | 6.1798 | 1.8585 | 2.6001 | 1.8311 | 2.5726
2161.1719 | 122.2 | -51.7227 | 70.3033 | -0.092392 | 0.019713 | 0.08414 | 0.088786 | 0.0029629 | 0.016829 | 0.010083 | -0.053192 | 0.023031 | 0.080301 | 0.08669 | -0.0016769 | 0.011478 | 0.010055 | -0.032043 | 70.1797 | -51.9745 | 70.1294 | -46.2662 | -49.6857 | -50.5417 | 0.054932 | 70.2209 | -51.7776 | 70.3445 | -45.6436 | -49.0997 | -49.9558 | -0.25177 | -0.1236 | -0.17395 | -0.050354 | 5.4565 | 5.7083 | 1.181 | 2.037 | 1.4328 | 2.2888 | -0.054932 | -0.082398 | 0.041198 | 0.1236 | 6.0791 | 6.134 | 1.767 | 2.623 | 1.8219 | 2.6779
2162.1875 | 122.3557 | -111.6714 | 10.2493 | -0.16869 | 0.015441 | 0.095248 | 0.083232 | -0.0021641 | 0.017684 | 0.0085876 | 0.0047913 | 0.07195 | 0.089059 | 0.076223 | -0.00018152 | 0.01276 | 0.014969 | -0.068665 | 10.2356 | -111.8866 | 10.1807 | -102.0813 | -108.9294 | -109.9915 | -0.091553 | 10.2631 | -111.7722 | 10.1944 | -101.825 | -108.2382 | -109.021 | -0.21515 | -0.013733 | -0.068665 | -0.054932 | 9.5901 | 9.8053 | 1.68 | 2.742 | 1.8951 | 2.9572 | -0.10071 | 0.013733 | -0.054932 | -0.068665 | 9.8465 | 9.9472 | 2.6505 | 3.4332 | 2.7512 | 3.5339
2163.1719 | 122.2961 | -113.6581 | 8.3633 | -0.12291 | 0.013304 | 0.092044 | 0.087718 | 0.0025357 | 0.01747 | 0.01286 | -0.053192 | 0.073232 | 0.091623 | 0.084768 | -0.00060892 | 0.016391 | 0.014755 | -0.022888 | 8.3542 | -113.8596 | 8.1665 | -103.9536 | -110.8887 | -111.9644 | -0.009155 | 8.3908 | -113.7039 | 8.2214 | -103.5965 | -110.3531 | -110.9573 | -0.20142 | -0.009155 | -0.19684 | -0.18768 | 9.7046 | 9.906 | 1.6937 | 2.7695 | 1.8951 | 2.9709 | -0.045776 | 0.027466 | -0.14191 | -0.16937 | 10.0616 | 10.1074 | 2.7008 | 3.3051 | 2.7466 | 3.3508
2164.1719 | 122.2412 | -113.9053 | 8.1299 | -0.13054 | 0.014586 | 0.090549 | 0.085368 | 0.0036039 | 0.019179 | 0.01051 | -0.080658 | 0.074514 | 0.091837 | 0.083272 | -0.00082252 | 0.013401 | 0.011123 | -0.10986 | 8.1207 | -114.0839 | 7.9102 | -104.1367 | -111.2274 | -112.2482 | -0.11902 | 8.1299 | -113.9557 | 7.9834 | -103.8757 | -110.7147 | -111.2091 | -0.17853 | -0.009155 | -0.21973 | -0.21057 | 9.7687 | 9.9472 | 1.6571 | 2.6779 | 1.8356 | 2.8564 | -0.050354 | 0 | -0.14648 | -0.14648 | 10.0296 | 10.08 | 2.6962 | 3.1906 | 2.7466 | 3.241
2165.1719 | 122.2733 | -114.1937 | 7.8369 | -0.10765 | 0.014586 | 0.094394 | 0.085154 | -0.00066869 | 0.018111 | 0.010724 | -0.074554 | 0.075155 | 0.089914 | 0.081136 | 3.2078e-05 | 0.014896 | 0.012832 | 0 | 7.8323 | -114.3723 | 7.663 | -104.3243 | -111.4288 | -112.4863 | 0.036621 | 7.8598 | -114.2212 | 7.7179 | -104.0726 | -110.9116 | -111.47 | -0.17853 | -0.004578 | -0.17395 | -0.16937 | 9.8694 | 10.0479 | 1.7075 | 2.7649 | 1.886 | 2.9434 | -0.027465 | 0.022888 | -0.11902 | -0.14191 | 10.1212 | 10.1486 | 2.7237 | 3.2822 | 2.7512 | 3.3096
2166.1719 | 122.2458 | -114.7797 | 7.3059 | -0.084763 | 0.0148 | 0.093112 | 0.085582 | 0.0014675 | 0.017684 | 0.012433 | -0.065399 | 0.076864 | 0.087778 | 0.081991 | -0.0010361 | 0.013187 | 0.011337 | 0 | 7.2464 | -114.9353 | 7.1136 | -104.8187 | -111.9919 | -113.0402 | 0.059509 | 7.3151 | -114.7705 | 7.164 | -104.5807 | -111.5341 | -112.001 | -0.15564 | -0.059509 | -0.19226 | -0.13275 | 9.9609 | 10.1166 | 1.7395 | 2.7878 | 1.8951 | 2.9434 | 0.009155 | 0.009155 | -0.14191 | -0.15106 | 10.199 | 10.1898 | 2.7786 | 3.2455 | 2.7695 | 3.2364
2167.1719 | 122.287 | -114.418 | 7.608 | -0.092392 | 0.01245 | 0.093539 | 0.084941 | 0.00039951 | 0.017684 | 0.011578 | -0.062347 | 0.076009 | 0.091623 | 0.082204 | 3.2078e-05 | 0.015323 | 0.016037 | -0.10071 | 7.5806 | -114.6423 | 7.3975 | -104.6082 | -111.7264 | -112.7518 | -0.09613 | 7.5943 | -114.473 | 7.5027 | -104.2648 | -111.2091 | -111.6623 | -0.2243 | -0.027466 | -0.21057 | -0.18311 | 9.8099 | 10.0342 | 1.6663 | 2.6917 | 1.8906 | 2.916 | -0.054931 | -0.013733 | -0.10528 | -0.091552 | 10.1532 | 10.2081 | 2.7557 | 3.2089 | 2.8107 | 3.2639
2168.1719 | 122.2366 | -114.6469 | 7.4203 | -0.10765 | 0.015013 | 0.094394 | 0.087718 | -0.00088229 | 0.017043 | 0.014996 | -0.092865 | 0.07665 | 0.089273 | 0.079213 | 0.00088648 | 0.012973 | 0.0083463 | -0.004578 | 7.3654 | -114.8529 | 7.2006 | -104.7043 | -111.9186 | -112.9486 | 0.036621 | 7.402 | -114.5828 | 7.2601 | -104.4388 | -111.4334 | -111.9141 | -0.20599 | -0.054932 | -0.21973 | -0.16479 | 9.9426 | 10.1486 | 1.6983 | 2.7283 | 1.9043 | 2.9343 | 0.064087 | -0.01831 | -0.16022 | -0.14191 | 10.2081 | 10.144 | 2.7328 | 3.2135 | 2.6688 | 3.1494
2169.1719 | 122.2046 | -114.8209 | 7.2876 | -0.13817 | 0.012877 | 0.093966 | 0.086436 | 0.0012539 | 0.017256 | 0.0090148 | -0.089813 | 0.075582 | 0.088419 | 0.084554 | 0.00088648 | 0.013187 | 0.017319 | -0.10529 | 7.2464 | -114.9902 | 6.9992 | -104.837 | -112.1429 | -113.1363 | -0.004578 | 7.2556 | -114.7705 | 7.0587 | -104.5715 | -111.6531 | -112.0926 | -0.16937 | -0.041199 | -0.28839 | -0.24719 | 9.9838 | 10.1532 | 1.6846 | 2.6779 | 1.8539 | 2.8473 | 0.050354 | -0.032044 | -0.22888 | -0.19684 | 10.2493 | 10.199 | 2.7283 | 3.1677 | 2.6779 | 3.1174
2170.1719 | 122.2366 | -114.743 | 7.3242 | -0.13817 | 0.014586 | 0.094607 | 0.083659 | -0.0015231 | 0.018965 | 0.011365 | -0.062347 | 0.075582 | 0.088205 | 0.083486 | -0.00082252 | 0.013187 | 0.017959 | -0.027466 | 7.338 | -114.8895 | 7.0816 | -104.8691 | -111.9827 | -113.0356 | 0.013733 | 7.3288 | -114.743 | 7.1777 | -104.5349 | -111.4746 | -111.9736 | -0.14648 | 0.013733 | -0.24262 | -0.25635 | 9.874 | 10.0204 | 1.7075 | 2.7603 | 1.8539 | 2.9068 | 0 | 0.004577 | -0.14649 | -0.15106 | 10.2081 | 10.2081 | 2.7695 | 3.2684 | 2.7695 | 3.2684
2171.1719 | 122.2321 | -114.8666 | 7.196 | -0.084763 | 0.014373 | 0.091403 | 0.085795 | -0.00024129 | 0.018325 | 0.011365 | -0.047089 | 0.076009 | 0.089059 | 0.08434 | -0.0018905 | 0.016391 | 0.01326 | -0.027466 | 7.1915 | -115.0589 | 6.9489 | -104.8233 | -112.1246 | -113.1592 | 0.004578 | 7.2098 | -114.8849 | 7.0267 | -104.6631 | -111.5799 | -112.1063 | -0.19226 | -0.004578 | -0.24719 | -0.24261 | 10.0433 | 10.2356 | 1.7075 | 2.742 | 1.8997 | 2.9343 | -0.018311 | 0.013733 | -0.16937 | -0.18311 | 10.2036 | 10.2219 | 2.7603 | 3.2867 | 2.7786 | 3.3051
2172.1719 | 122.2916 | -114.8621 | 7.1686 | -0.16106 | 0.011809 | 0.094394 | 0.085582 | -0.0013095 | 0.017256 | 0.0098694 | -0.040985 | 0.074941 | 0.088205 | 0.079 | -0.00039512 | 0.014896 | 0.016037 | -0.091553 | 7.1457 | -115.0314 | 6.9489 | -104.9286 | -112.1796 | -113.205 | -0.073242 | 7.1594 | -114.8895 | 7.0724 | -104.6082 | -111.6165 | -112.0834 | -0.16937 | -0.022888 | -0.21973 | -0.19684 | 9.9335 | 10.1028 | 1.6571 | 2.6825 | 1.8265 | 2.8519 | -0.027465 | -0.009155 | -0.09613 | -0.086975 | 10.2539 | 10.2814 | 2.7786 | 3.2455 | 2.8061 | 3.273
2173.1719 | 122.2321 | -114.798 | 7.2601 | -0.12291 | 0.014373 | 0.094607 | 0.084086 | 0.0057401 | 0.019179 | 0.012646 | -0.047089 | 0.074941 | 0.090341 | 0.079 | -0.0018905 | 0.013614 | 0.01326 | -0.009155 | 7.2372 | -114.9765 | 7.0541 | -104.7501 | -112.0651 | -113.0951 | 0.036621 | 7.2601 | -114.8117 | 7.1457 | -104.5761 | -111.6211 | -112.0514 | -0.17853 | -0.022888 | -0.20599 | -0.18311 | 10.0479 | 10.2264 | 1.7029 | 2.7328 | 1.8814 | 2.9114 | -0.013733 | 0 | -0.11444 | -0.11444 | 10.2219 | 10.2356 | 2.7466 | 3.1769 | 2.7603 | 3.1906
2174.1719 | 122.3145 | -114.8895 | 7.0953 | -0.1992 | 0.014159 | 0.094821 | 0.086436 | -0.0010959 | 0.016829 | 0.0096558 | -0.016571 | 0.073232 | 0.08671 | 0.085836 | -0.0014633 | 0.01276 | 0.013687 | -0.064087 | 7.045 | -115.1184 | 6.9763 | -104.6951 | -112.1841 | -113.1454 | -0.009155 | 7.0908 | -114.9078 | 7.0724 | -104.6539 | -111.6623 | -112.0651 | -0.22888 | -0.050354 | -0.11902 | -0.068665 | 10.1944 | 10.4233 | 1.7441 | 2.7054 | 1.973 | 2.9343 | -0.018311 | -0.004578 | -0.022888 | -0.01831 | 10.2356 | 10.2539 | 2.8244 | 3.2272 | 2.8427 | 3.2455
2175.1719 | 122.2366 | -114.9673 | 7.1182 | -0.054245 | 0.013091 | 0.09418 | 0.085154 | 0.0023221 | 0.014052 | 0.01051 | -0.08371 | 0.076223 | 0.089914 | 0.086049 | 0.0019547 | 0.015323 | 0.013687 | -0.018311 | 6.9992 | -115.1871 | 6.8436 | -104.9103 | -112.2849 | -113.2919 | 0.045776 | 7.045 | -115.0223 | 6.9122 | -104.8004 | -111.7676 | -112.2574 | -0.21973 | -0.11902 | -0.27466 | -0.15564 | 10.0571 | 10.2768 | 1.6754 | 2.6825 | 1.8951 | 2.9022 | -0.054932 | -0.073242 | -0.20599 | -0.13275 | 10.1669 | 10.2219 | 2.71 | 3.1998 | 2.7649 | 3.2547
2176.1719 | 122.2275 | -114.9673 | 7.132 | -0.16106 | 0.012877 | 0.091403 | 0.084086 | -0.00088229 | 0.017897 | 0.011151 | -0.074554 | 0.075582 | 0.08735 | 0.088613 | -0.00060892 | 0.01575 | 0.01561 | -0.009155 | 7.0129 | -115.1962 | 6.8436 | -104.7867 | -112.294 | -113.3102 | 0.022888 | 7.0541 | -115.0406 | 7.251 | -104.718 | -111.676 | -112.2391 | -0.22888 | -0.11902 | -0.28839 | -0.16937 | 10.1807 | 10.4095 | 1.6571 | 2.6733 | 1.886 | 2.9022 | -0.073242 | -0.07782 | 0.11902 | 0.19684 | 10.2493 | 10.3226 | 2.7283 | 3.2913 | 2.8015 | 3.3646
2177.1719 | 122.2321 | -114.9857 | 7.0724 | -0.10002 | 0.014373 | 0.092685 | 0.083873 | 0.0016813 | 0.018752 | 0.0090148 | -0.053192 | 0.075368 | 0.089914 | 0.085409 | -0.0027451 | 0.011905 | 0.016037 | -0.022888 | 7.0358 | -115.1688 | 6.8527 | -104.924 | -112.2803 | -113.2919 | 0.018311 | 7.077 | -114.9902 | 6.9077 | -104.7775 | -111.7355 | -112.2437 | -0.1831 | -0.036621 | -0.21973 | -0.18311 | 10.0616 | 10.2448 | 1.6937 | 2.7054 | 1.8768 | 2.8885 | -0.004577 | 0.004577 | -0.16479 | -0.16937 | 10.2081 | 10.2127 | 2.742 | 3.2501 | 2.7466 | 3.2547
2178.1719 | 122.2641 | -115.0497 | 7.0221 | -0.069504 | 0.013732 | 0.09418 | 0.088359 | -0.0030185 | 0.016616 | 0.011578 | -0.053192 | 0.075582 | 0.089487 | 0.081777 | -0.0018905 | 0.014255 | 0.013687 | -0.004578 | 6.9992 | -115.2283 | 6.7886 | -104.924 | -112.3535 | -113.3377 | 0.022888 | 7.0358 | -115.0452 | 6.8756 | -104.8416 | -111.8271 | -112.2849 | -0.17853 | -0.022888 | -0.23346 | -0.21057 | 10.1257 | 10.3043 | 1.712 | 2.6962 | 1.8906 | 2.8748 | 0.004578 | 0.013733 | -0.14649 | -0.16022 | 10.2081 | 10.2036 | 2.7649 | 3.2227 | 2.7603 | 3.2181
2179.1719 | 122.2504 | -115.0726 | 6.9489 | -0.13054 | 0.014159 | 0.093753 | 0.087504 | 0.0014675 | 0.018325 | 0.013715 | -0.077606 | 0.075368 | 0.090555 | 0.086904 | 0.00067288 | 0.014896 | 0.011978 | -0.11444 | 6.9489 | -115.2832 | 6.7291 | -104.9286 | -112.4451 | -113.4384 | -0.091553 | 6.9534 | -115.0955 | 6.8069 | -104.9011 | -111.8637 | -112.3535 | -0.21057 | 0 | -0.21973 | -0.21973 | 10.144 | 10.3546 | 1.6342 | 2.6276 | 1.8448 | 2.8381 | -0.022888 | 0.004577 | -0.14191 | -0.14648 | 10.1715 | 10.1944 | 2.7191 | 3.2089 | 2.742 | 3.2318
2180.1719 | 122.2412 | -114.9307 | 7.1457 | -0.077133 | 0.013518 | 0.093753 | 0.086009 | 0.0012539 | 0.018111 | 0.013715 | -0.080658 | 0.075582 | 0.08735 | 0.080922 | 0.00024568 | 0.016391 | 0.011978 | 0.009155 | 7.0267 | -115.1642 | 6.871 | -104.8187 | -112.3306 | -113.2965 | 0.004578 | 7.0587 | -115.0177 | 7.0038 | -104.6494 | -111.6302 | -112.2253 | -0.23346 | -0.11902 | -0.27466 | -0.15564 | 10.112 | 10.3455 | 1.6342 | 2.6001 | 1.8677 | 2.8336 | -0.086975 | -0.086975 | -0.14191 | -0.054932 | 10.2814 | 10.3683 | 2.7054 | 3.3005 | 2.7924 | 3.3875
2181.1719 | 122.2 | -115.1916 | 6.8481 | -0.10002 | 0.013732 | 0.095248 | 0.086009 | 0.0023221 | 0.01512 | 0.011792 | -0.092865 | 0.074727 | 0.0897 | 0.085622 | 0.00088648 | 0.015964 | 0.011978 | -0.045776 | 6.8253 | -115.4022 | 6.6467 | -104.9515 | -112.5412 | -113.5208 | 0.004578 | 6.839 | -115.2145 | 6.6879 | -105.0476 | -111.9965 | -112.5229 | -0.21057 | -0.022889 | -0.20142 | -0.17853 | 10.2402 | 10.4507 | 1.6708 | 2.6505 | 1.8814 | 2.861 | -0.022889 | -0.009156 | -0.16022 | -0.15106 | 10.144 | 10.1669 | 2.6688 | 3.1952 | 2.6917 | 3.2181
2182.1719 | 122.2137 | -115.1688 | 6.8985 | -0.12291 | 0.013945 | 0.092898 | 0.083232 | 0.0025357 | 0.014907 | 0.011578 | -0.092865 | 0.076223 | 0.089914 | 0.084554 | -0.00060892 | 0.014255 | 0.012405 | -0.054932 | 6.8939 | -115.3152 | 6.6879 | -105.0064 | -112.4542 | -113.4888 | -0.045776 | 6.9168 | -115.1825 | 6.7337 | -105.0385 | -111.85 | -112.4588 | -0.14648 | -0.004578 | -0.21057 | -0.20599 | 10.1624 | 10.3088 | 1.68 | 2.7145 | 1.8265 | 2.861 | -0.013733 | 0.01831 | -0.16479 | -0.1831 | 10.1303 | 10.144 | 2.71 | 3.3188 | 2.7237 | 3.3325
2183.1719 | 122.1725 | -115.1093 | 6.9717 | -0.12291 | 0.013091 | 0.093966 | 0.086436 | 0.0014675 | 0.019393 | 0.011365 | -0.08371 | 0.074941 | 0.086282 | 0.084768 | 3.2078e-05 | 0.01276 | 0.014755 | -0.022888 | 6.9031 | -115.2969 | 6.72 | -104.9881 | -112.4359 | -113.4247 | -0.027466 | 6.9351 | -115.1276 | 6.7657 | -104.9744 | -111.9781 | -112.4268 | -0.18768 | -0.068665 | -0.25177 | -0.18311 | 10.1212 | 10.3088 | 1.6846 | 2.6733 | 1.8723 | 2.861 | -0.01831 | -0.036621 | -0.20599 | -0.16937 | 10.1349 | 10.1532 | 2.6825 | 3.1311 | 2.7008 | 3.1494
2184.1719 | 122.2916 | -115.1321 | 6.8756 | -0.18394 | 0.016722 | 0.096744 | 0.082164 | 0.00018591 | 0.019179 | 0.01051 | 0.0017395 | 0.072805 | 0.088419 | 0.075582 | 0.00024568 | 0.012546 | 0.0098417 | -0.20142 | 6.7749 | -115.3748 | 6.7612 | -104.9515 | -112.3947 | -113.3835 | 0.022888 | 6.8481 | -115.1459 | 6.8253 | -104.9377 | -111.9232 | -112.3718 | -0.24262 | -0.10071 | -0.11444 | -0.013733 | 10.1807 | 10.4233 | 1.7487 | 2.7374 | 1.9913 | 2.98 | -0.013733 | -0.027465 | -0.050354 | -0.022889 | 10.1944 | 10.2081 | 2.7603 | 3.2089 | 2.774 | 3.2227
2185.1719 | 122.2046 | -115.1138 | 6.958 | -0.054245 | 0.013304 | 0.095889 | 0.085368 | -0.0017367 | 0.018325 | 0.012646 | -0.13864 | 0.075796 | 0.088205 | 0.079 | -0.0023177 | 0.014255 | 0.017532 | -0.018311 | 6.9214 | -115.274 | 6.7474 | -104.9286 | -112.413 | -113.3835 | 0.045776 | 7.045 | -115.0772 | 6.8298 | -104.9835 | -111.9507 | -112.3947 | -0.16022 | -0.036621 | -0.21057 | -0.17395 | 10.1852 | 10.3455 | 1.7303 | 2.7008 | 1.8906 | 2.861 | 0.036622 | 0.086975 | -0.12817 | -0.21515 | 10.1303 | 10.0937 | 2.7191 | 3.1631 | 2.6825 | 3.1265
2186.1719 | 122.1863 | -115.1138 | 6.9717 | -0.1458 | 0.013091 | 0.093112 | 0.085795 | 0.0023221 | 0.016188 | 0.01051 | -0.098968 | 0.074514 | 0.088419 | 0.083486 | 0.00088648 | 0.013828 | 0.013046 | -0.050354 | 6.9351 | -115.2832 | 6.7337 | -104.9286 | -112.4405 | -113.4201 | -0.050354 | 6.9534 | -115.1184 | 6.7886 | -105.0339 | -111.9873 | -112.4222 | -0.16937 | -0.036621 | -0.23804 | -0.20142 | 10.1852 | 10.3546 | 1.6937 | 2.6733 | 1.8631 | 2.8427 | -0.004577 | -0.018311 | -0.18311 | -0.16479 | 10.08 | 10.0845 | 2.6917 | 3.1265 | 2.6962 | 3.1311
2187.1719 | 122.168 | -115.1871 | 6.8985 | -0.1458 | 0.016082 | 0.093112 | 0.085368 | 0.0023221 | 0.016616 | 0.012646 | -0.068451 | 0.0743 | 0.088846 | 0.090963 | -0.00039512 | 0.01511 | 0.0096279 | -0.068665 | 6.8756 | -115.3793 | 6.6284 | -105.0522 | -112.6144 | -113.5345 | -0.16022 | 6.9031 | -115.2283 | 6.6971 | -105.1346 | -112.1292 | -112.5275 | -0.19226 | -0.022889 | -0.27008 | -0.24719 | 10.1349 | 10.3271 | 1.6525 | 2.5726 | 1.8448 | 2.7649 | -0.041198 | 0.004577 | -0.20142 | -0.20599 | 10.0525 | 10.0937 | 2.6596 | 3.0579 | 2.7008 | 3.0991
2188.1719 | 122.1771 | -115.1001 | 6.9809 | -0.20683 | 0.015441 | 0.095675 | 0.085795 | 0.0023221 | 0.016829 | 0.010937 | -0.11118 | 0.0743 | 0.088419 | 0.081777 | -0.0010361 | 0.015964 | 0.014541 | -0.086975 | 6.9397 | -115.274 | 6.6971 | -104.9515 | -112.5 | -113.3926 | -0.17395 | 6.9809 | -115.123 | 6.7795 | -105.0522 | -111.937 | -112.4451 | -0.17395 | -0.041199 | -0.28381 | -0.24261 | 10.1486 | 10.3226 | 1.7075 | 2.6001 | 1.8814 | 2.774 | -0.022888 | 0 | -0.20142 | -0.20142 | 10.0479 | 10.0708 | 2.655 | 3.1631 | 2.6779 | 3.186
2189.1719 | 122.2733 | -115.0223 | 6.9717 | -0.15343 | 0.011809 | 0.095462 | 0.083445 | -0.0010959 | 0.014693 | 0.013501 | -0.013519 | 0.073018 | 0.088205 | 0.080282 | 0.00024568 | 0.01276 | 0.011978 | -0.045776 | 6.9214 | -115.1871 | 6.8573 | -104.7363 | -112.2757 | -113.2828 | 0 | 6.9489 | -115.0543 | 6.9122 | -104.9194 | -111.7126 | -112.3169 | -0.16479 | -0.050354 | -0.11444 | -0.064087 | 10.2859 | 10.4507 | 1.7395 | 2.7466 | 1.9043 | 2.9114 | -0.032043 | -0.022888 | -0.05951 | -0.036622 | 10.1028 | 10.1349 | 2.7054 | 3.3096 | 2.7374 | 3.3417
2190.1719 | 122.1771 | -115.0635 | 7.0084 | -0.11528 | 0.014159 | 0.093326 | 0.089 | -0.00024129 | 0.014479 | 0.012006 | -0.095917 | 0.076864 | 0.0897 | 0.083059 | 0.00045928 | 0.012119 | 0.014114 | -0.10529 | 6.9534 | -115.2695 | 6.72 | -104.8233 | -112.5366 | -113.4521 | -0.14191 | 6.9901 | -115.0909 | 6.7978 | -105.0522 | -111.9965 | -112.413 | -0.20599 | -0.054932 | -0.28839 | -0.23346 | 10.2402 | 10.4462 | 1.6113 | 2.5269 | 1.8173 | 2.7328 | -0.027465 | -0.018311 | -0.21057 | -0.19226 | 10.0113 | 10.0388 | 2.6505 | 3.067 | 2.6779 | 3.0945
2191.1719 | 122.1542 | -115.036 | 7.0587 | -0.084763 | 0.014586 | 0.095248 | 0.086222 | -0.0010959 | 0.014907 | 0.011792 | -0.10202 | 0.074941 | 0.088632 | 0.087118 | -0.0012497 | 0.015323 | 0.014114 | -0.068665 | 6.9809 | -115.21 | 6.7932 | -104.8004 | -112.3993 | -113.3514 | -0.082397 | 7.0312 | -115.036 | 6.839 | -105.0568 | -111.9736 | -112.3947 | -0.17395 | -0.07782 | -0.2655 | -0.18768 | 10.2356 | 10.4095 | 1.6846 | 2.6367 | 1.8585 | 2.8107 | 0 | -0.027466 | -0.21973 | -0.19226 | 9.9792 | 9.9792 | 2.6413 | 3.0624 | 2.6413 | 3.0624
2192.1719 | 122.1725 | -115.0497 | 7.0267 | -0.12291 | 0.012664 | 0.092898 | 0.087077 | 0.0016813 | 0.014052 | 0.0098694 | -0.086761 | 0.07665 | 0.089059 | 0.080068 | -0.00039512 | 0.014041 | 0.010483 | -0.054932 | 6.9763 | -115.2328 | 6.7886 | -104.7363 | -112.4222 | -113.3698 | -0.027466 | 7.0312 | -115.0452 | 6.8344 | -105.043 | -111.9003 | -112.3947 | -0.1831 | -0.050354 | -0.23804 | -0.18768 | 10.3134 | 10.4965 | 1.68 | 2.6276 | 1.8631 | 2.8107 | 0.004578 | 0.004578 | -0.19226 | -0.19684 | 10.0067 | 10.0021 | 2.655 | 3.1494 | 2.6505 | 3.1448
2193.1719 | 122.168 | -115.1093 | 6.9534 | -0.20683 | 0.013091 | 0.094821 | 0.085368 | 0.0029629 | 0.015975 | 0.013715 | -0.095917 | 0.075368 | 0.087991 | 0.082631 | 0.00045928 | 0.015537 | 0.011551 | -0.10986 | 6.9168 | -115.3015 | 6.6742 | -104.7775 | -112.5229 | -113.4796 | -0.20599 | 6.9489 | -115.1505 | 6.752 | -105.13 | -112.1109 | -112.4817 | -0.19226 | -0.036621 | -0.27924 | -0.24261 | 10.3317 | 10.524 | 1.6296 | 2.5864 | 1.8219 | 2.7786 | -0.041199 | -0.004577 | -0.20142 | -0.19684 | 9.9792 | 10.0204 | 2.6276 | 2.9984 | 2.6688 | 3.0396
2194.1719 | 122.2321 | -115.1962 | 6.7932 | -0.22209 | 0.015227 | 0.094607 | 0.083873 | 0.0021085 | 0.018111 | 0.013715 | -0.028778 | 0.074727 | 0.0897 | 0.080282 | -0.00018152 | 0.014255 | 0.01326 | -0.059509 | 6.8253 | -115.3061 | 6.752 | -104.7546 | -112.4496 | -113.4201 | -0.032043 | 6.8573 | -115.1367 | 6.8253 | -105.0568 | -111.9141 | -112.5961 | -0.10986 | 0.032043 | -0.041199 | -0.073242 | 10.4416 | 10.5515 | 1.7761 | 2.7466 | 1.886 | 2.8564 | 0.059509 | 0.064087 | 0.032043 | -0.032044 | 10.1395 | 10.08 | 2.6001 | 3.2822 | 2.5406 | 3.2227
2195.1719 | 122.1771 | -115.1825 | 6.8985 | -0.13054 | 0.014373 | 0.093966 | 0.086009 | 0.0018949 | 0.014693 | 0.013287 | -0.092865 | 0.072591 | 0.087778 | 0.079 | -0.0014633 | 0.013828 | 0.013901 | -0.10529 | 6.8527 | -115.3473 | 6.6879 | -105.1025 | -112.5458 | -113.5071 | -0.12817 | 6.9351 | -115.1871 | 6.7062 | -105.1941 | -112.0377 | -112.5092 | -0.16479 | -0.045777 | -0.21057 | -0.16479 | 10.08 | 10.2448 | 1.6754 | 2.6367 | 1.8402 | 2.8015 | -0.004578 | 0.036621 | -0.19226 | -0.22888 | 9.9884 | 9.993 | 2.6733 | 3.1448 | 2.6779 | 3.1494
2196.1719 | 122.1863 | -115.1184 | 6.9351 | -0.10002 | 0.013091 | 0.093539 | 0.085582 | -0.0010959 | 0.019179 | 0.012646 | -0.12033 | 0.074727 | 0.087564 | 0.078786 | -0.00018152 | 0.012973 | 0.019455 | -0.11902 | 6.9031 | -115.3381 | 6.6147 | -104.7867 | -112.3444 | -113.5254 | -0.10986 | 6.9489 | -115.1413 | 6.7612 | -105.1804 | -112.0422 | -112.4817 | -0.21973 | -0.032044 | -0.32044 | -0.28839 | 10.3317 | 10.5515 | 1.593 | 2.774 | 1.8127 | 2.9938 | -0.022888 | 0.013733 | -0.17395 | -0.18768 | 9.938 | 9.9609 | 2.6367 | 3.0762 | 2.6596 | 3.0991
2197.1719 | 122.1588 | -115.1871 | 6.8756 | -0.10002 | 0.012236 | 0.092685 | 0.086436 | 0.00039951 | 0.015334 | 0.0092286 | -0.10812 | 0.072164 | 0.090128 | 0.081563 | -0.00018152 | 0.012332 | 0.012832 | -0.10529 | 6.8207 | -115.3656 | 6.6238 | -104.924 | -112.6465 | -113.6444 | 0.004578 | 6.9031 | -115.21 | 6.6971 | -105.2216 | -112.6099 | -112.5458 | -0.17853 | -0.054931 | -0.25177 | -0.19684 | 10.2631 | 10.4416 | 1.5427 | 2.5406 | 1.7212 | 2.7191 | -0.022888 | 0.027466 | -0.17853 | -0.20599 | 9.9655 | 9.9884 | 2.6413 | 2.5772 | 2.6642 | 2.6001
2198.1719 | 122.168 | -115.1871 | 6.8756 | -0.1458 | 0.014373 | 0.095035 | 0.086436 | -0.0032321 | 0.01747 | 0.014355 | -0.098968 | 0.073232 | 0.087137 | 0.081563 | 0.0013139 | 0.014041 | 0.011123 | -0.073242 | 6.8344 | -115.3976 | 6.6055 | -104.9973 | -112.6556 | -113.5666 | -0.1236 | 6.8802 | -115.2237 | 6.6696 | -105.2811 | -112.1155 | -112.5549 | -0.21057 | -0.041198 | -0.27008 | -0.22888 | 10.1898 | 10.4004 | 1.6205 | 2.5314 | 1.8311 | 2.742 | -0.036621 | 0.004578 | -0.20599 | -0.21057 | 9.906 | 9.9426 | 2.6321 | 3.0716 | 2.6688 | 3.1082
2199.1719 | 122.1588 | -115.1642 | 6.9031 | -0.13817 | 0.012877 | 0.093112 | 0.085582 | 0.0012539 | 0.017256 | 0.012006 | -0.10202 | 0.073446 | 0.088632 | 0.077291 | 0.00024568 | 0.018955 | 0.014328 | -0.10071 | 6.8619 | -115.3793 | 6.6147 | -104.9149 | -112.6556 | -113.6078 | -0.14648 | 6.8893 | -115.2191 | 6.7245 | -105.2307 | -111.9553 | -112.4496 | -0.21515 | -0.041199 | -0.28839 | -0.24719 | 10.2493 | 10.4645 | 1.5564 | 2.5085 | 1.7715 | 2.7237 | -0.054931 | -0.013733 | -0.17853 | -0.16479 | 9.9335 | 9.9884 | 2.7145 | 3.2089 | 2.7695 | 3.2639
2200.1719 | 122.1542 | -115.1367 | 6.9214 | -0.11528 | 0.0148 | 0.092685 | 0.087291 | 0.00039951 | 0.017897 | 0.01051 | -0.062347 | 0.073018 | 0.090128 | 0.080282 | -0.00039512 | 0.013187 | 0.011764 | -0.11444 | 6.839 | -115.3976 | 6.6284 | -104.9835 | -112.6511 | -113.5574 | 0.036621 | 6.9168 | -115.2008 | 6.6879 | -105.2444 | -111.731 | -112.5504 | -0.26093 | -0.082398 | -0.29297 | -0.21057 | 10.1532 | 10.4141 | 1.5793 | 2.4857 | 1.8402 | 2.7466 | -0.064087 | -0.004578 | -0.23346 | -0.22888 | 9.8923 | 9.9564 | 2.5864 | 3.4058 | 2.6505 | 3.4698
2201.1719 | 122.1359 | -115.2191 | 6.8619 | -0.13817 | 0.014586 | 0.094607 | 0.086222 | -0.0010959 | 0.015548 | 0.011578 | -0.095917 | 0.072164 | 0.08735 | 0.0837 | 3.2078e-05 | 0.012119 | 0.010269 | -0.073242 | 6.8298 | -115.3885 | 6.5826 | -105.043 | -112.6877 | -113.5757 | -0.15564 | 6.8756 | -115.2237 | 6.665 | -105.3131 | -112.1613 | -112.587 | -0.16937 | -0.032043 | -0.27923 | -0.24719 | 10.1761 | 10.3455 | 1.6434 | 2.5314 | 1.8127 | 2.7008 | -0.004578 | 0.013733 | -0.19684 | -0.21057 | 9.906 | 9.9106 | 2.6321 | 3.0579 | 2.6367 | 3.0624
2202.1719 | 122.2092 | -115.2695 | 6.7932 | -0.1458 | 0.014373 | 0.092257 | 0.087718 | 0.0021085 | 0.014907 | 0.0085876 | -0.034882 | 0.072164 | 0.090341 | 0.081991 | -0.00039512 | 0.012546 | 0.01561 | -0.10986 | 6.7657 | -115.4617 | 6.6284 | -105.0659 | -112.8067 | -113.6307 | -0.15564 | 6.7841 | -115.2832 | 6.601 | -105.3589 | -112.0605 | -112.6282 | -0.19226 | -0.027466 | -0.16479 | -0.13733 | 10.2036 | 10.3958 | 1.6388 | 2.4628 | 1.8311 | 2.655 | -0.013733 | -0.009155 | -0.19226 | -0.18311 | 9.9106 | 9.9243 | 2.6413 | 3.2089 | 2.655 | 3.2227
2203.1719 | 122.1588 | -115.3015 | 6.7886 | -0.11528 | 0.01245 | 0.09418 | 0.085368 | -0.00066869 | 0.017684 | 0.0094422 | -0.16916 | 0.074087 | 0.08735 | 0.083059 | 3.2078e-05 | 0.013401 | 0.012619 | -0.10986 | 6.7429 | -115.4755 | 6.5231 | -105.1437 | -112.7426 | -113.6856 | -0.018311 | 6.8115 | -115.2878 | 6.5872 | -105.3818 | -112.1933 | -112.6328 | -0.17395 | -0.045776 | -0.2655 | -0.21973 | 10.1578 | 10.3317 | 1.6159 | 2.5589 | 1.7899 | 2.7328 | 0.013733 | 0.022888 | -0.20142 | -0.2243 | 9.9197 | 9.906 | 2.6688 | 3.1082 | 2.655 | 3.0945
2204.1719 | 122.2916 | -115.0314 | 6.9534 | -0.19157 | 0.015227 | 0.094394 | 0.0843 | -0.0010959 | 0.019179 | 0.012219 | -0.12949 | 0.073018 | 0.088419 | 0.080068 | -0.0023177 | 0.011692 | 0.014755 | -0.064087 | 6.9031 | -115.2374 | 6.839 | -104.4617 | -112.4451 | -113.3331 | -0.050354 | 6.8802 | -115.1047 | 6.9077 | -105.0797 | -112.0605 | -112.3398 | -0.20599 | -0.050354 | -0.11444 | -0.064087 | 10.5698 | 10.7758 | 1.6983 | 2.5864 | 1.9043 | 2.7924 | -0.073242 | -0.073242 | -0.045776 | 0.027466 | 9.9518 | 10.025 | 2.6917 | 2.9709 | 2.7649 | 3.0441
2205.1719 | 122.1771 | -115.155 | 6.9489 | -0.1458 | 0.014159 | 0.093539 | 0.086436 | 0.0042447 | 0.015975 | 0.0096558 | -0.12338 | 0.074087 | 0.086069 | 0.08669 | 3.2078e-05 | 0.012973 | 0.013901 | -0.050354 | 6.8893 | -115.3336 | 6.6879 | -104.1916 | -112.5732 | -113.5254 | -0.054932 | 6.9031 | -115.155 | 6.7383 | -105.2536 | -112.1567 | -112.5092 | -0.17853 | -0.05951 | -0.26093 | -0.20142 | 10.9634 | 11.142 | 1.6296 | 2.5818 | 1.8082 | 2.7603 | 0 | -0.045777 | -0.21057 | -0.1648 | 9.9014 | 9.9014 | 2.6459 | 2.9984 | 2.6459 | 2.9984
2206.1719 | 122.1909 | -115.3656 | 6.7017 | -0.077133 | 0.012877 | 0.093753 | 0.0843 | 0.0036039 | 0.01512 | 0.01051 | -0.11423 | 0.072805 | 0.0897 | 0.082631 | -0.0010361 | 0.013614 | 0.011123 | -0.054932 | 6.6788 | -115.5212 | 6.4819 | -105.0934 | -112.7975 | -113.7131 | -0.045776 | 6.7017 | -115.3931 | 6.514 | -105.4642 | -112.3489 | -112.7151 | -0.15564 | -0.022888 | -0.21973 | -0.19684 | 10.2722 | 10.4279 | 1.6525 | 2.5681 | 1.8082 | 2.7237 | -0.027465 | 0 | -0.18768 | -0.18768 | 9.9014 | 9.9289 | 2.6505 | 3.0167 | 2.6779 | 3.0441
2207.1719 | 122.168 | -115.3839 | 6.6971 | -0.11528 | 0.015013 | 0.095035 | 0.087504 | 0.0029629 | 0.015548 | 0.011365 | -0.092865 | 0.075582 | 0.090128 | 0.085836 | 0.00088648 | 0.011905 | 0.010269 | -0.064087 | 6.6559 | -115.5807 | 6.427 | -104.9606 | -112.8708 | -113.7909 | -0.1236 | 6.6879 | -115.3839 | 6.4682 | -105.5237 | -112.3672 | -112.7701 | -0.19684 | -0.041199 | -0.27008 | -0.22888 | 10.4233 | 10.6201 | 1.593 | 2.5131 | 1.7899 | 2.71 | 0 | -0.009156 | -0.22888 | -0.21973 | 9.8602 | 9.8602 | 2.6138 | 3.0167 | 2.6138 | 3.0167
2208.1719 | 122.1725 | -115.5167 | 6.5643 | -0.069504 | 0.013518 | 0.095462 | 0.083873 | 0.00039951 | 0.016616 | 0.012006 | -0.092865 | 0.072164 | 0.088846 | 0.077718 | -0.00082252 | 0.013614 | 0.014114 | -0.059509 | 6.5186 | -115.7089 | 6.308 | -105.2582 | -112.9532 | -113.8824 | -0.086975 | 6.5506 | -115.5029 | 6.3538 | -105.661 | -112.468 | -112.8708 | -0.19226 | -0.045776 | -0.25635 | -0.21057 | 10.2585 | 10.4507 | 1.6342 | 2.5635 | 1.8265 | 2.7557 | 0.013733 | -0.013733 | -0.21057 | -0.19684 | 9.8557 | 9.8419 | 2.6459 | 3.0487 | 2.6321 | 3.035
2209.1719 | 122.2733 | -115.567 | 6.4362 | -0.092392 | 0.014586 | 0.092898 | 0.084086 | 0.0014675 | 0.01747 | 0.011151 | -0.016571 | 0.071737 | 0.090982 | 0.081136 | -0.0023177 | 0.015537 | 0.014755 | -0.050354 | 6.3766 | -115.7318 | 6.3217 | -105.2536 | -112.9715 | -113.8367 | -0.032043 | 6.4362 | -115.567 | 6.3812 | -105.6015 | -112.3672 | -112.8799 | -0.16479 | -0.059509 | -0.11444 | -0.054932 | 10.3134 | 10.4782 | 1.7303 | 2.5955 | 1.8951 | 2.7603 | 0 | 0 | -0.054931 | -0.054931 | 9.9655 | 9.9655 | 2.6871 | 3.1998 | 2.6871 | 3.1998
2210.1719 | 122.168 | -115.5121 | 6.5826 | -0.11528 | 0.011595 | 0.093112 | 0.088359 | 0.00039951 | 0.017256 | 0.013074 | -0.10507 | 0.073446 | 0.089273 | 0.089467 | 3.2078e-05 | 0.013614 | 0.011337 | -0.086975 | 6.514 | -115.7135 | 6.2714 | -105.3589 | -113.0219 | -113.9282 | -0.14191 | 6.5506 | -115.5579 | 6.3354 | -105.6793 | -112.5504 | -112.9211 | -0.20142 | -0.068665 | -0.31128 | -0.24261 | 10.1532 | 10.3546 | 1.5839 | 2.4902 | 1.7853 | 2.6916 | -0.045776 | -0.032044 | -0.24719 | -0.21515 | 9.8328 | 9.8785 | 2.5909 | 2.9617 | 2.6367 | 3.0075
2211.1719 | 122.1771 | -113.8733 | 8.2214 | -0.084763 | 0.013945 | 0.095675 | 0.085368 | -0.00066869 | 0.016616 | 0.014142 | -0.071503 | 0.073446 | 0.089059 | 0.086263 | -0.0018905 | 0.012973 | 0.016464 | -2.3071 | 7.0496 | -115.5121 | 6.2942 | -105.5237 | -112.9211 | -118.5516 | -0.76904 | 6.5598 | -115.4938 | 7.695 | -103.6514 | -110.7788 | -111.2183 | -1.6388 | -1.1719 | -1.9272 | -0.75531 | 8.3496 | 9.9884 | -4.6783 | 0.95215 | -3.0396 | 2.5909 | -1.6205 | -1.6617 | -0.52643 | 1.1353 | 10.2219 | 11.8423 | 2.655 | 3.0945 | 4.2755 | 4.715
2212.1719 | 122.2183 | -115.4068 | 6.6605 | -0.16869 | 0.012877 | 0.091617 | 0.085582 | -2.7692e-05 | 0.015548 | 0.010937 | -0.077606 | 0.074087 | 0.090341 | 0.079427 | 0.00024568 | 0.011478 | 0.013687 | 0.027466 | 6.7245 | -115.4892 | 6.4728 | -105.1758 | -112.8433 | -113.7314 | -0.17395 | 6.665 | -115.3885 | 6.4865 | -105.5283 | -112.7197 | -112.7563 | -0.082398 | 0.064087 | -0.18768 | -0.25177 | 10.231 | 10.3134 | 1.6754 | 2.5635 | 1.7578 | 2.6459 | 0.01831 | 0.004578 | -0.17395 | -0.17853 | 9.8785 | 9.8602 | 2.6505 | 2.6871 | 2.6321 | 2.6688
2213.1719 | 122.2137 | -115.4068 | 6.6147 | -0.13054 | 0.013518 | 0.095889 | 0.086009 | -0.00045509 | 0.016402 | 0.0088012 | -0.077606 | 0.073446 | 0.090341 | 0.087758 | -0.00039512 | 0.012973 | 0.015396 | -0.004578 | 6.6238 | -115.5853 | 6.6376 | -105.1117 | -112.7014 | -113.7588 | -0.032043 | 6.6925 | -115.4114 | 6.4819 | -105.5603 | -112.381 | -112.7609 | -0.17853 | 0.009155 | 0.022888 | 0.013733 | 10.2951 | 10.4736 | 1.6479 | 2.7054 | 1.8265 | 2.8839 | -0.004578 | 0.07782 | -0.13275 | -0.21057 | 9.8465 | 9.8511 | 2.6459 | 3.0258 | 2.6505 | 3.0304
2214.1719 | 122.2961 | -115.3748 | 6.6147 | -0.17632 | 0.014373 | 0.093753 | 0.084727 | -0.0032321 | 0.018111 | 0.01051 | -0.0013123 | 0.07195 | 0.088419 | 0.079854 | -0.0010361 | 0.011905 | 0.013687 | -0.045776 | 6.5964 | -115.5762 | 6.4957 | -105.098 | -112.793 | -113.6948 | 0.041199 | 6.6101 | -115.416 | 6.546 | -105.4688 | -112.2711 | -112.6831 | -0.20142 | -0.01831 | -0.11902 | -0.10071 | 10.2768 | 10.4782 | 1.68 | 2.5818 | 1.8814 | 2.7832 | -0.041199 | -0.004578 | -0.068664 | -0.064086 | 9.906 | 9.9472 | 2.6917 | 3.1036 | 2.7328 | 3.1448
2215.1719 | 122.1588 | -115.4114 | 6.6925 | -0.10765 | 0.013304 | 0.094821 | 0.084513 | 0.00018591 | 0.015548 | 0.011365 | -0.08371 | 0.072164 | 0.088205 | 0.083913 | -0.00082252 | 0.01276 | 0.011551 | -0.045776 | 6.6238 | -115.5945 | 6.4316 | -105.1575 | -112.9349 | -113.8092 | -0.10529 | 6.6696 | -115.4205 | 6.4728 | -105.5786 | -112.3444 | -112.7747 | -0.1831 | -0.068665 | -0.26092 | -0.19226 | 10.2539 | 10.437 | 1.6022 | 2.4765 | 1.7853 | 2.6596 | -0.009155 | -0.022888 | -0.21973 | -0.19684 | 9.8328 | 9.8419 | 2.6367 | 3.067 | 2.6459 | 3.0762
2216.1719 | 122.1725 | -115.4755 | 6.5872 | -0.11528 | 0.013091 | 0.094394 | 0.086222 | 0.00039951 | 0.016188 | 0.01051 | -0.080658 | 0.072805 | 0.090128 | 0.081136 | -0.0010361 | 0.012546 | 0.015396 | -0.054932 | 6.5781 | -115.6586 | 6.3583 | -105.0797 | -112.9395 | -113.8184 | -0.07782 | 6.5872 | -115.4709 | 6.395 | -105.6473 | -112.4634 | -112.8342 | -0.1831 | -0.009155 | -0.22888 | -0.21973 | 10.3958 | 10.5789 | 1.6571 | 2.536 | 1.8402 | 2.7191 | 0.004578 | 0 | -0.19226 | -0.19226 | 9.8282 | 9.8236 | 2.6413 | 3.0121 | 2.6367 | 3.0075
2217.1719 | 122.1771 | -115.5304 | 6.5369 | -0.10765 | 0.014373 | 0.091617 | 0.086222 | -0.00024129 | 0.014907 | 0.011578 | -0.092865 | 0.071737 | 0.091409 | 0.079213 | 0.00088648 | 0.012546 | 0.012619 | -0.009155 | 6.5048 | -115.6906 | 6.308 | -105.1666 | -113.0127 | -113.9236 | -0.018311 | 6.5094 | -115.5624 | 6.3309 | -105.6976 | -112.5687 | -112.9074 | -0.16022 | -0.032043 | -0.22888 | -0.19684 | 10.3638 | 10.524 | 1.6068 | 2.5177 | 1.767 | 2.6779 | -0.032043 | -0.027466 | -0.20599 | -0.17853 | 9.8328 | 9.8648 | 2.623 | 2.9617 | 2.655 | 2.9938
2218.1719 | 122.2824 | -115.4068 | 6.6193 | -0.092392 | 0.013945 | 0.093539 | 0.082591 | -0.0043003 | 0.017043 | 0.0098694 | 0.041412 | 0.071096 | 0.087778 | 0.079641 | 0.0019547 | 0.010196 | 0.016678 | -0.073242 | 6.5552 | -115.5807 | 6.4499 | -104.9011 | -112.8387 | -113.7772 | -0.091553 | 6.5506 | -115.48 | 6.4636 | -105.5511 | -112.4176 | -112.7655 | -0.17395 | -0.064087 | -0.16937 | -0.10529 | 10.5057 | 10.6796 | 1.6296 | 2.5681 | 1.8036 | 2.742 | -0.073243 | -0.068665 | -0.15564 | -0.086975 | 9.8557 | 9.9289 | 2.6413 | 2.9892 | 2.7145 | 3.0624
2219.1719 | 122.2961 | -115.5075 | 6.4545 | -0.16869 | 0.013945 | 0.093539 | 0.084941 | -0.00066869 | 0.016402 | 0.011578 | -0.040985 | 0.069173 | 0.0897 | 0.084768 | -0.0021041 | 0.013187 | 0.011551 | 0.022888 | 6.459 | -115.6769 | 6.3995 | -105.1987 | -112.8159 | -113.8687 | 0.013733 | 6.4911 | -115.5212 | 6.3217 | -105.6564 | -112.2482 | -112.7472 | -0.16937 | 0.004577 | -0.054932 | -0.059509 | 10.3088 | 10.4782 | 1.6388 | 2.6917 | 1.8082 | 2.861 | -0.013733 | 0.036621 | -0.13275 | -0.16937 | 9.8511 | 9.8648 | 2.7603 | 3.2593 | 2.774 | 3.273
2220.1719 | 122.2412 | -116.1163 | 5.9463 | -0.077133 | 0.014586 | 0.094821 | 0.08665 | -0.00045509 | 0.016188 | 0.0096558 | -0.089813 | 0.073659 | 0.08735 | 0.080068 | 3.2078e-05 | 0.012119 | 0.014114 | 0 | 5.9235 | -116.3086 | 5.7312 | -105.6931 | -113.5483 | -114.4684 | 0.073242 | 5.9647 | -116.1118 | 5.8044 | -106.279 | -113.0127 | -113.4476 | -0.19226 | -0.022888 | -0.21515 | -0.19226 | 10.4233 | 10.6155 | 1.6479 | 2.5681 | 1.8402 | 2.7603 | 0.004578 | 0.018311 | -0.14191 | -0.16022 | 9.8373 | 9.8328 | 2.6688 | 3.1036 | 2.6642 | 3.0991
2221.1719 | 122.2412 | -115.7913 | 6.2439 | -0.10002 | 0.014373 | 0.091617 | 0.086222 | -0.0028049 | 0.01512 | 0.014569 | -0.10507 | 0.073232 | 0.089487 | 0.084127 | 0.00045928 | 0.012332 | 0.012619 | 0.018311 | 6.1478 | -116.0843 | 5.9784 | -105.2948 | -113.3514 | -114.1525 | 0.036621 | 6.1981 | -115.8875 | 6.0196 | -105.8441 | -112.6877 | -113.2141 | -0.29297 | -0.09613 | -0.2655 | -0.16937 | 10.4965 | 10.7895 | 1.6388 | 2.4399 | 1.9318 | 2.7328 | -0.09613 | -0.045776 | -0.2243 | -0.17853 | 9.9472 | 10.0433 | 2.5772 | 3.1036 | 2.6733 | 3.1998
2222.1719 | 122.2366 | -115.9241 | 6.1249 | -0.092392 | 0.012877 | 0.093112 | 0.083659 | -0.0032321 | 0.017043 | 0.012006 | -0.10202 | 0.073018 | 0.088632 | 0.080709 | 0.00024568 | 0.01276 | 0.01091 | 0.07782 | 6.1478 | -116.0248 | 5.9418 | -105.4047 | -113.4018 | -114.2761 | -0.036621 | 6.1295 | -115.947 | 5.9463 | -106.1325 | -112.9349 | -113.3102 | -0.10071 | 0.022888 | -0.18311 | -0.20599 | 10.5194 | 10.6201 | 1.6479 | 2.5223 | 1.7487 | 2.623 | -0.022888 | 0.004578 | -0.17853 | -0.18311 | 9.7916 | 9.8145 | 2.6138 | 2.9892 | 2.6367 | 3.0121
2223.1875 | 122.2504 | -115.8646 | 6.1707 | -0.13054 | 0.013304 | 0.092471 | 0.082804 | -0.0010959 | 0.014052 | 0.0088012 | 0.023102 | 0.071523 | 0.090341 | 0.081777 | 0.0017411 | 0.012332 | 0.015182 | -0.059509 | 6.134 | -116.011 | 6.0104 | -105.2032 | -113.3286 | -114.18 | -0.027466 | 6.1478 | -115.8508 | 6.0425 | -106.0135 | -112.8113 | -113.1958 | -0.14648 | -0.036621 | -0.16022 | -0.1236 | 10.6613 | 10.8078 | 1.6846 | 2.536 | 1.8311 | 2.6825 | 0.013733 | -0.022888 | -0.12817 | -0.10529 | 9.8511 | 9.8373 | 2.6688 | 3.0533 | 2.655 | 3.0396
2224.1719 | 122.2366 | -115.8554 | 6.1935 | -0.10002 | 0.013732 | 0.094821 | 0.084727 | 0.0033901 | 0.018111 | 0.011365 | -0.08371 | 0.072591 | 0.090341 | 0.081563 | -0.0025315 | 0.013614 | 0.014969 | -0.10529 | 6.1569 | -116.0477 | 5.9372 | -105.5283 | -113.4384 | -114.2395 | -0.09613 | 6.1615 | -115.892 | 6.0562 | -106.0501 | -112.738 | -113.2324 | -0.19226 | -0.036621 | -0.25635 | -0.21973 | 10.3271 | 10.5194 | 1.6159 | 2.417 | 1.8082 | 2.6093 | -0.036621 | -0.032043 | -0.13733 | -0.10529 | 9.8053 | 9.8419 | 2.623 | 3.1174 | 2.6596 | 3.154
2225.1719 | 122.2321 | -115.3656 | 6.6605 | -0.061874 | 0.012877 | 0.13434 | 0.085582 | 0.00039951 | 0.018538 | 0.013501 | -0.12338 | 0.11724 | 0.089914 | 0.081563 | -0.0016769 | 0.01511 | 0.01326 | 0 | 6.6193 | -115.5579 | 6.4911 | -38.3789 | -112.8113 | -113.0081 | 0.050354 | 6.6788 | -115.4114 | 6.5186 | -105.6015 | -112.6007 | -112.7609 | -0.19226 | -0.041198 | -0.16937 | -0.12817 | 76.9867 | 77.179 | 2.3575 | 2.5543 | 2.5497 | 2.7466 | -0.045776 | 0.018311 | -0.14191 | -0.16022 | 9.7641 | 9.8099 | 2.6047 | 2.7649 | 2.6505 | 2.8107
2226.1719 | 122.2046 | -117.4667 | 4.6051 | -0.054245 | 0.014159 | 0.097384 | 0.085582 | 0.00082671 | 0.016402 | 0.010724 | -0.047089 | 0.074087 | 0.089487 | 0.079 | -0.0012497 | 0.01276 | 0.0066373 | 0.022888 | 4.6005 | -117.5308 | 4.3579 | -106.5079 | -114.9582 | -116.4505 | -0.050354 | 4.6417 | -117.4759 | 4.4037 | -107.6065 | -114.3539 | -114.8117 | -0.064087 | -0.004578 | -0.24719 | -0.24261 | 10.9589 | 11.0229 | 1.0162 | 2.5085 | 1.0803 | 2.5726 | -0.009155 | 0.036621 | -0.20142 | -0.23804 | 9.8602 | 9.8694 | 2.655 | 3.1128 | 2.6642 | 3.1219
2227.1719 | 122.2046 | -117.5262 | 4.5731 | 0.01442 | 0.011595 | 0.092898 | 0.085368 | -0.0013095 | 0.016616 | 0.012006 | -0.071503 | 0.073018 | 0.090128 | 0.080922 | 0.00088648 | 0.012119 | 0.012405 | -0.07782 | 4.747 | -117.5583 | 4.4037 | -107.4921 | -114.8895 | -116.0614 | -0.11902 | 4.6005 | -117.5171 | 4.3579 | -107.6889 | -114.5691 | -114.8849 | -0.032044 | 0.17395 | -0.16937 | -0.34332 | 10.0342 | 10.0662 | 1.4648 | 2.6367 | 1.4969 | 2.6688 | 0.009155 | 0.027466 | -0.21515 | -0.24261 | 9.8373 | 9.8282 | 2.6413 | 2.9572 | 2.6321 | 2.948
2228.1719 | 122.2046 | -117.3569 | 4.7516 | -0.11528 | 0.015013 | 0.095035 | 0.085795 | -0.00088229 | 0.015548 | 0.012433 | -0.092865 | 0.07195 | 0.089487 | 0.083272 | -0.0018905 | 0.014896 | 0.0081327 | -0.054932 | 4.6829 | -117.5766 | 4.4449 | -107.016 | -114.8758 | -115.773 | -0.054932 | 4.7241 | -117.3843 | 4.5044 | -107.5607 | -114.5325 | -114.7202 | -0.21973 | -0.068665 | -0.3067 | -0.23804 | 10.3409 | 10.5606 | 1.5839 | 2.4811 | 1.8036 | 2.7008 | -0.027465 | -0.027466 | -0.24719 | -0.21973 | 9.7961 | 9.8236 | 2.6367 | 2.8244 | 2.6642 | 2.8519
2229.1719 | 122.168 | -117.1875 | 4.8935 | -0.13817 | 0.013091 | 0.093326 | 0.085582 | 0.0010403 | 0.016616 | 0.010937 | -0.12033 | 0.072805 | 0.088632 | 0.076864 | 0.0015275 | 0.012973 | 0.0074917 | -0.07782 | 4.8523 | -117.3843 | 4.6051 | -106.4255 | -114.7339 | -115.6082 | -0.17395 | 4.8889 | -117.215 | 4.6783 | -107.3776 | -114.2624 | -114.592 | -0.19684 | -0.041199 | -0.28839 | -0.24719 | 10.762 | 10.9589 | 1.5793 | 2.4536 | 1.7761 | 2.6505 | -0.027466 | -0.004578 | -0.21515 | -0.21057 | 9.8099 | 9.8373 | 2.5955 | 2.9251 | 2.623 | 2.9526
2230.1719 | 122.2183 | -117.0822 | 4.9805 | -0.077133 | 0.015013 | 0.092471 | 0.084727 | 0.00061311 | 0.016829 | 0.011151 | -0.047089 | 0.073018 | 0.090555 | 0.086263 | -0.00039512 | 0.013401 | 0.014328 | 0.027466 | 5.0217 | -117.2241 | 4.7791 | -106.2698 | -114.505 | -115.4846 | 0.004578 | 5.0034 | -117.0593 | 4.8569 | -107.2037 | -114.0976 | -114.4226 | -0.14191 | 0.041198 | -0.20142 | -0.24261 | 10.8124 | 10.9543 | 1.5976 | 2.5772 | 1.7395 | 2.7191 | 0.022888 | 0.022888 | -0.1236 | -0.14648 | 9.8785 | 9.8557 | 2.6596 | 2.9846 | 2.6367 | 2.9617
2231.1719 | 122.2366 | -117.1646 | 4.9301 | -0.11528 | 0.010741 | 0.09418 | 0.0843 | 0.0016813 | 0.016188 | 0.0075196 | -0.044037 | 0.073873 | 0.086496 | 0.079427 | -0.00082252 | 0.013614 | 0.015823 | -0.027466 | 4.8935 | -117.3569 | 4.6829 | -106.4438 | -114.6011 | -115.4892 | 0.009155 | 4.9438 | -117.1509 | 4.7607 | -107.3456 | -114.1251 | -114.4913 | -0.19226 | -0.036621 | -0.24719 | -0.21057 | 10.7208 | 10.9131 | 1.6754 | 2.5635 | 1.8677 | 2.7557 | 0.013733 | 0.013733 | -0.16937 | -0.18311 | 9.819 | 9.8053 | 2.6733 | 3.0396 | 2.6596 | 3.0258
2232.1719 | 122.2275 | -117.1921 | 4.8752 | -0.061874 | 0.014373 | 0.093112 | 0.085795 | 0.0021085 | 0.014479 | 0.010083 | -0.08371 | 0.072378 | 0.090128 | 0.077932 | 0.0013139 | 0.011692 | 0.016464 | 0.14648 | 4.8798 | -117.2516 | 4.6875 | -106.4758 | -114.537 | -115.535 | 0.022888 | 4.8798 | -117.2058 | 4.7058 | -107.3822 | -113.9648 | -114.5508 | -0.059509 | 0.004578 | -0.18768 | -0.19226 | 10.7162 | 10.7758 | 1.6571 | 2.655 | 1.7166 | 2.7145 | -0.013733 | 0.004578 | -0.16937 | -0.17395 | 9.8099 | 9.8236 | 2.6413 | 3.2272 | 2.655 | 3.241
2233.1719 | 122.1588 | -116.8213 | 5.3055 | -0.077133 | 0.011382 | 0.090335 | 0.083873 | 0.00061311 | 0.014266 | 0.010083 | -0.03183 | 0.072805 | 0.091409 | 0.073873 | -0.0010361 | 0.013401 | 0.011764 | -0.036621 | 5.2368 | -117.009 | 4.9942 | -106.0959 | -114.3631 | -115.1459 | -0.013733 | 5.2551 | -116.8121 | 5.0217 | -106.8741 | -113.7222 | -114.2441 | -0.18768 | -0.068665 | -0.31128 | -0.24261 | 10.7254 | 10.9131 | 1.6754 | 2.4582 | 1.8631 | 2.6459 | 0.009155 | -0.050354 | -0.28381 | -0.23346 | 9.9472 | 9.9381 | 2.5772 | 3.0991 | 2.5681 | 3.0899
2234.1719 | 122.2183 | -116.6702 | 5.4062 | -0.11528 | 0.013518 | 0.093539 | 0.085795 | 0.0018949 | 0.014907 | 0.013928 | -0.092865 | 0.071737 | 0.0897 | 0.08434 | -0.0014633 | 0.014255 | 0.014328 | -0.041199 | 5.4245 | -116.8121 | 5.1315 | -106.3614 | -114.1617 | -115.0085 | -0.032043 | 5.4062 | -116.7068 | 5.1773 | -106.9382 | -113.7085 | -114.0472 | -0.14191 | 0.018311 | -0.27466 | -0.29297 | 10.3088 | 10.4507 | 1.6617 | 2.5085 | 1.8036 | 2.6505 | -0.036621 | 0 | -0.22888 | -0.22888 | 9.7321 | 9.7687 | 2.623 | 2.9617 | 2.6596 | 2.9984
2235.1719 | 122.1863 | -116.629 | 5.452 | -0.10765 | 0.01245 | 0.093753 | 0.084513 | -0.00045509 | 0.015975 | 0.011365 | -0.028778 | 0.070028 | 0.090555 | 0.080282 | 0.00088648 | 0.011692 | 0.016037 | -0.045776 | 5.4428 | -116.7618 | 5.1636 | -106.3248 | -114.0884 | -115.0452 | -0.045776 | 5.4932 | -116.6336 | 5.2231 | -106.8695 | -113.6307 | -114.0289 | -0.13275 | -0.009155 | -0.28839 | -0.27924 | 10.3043 | 10.437 | 1.5839 | 2.5406 | 1.7166 | 2.6733 | -0.004578 | 0.041199 | -0.22888 | -0.27008 | 9.7595 | 9.7641 | 2.6001 | 2.9984 | 2.6047 | 3.0029
2236.1719 | 122.2229 | -116.7023 | 5.3741 | -0.10002 | 0.012236 | 0.092257 | 0.084727 | -0.0032321 | 0.016616 | 0.010937 | -0.040985 | 0.071309 | 0.090341 | 0.081136 | -0.0023177 | 0.013187 | 0.012192 | -0.027466 | 5.3284 | -116.8991 | 5.1224 | -106.3705 | -114.2395 | -115.0497 | 0.009155 | 5.3741 | -116.7114 | 5.1636 | -106.929 | -113.5666 | -114.0656 | -0.19684 | -0.045777 | -0.25177 | -0.20599 | 10.3317 | 10.5286 | 1.6525 | 2.4628 | 1.8494 | 2.6596 | -0.009155 | 0 | -0.21057 | -0.21057 | 9.7733 | 9.7824 | 2.6367 | 3.1357 | 2.6459 | 3.1448
2237.1719 | 122.2549 | -116.6702 | 5.4245 | -0.061874 | 0.012664 | 0.091617 | 0.083873 | -0.0036593 | 0.014907 | 0.011578 | -0.028778 | 0.070669 | 0.087564 | 0.081136 | 0.00067288 | 0.013187 | 0.012405 | -0.036621 | 5.4245 | -116.8396 | 5.1544 | -106.4575 | -114.2258 | -115.0406 | -0.027466 | 5.4153 | -116.6702 | 5.1865 | -106.9016 | -113.6078 | -114.0427 | -0.16937 | 0 | -0.27008 | -0.27008 | 10.2127 | 10.3821 | 1.6296 | 2.4445 | 1.799 | 2.6138 | 0 | -0.009156 | -0.23804 | -0.22888 | 9.7687 | 9.7687 | 2.6276 | 3.0624 | 2.6276 | 3.0624
2238.1719 | 122.2183 | -116.7435 | 5.3284 | -0.038986 | 0.013304 | 0.092471 | 0.085368 | 0.00039951 | 0.013198 | 0.013928 | -0.034882 | 0.071096 | 0.0897 | 0.082631 | 0.00045928 | 0.013401 | 0.0072781 | -0.054932 | 5.3329 | -116.922 | 5.0674 | -106.5445 | -114.267 | -115.1001 | 0.068665 | 5.3238 | -116.7343 | 5.0995 | -107.0068 | -113.9145 | -114.1571 | -0.17853 | 0.004578 | -0.26093 | -0.2655 | 10.199 | 10.3775 | 1.6434 | 2.4765 | 1.8219 | 2.655 | 0.009155 | -0.004577 | -0.22888 | -0.22431 | 9.7366 | 9.7275 | 2.5864 | 2.829 | 2.5772 | 2.8198
2239.1719 | 122.1771 | -116.9037 | 5.191 | -0.077133 | 0.0148 | 0.093539 | 0.085368 | -0.00088229 | 0.014907 | 0.0092286 | -0.089813 | 0.073232 | 0.089273 | 0.087972 | -0.0018905 | 0.011478 | 0.013046 | -0.082397 | 5.1132 | -117.0868 | 4.9118 | -106.8375 | -114.5279 | -115.2878 | -0.091553 | 5.1727 | -116.9128 | 4.9484 | -107.1899 | -113.5391 | -114.3127 | -0.1831 | -0.07782 | -0.27924 | -0.20142 | 10.0662 | 10.2493 | 1.6159 | 2.3758 | 1.799 | 2.5589 | -0.009155 | -0.018311 | -0.24261 | -0.2243 | 9.7137 | 9.7229 | 2.5909 | 3.3646 | 2.6001 | 3.3737
2240.1719 | 122.1817 | -116.7068 | 5.4062 | -0.084763 | 0.01245 | 0.093112 | 0.082377 | -0.00024129 | 0.016829 | 0.010937 | 0.0047913 | 0.072378 | 0.088632 | 0.077718 | 0.0011001 | 0.013187 | 0.013473 | -0.004578 | 5.3833 | -116.8121 | 5.1315 | -106.2012 | -114.2075 | -115.0909 | 0.013733 | 5.4062 | -116.6977 | 5.1361 | -106.9702 | -113.6673 | -114.093 | -0.10529 | -0.022888 | -0.27466 | -0.25177 | 10.5057 | 10.611 | 1.6159 | 2.4994 | 1.7212 | 2.6047 | 0.009155 | 0 | -0.27008 | -0.27008 | 9.7366 | 9.7275 | 2.6138 | 3.0396 | 2.6047 | 3.0304
2241.1719 | 122.1863 | -116.597 | 5.5069 | -0.13817 | 0.013518 | 0.092898 | 0.084086 | -0.0015231 | 0.012984 | 0.007306 | -0.044037 | 0.069173 | 0.087564 | 0.081136 | 0.00045928 | 0.012332 | 0.013901 | -0.059509 | 5.4703 | -116.7664 | 5.2231 | -105.4001 | -114.1891 | -114.9307 | 0.21973 | 5.5344 | -116.5329 | 5.2689 | -106.929 | -113.5895 | -113.9969 | -0.16937 | -0.036621 | -0.28381 | -0.24719 | 11.1969 | 11.3663 | 1.6663 | 2.4078 | 1.8356 | 2.5772 | 0.064087 | 0.027466 | -0.23804 | -0.2655 | 9.668 | 9.6039 | 2.6001 | 3.0075 | 2.536 | 2.9434
2242.1719 | 122.1817 | -116.4413 | 5.6763 | -0.10002 | 0.01245 | 0.092257 | 0.084727 | -0.0034457 | 0.015548 | 0.0094422 | -0.010467 | 0.072805 | 0.090341 | 0.083486 | -0.0018905 | 0.0084873 | 0.011978 | -0.045776 | 5.5847 | -116.6565 | 5.3467 | -105.8029 | -114.0198 | -114.8163 | 0.022888 | 5.6076 | -116.4688 | 5.5161 | -106.7963 | -113.3789 | -113.8962 | -0.21515 | -0.091553 | -0.32959 | -0.23804 | 10.6384 | 10.8536 | 1.6251 | 2.4216 | 1.8402 | 2.6367 | -0.027466 | -0.068665 | -0.16022 | -0.091553 | 9.6451 | 9.6725 | 2.5452 | 3.0624 | 2.5726 | 3.0899
2243.1719 | 122.2137 | -116.7023 | 5.3696 | -0.19157 | 0.014159 | 0.09183 | 0.085154 | 0.00018591 | 0.014907 | 0.010297 | -0.040985 | 0.070455 | 0.0897 | 0.081136 | -0.0014633 | 0.012546 | 0.011337 | -0.004578 | 5.3833 | -116.8625 | 5.1041 | -105.2216 | -114.3173 | -115.2008 | -0.009155 | 5.3513 | -116.7435 | 5.1407 | -107.0984 | -113.7588 | -114.1113 | -0.16022 | 0.013733 | -0.2655 | -0.27924 | 11.4807 | 11.6409 | 1.5015 | 2.3849 | 1.6617 | 2.5452 | -0.041198 | -0.018311 | -0.22888 | -0.21057 | 9.6039 | 9.6451 | 2.5909 | 2.9434 | 2.6321 | 2.9846
2244.1719 | 122.2092 | -116.423 | 5.6259 | -0.016098 | 0.013732 | 0.095462 | 0.086436 | 0.0053129 | 0.013839 | 0.014355 | -0.08371 | 0.071737 | 0.088419 | 0.08434 | -0.00060892 | 0.01276 | 0.011123 | -0.013733 | 5.6671 | -116.5924 | 5.4337 | -105.6244 | -114.0152 | -114.83 | -0.027466 | 5.6351 | -116.4139 | 5.4565 | -106.81 | -113.4521 | -113.8367 | -0.16937 | 0.041198 | -0.19226 | -0.23346 | 10.7986 | 10.968 | 1.593 | 2.4078 | 1.7624 | 2.5772 | 0.009156 | 0.009155 | -0.16937 | -0.17853 | 9.613 | 9.6039 | 2.5864 | 2.9709 | 2.5772 | 2.9617
2245.1719 | 122.2 | -116.6107 | 5.4749 | -0.12291 | 0.013091 | 0.093326 | 0.083445 | -0.0013095 | 0.014052 | 0.0096558 | 0.023102 | 0.071737 | 0.090982 | 0.076864 | -0.0021041 | 0.011692 | 0.0096279 | -0.036621 | 5.4428 | -116.8304 | 5.1727 | -105.7068 | -114.2029 | -114.7202 | -0.10986 | 5.4932 | -116.6153 | 5.2505 | -107.0297 | -113.5986 | -114.0289 | -0.21973 | -0.032044 | -0.30212 | -0.27008 | 10.9039 | 11.1237 | 1.8906 | 2.4078 | 2.1103 | 2.6276 | -0.004577 | 0.01831 | -0.2243 | -0.24261 | 9.581 | 9.5856 | 2.5818 | 3.0121 | 2.5864 | 3.0167
2246.1719 | 122.1954 | -116.4001 | 5.7266 | -0.077133 | 0.013304 | 0.091617 | 0.082377 | -0.0013095 | 0.015761 | 0.011578 | 0.013947 | 0.071309 | 0.089487 | 0.084981 | -0.0012497 | 0.0099827 | 0.011123 | 0.004578 | 5.658 | -116.5375 | 5.5161 | -105.1025 | -113.9053 | -114.6561 | 0.013733 | 5.7083 | -116.3818 | 5.5069 | -106.7551 | -113.4567 | -113.7955 | -0.13733 | -0.068665 | -0.21057 | -0.14191 | 11.2976 | 11.4349 | 1.7441 | 2.4948 | 1.8814 | 2.6321 | 0.01831 | -0.018311 | -0.21973 | -0.20142 | 9.6451 | 9.6268 | 2.6047 | 2.9434 | 2.5864 | 2.9251
2247.1719 | 122.2183 | -116.4276 | 5.6625 | -0.16106 | 0.016509 | 0.091403 | 0.082164 | -0.0043003 | 0.013198 | 0.0098694 | 0.0017395 | 0.070028 | 0.091623 | 0.078145 | 3.2078e-05 | 0.01041 | 0.014541 | 0.054932 | 5.6168 | -116.5283 | 5.484 | -105.5237 | -113.8733 | -114.592 | 0.082397 | 5.6808 | -116.4185 | 5.4749 | -106.8054 | -113.4155 | -113.8 | -0.10071 | -0.045777 | -0.17853 | -0.13275 | 10.9039 | 11.0046 | 1.8356 | 2.5543 | 1.9363 | 2.655 | 0.009155 | 0.01831 | -0.18768 | -0.20599 | 9.6222 | 9.613 | 2.6276 | 3.0121 | 2.6184 | 3.0029
2248.1719 | 122.1451 | -116.4459 | 5.6534 | -0.077133 | 0.01245 | 0.089908 | 0.083018 | 0.0014675 | 0.016402 | 0.0096558 | -0.034882 | 0.071309 | 0.091623 | 0.081991 | -0.00039512 | 0.01041 | 0.01561 | -0.036621 | 5.6213 | -116.5833 | 5.4062 | -105.2032 | -113.9969 | -114.7659 | 0.10529 | 5.6351 | -116.423 | 5.4016 | -106.8787 | -113.4109 | -113.8962 | -0.13733 | -0.032043 | -0.24719 | -0.21515 | 11.2427 | 11.38 | 1.68 | 2.449 | 1.8173 | 2.5864 | 0.022888 | -0.01831 | -0.25177 | -0.23346 | 9.5673 | 9.5444 | 2.5497 | 3.035 | 2.5269 | 3.0121
2249.1719 | 122.168 | -116.5466 | 5.484 | -0.10765 | 0.014159 | 0.093326 | 0.084941 | 0.0029629 | 0.014266 | 0.011151 | -0.12033 | 0.07195 | 0.089273 | 0.080709 | -0.0010361 | 0.012119 | 0.011337 | -0.013733 | 5.4062 | -116.8076 | 5.2551 | -105.1117 | -114.18 | -114.9078 | -0.045776 | 5.484 | -116.6153 | 5.2551 | -107.0709 | -113.6902 | -113.9969 | -0.26093 | -0.07782 | -0.22888 | -0.15106 | 11.4349 | 11.6959 | 1.6388 | 2.3666 | 1.8997 | 2.6276 | -0.068664 | 0 | -0.22888 | -0.22888 | 9.4757 | 9.5444 | 2.5497 | 2.8564 | 2.6184 | 2.9251
2250.1719 | 122.2366 | -116.4551 | 5.6168 | -0.12291 | 0.014586 | 0.090762 | 0.085154 | 0.0014675 | 0.015761 | 0.010937 | -0.044037 | 0.070028 | 0.088846 | 0.084768 | 0.0011001 | 0.013187 | 0.011764 | -0.036621 | 5.5389 | -116.6931 | 5.2826 | -105.336 | -114.093 | -114.917 | -0.004578 | 5.5527 | -116.5421 | 5.3146 | -106.9931 | -113.5849 | -113.8092 | -0.23804 | -0.07782 | -0.33417 | -0.25635 | 11.1191 | 11.3571 | 1.5381 | 2.3621 | 1.7761 | 2.6001 | -0.086975 | -0.064087 | -0.30212 | -0.23804 | 9.462 | 9.5489 | 2.6459 | 2.8702 | 2.7328 | 2.9572
2251.1719 | 122.1863 | -116.6153 | 5.5069 | -0.038986 | 0.013304 | 0.090549 | 0.084086 | 0.0018949 | 0.01747 | 0.0098694 | -0.013519 | 0.0696 | 0.089059 | 0.082418 | 0.0017411 | 0.011264 | 0.012192 | -0.018311 | 5.452 | -116.7618 | 5.2094 | -106.1691 | -114.1891 | -114.9994 | -0.009155 | 5.4749 | -116.6061 | 5.246 | -107.0435 | -113.6307 | -114.061 | -0.14648 | -0.054932 | -0.29755 | -0.24261 | 10.4462 | 10.5927 | 1.6159 | 2.4261 | 1.7624 | 2.5726 | 0.009155 | -0.032043 | -0.26093 | -0.22888 | 9.5718 | 9.5627 | 2.5543 | 2.9846 | 2.5452 | 2.9755
2252.1719 | 122.1634 | -116.391 | 5.7358 | -0.069504 | 0.012877 | 0.09183 | 0.084086 | 0.00061311 | 0.014479 | 0.01051 | -0.034882 | 0.069173 | 0.087778 | 0.078786 | 0.00024568 | 0.012332 | 0.013046 | 0.018311 | 5.6259 | -116.5787 | 5.4474 | -105.4504 | -113.9603 | -114.7293 | 0.036621 | 5.6625 | -116.3681 | 5.484 | -106.7047 | -113.3881 | -113.8412 | -0.18768 | -0.10986 | -0.28839 | -0.17853 | 10.9406 | 11.1282 | 1.6617 | 2.4307 | 1.8494 | 2.6184 | 0.022888 | -0.073242 | -0.25177 | -0.17853 | 9.6863 | 9.6634 | 2.5497 | 3.0029 | 2.5269 | 2.98
2253.1719 | 122.2687 | -116.2216 | 5.809 | -0.00083923 | 0.016722 | 0.094821 | 0.084086 | -2.7692e-05 | 0.015761 | 0.010297 | -0.0013123 | 0.071523 | 0.088205 | 0.081136 | 3.2078e-05 | 0.011905 | 0.01091 | -0.13733 | 5.7678 | -116.4001 | 5.7037 | -105.1895 | -113.7405 | -114.5554 | 0.11444 | 5.809 | -116.2216 | 5.7175 | -106.5857 | -113.2233 | -113.5986 | -0.17853 | -0.041199 | -0.10529 | -0.064087 | 11.0321 | 11.2106 | 1.6663 | 2.4811 | 1.8448 | 2.6596 | 0 | 0 | -0.091553 | -0.091553 | 9.6359 | 9.6359 | 2.623 | 2.9984 | 2.623 | 2.9984
2254.1719 | 122.2183 | -116.1255 | 5.9463 | -0.12291 | 0.011595 | 0.092898 | 0.087718 | -0.0019503 | 0.015548 | 0.012006 | -0.086761 | 0.070882 | 0.087564 | 0.080068 | -0.0031723 | 0.012332 | 0.012619 | 0.009155 | 5.9097 | -116.3223 | 5.7312 | -104.8965 | -113.7177 | -114.3311 | 0.022888 | 5.9463 | -116.1346 | 5.7632 | -106.5491 | -113.118 | -113.4979 | -0.19684 | -0.036621 | -0.21515 | -0.17853 | 11.2289 | 11.4258 | 1.7944 | 2.4078 | 1.9913 | 2.6047 | -0.009156 | 0 | -0.1831 | -0.1831 | 9.5764 | 9.5856 | 2.6276 | 3.0075 | 2.6367 | 3.0167
2255.1719 | 122.2778 | -116.5375 | 5.5023 | -0.11528 | 0.014373 | 0.090549 | 0.085154 | 0.0018949 | 0.016188 | 0.0085876 | 0.0017395 | 0.067891 | 0.087137 | 0.07195 | -0.0016769 | 0.013401 | 0.0098417 | 0.032043 | 5.452 | -116.7114 | 5.365 | -106.3889 | -114.0518 | -114.8621 | 0.1236 | 5.4932 | -116.5558 | 5.365 | -106.9565 | -113.5391 | -113.9145 | -0.17395 | -0.050354 | -0.13733 | -0.086975 | 10.1486 | 10.3226 | 1.6754 | 2.4857 | 1.8494 | 2.6596 | -0.01831 | -0.009155 | -0.13733 | -0.12817 | 9.581 | 9.5993 | 2.623 | 2.9984 | 2.6413 | 3.0167
2256.1719 | 122.2778 | -116.7664 | 5.2094 | -0.12291 | 0.012664 | 0.092044 | 0.084941 | -0.00088229 | 0.014052 | 0.010297 | -0.028778 | 0.06896 | 0.086923 | 0.074727 | -0.00018152 | 0.01041 | 0.01561 | -0.041199 | 5.1636 | -117.041 | 5.1132 | -106.3797 | -114.2441 | -115.0726 | 0.21973 | 5.2048 | -116.8076 | 5.1453 | -107.1762 | -113.7131 | -114.1754 | -0.27466 | -0.045777 | -0.096131 | -0.050354 | 10.3867 | 10.6613 | 1.6937 | 2.5223 | 1.9684 | 2.7969 | -0.041199 | -0.004578 | -0.064087 | -0.059509 | 9.5901 | 9.6313 | 2.5909 | 3.0533 | 2.6321 | 3.0945
2257.1719 | 122.1954 | -116.9312 | 5.1544 | -0.10765 | 0.01245 | 0.091403 | 0.086009 | 0.0012539 | 0.016188 | 0.0096558 | -0.03183 | 0.068105 | 0.086923 | 0.080495 | -0.00082252 | 0.01041 | 0.012619 | 0 | 5.0903 | -117.0959 | 4.8981 | -106.1737 | -114.5279 | -115.2969 | 0.009155 | 5.1407 | -116.9495 | 4.9255 | -107.4188 | -114.0289 | -114.3906 | -0.16479 | -0.064087 | -0.25635 | -0.19226 | 10.7574 | 10.9222 | 1.6342 | 2.4033 | 1.799 | 2.5681 | -0.018311 | -0.013733 | -0.22888 | -0.21515 | 9.5123 | 9.5306 | 2.5406 | 2.9022 | 2.5589 | 2.9205
2258.1719 | 122.2641 | -116.7618 | 5.2963 | -0.18394 | 0.015227 | 0.091403 | 0.083232 | 0.0018949 | 0.014693 | 0.0077332 | -0.016571 | 0.0696 | 0.08735 | 0.085195 | 0.00045928 | 0.01511 | 0.010483 | 0.068665 | 5.2322 | -116.9174 | 5.1636 | -106.5262 | -114.2944 | -115.0726 | 0.086975 | 5.2734 | -116.7664 | 5.1636 | -107.1579 | -113.7497 | -114.1251 | -0.15564 | -0.064087 | -0.13275 | -0.068665 | 10.2356 | 10.3912 | 1.6891 | 2.4673 | 1.8448 | 2.623 | -0.004577 | -0.022889 | -0.13275 | -0.10986 | 9.6039 | 9.6085 | 2.6367 | 3.0121 | 2.6413 | 3.0167
2259.1719 | 122.1725 | -116.8076 | 5.2963 | -0.069504 | 0.012236 | 0.092471 | 0.083445 | -2.7692e-05 | 0.016402 | 0.013287 | -0.053192 | 0.0696 | 0.088205 | 0.079213 | 0.00088648 | 0.011264 | 0.010055 | -0.018311 | 5.1407 | -117.0731 | 4.921 | -106.778 | -114.5096 | -115.2695 | 0.059509 | 5.2094 | -116.835 | 5.0446 | -107.3135 | -113.8824 | -114.2853 | -0.2655 | -0.15564 | -0.37537 | -0.21973 | 10.0296 | 10.2951 | 1.5381 | 2.298 | 1.8036 | 2.5635 | -0.027466 | -0.086975 | -0.25177 | -0.16479 | 9.494 | 9.5215 | 2.5223 | 2.9251 | 2.5497 | 2.9526
2260.1719 | 122.319 | -116.6473 | 5.365 | -0.10765 | 0.010741 | 0.091617 | 0.083232 | -0.0010959 | 0.014693 | 0.0085876 | -0.010467 | 0.067251 | 0.088419 | 0.084127 | 0.00088648 | 0.0095555 | 0.013473 | 0 | 5.4291 | -116.7526 | 5.2597 | -106.6223 | -114.1296 | -114.9033 | 0.10071 | 5.4016 | -116.6061 | 5.3238 | -107.1213 | -113.7405 | -113.9145 | -0.10529 | 0.064087 | -0.10528 | -0.16937 | 10.025 | 10.1303 | 1.7441 | 2.5177 | 1.8494 | 2.623 | 0.041199 | 0.036621 | -0.041198 | -0.077819 | 9.5261 | 9.4849 | 2.7328 | 2.9068 | 2.6916 | 2.8656
2261.1719 | 122.2687 | -116.5878 | 5.4291 | -0.1458 | 0.013732 | 0.09183 | 0.084086 | -0.0010959 | 0.01512 | 0.0085876 | -0.0074159 | 0.0696 | 0.086282 | 0.084768 | 0.00067288 | 0.010623 | 0.013046 | -0.032043 | 5.3604 | -116.7801 | 5.2689 | -106.7001 | -114.1708 | -114.917 | 0.068665 | 5.4062 | -116.6153 | 5.3467 | -107.1304 | -113.6261 | -113.9969 | -0.19226 | -0.068664 | -0.16022 | -0.091553 | 9.8877 | 10.08 | 1.6708 | 2.417 | 1.8631 | 2.6093 | -0.027465 | -0.022888 | -0.082397 | -0.059509 | 9.4574 | 9.4849 | 2.5909 | 2.9617 | 2.6184 | 2.9892
2262.1719 | 122.3099 | -116.6061 | 5.397 | -0.12291 | 0.012877 | 0.092685 | 0.084086 | -0.0043003 | 0.017043 | 0.0090148 | -0.022675 | 0.066823 | 0.085428 | 0.081777 | -0.0016769 | 0.011905 | 0.011764 | -0.009155 | 5.3604 | -116.7892 | 5.3009 | -106.5445 | -114.1388 | -114.9216 | 0.15564 | 5.397 | -116.6107 | 5.3375 | -107.0938 | -113.6261 | -114.0106 | -0.18311 | -0.036621 | -0.096131 | -0.05951 | 10.0616 | 10.2448 | 1.6846 | 2.4673 | 1.8677 | 2.6505 | -0.004578 | 0 | -0.05951 | -0.05951 | 9.5123 | 9.5169 | 2.5955 | 2.98 | 2.6001 | 2.9846
2263.1719 | 122.3053 | -116.6519 | 5.3696 | -0.15343 | 0.012877 | 0.090335 | 0.085795 | 0.0014675 | 0.014479 | 0.011792 | -0.025726 | 0.068532 | 0.086496 | 0.077504 | -0.00018152 | 0.0093417 | 0.011551 | 0 | 5.3192 | -116.8533 | 5.2277 | -106.7734 | -114.2166 | -115.0177 | 0.07782 | 5.3558 | -116.6656 | 5.2505 | -107.2174 | -113.7222 | -114.1022 | -0.20142 | -0.050354 | -0.14191 | -0.091553 | 9.8785 | 10.08 | 1.6342 | 2.4353 | 1.8356 | 2.6367 | -0.013732 | -0.013733 | -0.11902 | -0.10529 | 9.4345 | 9.4482 | 2.5497 | 2.9297 | 2.5635 | 2.9434
2264.1875 | 122.2961 | -116.629 | 5.3558 | -0.1458 | 0.014159 | 0.092044 | 0.086009 | 0.0031765 | 0.014479 | 0.012006 | -0.080658 | 0.067891 | 0.085214 | 0.082631 | -0.0016769 | 0.012546 | 0.0083463 | -0.041199 | 5.3146 | -116.8396 | 5.2643 | -106.5125 | -114.212 | -114.9673 | -0.032043 | 5.3284 | -116.7206 | 5.2872 | -107.1625 | -113.7817 | -114.0656 | -0.21057 | -0.041199 | -0.091553 | -0.050354 | 10.1166 | 10.3271 | 1.6617 | 2.417 | 1.8723 | 2.6276 | -0.091553 | -0.027466 | -0.068665 | -0.041199 | 9.4666 | 9.5581 | 2.5635 | 2.8473 | 2.655 | 2.9388
2265.1719 | 122.2687 | -116.7114 | 5.3238 | 0.029678 | 0.010741 | 0.093966 | 0.085368 | -0.0051547 | 0.014479 | 0.0094422 | -0.034882 | 0.067891 | 0.087991 | 0.079641 | -0.00082252 | 0.010196 | 0.010055 | -0.073242 | 5.2094 | -116.9312 | 5.1361 | -106.6086 | -114.3311 | -115.0772 | 0.13275 | 5.3055 | -116.7297 | 5.2139 | -107.309 | -113.9145 | -114.1617 | -0.21973 | -0.11444 | -0.18768 | -0.073243 | 10.1028 | 10.3226 | 1.6342 | 2.3804 | 1.8539 | 2.6001 | -0.01831 | -0.018311 | -0.10986 | -0.091553 | 9.4025 | 9.4208 | 2.5497 | 2.7969 | 2.5681 | 2.8152
2266.1719 | 122.3007 | -116.7435 | 5.2689 | -0.10765 | 0.013091 | 0.091403 | 0.085368 | -0.0040867 | 0.014266 | 0.0092286 | -0.03183 | 0.067891 | 0.08735 | 0.077932 | 0.00024568 | 0.013614 | 0.011337 | -0.022888 | 5.2277 | -116.9632 | 5.1224 | -105.9402 | -114.3356 | -115.1093 | 0.013733 | 5.278 | -116.7755 | 5.1682 | -107.2815 | -113.6124 | -114.1891 | -0.21973 | -0.041199 | -0.14649 | -0.10529 | 10.8032 | 11.0229 | 1.6342 | 2.4078 | 1.8539 | 2.6276 | -0.032044 | 0.009155 | -0.10071 | -0.10986 | 9.462 | 9.494 | 2.5543 | 3.1311 | 2.5864 | 3.1631
2267.1719 | 122.113 | -116.8579 | 5.191 | -0.15343 | 0.013091 | 0.091617 | 0.084727 | 0.00018591 | 0.01512 | 0.0098694 | -0.062347 | 0.068105 | 0.088205 | 0.079427 | -0.0010361 | 0.012332 | 0.011551 | -0.086975 | 5.1636 | -117.0227 | 5.0446 | -106.6269 | -114.4455 | -115.1825 | -0.073242 | 5.2048 | -116.8396 | 5.0812 | -107.428 | -113.8687 | -114.267 | -0.16479 | -0.027466 | -0.14648 | -0.11902 | 10.231 | 10.3958 | 1.6754 | 2.4124 | 1.8402 | 2.5772 | 0.01831 | 0.013733 | -0.10986 | -0.1236 | 9.4299 | 9.4116 | 2.5909 | 2.9892 | 2.5726 | 2.9709
2268.1719 | 122.2595 | -116.7618 | 5.2551 | -0.16869 | 0.014159 | 0.093112 | 0.086009 | -0.0036593 | 0.014266 | 0.010297 | -0.071503 | 0.065969 | 0.086282 | 0.084768 | -0.0016769 | 0.011478 | 0.012405 | -0.10071 | 5.2048 | -116.9495 | 5.0903 | -106.7917 | -114.4318 | -115.1459 | -0.050354 | 5.2368 | -116.7938 | 5.1407 | -107.373 | -113.8275 | -114.2395 | -0.18768 | -0.050354 | -0.16479 | -0.11444 | 9.9701 | 10.1578 | 1.6159 | 2.33 | 1.8036 | 2.5177 | -0.032043 | -0.018311 | -0.11444 | -0.09613 | 9.3887 | 9.4208 | 2.5223 | 2.9343 | 2.5543 | 2.9663
2269.1719 | 122.3099 | -116.6519 | 5.3375 | -0.10765 | 0.012877 | 0.090762 | 0.086436 | -0.0021641 | 0.013625 | 0.011365 | -0.056244 | 0.067678 | 0.08735 | 0.090963 | -0.0018905 | 0.010196 | 0.013046 | -0.013733 | 5.3055 | -116.8579 | 5.2094 | -106.3843 | -114.2029 | -115.0131 | 0.027466 | 5.3238 | -116.6885 | 5.2826 | -107.2403 | -113.8 | -114.0839 | -0.20599 | -0.032043 | -0.12817 | -0.09613 | 10.2676 | 10.4736 | 1.6388 | 2.449 | 1.8448 | 2.655 | -0.036621 | -0.013732 | -0.054931 | -0.041199 | 9.4116 | 9.4482 | 2.5681 | 2.8519 | 2.6047 | 2.8885
2270.1719 | 122.2504 | -116.7389 | 5.3192 | -0.13054 | 0.009032 | 0.09183 | 0.0843 | -0.00045509 | 0.010848 | 0.0096558 | -0.077606 | 0.066396 | 0.088419 | 0.085622 | -0.00060892 | 0.012332 | 0.012405 | -0.091553 | 5.2872 | -116.8762 | 5.1636 | -106.5628 | -114.3356 | -115.0772 | -0.036621 | 5.3284 | -116.716 | 5.1956 | -107.3227 | -113.8184 | -114.1937 | -0.13733 | -0.032044 | -0.15564 | -0.1236 | 10.1761 | 10.3134 | 1.6617 | 2.4033 | 1.799 | 2.5406 | 0.022889 | 0.009155 | -0.1236 | -0.13275 | 9.4162 | 9.3933 | 2.5452 | 2.9205 | 2.5223 | 2.8976
2271.1719 | 122.2641 | -116.6565 | 5.3558 | -0.24498 | 0.012023 | 0.091189 | 0.085582 | 0.0010403 | 0.015761 | 0.012219 | -0.059296 | 0.070455 | 0.086496 | 0.085409 | -0.0014633 | 0.012332 | 0.0064237 | -0.059509 | 5.3329 | -116.7892 | 5.2185 | -106.3934 | -114.2853 | -114.9628 | 0 | 5.4016 | -116.7114 | 5.2597 | -107.2586 | -113.8733 | -114.1296 | -0.13275 | -0.022888 | -0.13733 | -0.11444 | 10.2631 | 10.3958 | 1.6937 | 2.3712 | 1.8265 | 2.504 | -0.054932 | 0.045776 | -0.09613 | -0.14191 | 9.3979 | 9.4528 | 2.5269 | 2.7832 | 2.5818 | 2.8381
2272.1719 | 122.2458 | -116.7572 | 5.2551 | -0.15343 | 0.012877 | 0.092471 | 0.086436 | -0.0015231 | 0.014479 | 0.011578 | -0.077606 | 0.069173 | 0.086923 | 0.078145 | -0.0025315 | 0.010837 | 0.01091 | -0.086975 | 5.2048 | -116.922 | 5.1178 | -106.6544 | -114.3814 | -115.1459 | -0.073242 | 5.2368 | -116.803 | 5.1315 | -107.3776 | -113.7222 | -114.2441 | -0.16479 | -0.050354 | -0.13733 | -0.086975 | 10.1028 | 10.2676 | 1.6113 | 2.3758 | 1.7761 | 2.5406 | -0.045777 | -0.018311 | -0.1236 | -0.10529 | 9.3796 | 9.4254 | 2.5131 | 3.035 | 2.5589 | 3.0807
2273.1719 | 122.2504 | -116.6107 | 5.4245 | -0.16869 | 0.011809 | 0.093112 | 0.088145 | -0.0025913 | 0.016402 | 0.011151 | -0.08371 | 0.066182 | 0.08671 | 0.076436 | -0.0018905 | 0.013187 | 0.012619 | -0.07782 | 5.3833 | -116.7572 | 5.2689 | -106.5948 | -114.2166 | -114.9902 | -0.073242 | 5.4245 | -116.6245 | 5.3238 | -107.2037 | -113.6719 | -114.5874 | -0.14648 | -0.041199 | -0.15564 | -0.11444 | 10.0159 | 10.1624 | 1.6205 | 2.3941 | 1.767 | 2.5406 | -0.013733 | 0 | -0.10071 | -0.10071 | 9.407 | 9.4208 | 2.0233 | 2.9388 | 2.037 | 2.9526
2274.1719 | 122.2549 | -116.7572 | 5.2963 | -0.15343 | 0.014586 | 0.092685 | 0.084941 | -0.0051547 | 0.015548 | 0.010724 | -0.074554 | 0.06661 | 0.086496 | 0.080282 | -0.0033859 | 0.011905 | 0.013901 | 0.036621 | 5.246 | -116.9174 | 5.1636 | -106.485 | -114.299 | -115.0589 | 0.10071 | 5.2826 | -116.7664 | 5.1727 | -107.3914 | -113.7909 | -114.2212 | -0.16022 | -0.050354 | -0.13275 | -0.082398 | 10.2722 | 10.4324 | 1.6983 | 2.4582 | 1.8585 | 2.6184 | -0.009155 | -0.013733 | -0.1236 | -0.10986 | 9.3658 | 9.375 | 2.536 | 2.9663 | 2.5452 | 2.9755
2275.1719 | 122.2916 | -116.6702 | 5.3284 | -0.11528 | 0.014159 | 0.092257 | 0.085795 | 0.00018591 | 0.015548 | 0.010937 | -0.13254 | 0.067251 | 0.0897 | 0.086904 | -0.0012497 | 0.013187 | 0.014328 | -0.032043 | 5.3009 | -116.8854 | 5.246 | -106.5765 | -114.2807 | -115.0223 | -0.21973 | 5.3055 | -116.7755 | 5.2689 | -107.2998 | -113.887 | -114.3906 | -0.21515 | -0.027466 | -0.082397 | -0.054931 | 10.0937 | 10.3088 | 1.6479 | 2.3895 | 1.8631 | 2.6047 | -0.10529 | -0.022888 | -0.059509 | -0.036621 | 9.3704 | 9.4757 | 2.2797 | 2.7832 | 2.3849 | 2.8885
2276.1719 | 122.2778 | -116.6611 | 5.3375 | -0.061874 | 0.013091 | 0.094394 | 0.086436 | -0.0017367 | 0.013411 | 0.0079468 | -0.10202 | 0.06661 | 0.083719 | 0.08135 | -0.0012497 | 0.011051 | 0.0096279 | -0.059509 | 5.2597 | -116.8442 | 5.2322 | -106.5765 | -114.2624 | -115.0223 | 0.027466 | 5.3284 | -116.7068 | 5.2643 | -107.3044 | -113.8 | -114.1388 | -0.1831 | -0.077819 | -0.10529 | -0.027466 | 10.0845 | 10.2676 | 1.6388 | 2.3987 | 1.8219 | 2.5818 | -0.045776 | -0.009155 | -0.073242 | -0.064087 | 9.3567 | 9.4025 | 2.5223 | 2.861 | 2.5681 | 2.9068
2277.1719 | 122.2046 | -116.5649 | 5.452 | -0.10765 | 0.014159 | 0.093539 | 0.084727 | -0.00045509 | 0.018325 | 0.011151 | -0.065399 | 0.065969 | 0.088632 | 0.082418 | -0.0010361 | 0.011264 | 0.01091 | 0.032043 | 5.4062 | -116.748 | 5.3741 | -106.6177 | -114.1159 | -114.8712 | 0.13275 | 5.4382 | -116.5695 | 5.3833 | -107.1487 | -113.6032 | -114.0198 | -0.18311 | -0.045776 | -0.077819 | -0.032043 | 9.9472 | 10.1303 | 1.6937 | 2.449 | 1.8768 | 2.6321 | -0.004578 | -0.013733 | -0.068664 | -0.054931 | 9.4162 | 9.4208 | 2.5452 | 2.9617 | 2.5497 | 2.9663
2278.1719 | 122.2412 | -116.5787 | 5.4199 | -0.16869 | 0.01245 | 0.093326 | 0.085582 | 0.0025357 | 0.015975 | 0.01286 | -0.092865 | 0.066182 | 0.086282 | 0.086263 | -0.0016769 | 0.012332 | 0.01326 | -0.004578 | 5.3787 | -116.7664 | 5.3101 | -106.4804 | -114.1937 | -114.9582 | 0.059509 | 5.4428 | -116.5924 | 5.365 | -107.222 | -113.4247 | -114.0656 | -0.18768 | -0.041199 | -0.10986 | -0.068664 | 10.0983 | 10.2859 | 1.6205 | 2.3849 | 1.8082 | 2.5726 | -0.013733 | 0.022888 | -0.054932 | -0.07782 | 9.3567 | 9.3704 | 2.5131 | 3.154 | 2.5269 | 3.1677
2279.1875 | 122.1542 | -116.684 | 5.4382 | -0.092392 | 0.012877 | 0.093326 | 0.089 | 0.0014675 | 0.013625 | 0.0079468 | -0.18442 | 0.068746 | 0.087778 | 0.088613 | -0.00018152 | 0.013187 | 0.0081327 | -0.059509 | 5.3741 | -116.8625 | 5.1956 | -106.4301 | -114.3036 | -115.0223 | -0.07782 | 5.4199 | -116.6656 | 5.2414 | -107.1716 | -113.8092 | -114.3723 | -0.17853 | -0.064086 | -0.24261 | -0.17853 | 10.2539 | 10.4324 | 1.6617 | 2.3804 | 1.8402 | 2.5589 | 0.018311 | -0.01831 | -0.19684 | -0.17853 | 9.5123 | 9.494 | 2.3117 | 2.8748 | 2.2934 | 2.8564
2280.1719 | 122.3007 | -116.51 | 5.4886 | -0.12291 | 0.01245 | 0.092898 | 0.086222 | 0.00018591 | 0.016402 | 0.0085876 | -0.17831 | 0.067251 | 0.08671 | 0.08135 | -0.00039512 | 0.013401 | 0.012405 | -0.068665 | 5.4291 | -116.6931 | 5.4016 | -105.9952 | -114.0976 | -114.798 | -0.10529 | 5.4291 | -116.5924 | 5.4337 | -107.1716 | -113.7131 | -113.9694 | -0.18311 | -0.059509 | -0.086975 | -0.027466 | 10.5148 | 10.6979 | 1.712 | 2.4124 | 1.8951 | 2.5955 | -0.082397 | -0.059509 | -0.054931 | 0.004578 | 9.3384 | 9.4208 | 2.5406 | 2.7969 | 2.623 | 2.8793
2281.1719 | 122.2824 | -116.4917 | 5.5161 | -0.1458 | 0.012664 | 0.09183 | 0.087291 | 0.0021085 | 0.014266 | 0.0092286 | -0.14169 | 0.067037 | 0.087137 | 0.081136 | -0.0012497 | 0.012973 | 0.013901 | -0.10986 | 5.4794 | -116.7389 | 5.3101 | -106.2927 | -114.2212 | -114.9261 | -0.073242 | 5.5023 | -116.5466 | 5.4382 | -107.0938 | -113.6307 | -113.3286 | -0.24719 | -0.036621 | -0.20599 | -0.16937 | 10.199 | 10.4462 | 1.5656 | 2.2705 | 1.8127 | 2.5177 | -0.054932 | -0.013733 | -0.07782 | -0.064087 | 9.3979 | 9.4528 | 3.1631 | 2.861 | 3.2181 | 2.916
2282.1719 | 122.2412 | -116.5192 | 5.5252 | -0.1458 | 0.014373 | 0.092898 | 0.084727 | -0.0034457 | 0.017043 | 0.012006 | -0.068451 | 0.067464 | 0.085214 | 0.082204 | -0.00018152 | 0.01276 | 0.0098417 | -0.041199 | 5.4657 | -116.6656 | 5.365 | -106.2149 | -114.1296 | -114.8621 | -0.045776 | 5.5023 | -116.51 | 5.4199 | -107.1671 | -113.3972 | -114.006 | -0.14648 | -0.05951 | -0.16022 | -0.10071 | 10.3043 | 10.4507 | 1.6571 | 2.3895 | 1.8036 | 2.536 | 0.009155 | -0.022889 | -0.10529 | -0.082397 | 9.3521 | 9.343 | 2.5131 | 3.1219 | 2.504 | 3.1128
2283.1719 | 122.2733 | -116.9174 | 5.0858 | -0.18394 | 0.013732 | 0.093112 | 0.085795 | 0.0029629 | 0.01512 | 0.0085876 | -0.034882 | 0.068532 | 0.089273 | 0.081136 | -0.0023177 | 0.011478 | 0.0096279 | -0.068665 | 5.0629 | -117.1143 | 5.0079 | -106.7505 | -114.7751 | -115.2237 | -0.041199 | 5.0674 | -116.922 | 5.0537 | -107.5562 | -114.0884 | -114.3723 | -0.19684 | -0.022888 | -0.077819 | -0.054931 | 10.1669 | 10.3638 | 1.6937 | 2.1423 | 1.8906 | 2.3392 | -0.004578 | -0.01831 | -0.032043 | -0.013733 | 9.3613 | 9.3658 | 2.5452 | 2.829 | 2.5497 | 2.8336
2284.1875 | 122.1405 | -116.6931 | 5.4062 | -0.061874 | 0.012236 | 0.092044 | 0.083232 | -0.0023777 | 0.014052 | 0.01051 | -0.071503 | 0.068319 | 0.087137 | 0.081777 | -0.00039512 | 0.010196 | 0.012405 | -0.041199 | 5.3375 | -116.8579 | 5.1361 | -106.366 | -114.3219 | -115.1184 | 0 | 5.3833 | -116.6931 | 5.191 | -107.4097 | -113.8 | -114.2258 | -0.16479 | -0.068665 | -0.27008 | -0.20142 | 10.3271 | 10.4919 | 1.5747 | 2.3712 | 1.7395 | 2.536 | 0 | -0.022888 | -0.21515 | -0.19226 | 9.2834 | 9.2834 | 2.4673 | 2.8931 | 2.4673 | 2.8931
2285.1719 | 122.2961 | -116.7297 | 5.2917 | -0.17632 | 0.011809 | 0.093753 | 0.088145 | 0.0012539 | 0.014266 | 0.0098694 | -0.08371 | 0.068746 | 0.086496 | 0.07665 | -0.0021041 | 0.010623 | 0.011551 | -0.068665 | 5.2826 | -116.8762 | 5.2231 | -106.3156 | -114.3082 | -115.0497 | -0.022888 | 5.3375 | -116.6611 | 5.2963 | -107.3227 | -113.7955 | -114.1068 | -0.14648 | -0.009155 | -0.068665 | -0.05951 | 10.4141 | 10.5606 | 1.68 | 2.4216 | 1.8265 | 2.5681 | 0.068664 | 0.045776 | 0.004578 | -0.041198 | 9.407 | 9.3384 | 2.623 | 2.9343 | 2.5543 | 2.8656
2286.1719 | 122.3053 | -116.6656 | 5.3375 | -0.15343 | 0.013518 | 0.092471 | 0.087077 | -0.00088229 | 0.015334 | 0.011578 | -0.11423 | 0.066823 | 0.084787 | 0.081563 | 0.0011001 | 0.01041 | 0.010696 | -0.27008 | 5.3101 | -116.8854 | 5.2185 | -106.3248 | -114.5142 | -115.0497 | 0.004578 | 5.3329 | -116.7114 | 5.2917 | -107.3364 | -113.7955 | -114.1296 | -0.21973 | -0.027465 | -0.11902 | -0.091553 | 10.3409 | 10.5606 | 1.6159 | 2.1515 | 1.8356 | 2.3712 | -0.045777 | -0.004577 | -0.045776 | -0.041199 | 9.3292 | 9.375 | 2.536 | 2.8702 | 2.5818 | 2.916
2287.1719 | 122.3053 | -116.6519 | 5.3329 | -0.16106 | 0.01245 | 0.094607 | 0.086863 | 0.00061311 | 0.014907 | 0.011578 | -0.11728 | 0.068105 | 0.086069 | 0.08135 | 0.00024568 | 0.011692 | 0.010269 | -0.054932 | 5.3101 | -116.8533 | 5.2277 | -106.192 | -114.3723 | -115.0635 | -0.050354 | 5.3513 | -116.6656 | 5.2963 | -107.3502 | -113.8412 | -114.1342 | -0.20142 | -0.022888 | -0.10529 | -0.082398 | 10.4599 | 10.6613 | 1.5884 | 2.2797 | 1.7899 | 2.4811 | -0.013732 | 0.01831 | -0.036621 | -0.054931 | 9.3018 | 9.3155 | 2.5177 | 2.8107 | 2.5314 | 2.8244
2288.1719 | 122.2549 | -116.684 | 5.3421 | -0.15343 | 0.01245 | 0.095248 | 0.087291 | -0.0040867 | 0.013625 | 0.010297 | -0.12338 | 0.067251 | 0.08671 | 0.083272 | 0.0015275 | 0.012332 | 0.0096279 | -0.10071 | 5.3101 | -116.89 | 5.1498 | -106.1966 | -114.3814 | -115.1321 | -0.064087 | 5.3238 | -116.6977 | 5.2872 | -107.3273 | -113.7314 | -114.1388 | -0.20599 | -0.032043 | -0.19226 | -0.16022 | 10.4874 | 10.6934 | 1.5518 | 2.3026 | 1.7578 | 2.5085 | -0.013733 | -0.01831 | -0.054932 | -0.036622 | 9.3567 | 9.3704 | 2.5452 | 2.9526 | 2.5589 | 2.9663
2289.1719 | 122.2595 | -116.7389 | 5.3009 | -0.12291 | 0.015441 | 0.09418 | 0.084941 | -0.0019503 | 0.01747 | 0.0092286 | -0.010467 | 0.065542 | 0.08735 | 0.084127 | 0.0011001 | 0.011478 | 0.013687 | -0.082397 | 5.246 | -116.9312 | 5.1315 | -105.9128 | -114.3723 | -115.0909 | -0.036621 | 5.2643 | -116.7435 | 5.1727 | -107.4051 | -113.8412 | -114.1891 | -0.19226 | -0.054931 | -0.16937 | -0.11444 | 10.8261 | 11.0184 | 1.648 | 2.3666 | 1.8402 | 2.5589 | -0.004577 | -0.036621 | -0.12817 | -0.091553 | 9.3338 | 9.3384 | 2.5497 | 2.8976 | 2.5543 | 2.9022
2290.1719 | 122.2275 | -116.5512 | 5.484 | -0.16106 | 0.011382 | 0.092257 | 0.087291 | 0.00061311 | 0.016188 | 0.0092286 | -0.12338 | 0.067251 | 0.086069 | 0.082631 | -0.0014633 | 0.011692 | 0.012405 | -0.10529 | 5.4794 | -116.7435 | 5.2689 | -105.5969 | -114.2441 | -115.0635 | -0.11444 | 5.4886 | -116.597 | 5.3604 | -107.2723 | -113.736 | -114.0518 | -0.19226 | -0.004578 | -0.21515 | -0.21057 | 10.9543 | 11.1465 | 1.4877 | 2.3071 | 1.68 | 2.4994 | -0.045777 | 0.004577 | -0.1236 | -0.12817 | 9.2789 | 9.3246 | 2.4994 | 2.8152 | 2.5452 | 2.861
2291.1719 | 122.319 | -116.6565 | 5.3421 | -0.1992 | 0.012236 | 0.091403 | 0.085582 | -0.0013095 | 0.014479 | 0.010083 | -0.10202 | 0.066396 | 0.087991 | 0.083486 | 3.2078e-05 | 0.013187 | 0.011123 | -0.16022 | 5.3055 | -116.8716 | 5.1956 | -106.2515 | -114.5233 | -115.0635 | 0.009155 | 5.3421 | -116.7023 | 5.2917 | -107.3044 | -113.8046 | -114.1205 | -0.21515 | -0.036621 | -0.14648 | -0.10986 | 10.405 | 10.6201 | 1.593 | 2.1332 | 1.8082 | 2.3483 | -0.045777 | 0 | -0.050354 | -0.050354 | 9.3521 | 9.3979 | 2.536 | 2.8519 | 2.5818 | 2.8976
2292.1719 | 122.2595 | -116.7114 | 5.3284 | -0.10765 | 0.011382 | 0.09183 | 0.083018 | -0.0023777 | 0.01277 | 0.0094422 | -0.086761 | 0.065755 | 0.08735 | 0.077718 | 0.0038773 | 0.010623 | 0.011978 | -0.07782 | 5.2872 | -116.8762 | 5.1361 | -106.0318 | -114.5004 | -115.2145 | -0.022888 | 5.2551 | -116.7297 | 5.2094 | -107.3593 | -114.061 | -114.0884 | -0.16479 | -0.041199 | -0.19226 | -0.15106 | 10.6796 | 10.8444 | 1.4969 | 2.211 | 1.6617 | 2.3758 | -0.01831 | -0.073242 | -0.11902 | -0.045776 | 9.3521 | 9.3704 | 2.623 | 2.6505 | 2.6413 | 2.6688
2293.1719 | 122.287 | -116.7892 | 5.2185 | -0.16106 | 0.011382 | 0.094394 | 0.08665 | 0.0044583 | 0.017256 | 0.0075196 | -0.08371 | 0.067037 | 0.087137 | 0.079854 | -0.0029587 | 0.011905 | 0.01625 | -0.12817 | 5.1727 | -117.009 | 5.0308 | -106.2012 | -114.6378 | -115.2649 | -0.068665 | 5.2231 | -116.8167 | 5.1407 | -107.4646 | -113.8733 | -114.2212 | -0.21973 | -0.045777 | -0.18768 | -0.14191 | 10.5881 | 10.8078 | 1.5244 | 2.1515 | 1.7441 | 2.3712 | -0.027465 | 0.004577 | -0.07782 | -0.082397 | 9.3246 | 9.3521 | 2.5681 | 2.916 | 2.5955 | 2.9434
2294.1719 | 122.2595 | -116.9128 | 5.1041 | -0.13817 | 0.011382 | 0.093112 | 0.082377 | -0.0019503 | 0.015761 | 0.012006 | -0.040985 | 0.065542 | 0.08735 | 0.082845 | -0.00039512 | 0.010837 | 0.010269 | -0.09613 | 5.1086 | -117.0639 | 4.9713 | -106.1325 | -114.5187 | -115.2466 | -0.082397 | 5.1498 | -116.9083 | 4.9438 | -107.6569 | -114.0884 | -114.386 | -0.15106 | 0.004578 | -0.13275 | -0.13733 | 10.7803 | 10.9314 | 1.6663 | 2.3941 | 1.8173 | 2.5452 | 0.004578 | 0.045776 | -0.16022 | -0.20599 | 9.256 | 9.2514 | 2.5269 | 2.8244 | 2.5223 | 2.8198
2295.1719 | 122.2092 | -117.1234 | 4.9347 | -0.11528 | 0.01245 | 0.094394 | 0.090281 | 0.0010403 | 0.016402 | 0.011365 | -0.1539 | 0.066823 | 0.086069 | 0.081991 | 0.00088648 | 0.0097691 | 0.01091 | -0.091553 | 4.9118 | -117.2974 | 4.628 | -106.8466 | -114.6606 | -115.6219 | -0.07782 | 4.9622 | -117.1188 | 4.7607 | -107.8674 | -114.299 | -114.6469 | -0.17395 | -0.022888 | -0.3067 | -0.28381 | 10.2768 | 10.4507 | 1.5015 | 2.4628 | 1.6754 | 2.6367 | 0.004578 | 0.027466 | -0.17395 | -0.20142 | 9.256 | 9.2514 | 2.4765 | 2.8244 | 2.4719 | 2.8198
2296.1719 | 122.168 | -117.1555 | 4.9164 | -0.10765 | 0.015441 | 0.092685 | 0.08964 | -0.0010959 | 0.014479 | 0.0096558 | -0.19968 | 0.068105 | 0.085855 | 0.085836 | -0.00018152 | 0.014255 | 0.011551 | -0.091553 | 4.9072 | -117.3431 | 4.6417 | -107.0572 | -114.8987 | -115.7364 | -0.032043 | 4.9393 | -117.1371 | 4.7241 | -107.8903 | -114.3036 | -114.711 | -0.18768 | -0.009155 | -0.27466 | -0.2655 | 10.0983 | 10.2859 | 1.4191 | 2.2568 | 1.6068 | 2.4445 | 0.018311 | 0.022888 | -0.19226 | -0.21515 | 9.2651 | 9.2468 | 2.4445 | 2.8519 | 2.4261 | 2.8336
2297.1719 | 122.2183 | -117.3065 | 4.7516 | -0.11528 | 0.010527 | 0.092685 | 0.087718 | -0.00088229 | 0.015761 | 0.0098694 | -0.12033 | 0.067251 | 0.089273 | 0.084981 | 0.00067288 | 0.012973 | 0.011123 | -0.10071 | 4.7287 | -117.4942 | 4.4998 | -106.7413 | -115.0131 | -115.7547 | -0.073242 | 4.7379 | -117.3386 | 4.5914 | -108.0322 | -114.4363 | -114.8438 | -0.18768 | -0.022888 | -0.25177 | -0.22888 | 10.5652 | 10.7529 | 1.5518 | 2.2934 | 1.7395 | 2.4811 | -0.032043 | -0.013733 | -0.16022 | -0.14648 | 9.2743 | 9.3063 | 2.4628 | 2.8702 | 2.4948 | 2.9022
2298.1719 | 122.2916 | -117.3248 | 4.6875 | -0.15343 | 0.01245 | 0.094821 | 0.087718 | 0.00061311 | 0.015334 | 0.010083 | -0.098968 | 0.066823 | 0.08671 | 0.08434 | -0.00082252 | 0.011905 | 0.013473 | -0.07782 | 4.6692 | -117.5079 | 4.5273 | -107.103 | -114.9124 | -115.7547 | -0.20142 | 4.6692 | -117.3523 | 4.6326 | -108.0093 | -114.6286 | -114.8163 | -0.18311 | -0.018311 | -0.16022 | -0.14191 | 10.2219 | 10.405 | 1.5701 | 2.4124 | 1.7532 | 2.5955 | -0.027466 | -0.018311 | -0.054932 | -0.036621 | 9.3155 | 9.343 | 2.5085 | 2.6962 | 2.536 | 2.7237
2299.1719 | 122.2824 | -117.3111 | 4.7287 | -0.12291 | 0.013518 | 0.093753 | 0.083873 | -0.0036593 | 0.015975 | 0.011365 | 0.010895 | 0.065328 | 0.085855 | 0.078145 | -0.0021041 | 0.0093417 | 0.0079189 | -0.045776 | 4.6646 | -117.5034 | 4.5456 | -107.0709 | -115.036 | -115.686 | 0.004578 | 4.6829 | -117.3248 | 4.6051 | -107.9819 | -114.1205 | -114.798 | -0.19226 | -0.064087 | -0.18311 | -0.11902 | 10.2402 | 10.4324 | 1.6251 | 2.2751 | 1.8173 | 2.4673 | -0.013733 | -0.045777 | -0.1236 | -0.077819 | 9.3292 | 9.343 | 2.5131 | 3.1906 | 2.5269 | 3.2043
2300.1719 | 122.2961 | -117.3798 | 4.6509 | -0.092392 | 0.012023 | 0.094821 | 0.087077 | -0.00024129 | 0.016188 | 0.013501 | -0.092865 | 0.066823 | 0.087564 | 0.082418 | -0.00039512 | 0.013828 | 0.0089871 | -0.10986 | 4.6417 | -117.5949 | 4.4495 | -106.8787 | -115.0452 | -115.8875 | -0.027466 | 4.6646 | -117.3981 | 4.5731 | -108.11 | -114.4592 | -114.8575 | -0.21515 | -0.009155 | -0.20142 | -0.19226 | 10.5011 | 10.7162 | 1.4923 | 2.3346 | 1.7075 | 2.5497 | -0.01831 | 0.013733 | -0.07782 | -0.091553 | 9.2697 | 9.288 | 2.5223 | 2.9205 | 2.5406 | 2.9388
2301.1719 | 122.2595 | -117.3157 | 4.6921 | -0.084763 | 0.0096728 | 0.093539 | 0.088786 | -0.0049411 | 0.014266 | 0.010297 | -0.12643 | 0.065969 | 0.087564 | 0.083059 | -0.00082252 | 0.013401 | 0.012619 | -0.12817 | 4.6188 | -117.5995 | 4.4769 | -107.3593 | -115.0269 | -115.8325 | -0.059509 | 4.6738 | -117.366 | 4.5822 | -108.0643 | -114.5508 | -114.8758 | -0.28381 | -0.073243 | -0.21515 | -0.14191 | 9.9564 | 10.2402 | 1.4832 | 2.2888 | 1.767 | 2.5726 | -0.050354 | -0.018311 | -0.10986 | -0.091553 | 9.2514 | 9.3018 | 2.4399 | 2.7649 | 2.4902 | 2.8152
2302.1719 | 122.2321 | -117.1829 | 4.866 | -0.1458 | 0.013091 | 0.094607 | 0.083018 | -0.0010959 | 0.013625 | 0.0094422 | -0.025726 | 0.06661 | 0.085855 | 0.079641 | -0.00039512 | 0.010837 | 0.012192 | -0.10986 | 4.834 | -117.3569 | 4.6783 | -106.572 | -114.7705 | -115.567 | -0.068665 | 4.8523 | -117.1692 | 4.7241 | -107.9086 | -114.2441 | -114.7018 | -0.17395 | -0.032044 | -0.18768 | -0.15564 | 10.611 | 10.7849 | 1.6159 | 2.4124 | 1.7899 | 2.5864 | 0.013733 | -0.013733 | -0.14191 | -0.12817 | 9.2743 | 9.2606 | 2.4811 | 2.9388 | 2.4673 | 2.9251
2303.1719 | 122.1909 | -116.6794 | 5.4291 | -0.16106 | 0.011809 | 0.093966 | 0.087931 | 0.0046719 | 0.016829 | 0.012006 | -0.13864 | 0.067037 | 0.087778 | 0.084768 | -0.0010361 | 0.015964 | 0.012405 | -0.12817 | 5.3925 | -116.8762 | 5.1315 | -106.2332 | -114.4547 | -115.1962 | -0.18768 | 5.4062 | -116.6748 | 5.1498 | -107.309 | -113.9465 | -114.3173 | -0.19684 | -0.036621 | -0.29755 | -0.26093 | 10.4462 | 10.643 | 1.4832 | 2.2247 | 1.68 | 2.4216 | 0.004577 | -0.022888 | -0.27924 | -0.25635 | 9.3704 | 9.3658 | 2.3621 | 2.7328 | 2.3575 | 2.7283
2304.1719 | 122.2549 | -116.684 | 5.3558 | -0.19157 | 0.014159 | 0.092044 | 0.084086 | 0.0012539 | 0.016616 | 0.0075196 | -0.019623 | 0.067891 | 0.089059 | 0.076864 | -0.00018152 | 0.0095555 | 0.011551 | -0.09613 | 5.1544 | -116.9998 | 5.0629 | -105.8899 | -114.3082 | -115.242 | -0.068665 | 5.2963 | -116.7618 | 5.2048 | -107.4417 | -113.8458 | -114.2303 | -0.31586 | -0.20142 | -0.29297 | -0.091553 | 10.7941 | 11.1099 | 1.442 | 2.3758 | 1.7578 | 2.6916 | -0.07782 | -0.059509 | -0.15106 | -0.091553 | 9.2422 | 9.3201 | 2.4536 | 2.8381 | 2.5314 | 2.916
2305.1719 | 122.2641 | -117.1371 | 4.8569 | -0.12291 | 0.012664 | 0.092898 | 0.085154 | -0.0030185 | 0.015548 | 0.0096558 | -0.11118 | 0.066396 | 0.086282 | 0.08135 | 3.2078e-05 | 0.013614 | 0.0083463 | -0.12817 | 4.834 | -117.366 | 4.6555 | -106.9885 | -114.8804 | -115.7318 | -0.14648 | 4.8752 | -117.1875 | 4.7562 | -107.8903 | -114.1571 | -114.7018 | -0.22888 | -0.022889 | -0.20142 | -0.17853 | 10.1486 | 10.3775 | 1.4053 | 2.2568 | 1.6342 | 2.4857 | -0.050354 | 0.01831 | -0.10071 | -0.11902 | 9.2468 | 9.2972 | 2.4353 | 2.98 | 2.4857 | 3.0304
2306.1719 | 122.2366 | -116.8671 | 5.1727 | -0.11528 | 0.011382 | 0.092257 | 0.086436 | 0.00082671 | 0.015548 | 0.0094422 | -0.10812 | 0.068746 | 0.086496 | 0.082418 | -0.0016769 | 0.014255 | 0.01091 | -0.11444 | 5.1636 | -117.0685 | 4.9576 | -106.7596 | -114.5554 | -115.3564 | -0.1831 | 5.1773 | -116.8625 | 5.0446 | -107.6843 | -114.0518 | -114.4409 | -0.20142 | -0.009155 | -0.21515 | -0.20599 | 10.1074 | 10.3088 | 1.5106 | 2.3117 | 1.712 | 2.5131 | 0.004577 | 0.004578 | -0.12817 | -0.13275 | 9.1827 | 9.1782 | 2.4261 | 2.8152 | 2.4216 | 2.8107
2307.1719 | 122.2183 | -116.8304 | 5.2139 | -0.12291 | 0.012236 | 0.092471 | 0.085154 | -0.0013095 | 0.014052 | 0.011151 | -0.11118 | 0.066823 | 0.088205 | 0.082631 | 0.0011001 | 0.012546 | 0.010055 | -0.13733 | 5.159 | -117.0731 | 4.9438 | -106.517 | -114.5554 | -115.4114 | -0.091553 | 5.1956 | -116.8854 | 5.0354 | -107.6843 | -114.0152 | -114.4547 | -0.24262 | -0.054931 | -0.27008 | -0.21515 | 10.3134 | 10.556 | 1.4191 | 2.2751 | 1.6617 | 2.5177 | -0.054932 | -0.01831 | -0.17853 | -0.16022 | 9.1461 | 9.201 | 2.3758 | 2.8152 | 2.4307 | 2.8702
2308.1719 | 122.1909 | -119.0781 | 2.9755 | -0.031357 | 0.012023 | 0.095462 | 0.085154 | -0.0025913 | 0.014479 | 0.010937 | -0.13864 | 0.14181 | 0.0944 | -0.0002543 | 0.07715 | 0.08411 | 0.0036465 | -0.17395 | 2.98 | -119.2795 | 2.71 | -107.048 | -116.7847 | -117.5812 | -0.10529 | 2.9709 | -119.0872 | -96.7163 | 2.8519 | -9.5489 | -2.8519 | -0.20142 | 0.004578 | -0.2655 | -0.27008 | 12.03 | 12.2314 | 1.4969 | 2.2934 | 1.6983 | 2.4948 | -0.009155 | -0.004578 | -99.6918 | -99.6872 | 121.9299 | 121.9391 | 116.2262 | 109.5291 | 116.2354 | 109.5383
2309.1719 | 122.3007 | -119.0094 | 2.9846 | -0.18394 | 0.013518 | 0.093966 | 0.083659 | -0.0025913 | 0.015334 | 0.011151 | -0.056244 | 0.14565 | 0.095682 | -0.0028177 | 0.074159 | 0.089237 | 0.0023647 | -0.064087 | 2.948 | -119.2017 | 2.8564 | -108.8791 | -116.6519 | -117.4438 | 0.013733 | 2.98 | -119.046 | -95.9152 | 3.0029 | -6.2393 | -2.4124 | -0.19226 | -0.036621 | -0.12817 | -0.091553 | 10.1303 | 10.3226 | 1.5656 | 2.3575 | 1.7578 | 2.5497 | -0.036622 | -0.004577 | -98.8998 | -98.8953 | 122.0123 | 122.049 | 116.597 | 112.7701 | 116.6336 | 112.8067
2310.1719 | 122.1725 | -119.0918 | 3.0121 | -0.054245 | 0.012877 | 0.091403 | 0.086222 | -0.0013095 | 0.014693 | 0.010083 | -0.15085 | 0.14843 | 0.096536 | 0.0033773 | 0.076936 | 0.089237 | 0.0012967 | -0.064087 | 2.9572 | -119.2841 | 2.7145 | -108.8196 | -116.7709 | -117.5812 | -0.1236 | 2.9755 | -119.0689 | -96.2128 | 2.8473 | -6.4728 | -2.5223 | -0.19226 | -0.054932 | -0.29755 | -0.24261 | 10.2722 | 10.4645 | 1.5106 | 2.3209 | 1.7029 | 2.5131 | 0.022888 | -0.036621 | -99.2249 | -99.1882 | 121.9391 | 121.9162 | 116.5695 | 112.619 | 116.5466 | 112.5961
2311.1719 | 132.6233 | -73.6542 | 40.2328 | -0.35942 | 0.024626 | 0.12024 | 0.023417 | 0.041415 | 0.061263 | 0.050458 | -0.26071 | 0.039907 | 0.095468 | -0.0028177 | 0.072664 | 0.046085 | -0.0081027 | 0.091553 | 20.3522 | -46.1746 | 26.9485 | 0.09613 | -47.4426 | -50.4501 | 0.10071 | 22.4899 | -52.8717 | -56.0577 | 40.8096 | 0.12817 | 20.5261 | 27.4796 | -19.8807 | -13.2843 | 6.5964 | 73.7503 | 46.2708 | 23.204 | 26.2115 | -4.2755 | -1.268 | 20.7825 | -17.7429 | -96.2906 | -78.5477 | 114.4638 | 93.6813 | 94.1803 | 73.7823 | 73.3978 | 52.9999
2312.1719 | 122.4472 | -116.0385 | 5.8273 | -0.15343 | 0.01245 | 0.092471 | 0.085582 | -0.00088229 | 0.015761 | 0.0085876 | -0.080658 | 0.14415 | 0.096963 | -0.0019631 | 0.075014 | 0.089237 | 0.0010831 | -0.10986 | 5.8548 | -115.4846 | 5.7816 | -105.542 | -113.3698 | -114.1388 | -0.013733 | 5.6213 | -115.4205 | -93.5211 | 5.6717 | -3.6301 | 0.036621 | 0.55389 | 0.027465 | -0.045777 | -0.073242 | 10.4965 | 9.9426 | 1.8997 | 2.6688 | 1.3458 | 2.1149 | 0.61798 | -0.20599 | -99.3484 | -99.1425 | 121.7102 | 121.0922 | 116.0751 | 112.4084 | 115.4572 | 111.7905
2313.1719 | 122.2778 | -116.0248 | 5.9509 | -0.41283 | 0.0092456 | 0.087558 | 0.083445 | -0.00045509 | 0.014052 | 0.007306 | -0.18442 | 0.14116 | 0.096963 | 0.0016685 | 0.075441 | 0.088382 | 0.0079189 | -0.39368 | 5.722 | -116.011 | 5.6992 | -105.954 | -113.6261 | -114.3585 | -0.1831 | 5.7632 | -116.1713 | -93.8782 | 5.8273 | -3.9459 | -0.2243 | 0.013733 | -0.22888 | -0.25177 | -0.022888 | 10.0708 | 10.0571 | 1.6663 | 2.3987 | 1.6525 | 2.3849 | -0.14648 | -0.18768 | -99.8291 | -99.6414 | 121.8521 | 121.9986 | 115.8005 | 112.0789 | 115.947 | 112.2253
2314.1719 | 122.287 | -116.0889 | 5.8914 | -0.12291 | 0.0098864 | 0.090549 | 0.0843 | 0.00018591 | 0.015334 | 0.0090148 | -0.0043641 | 0.14565 | 0.09675 | -0.0011087 | 0.076936 | 0.086887 | 0.0045011 | -0.10529 | 5.8548 | -116.098 | 5.7312 | -105.9036 | -113.7772 | -114.4775 | -0.054932 | 5.8777 | -115.6631 | -93.8049 | 5.8823 | -3.7079 | 0.009155 | -0.009155 | -0.036621 | -0.16022 | -0.1236 | 10.1852 | 10.1944 | 1.6113 | 2.3117 | 1.6205 | 2.3209 | 0.42572 | -0.013732 | -99.6963 | -99.6826 | 121.9711 | 121.5454 | 116.098 | 112.381 | 115.6723 | 111.9553
2315.1719 | 122.2183 | -115.9973 | 6.0562 | -0.29839 | 0.007323 | 0.087131 | 0.074473 | -0.016477 | 0.011702 | 0.005597 | -0.30344 | 0.13155 | 0.091196 | 0.0010275 | 0.078645 | 0.08411 | -0.0025485 | -0.13733 | 5.9418 | -115.4388 | 5.8685 | -104.8325 | -113.2553 | -114.3585 | -0.15106 | 6.102 | -115.3519 | -93.7729 | 5.7632 | -3.6163 | 0 | 0.55847 | -0.11444 | -0.18768 | -0.073242 | 11.1649 | 10.6064 | 1.6388 | 2.742 | 1.0803 | 2.1835 | 0.64545 | 0.045777 | -99.8291 | -99.8749 | 121.7606 | 121.1151 | 115.9973 | 112.381 | 115.3519 | 111.7355
2316.1719 | 122.1863 | -116.0522 | 6.0104 | -0.13817 | 0.012236 | 0.083713 | 0.073191 | -0.014554 | 0.0080708 | -0.0067932 | -0.4072 | 0.14245 | 0.082223 | -0.011149 | 0.067964 | 0.073856 | -0.0055393 | -0.15564 | 5.8823 | -115.5624 | 5.7907 | -105.6244 | -113.475 | -114.5187 | -0.064087 | 5.9784 | -115.7822 | -93.8049 | 5.8914 | -3.8132 | -0.17395 | 0.48981 | -0.12817 | -0.21973 | -0.091553 | 10.4279 | 9.938 | 1.5335 | 2.5772 | 1.0437 | 2.0874 | 0.27008 | -0.032043 | -99.8154 | -99.7833 | 121.9437 | 121.6736 | 115.8783 | 112.2391 | 115.6082 | 111.969
2317.1719 | 122.1863 | -117.247 | 4.7928 | -0.069504 | 0.013091 | 0.087344 | 0.078105 | -0.019681 | 0.010207 | -0.00081179 | -0.30954 | 0.12941 | 0.085001 | -0.0083719 | 0.069459 | 0.071933 | -0.0078891 | -0.086975 | 4.6921 | -116.7984 | 4.5776 | -106.5399 | -114.505 | -115.86 | -0.045776 | 4.7882 | -116.6931 | -95.05 | 4.6783 | -4.422 | -1.1124 | 0.44861 | -0.10071 | -0.21515 | -0.11444 | 10.7071 | 10.2585 | 1.387 | 2.742 | 0.93842 | 2.2934 | 0.55389 | -0.004578 | -99.8428 | -99.8383 | 121.9254 | 121.3715 | 116.1346 | 112.825 | 115.5807 | 112.2711
2318.1719 | 122.168 | -117.6727 | 4.3854 | -0.11528 | 0.012023 | 0.089908 | 0.081523 | -0.012418 | 0.013411 | 0.0036744 | -0.23325 | 0.18902 | 0.14439 | 0.0063681 | 0.075014 | 0.081974 | -0.0012667 | -0.073242 | 4.3396 | -117.0731 | 4.0878 | -107.2952 | -114.9445 | -115.8737 | -0.11902 | 4.3533 | -116.8854 | -5.2277 | 4.2252 | -4.6051 | 0.29755 | 0.59967 | -0.045776 | -0.29755 | -0.25177 | 10.3775 | 9.7778 | 1.799 | 2.7283 | 1.1993 | 2.1286 | 0.78735 | -0.032043 | -9.613 | -9.581 | 121.8979 | 121.1105 | 117.9703 | 113.0676 | 117.1829 | 112.2803
2319.1719 | 122.2549 | -117.9565 | 4.0421 | -0.16869 | 0.012877 | 0.091403 | 0.084941 | -0.0019503 | 0.014479 | 0.011578 | -0.037933 | 0.20226 | 0.15058 | -0.0040995 | 0.075655 | 0.088809 | 0.0047147 | -0.073242 | 4.0237 | -118.1259 | 3.8635 | -107.6202 | -115.6357 | -116.3818 | -0.045776 | 4.0283 | -117.8513 | -4.6829 | 4.0146 | -4.4174 | 0.16937 | -0.16937 | -0.01831 | -0.17853 | -0.16022 | 10.3363 | 10.5057 | 1.5747 | 2.3209 | 1.7441 | 2.4902 | 0.10529 | -0.013733 | -8.725 | -8.7112 | 121.9711 | 121.8658 | 118.1259 | 113.5391 | 118.0206 | 113.4338
2320.1719 | 122.1405 | -117.2607 | 4.8111 | -0.084763 | 0.013945 | 0.14801 | 0.085154 | -0.0043003 | 0.014266 | 0.013715 | -0.25766 | 0.25503 | 0.14823 | -0.0068765 | 0.073732 | 0.086246 | 0.0015103 | -0.073242 | 4.7836 | -117.0639 | 4.5273 | -14.9231 | -114.5966 | -114.4089 | -0.14191 | 4.7974 | -116.803 | -4.834 | 4.6509 | -4.4998 | 0.72327 | 0.19684 | -0.027466 | -0.28381 | -0.25635 | 102.3376 | 102.1408 | 2.8519 | 2.6642 | 2.655 | 2.4673 | 0.45776 | -0.013733 | -9.6451 | -9.6313 | 121.9116 | 121.4539 | 117.984 | 112.7609 | 117.5262 | 112.3032
2321.1719 | 122.1222 | -117.3889 | 4.6783 | -0.10765 | 0.015227 | 0.14972 | 0.082377 | -0.0040867 | 0.013839 | 0.011151 | -0.22104 | 0.25695 | 0.1493 | -0.0040995 | 0.078432 | 0.086673 | -0.00083948 | -0.086975 | 4.6097 | -117.1234 | 4.3854 | -11.2015 | -114.7751 | -114.505 | -0.13275 | 4.6646 | -116.954 | -4.6921 | 4.509 | -4.5639 | 0.59052 | 0.2655 | -0.068665 | -0.29297 | -0.2243 | 106.1874 | 105.9219 | 2.8839 | 2.6138 | 2.6184 | 2.3483 | 0.43487 | -0.013733 | -9.3704 | -9.3567 | 121.8979 | 121.463 | 117.9794 | 112.825 | 117.5446 | 112.3901
2322.1719 | 122.2733 | -117.4118 | 4.5914 | -0.15343 | 0.013518 | 0.15122 | 0.084727 | -0.00045509 | 0.014907 | 0.016919 | -0.20273 | 0.26528 | 0.14866 | -0.0036721 | 0.074586 | 0.089664 | 0.0021511 | -0.050354 | 4.5502 | -117.4896 | 4.4769 | -10.1669 | -114.9582 | -114.7614 | -0.059509 | 4.5136 | -117.3615 | -4.5044 | 4.5959 | -4.3579 | 0.66376 | -0.07782 | -0.041199 | -0.11444 | -0.073242 | 107.2449 | 107.3227 | 2.6505 | 2.4536 | 2.7283 | 2.5314 | 0.050354 | -0.07782 | -9.0958 | -9.0179 | 122.0078 | 121.9574 | 118.0756 | 113.0539 | 118.0252 | 113.0035
2323.1719 | 122.1634 | -117.3615 | 4.6234 | -0.092392 | 0.0101 | 0.15122 | 0.086222 | 0.0025357 | 0.015975 | 0.014569 | -0.26376 | 0.25738 | 0.1461 | -0.0079447 | 0.0701 | 0.078128 | 0.0023647 | -0.032043 | 4.6097 | -117.2607 | 4.4128 | -8.9996 | -114.7156 | -114.418 | 0.004578 | 4.6326 | -117.0273 | -4.5593 | 4.5456 | -4.39 | 0.59967 | 0.10071 | -0.013733 | -0.21057 | -0.19684 | 108.3618 | 108.2611 | 2.9434 | 2.6459 | 2.8427 | 2.5452 | 0.33417 | 0.009155 | -9.1827 | -9.1919 | 121.907 | 121.5729 | 117.9611 | 112.9715 | 117.627 | 112.6373
2324.1719 | 122.2549 | -117.4072 | 4.5319 | -0.19157 | 0.010741 | 0.15229 | 0.087077 | 0.0027493 | 0.015548 | 0.017133 | -0.13254 | 0.26656 | 0.15101 | -0.0051675 | 0.071596 | 0.089664 | 0.0017239 | -0.054932 | 4.5181 | -117.5812 | 4.4174 | -9.0912 | -115.0269 | -114.8163 | 0.009155 | 4.509 | -117.3981 | -4.4403 | 4.5639 | -4.509 | 0.63171 | -0.17395 | -0.013733 | -0.11444 | -0.10071 | 108.316 | 108.49 | 2.5909 | 2.3804 | 2.7649 | 2.5543 | 0.009156 | -0.022888 | -8.9722 | -8.9493 | 121.9711 | 121.962 | 118.0389 | 112.8983 | 118.0298 | 112.8891
2325.1719 | 122.1359 | -117.4805 | 4.5456 | -0.12291 | 0.010955 | 0.14951 | 0.087291 | -0.00066869 | 0.014052 | 0.013928 | -0.26682 | 0.25887 | 0.14482 | -0.0068765 | 0.070955 | 0.076419 | 0.0015103 | -0.036621 | 4.5227 | -117.5491 | 4.335 | -8.7799 | -114.9216 | -114.6194 | 0 | 4.5273 | -117.2974 | -4.5136 | 4.4312 | -4.509 | 0.49896 | -0.068664 | -0.022888 | -0.21057 | -0.18768 | 108.7006 | 108.7692 | 2.861 | 2.5589 | 2.9297 | 2.6276 | 0.18311 | -0.01831 | -9.0591 | -9.0408 | 121.9116 | 121.7285 | 117.9794 | 112.9715 | 117.7963 | 112.7884
2326.1719 | 122.0947 | -117.4896 | 4.5685 | -0.17632 | 0.012023 | 0.15378 | 0.08665 | -0.0030185 | 0.015334 | 0.01756 | -0.20273 | 0.2625 | 0.14909 | -0.0030313 | 0.076082 | 0.083896 | 0.0019375 | -0.068665 | 4.509 | -117.5217 | 4.2847 | -8.5693 | -115.0497 | -114.6973 | -0.10529 | 4.5227 | -117.4164 | -4.509 | 4.4037 | -4.5364 | 0.48981 | -0.032043 | -0.059509 | -0.28381 | -0.2243 | 108.9203 | 108.9523 | 2.7924 | 2.4399 | 2.8244 | 2.4719 | 0.073242 | -0.045776 | -9.0775 | -9.0317 | 121.8933 | 121.8201 | 117.9794 | 112.9532 | 117.9062 | 112.8799
2327.1719 | 122.1268 | -117.4622 | 4.5639 | -0.10002 | 0.011809 | 0.15143 | 0.085368 | 0.0016813 | 0.017043 | 0.018201 | -0.18747 | 0.26357 | 0.14845 | 0.0016685 | 0.075227 | 0.083683 | 0.0010831 | -0.050354 | 4.5456 | -117.453 | 4.3304 | -8.4869 | -115.0131 | -114.6835 | -0.045776 | 4.541 | -117.3523 | -4.4861 | 4.4449 | -4.4907 | 0.50812 | 0.009155 | -0.018311 | -0.23346 | -0.21515 | 108.9752 | 108.9661 | 2.7786 | 2.449 | 2.7695 | 2.4399 | 0.10986 | -0.022888 | -9.05 | -9.0271 | 121.907 | 121.7972 | 117.9703 | 112.9715 | 117.8604 | 112.8616
2328.1719 | 122.1222 | -117.4255 | 4.5914 | -0.11528 | 0.010955 | 0.15058 | 0.084941 | 0.0025357 | 0.013625 | 0.017346 | -0.17831 | 0.26378 | 0.14951 | -0.0017495 | 0.076936 | 0.085819 | 0.0025785 | -0.036621 | 4.5776 | -117.3752 | 4.3488 | -8.4 | -115.0085 | -114.6652 | -0.050354 | 4.5547 | -117.3065 | -4.4357 | 4.4632 | -4.4678 | 0.54474 | 0.050354 | -0.013733 | -0.24262 | -0.22888 | 109.0256 | 108.9752 | 2.7603 | 2.417 | 2.71 | 2.3666 | 0.11902 | -0.036621 | -9.0271 | -8.9905 | 121.8887 | 121.7697 | 117.9703 | 112.9578 | 117.8513 | 112.8387
2329.1719 | 122.2412 | -117.3431 | 4.6326 | -0.16869 | 0.013304 | 0.15293 | 0.08195 | -0.0015231 | 0.017684 | 0.017346 | -0.068451 | 0.26784 | 0.14994 | 0.0006003 | 0.076082 | 0.087314 | -0.0010531 | -0.059509 | 4.5502 | -117.5171 | 4.4312 | -8.02 | -114.9811 | -114.7842 | 0.050354 | 4.5364 | -117.334 | -4.3167 | 4.6371 | -4.2068 | 0.66376 | -0.17395 | -0.082397 | -0.20142 | -0.11902 | 109.3231 | 109.4971 | 2.5589 | 2.3621 | 2.7328 | 2.536 | 0.009156 | -0.09613 | -8.9493 | -8.8531 | 121.9803 | 121.9711 | 118.0069 | 113.1363 | 117.9977 | 113.1271
2330.1875 | 122.2229 | -117.3431 | 4.6417 | -0.19157 | 0.011595 | 0.15015 | 0.084727 | 0.0012539 | 0.017043 | 0.01286 | -0.044037 | 0.26699 | 0.14909 | -0.0034585 | 0.073305 | 0.089023 | 0.0030057 | -0.091553 | 4.5639 | -117.5446 | 4.3533 | -8.1619 | -114.9994 | -114.7614 | 0.027466 | 4.541 | -117.3706 | -4.3625 | 4.6005 | -4.4678 | 0.63629 | -0.20142 | -0.07782 | -0.28839 | -0.21057 | 109.1812 | 109.3826 | 2.5818 | 2.3438 | 2.7832 | 2.5452 | -0.027465 | -0.10071 | -9.0042 | -8.9035 | 121.9437 | 121.9711 | 117.9794 | 112.8754 | 118.0069 | 112.9028
2331.1719 | 122.1222 | -117.247 | 4.7791 | -0.10002 | 0.010314 | 0.15335 | 0.085795 | -0.0055819 | 0.013625 | 0.012433 | -0.20883 | 0.25823 | 0.14588 | -0.0053811 | 0.075014 | 0.080692 | 0.0015103 | 0.045776 | 4.7241 | -117.366 | 4.5227 | -8.1207 | -114.9307 | -114.6149 | 0.013733 | 4.715 | -117.247 | -4.3625 | 4.7012 | -4.4632 | 0.73242 | -0.11902 | -0.054932 | -0.25635 | -0.20142 | 109.1263 | 109.2453 | 2.6321 | 2.3163 | 2.7512 | 2.4353 | 0 | -0.064087 | -9.1415 | -9.0775 | 121.9482 | 121.9482 | 117.9794 | 112.7838 | 117.9794 | 112.7838
2332.1719 | 122.1313 | -117.2836 | 4.7333 | -0.37468 | 0.011809 | 0.14929 | 0.08665 | -0.0038729 | 0.01512 | 0.017133 | -0.19052 | 0.26912 | 0.15165 | 0.0023093 | 0.075014 | 0.089023 | -0.0040439 | -0.013733 | 4.6783 | -117.4805 | 4.4586 | -8.075 | -114.9811 | -114.7751 | -0.1236 | 4.7058 | -117.3019 | -4.4083 | 4.5959 | -4.4769 | 0.64087 | -0.19684 | -0.054931 | -0.27466 | -0.21973 | 109.2087 | 109.4055 | 2.5085 | 2.3026 | 2.7054 | 2.4994 | -0.018311 | -0.027465 | -9.1415 | -9.1141 | 121.8796 | 121.8979 | 117.9245 | 112.8067 | 117.9428 | 112.825
2333.1719 | 122.1176 | -117.2516 | 4.747 | -0.10002 | 0.012236 | 0.15229 | 0.087077 | 0.0012539 | 0.014479 | 0.014783 | -0.17831 | 0.2625 | 0.14973 | 0.0046591 | 0.075014 | 0.08411 | -0.0034031 | -0.004578 | 4.7379 | -117.3981 | 4.5547 | -8.0521 | -114.8346 | -114.6698 | -0.059509 | 4.747 | -117.2333 | -4.3945 | 4.628 | -4.4586 | 0.67749 | -0.14648 | -0.009155 | -0.19226 | -0.1831 | 109.1995 | 109.346 | 2.5818 | 2.417 | 2.7283 | 2.5635 | 0.018311 | 0 | -9.1415 | -9.1415 | 121.8796 | 121.8613 | 117.9291 | 112.793 | 117.9108 | 112.7747
2334.1875 | 122.1771 | -117.1967 | 4.7424 | -0.12291 | 0.012877 | 0.15229 | 0.083873 | -0.0017367 | 0.015761 | 0.015851 | -0.08371 | 0.26912 | 0.14887 | -0.0066629 | 0.073732 | 0.088596 | 0.0047147 | -0.07782 | 4.6921 | -117.3889 | 4.5914 | -7.8827 | -114.8621 | -114.6011 | 0 | 4.7104 | -117.1692 | -4.3213 | 4.7287 | -4.4449 | 0.76904 | -0.19226 | -0.050354 | -0.15106 | -0.10071 | 109.314 | 109.5062 | 2.5955 | 2.3346 | 2.7878 | 2.5269 | 0.027466 | -0.032044 | -9.0637 | -9.0317 | 121.9254 | 121.8979 | 117.9657 | 112.7518 | 117.9382 | 112.7243
2335.1719 | 122.1771 | -117.2791 | 4.6692 | -0.13817 | 0.011809 | 0.15164 | 0.085795 | -0.00066869 | 0.01277 | 0.015424 | -0.086761 | 0.26549 | 0.1493 | -0.0049539 | 0.072237 | 0.087314 | 0.0015103 | -0.064087 | 4.6555 | -117.4576 | 4.4952 | -7.9376 | -114.8987 | -114.6881 | -0.11902 | 4.6371 | -117.3752 | -4.3533 | 4.6463 | -4.422 | 0.72784 | -0.17853 | -0.013732 | -0.17395 | -0.16022 | 109.3414 | 109.52 | 2.5909 | 2.3804 | 2.7695 | 2.5589 | -0.09613 | -0.032043 | -9.0225 | -8.9905 | 121.9254 | 122.0215 | 118.0069 | 112.8571 | 118.103 | 112.9532
2336.1719 | 122.1268 | -117.2653 | 4.7287 | -0.10765 | 0.013732 | 0.15058 | 0.084513 | -0.0028049 | 0.01512 | 0.017987 | -0.21799 | 0.26079 | 0.14567 | -0.0083719 | 0.075014 | 0.080265 | -0.00062588 | 0.004578 | 4.7379 | -117.3798 | 4.5502 | -7.9102 | -114.9399 | -114.7247 | -0.004578 | 4.7058 | -117.2607 | -4.3488 | 4.628 | -4.5227 | 0.70038 | -0.11444 | 0.009155 | -0.17853 | -0.18768 | 109.3552 | 109.4696 | 2.5406 | 2.3254 | 2.655 | 2.4399 | 0.004578 | -0.022888 | -9.0775 | -9.0546 | 121.8933 | 121.8887 | 117.9657 | 112.7426 | 117.9611 | 112.738
2337.1719 | 122.1359 | -117.1692 | 4.8248 | -0.069504 | 0.012236 | 0.15164 | 0.086222 | -0.0013095 | 0.014907 | 0.017987 | -0.18747 | 0.26763 | 0.1508 | 0.0010275 | 0.073946 | 0.085391 | -0.0027621 | -0.054932 | 4.7653 | -117.3843 | 4.6051 | -7.8415 | -114.8895 | -114.6286 | 0.059509 | 4.7974 | -117.1875 | -4.3259 | 4.7333 | -4.4128 | 0.79193 | -0.21515 | -0.059509 | -0.21973 | -0.16022 | 109.3277 | 109.5428 | 2.5406 | 2.2797 | 2.7557 | 2.4948 | -0.018311 | -0.027466 | -9.1507 | -9.1232 | 121.9025 | 121.9208 | 117.9611 | 112.7563 | 117.9794 | 112.7747
2338.1719 | 122.168 | -117.2745 | 4.7012 | -0.069504 | 0.011809 | 0.15293 | 0.08665 | -0.00045509 | 0.01512 | 0.017346 | -0.13864 | 0.2672 | 0.14973 | 0.0014547 | 0.073305 | 0.084751 | 0.0015103 | 0.004578 | 4.6555 | -117.4393 | 4.509 | -7.8094 | -114.7751 | -114.5599 | 0.009155 | 4.7653 | -117.2333 | -4.303 | 4.6097 | -4.5456 | 0.68207 | -0.16479 | -0.045776 | -0.19226 | -0.14649 | 109.465 | 109.6298 | 2.7145 | 2.4994 | 2.8793 | 2.6642 | 0.041199 | 0.064087 | -9.0042 | -9.0683 | 121.8842 | 121.843 | 117.9565 | 112.7289 | 117.9153 | 112.6877
2339.1719 | 122.0947 | -117.2653 | 4.7424 | -0.11528 | 0.011168 | 0.15186 | 0.087504 | -0.0023777 | 0.016616 | 0.019482 | -0.19052 | 0.26934 | 0.14973 | 0.0010275 | 0.073946 | 0.090305 | -0.00083948 | -0.082397 | 4.6646 | -117.4667 | 4.4495 | -7.9468 | -115.0131 | -114.7888 | -0.14191 | 4.7104 | -117.2607 | -4.3579 | 4.5776 | -4.541 | 0.69122 | -0.20142 | -0.07782 | -0.29297 | -0.21515 | 109.3185 | 109.52 | 2.4765 | 2.2522 | 2.6779 | 2.4536 | 0.004578 | -0.032044 | -9.1003 | -9.0683 | 121.843 | 121.8384 | 117.9565 | 112.7243 | 117.952 | 112.7197
2340.1719 | 122.2183 | -117.2241 | 4.6692 | -0.11528 | 0.01245 | 0.15335 | 0.084941 | 0.0040311 | 0.016188 | 0.014569 | -0.14169 | 0.26699 | 0.15058 | -4.0504e-05 | 0.073732 | 0.089237 | 0.00065592 | -0.054932 | 4.628 | -117.4438 | 4.5868 | -7.6813 | -114.9307 | -114.6515 | -0.018311 | 4.6829 | -117.2516 | -4.2068 | 4.715 | -4.3442 | 0.75531 | -0.21973 | -0.041198 | -0.082397 | -0.041199 | 109.5428 | 109.7626 | 2.5726 | 2.2934 | 2.7924 | 2.5131 | -0.027466 | 0.013733 | -8.876 | -8.8898 | 121.9391 | 121.9666 | 117.9794 | 112.8799 | 118.0069 | 112.9074
2341.1719 | 122.1268 | -117.2607 | 4.7012 | -0.077133 | 0.012236 | 0.15271 | 0.086009 | 0.0023221 | 0.014266 | 0.015851 | -0.18747 | 0.2672 | 0.14866 | 0.0027365 | 0.072877 | 0.086032 | -0.00019868 | -0.032043 | 4.7012 | -117.4484 | 4.5181 | -7.7911 | -114.9216 | -114.7247 | -0.018311 | 4.6875 | -117.3065 | -4.2984 | 4.6143 | -4.6051 | 0.68207 | -0.18768 | 0 | -0.18311 | -0.18311 | 109.4696 | 109.6573 | 2.536 | 2.3392 | 2.7237 | 2.5269 | -0.045777 | -0.013733 | -8.9996 | -8.9859 | 121.875 | 121.9208 | 117.9428 | 112.6556 | 117.9886 | 112.7014
2342.1719 | 122.2229 | -117.3019 | 4.6326 | -0.1458 | 0.011382 | 0.15229 | 0.085368 | -0.00066869 | 0.017043 | 0.017346 | -0.13559 | 0.26677 | 0.14909 | 0.0031637 | 0.075014 | 0.086887 | -0.0023349 | -0.10071 | 4.628 | -117.485 | 4.4907 | -7.6859 | -115.0085 | -114.4775 | -0.009155 | 4.6143 | -117.3065 | -4.2252 | 4.66 | -4.4037 | 0.75531 | -0.1831 | -0.004577 | -0.14191 | -0.13733 | 109.6161 | 109.7992 | 2.8244 | 2.2934 | 3.0075 | 2.4765 | -0.004578 | -0.01831 | -8.8577 | -8.8394 | 121.962 | 121.9666 | 118.0573 | 112.8983 | 118.0618 | 112.9028
2343.1719 | 122.1359 | -117.3386 | 4.6509 | -0.077133 | 0.013091 | 0.15271 | 0.084941 | -0.0023777 | 0.014052 | 0.018201 | -0.19052 | 0.26912 | 0.14845 | -0.0030313 | 0.071809 | 0.089023 | -0.0034031 | -0.027466 | 4.6921 | -117.485 | 4.4449 | -7.7271 | -114.9628 | -114.7842 | -0.022888 | 4.6417 | -117.3157 | -4.2984 | 4.5227 | -4.5914 | 0.66376 | -0.14648 | 0.041199 | -0.20599 | -0.24719 | 109.6115 | 109.758 | 2.5543 | 2.3758 | 2.7008 | 2.5223 | 0.022888 | -0.009155 | -8.9493 | -8.9401 | 121.8613 | 121.8384 | 118.0023 | 112.7472 | 117.9794 | 112.7243
2344.1719 | 122.1634 | -117.3203 | 4.6463 | -0.13054 | 0.0088184 | 0.151 | 0.081309 | -0.0038729 | 0.013625 | 0.019055 | -0.059296 | 0.26571 | 0.15101 | -0.0036721 | 0.071809 | 0.089023 | 0.00065592 | -0.073242 | 4.6417 | -117.4347 | 4.509 | -7.6538 | -114.9216 | -114.6835 | -0.036621 | 4.6692 | -117.2653 | -4.2847 | 4.6692 | -4.4357 | 0.71411 | -0.11444 | -0.004577 | -0.13733 | -0.13275 | 109.6664 | 109.7809 | 2.6367 | 2.3987 | 2.7512 | 2.5131 | 0.054931 | 0.022888 | -8.931 | -8.9539 | 121.9894 | 121.9345 | 118.0344 | 112.8845 | 117.9794 | 112.8296
2345.1719 | 122.1222 | -117.3294 | 4.6783 | -0.10765 | 0.012023 | 0.15036 | 0.087077 | 0.0021085 | 0.01747 | 0.018628 | -0.19968 | 0.2672 | 0.15101 | -0.0043131 | 0.07245 | 0.086246 | -0.0040439 | -0.018311 | 4.7012 | -117.453 | 4.4632 | -7.6767 | -114.9536 | -114.743 | 0.004578 | 4.6509 | -117.2928 | -4.2892 | 4.5731 | -4.5044 | 0.68665 | -0.1236 | 0.022888 | -0.21515 | -0.23804 | 109.6527 | 109.7763 | 2.5864 | 2.3758 | 2.71 | 2.4994 | 0.036621 | -0.027466 | -8.9676 | -8.9401 | 121.9025 | 121.8658 | 118.0161 | 112.825 | 117.9794 | 112.7884
2346.1719 | 122.1268 | -117.2745 | 4.6738 | -0.29076 | 0.011595 | 0.1525 | 0.0843 | -0.0032321 | 0.012343 | 0.016064 | -0.18442 | 0.26699 | 0.14717 | 0.0003867 | 0.073305 | 0.08945 | -0.0023349 | 0.009155 | 4.6967 | -117.4393 | 4.4861 | -7.6721 | -114.9857 | -114.7476 | -0.009155 | 4.6646 | -117.2882 | -4.3076 | 4.6005 | -4.4632 | 0.72327 | -0.16479 | 0.022888 | -0.18768 | -0.21057 | 109.6024 | 109.7672 | 2.5269 | 2.2888 | 2.6916 | 2.4536 | -0.013733 | -0.009155 | -8.9813 | -8.9722 | 121.875 | 121.8887 | 117.9977 | 112.8113 | 118.0115 | 112.825
2347.1719 | 122.3145 | -117.3157 | 4.6783 | -0.13817 | 0.010955 | 0.15271 | 0.085368 | 0.0010403 | 0.014479 | 0.016064 | -0.19052 | 0.26806 | 0.14717 | -0.0011087 | 0.075441 | 0.086032 | 0.00086952 | -0.07782 | 4.628 | -117.5217 | 4.3716 | -7.7637 | -115.0772 | -114.8392 | -0.12817 | 4.6417 | -117.3477 | -4.3213 | 4.5181 | -4.509 | 0.6958 | -0.20599 | -0.050354 | -0.3067 | -0.25635 | 109.552 | 109.758 | 2.4765 | 2.2385 | 2.6825 | 2.4445 | -0.032043 | -0.036621 | -8.9996 | -8.963 | 121.8338 | 121.8658 | 118.0115 | 112.8067 | 118.0435 | 112.8387
2348.1719 | 122.081 | -117.3477 | 4.6646 | -0.069504 | 0.013518 | 0.15229 | 0.088359 | 0.0021085 | 0.012984 | 0.016492 | -0.17831 | 0.26635 | 0.14887 | -4.0504e-05 | 0.073305 | 0.088169 | 0.00065592 | -0.064087 | 4.6143 | -117.5125 | 4.39 | -7.6492 | -115.0223 | -114.8254 | -0.07782 | 4.6326 | -117.3706 | -4.3076 | 4.5227 | -4.5136 | 0.68665 | -0.16479 | -0.050354 | -0.27466 | -0.2243 | 109.6985 | 109.8633 | 2.5223 | 2.3254 | 2.6871 | 2.4902 | -0.022888 | -0.032044 | -8.9722 | -8.9401 | 121.8704 | 121.8933 | 118.0344 | 112.8342 | 118.0573 | 112.8571
2349.1719 | 122.2046 | -117.4438 | 4.4586 | -0.1458 | 0.012664 | 0.151 | 0.083659 | -0.0045139 | 0.01277 | 0.019055 | -0.12033 | 0.26549 | 0.14866 | -0.0034585 | 0.073946 | 0.089664 | -0.0019077 | -0.036621 | 4.4403 | -117.6315 | 4.3625 | -7.5851 | -115.1093 | -115.0131 | -0.009155 | 4.4632 | -117.4622 | -4.1885 | 4.4907 | -4.4495 | 0.65918 | -0.18768 | -0.01831 | -0.09613 | -0.07782 | 109.8587 | 110.0464 | 2.4307 | 2.3346 | 2.6184 | 2.5223 | -0.01831 | 0.004578 | -8.6472 | -8.6517 | 121.9345 | 121.9528 | 118.103 | 112.9944 | 118.1213 | 113.0127
2350.1719 | 122.0856 | -117.4301 | 4.5456 | -0.084763 | 0.012877 | 0.15164 | 0.087931 | 0.0012539 | 0.016188 | 0.018414 | -0.17526 | 0.26592 | 0.14845 | -0.0015359 | 0.070955 | 0.088169 | -0.0027621 | -0.059509 | 4.5181 | -117.5949 | 4.335 | -7.6263 | -115.1184 | -114.7476 | -0.045776 | 4.5547 | -117.4255 | -4.3259 | 4.4403 | -4.628 | 0.58136 | -0.16479 | -0.027466 | -0.21057 | -0.1831 | 109.8038 | 109.9686 | 2.6825 | 2.3117 | 2.8473 | 2.4765 | 0.004578 | 0.009156 | -8.8715 | -8.8806 | 121.8704 | 121.8658 | 118.0115 | 112.8021 | 118.0069 | 112.7975
2351.1719 | 122.1542 | -117.3111 | 4.6921 | -0.10002 | 0.013091 | 0.15079 | 0.084727 | 0.0027493 | 0.015334 | 0.016278 | -0.17221 | 0.26806 | 0.14631 | 0.0055137 | 0.075441 | 0.089237 | 0.0025785 | 0.004578 | 4.6143 | -117.5537 | 4.4312 | -7.5348 | -115.0406 | -114.7705 | 0.045776 | 4.6097 | -117.3615 | -4.2526 | 4.5868 | -4.5456 | 0.72327 | -0.24262 | -0.07782 | -0.26093 | -0.18311 | 109.7763 | 110.0189 | 2.5406 | 2.2705 | 2.7832 | 2.5131 | -0.050354 | -0.082398 | -8.9447 | -8.8623 | 121.8979 | 121.9482 | 118.0344 | 112.7655 | 118.0847 | 112.8159
2352.1719 | 122.1268 | -117.366 | 4.6051 | -0.11528 | 0.011168 | 0.15293 | 0.085795 | 0.0021085 | 0.015548 | 0.015851 | -0.16611 | 0.26827 | 0.14887 | 0.0020957 | 0.073091 | 0.087528 | -0.0042575 | -0.10529 | 4.5822 | -117.5674 | 4.3579 | -7.5897 | -115.0635 | -114.8483 | -0.11902 | 4.5914 | -117.3798 | -4.3167 | 4.5181 | -4.5639 | 0.6958 | -0.20142 | -0.022889 | -0.24719 | -0.2243 | 109.7763 | 109.9777 | 2.5177 | 2.3026 | 2.7191 | 2.504 | -0.013733 | -0.013733 | -8.9218 | -8.9081 | 121.8842 | 121.8979 | 118.0618 | 112.8021 | 118.0756 | 112.8159
2353.1719 | 122.1313 | -117.3386 | 4.6371 | -0.13054 | 0.012023 | 0.15186 | 0.088145 | 0.00018591 | 0.015548 | 0.013501 | -0.16 | 0.26827 | 0.14951 | -0.0004679 | 0.073946 | 0.089023 | 0.00086952 | -0.036621 | 4.6967 | -117.485 | 4.4266 | -7.4936 | -115.0269 | -114.7934 | 0.009155 | 4.6097 | -117.3477 | -4.3304 | 4.5227 | -4.5044 | 0.70496 | -0.14648 | 0.059509 | -0.21057 | -0.27008 | 109.845 | 109.9915 | 2.5452 | 2.3117 | 2.6916 | 2.4582 | -0.009155 | -0.027466 | -8.9676 | -8.9401 | 121.8613 | 121.8704 | 118.0435 | 112.8342 | 118.0527 | 112.8433
2354.1719 | 122.1771 | -117.2745 | 4.6509 | -0.16106 | 0.012877 | 0.1525 | 0.084941 | 0.00082671 | 0.015761 | 0.019055 | -0.080658 | 0.26656 | 0.15016 | 0.0020957 | 0.073946 | 0.0871 | -0.0010531 | -0.032043 | 4.5547 | -117.4942 | 4.4815 | -7.4158 | -114.9399 | -114.7156 | 0.032043 | 4.5868 | -117.3111 | -4.2618 | 4.6555 | -4.3762 | 0.80566 | -0.21973 | -0.09613 | -0.16937 | -0.073243 | 109.8587 | 110.0784 | 2.5589 | 2.3346 | 2.7786 | 2.5543 | -0.036621 | -0.064087 | -8.9127 | -8.8486 | 121.9299 | 121.9666 | 118.0801 | 112.8983 | 118.1168 | 112.9349
2355.1719 | 122.0535 | -117.4026 | 4.5776 | -0.092392 | 0.01245 | 0.1525 | 0.086863 | -0.00024129 | 0.015548 | 0.018201 | -0.18442 | 0.26784 | 0.14717 | -0.0032449 | 0.073732 | 0.089664 | -0.0012667 | -0.045776 | 4.5456 | -117.5995 | 4.3213 | -7.5165 | -115.0726 | -114.8666 | -0.032043 | 4.5731 | -117.3889 | -4.3488 | 4.454 | -4.4952 | 0.65002 | -0.19684 | -0.032044 | -0.25635 | -0.2243 | 109.8862 | 110.083 | 2.536 | 2.33 | 2.7328 | 2.5269 | 0.013733 | -0.004578 | -8.9264 | -8.9218 | 121.8567 | 121.843 | 118.0527 | 112.9074 | 118.0389 | 112.8937
2356.1719 | 122.113 | -117.3889 | 4.5822 | -0.084763 | 0.01245 | 0.15164 | 0.084513 | 0.0018949 | 0.015761 | 0.016492 | -0.16611 | 0.26613 | 0.14717 | 0.0012411 | 0.074373 | 0.088382 | 0.0010831 | -0.036621 | 4.5914 | -117.5537 | 4.3762 | -7.4661 | -115.0497 | -114.8163 | -0.018311 | 4.5868 | -117.3615 | -4.3625 | 4.4815 | -4.4632 | 0.65918 | -0.1648 | 0.009156 | -0.20599 | -0.21515 | 109.9228 | 110.0876 | 2.5726 | 2.3392 | 2.7374 | 2.504 | 0.027466 | 0.004578 | -8.9447 | -8.9493 | 121.8704 | 121.843 | 118.0481 | 112.9257 | 118.0206 | 112.8983
2357.1875 | 122.3053 | -117.3889 | 4.5776 | -0.084763 | 0.0101 | 0.15186 | 0.082164 | -0.0038729 | 0.015975 | 0.015637 | -0.022675 | 0.26442 | 0.14994 | 0.0031637 | 0.073518 | 0.085819 | 0.00044212 | 0.013733 | 4.5364 | -117.4896 | 4.4815 | -7.3425 | -114.8849 | -114.7156 | 0.12817 | 4.5456 | -117.3248 | -4.2801 | 4.6371 | -4.5731 | 0.59052 | -0.10071 | -0.041199 | -0.096131 | -0.054932 | 110.0464 | 110.1471 | 2.6733 | 2.504 | 2.774 | 2.6047 | 0.064087 | -0.032044 | -8.8577 | -8.8257 | 122.0261 | 121.962 | 117.9794 | 112.8159 | 117.9153 | 112.7518
2358.1719 | 122.1405 | -117.3615 | 4.628 | -0.077133 | 0.013304 | 0.15015 | 0.087718 | 0.0016813 | 0.017256 | 0.015424 | -0.16306 | 0.26806 | 0.14909 | -0.0008951 | 0.072023 | 0.088169 | 0.0021511 | 0.027466 | 4.5685 | -117.5446 | 4.3991 | -7.489 | -115.0131 | -114.8209 | 0.041199 | 4.5914 | -117.3752 | -4.4037 | 4.541 | -4.5593 | 0.6958 | -0.18311 | -0.05951 | -0.22888 | -0.16937 | 109.8724 | 110.0555 | 2.5406 | 2.3483 | 2.7237 | 2.5314 | -0.013733 | -0.036621 | -9.0317 | -8.9951 | 121.9025 | 121.9162 | 118.0573 | 112.8021 | 118.071 | 112.8159
2359.1719 | 122.0718 | -117.3294 | 4.628 | -0.16106 | 0.010955 | 0.15229 | 0.087504 | 0.00039951 | 0.01277 | 0.016705 | -0.16916 | 0.26827 | 0.1508 | 0.0010275 | 0.074159 | 0.088382 | -0.0016941 | -0.045776 | 4.6097 | -117.5308 | 4.4128 | -7.4936 | -115.0635 | -114.8163 | -0.050354 | 4.6234 | -117.3569 | -4.4861 | 4.5044 | -4.5731 | 0.63171 | -0.20142 | -0.018311 | -0.21515 | -0.19684 | 109.8358 | 110.0372 | 2.5131 | 2.2659 | 2.7145 | 2.4673 | -0.027466 | -0.004578 | -9.1141 | -9.1095 | 121.8338 | 121.8613 | 117.9611 | 112.7563 | 117.9886 | 112.7838
2360.1719 | 122.1222 | -117.2836 | 4.6692 | -0.054245 | 0.013732 | 0.151 | 0.087718 | 0.0016813 | 0.012343 | 0.016705 | -0.17526 | 0.2672 | 0.14652 | 0.0006003 | 0.073305 | 0.087955 | 0.0023647 | -0.018311 | 4.6509 | -117.4805 | 4.5776 | -7.3013 | -114.8483 | -114.6469 | 0 | 4.7379 | -117.2791 | -4.4495 | 4.5914 | -4.5502 | 0.72784 | -0.19684 | -0.01831 | -0.091552 | -0.073242 | 109.9823 | 110.1791 | 2.6367 | 2.4353 | 2.8336 | 2.6321 | 0.004577 | 0.068665 | -9.1187 | -9.1873 | 121.875 | 121.8704 | 118.0115 | 112.7335 | 118.0069 | 112.7289
2361.1719 | 122.0947 | -117.3431 | 4.6371 | -0.10002 | 0.010955 | 0.15271 | 0.086863 | 0.00061311 | 0.014052 | 0.019482 | -0.1539 | 0.26613 | 0.14717 | 0.0052999 | 0.072023 | 0.086887 | 0.0021511 | -0.050354 | 4.5868 | -117.5629 | 4.3808 | -7.4203 | -114.2899 | -114.83 | -0.036621 | 4.6005 | -117.366 | -4.4495 | 4.5227 | -4.5364 | 0.62714 | -0.21973 | -0.050354 | -0.25635 | -0.20599 | 109.9228 | 110.1425 | 2.5131 | 3.0533 | 2.7328 | 3.273 | -0.022888 | -0.036621 | -9.0866 | -9.05 | 121.8658 | 121.8887 | 117.9703 | 112.8067 | 117.9932 | 112.8296
2362.1875 | 122.1039 | -117.3431 | 4.6234 | -0.077133 | 0.013518 | 0.1525 | 0.085582 | -0.0017367 | 0.015334 | 0.017346 | -0.092865 | 0.26378 | 0.14802 | -0.0070901 | 0.07245 | 0.089664 | 0.0030057 | -0.022888 | 4.5731 | -117.5354 | 4.4083 | -7.3837 | -114.8666 | -114.6515 | -0.004578 | 4.6783 | -117.3065 | -4.3808 | 4.5502 | -4.4495 | 0.65918 | -0.19226 | -0.050354 | -0.21515 | -0.16479 | 109.9594 | 110.1517 | 2.6917 | 2.4765 | 2.8839 | 2.6688 | 0.036621 | 0.054932 | -9.0042 | -9.0591 | 121.8933 | 121.8567 | 118.0023 | 112.8937 | 117.9657 | 112.8571
2363.1719 | 122.1085 | -117.2379 | 4.7379 | -0.077133 | 0.012877 | 0.15271 | 0.085795 | 0.00061311 | 0.014266 | 0.019696 | -0.1539 | 0.26656 | 0.14887 | 0.0012411 | 0.075227 | 0.08945 | 1.4916e-05 | -0.036621 | 4.7516 | -117.4438 | 4.5593 | -7.4661 | -114.8895 | -114.7064 | 0 | 4.7424 | -117.1875 | -4.2984 | 4.66 | -4.4769 | 0.74158 | -0.20599 | 0.013733 | -0.17853 | -0.19226 | 109.7717 | 109.9777 | 2.5314 | 2.3483 | 2.7374 | 2.5543 | 0.050354 | 0.004578 | -9.0363 | -9.0408 | 121.8979 | 121.8475 | 117.9794 | 112.7609 | 117.9291 | 112.7106
2364.1719 | 122.0764 | -78.8452 | 43.0252 | -0.33653 | 0.00091438 | 0.12558 | 0.086863 | -0.0019503 | 0.012984 | 0.013715 | -0.14474 | 0.1935 | 0.13926 | -0.0030313 | 0.075441 | 0.061038 | 0.0038601 | 0.10986 | 43.0115 | -78.8315 | 43.0161 | -5.6396 | -76.767 | -76.9272 | -0.022888 | 43.1625 | -78.4698 | -1.2955 | 42.2379 | -3.3508 | 32.9041 | 0.013733 | -0.013733 | -0.009156 | 0.004577 | 73.2056 | 73.1918 | 1.918 | 2.0782 | 1.9043 | 2.0645 | 0.37537 | 0.13733 | -44.3207 | -44.458 | 121.0831 | 120.7077 | 111.7493 | 75.4944 | 111.3739 | 75.119
2365.1719 | 122.2046 | -115.8325 | 6.0425 | -0.077133 | 0.013518 | 0.15164 | 0.086436 | 0.0014675 | 0.015761 | 0.017987 | -0.1539 | 0.27746 | 0.16447 | 0.0057273 | 0.073518 | 0.085605 | 0.0017239 | 0.013733 | 6.0104 | -116.1118 | 5.8182 | -7.4432 | -113.5483 | -113.3606 | 0.045776 | 6.0013 | -115.9058 | -1.181 | 5.9555 | -4.5776 | 1.8906 | -0.27923 | -0.032043 | -0.2243 | -0.19226 | 108.3893 | 108.6685 | 2.4719 | 2.2842 | 2.7512 | 2.5635 | -0.073242 | -0.041198 | -7.2235 | -7.1823 | 121.788 | 121.8613 | 117.7231 | 111.2549 | 117.7963 | 111.3281
2366.1719 | 122.1497 | -116.1209 | 5.7999 | -0.084763 | 0.012877 | 0.15378 | 0.087931 | -0.00024129 | 0.014052 | 0.017133 | -0.17221 | 0.28109 | 0.16468 | -0.0083719 | 0.072877 | 0.088809 | 0.0025785 | -0.036621 | 5.7724 | -116.3177 | 5.6168 | -7.4203 | -113.855 | -113.6673 | -0.027466 | 5.7953 | -116.153 | -1.236 | 5.7037 | -4.6921 | 1.6708 | -0.19684 | -0.027466 | -0.18311 | -0.15564 | 108.7006 | 108.8974 | 2.4536 | 2.2659 | 2.6505 | 2.4628 | -0.032043 | -0.004578 | -7.0358 | -7.0312 | 121.8246 | 121.8567 | 117.7917 | 111.4288 | 117.8238 | 111.4609
2367.1875 | 122.1222 | -116.3086 | 5.6534 | -0.16869 | 0.011382 | 0.15207 | 0.086863 | 0.0014675 | 0.015761 | 0.015637 | -0.10812 | 0.27938 | 0.16468 | -0.0034585 | 0.073305 | 0.085605 | 0.0032193 | -0.11444 | 5.6168 | -116.51 | 5.4382 | -7.3105 | -114.0518 | -113.9099 | -0.12817 | 5.6534 | -116.3086 | -1.2726 | 5.5206 | -4.7058 | 1.5381 | -0.20142 | -0.036621 | -0.21515 | -0.17853 | 108.9981 | 109.1995 | 2.3987 | 2.2568 | 2.6001 | 2.4582 | 0 | 0 | -6.926 | -6.926 | 121.8292 | 121.8292 | 117.8467 | 111.6028 | 117.8467 | 111.6028
2368.1719 | 122.0306 | -116.3635 | 5.5527 | -0.10002 | 0.013518 | 0.15335 | 0.08665 | 0.0012539 | 0.01747 | 0.018842 | -0.16916 | 0.28386 | 0.16425 | 0.0016685 | 0.072664 | 0.086887 | 0.00044212 | -0.1236 | 5.5481 | -116.6016 | 5.2872 | -6.2073 | -114.1754 | -114.537 | -0.14648 | 5.5481 | -116.3864 | -1.2909 | 5.4108 | -4.7745 | 1.5106 | -0.23804 | -0.004577 | -0.2655 | -0.26093 | 110.1562 | 110.3943 | 1.8265 | 2.1881 | 2.0645 | 2.4261 | -0.022889 | -0.004577 | -6.8436 | -6.839 | 121.7743 | 121.7972 | 117.8741 | 111.5891 | 117.897 | 111.6119
2369.1719 | 122.113 | -116.5009 | 5.4794 | -0.17632 | 0.013304 | 0.15485 | 0.088786 | 0.0012539 | 0.015548 | 0.018628 | -0.15695 | 0.28194 | 0.1681 | -4.0504e-05 | 0.070741 | 0.086246 | -0.0010531 | -0.07782 | 5.4703 | -116.6428 | 5.2689 | -6.2576 | -114.2487 | -113.9694 | -0.050354 | 5.4749 | -116.4871 | -1.3092 | 5.3833 | -4.66 | 1.387 | -0.14191 | -0.009155 | -0.21057 | -0.20142 | 110.2432 | 110.3851 | 2.5314 | 2.2522 | 2.6733 | 2.3941 | 0.013732 | -0.004577 | -6.7886 | -6.7841 | 121.8842 | 121.8704 | 117.8879 | 111.8408 | 117.8741 | 111.8271
2370.1719 | 122.1359 | -116.5375 | 5.4703 | -0.12291 | 0.011595 | 0.15442 | 0.087077 | 0.0014675 | 0.013411 | 0.014142 | -0.12643 | 0.28109 | 0.16554 | -0.0034585 | 0.072877 | 0.088169 | 0.0017239 | -0.004578 | 5.4016 | -116.7023 | 5.2551 | -6.2576 | -114.1479 | -113.9648 | 0.26093 | 5.4749 | -116.3818 | -1.236 | 5.3925 | -4.5593 | 1.3962 | -0.16479 | -0.068665 | -0.21515 | -0.14648 | 110.2798 | 110.4446 | 2.5726 | 2.3895 | 2.7374 | 2.5543 | 0.15564 | 0.004578 | -6.7062 | -6.7108 | 121.9299 | 121.7743 | 117.9337 | 111.9781 | 117.778 | 111.8225
2371.1719 | 122.1359 | -116.5512 | 5.4016 | -0.084763 | 0.013304 | 0.15528 | 0.087931 | 0.00082671 | 0.015334 | 0.018201 | -0.16306 | 0.28023 | 0.16575 | 0.0014547 | 0.071382 | 0.086246 | 0.0019375 | -0.009155 | 5.4245 | -116.716 | 5.2094 | -6.2897 | -114.1571 | -114.0152 | -0.004578 | 5.3925 | -116.5421 | -1.2955 | 5.3375 | -4.4586 | 1.3412 | -0.16479 | 0.022889 | -0.19226 | -0.21515 | 110.2615 | 110.4263 | 2.536 | 2.3941 | 2.7008 | 2.5589 | 0.009155 | -0.009155 | -6.6971 | -6.6879 | 121.8887 | 121.8796 | 117.8925 | 112.0926 | 117.8833 | 112.0834
2372.1875 | 122.1451 | -116.5558 | 5.3879 | -0.13054 | 0.010955 | 0.15293 | 0.086863 | 0.00082671 | 0.015761 | 0.018414 | -0.12643 | 0.27959 | 0.16703 | 0.0008139 | 0.075014 | 0.088169 | 0.0025785 | -0.40741 | 5.3787 | -116.748 | 5.1407 | -6.3721 | -114.4913 | -114.0656 | -0.11444 | 5.3741 | -116.5787 | -1.2955 | 5.3146 | -4.4769 | 1.3458 | -0.19226 | -0.009155 | -0.24719 | -0.23804 | 110.1837 | 110.376 | 2.4902 | 2.0645 | 2.6825 | 2.2568 | -0.022888 | -0.013732 | -6.6833 | -6.6696 | 121.8704 | 121.8933 | 117.9016 | 112.0789 | 117.9245 | 112.1017
2373.1719 | 122.1542 | -116.629 | 5.3558 | -0.15343 | 0.010955 | 0.14908 | 0.087077 | 0.00039951 | 0.014479 | 0.016064 | -0.11728 | 0.28045 | 0.16767 | 0.0035911 | 0.071809 | 0.085605 | 0.00086952 | -0.009155 | 5.3101 | -116.8213 | 5.1224 | -6.2576 | -114.2212 | -114.0793 | 0.054932 | 5.3558 | -116.5695 | -1.3229 | 5.2826 | -4.6783 | 1.3229 | -0.19226 | -0.045776 | -0.23346 | -0.18768 | 110.3714 | 110.5637 | 2.5497 | 2.4078 | 2.742 | 2.6001 | 0.059509 | 0 | -6.6788 | -6.6788 | 121.9116 | 121.8521 | 117.952 | 111.9507 | 117.8925 | 111.8912
2374.1719 | 122.1405 | -116.6428 | 5.3513 | -0.16106 | 0.013091 | 0.15207 | 0.08665 | 0.0014675 | 0.01512 | 0.016919 | -0.13864 | 0.28087 | 0.16511 | 0.0055137 | 0.0701 | 0.085178 | 0.0038601 | -0.13733 | 5.3329 | -116.7938 | 5.1041 | -6.2668 | -114.3082 | -114.1388 | -0.09613 | 5.3055 | -116.6153 | -1.3046 | 5.2734 | -4.9164 | 1.2955 | -0.15106 | -0.01831 | -0.24719 | -0.22888 | 110.376 | 110.527 | 2.504 | 2.3346 | 2.655 | 2.4857 | 0.027466 | -0.045776 | -6.6559 | -6.6101 | 121.9162 | 121.8887 | 117.9382 | 111.7264 | 117.9108 | 111.6989
2375.1719 | 122.0947 | -116.6336 | 5.3375 | -0.10002 | 0.01245 | 0.15421 | 0.087077 | 0.0027493 | 0.012557 | 0.019696 | -0.15085 | 0.28194 | 0.16532 | 0.0050863 | 0.073091 | 0.08646 | 0.0015103 | -0.032043 | 5.3238 | -116.7984 | 5.1498 | -4.8019 | -114.299 | -114.0564 | -0.027466 | 5.3192 | -116.6245 | -1.3321 | 5.2689 | -4.8981 | 1.2772 | -0.16479 | -0.013732 | -0.18768 | -0.17395 | 111.8317 | 111.9965 | 2.5772 | 2.3346 | 2.742 | 2.4994 | 0.009155 | -0.01831 | -6.6696 | -6.6513 | 121.9025 | 121.8933 | 117.9108 | 111.7355 | 117.9016 | 111.7264
2376.1719 | 122.0673 | -116.7389 | 5.2551 | -0.092392 | 0.011809 | 0.15293 | 0.087931 | 0.00082671 | 0.015761 | 0.01521 | -0.19052 | 0.28087 | 0.16297 | -0.0017495 | 0.072237 | 0.088809 | 0.0012967 | -0.064087 | 5.2322 | -116.9037 | 5.0171 | -4.8065 | -114.2349 | -114.2166 | -0.082397 | 5.2368 | -116.7114 | -1.3321 | 5.1132 | -4.9164 | 1.1948 | -0.16479 | -0.022888 | -0.23804 | -0.21515 | 111.9324 | 112.0972 | 2.5223 | 2.504 | 2.6871 | 2.6688 | 0.027466 | -0.018311 | -6.5872 | -6.5689 | 121.8521 | 121.8246 | 117.9337 | 111.8225 | 117.9062 | 111.795
2377.1875 | 122.0947 | -116.6885 | 5.3192 | -0.10765 | 0.01245 | 0.15229 | 0.084513 | -0.00024129 | 0.01512 | 0.016278 | -0.077606 | 0.28301 | 0.16618 | -0.0075173 | 0.072023 | 0.084323 | -0.0014805 | -0.022888 | 5.278 | -116.8167 | 5.0537 | -4.8065 | -114.3082 | -114.1251 | 0.013733 | 5.2872 | -116.6702 | -1.3458 | 5.191 | -4.7882 | 1.2177 | -0.12817 | -0.041199 | -0.2655 | -0.2243 | 111.882 | 112.0102 | 2.5635 | 2.3804 | 2.6916 | 2.5085 | 0.018311 | -0.032044 | -6.665 | -6.633 | 121.8796 | 121.8613 | 117.9062 | 111.9003 | 117.8879 | 111.882
2378.1719 | 122.1268 | -116.6748 | 5.3055 | -0.10002 | 0.010955 | 0.15442 | 0.086009 | -0.0036593 | 0.014907 | 0.01756 | -0.12033 | 0.28301 | 0.16596 | 0.0070089 | 0.072877 | 0.087528 | 0.00022852 | 0.004578 | 5.246 | -116.8716 | 5.0766 | -4.7974 | -114.3173 | -114.1113 | 0.07782 | 5.2551 | -116.6519 | -1.2909 | 5.2322 | -4.6692 | 1.2497 | -0.19684 | -0.059509 | -0.22888 | -0.16937 | 111.8774 | 112.0743 | 2.5635 | 2.3575 | 2.7603 | 2.5543 | 0.022888 | -0.050354 | -6.5964 | -6.546 | 121.907 | 121.8842 | 117.9245 | 112.0056 | 117.9016 | 111.9827
2379.1719 | 122.1451 | -116.6977 | 5.2277 | -0.13054 | 0.012877 | 0.15229 | 0.086863 | -0.00088229 | 0.015548 | 0.017133 | -0.17221 | 0.28109 | 0.16404 | 0.0016685 | 0.071168 | 0.085178 | 0.0030057 | -0.13733 | 5.2139 | -116.9174 | 5.0171 | -4.8523 | -114.4913 | -114.2395 | -0.082397 | 5.2139 | -116.7252 | -1.2863 | 5.1544 | -4.7333 | 1.2085 | -0.21973 | -0.013733 | -0.21057 | -0.19684 | 111.8454 | 112.0651 | 2.4582 | 2.2064 | 2.6779 | 2.4261 | -0.027466 | -0.013733 | -6.514 | -6.5002 | 121.8521 | 121.8796 | 117.9062 | 111.9644 | 117.9337 | 111.9919
2380.1719 | 122.0901 | -116.7343 | 5.2185 | -0.1458 | 0.011809 | 0.15357 | 0.089 | 0.0029629 | 0.011489 | 0.017133 | -0.18137 | 0.28216 | 0.16596 | -0.0023905 | 0.073305 | 0.088596 | 0.0019375 | -0.032043 | 5.2094 | -116.89 | 5.0262 | -4.8523 | -114.4043 | -114.2166 | -0.027466 | 5.1956 | -116.7343 | -1.3458 | 5.1224 | -4.8111 | 1.1673 | -0.15564 | -0.009155 | -0.19226 | -0.18311 | 111.882 | 112.0377 | 2.5177 | 2.33 | 2.6733 | 2.4857 | 0 | -0.022888 | -6.5643 | -6.5414 | 121.8567 | 121.8567 | 117.9016 | 111.9232 | 117.9016 | 111.9232
2381.1719 | 122.1176 | -116.7206 | 5.2139 | -0.12291 | 0.012236 | 0.15293 | 0.08665 | 0.0023221 | 0.013198 | 0.018414 | -0.12949 | 0.28066 | 0.16361 | -0.0030313 | 0.073518 | 0.084323 | 0.00065592 | -0.004578 | 5.2002 | -116.9357 | 4.985 | -4.8569 | -114.4089 | -114.2075 | 0.013733 | 5.2139 | -116.7435 | -1.355 | 5.159 | -4.7791 | 1.1948 | -0.21515 | -0.013733 | -0.22888 | -0.21515 | 111.8637 | 112.0789 | 2.5131 | 2.3117 | 2.7283 | 2.5269 | -0.022888 | 0 | -6.5689 | -6.5689 | 121.8796 | 121.9025 | 117.9153 | 111.9415 | 117.9382 | 111.9644
2382.1875 | 122.2046 | -116.7343 | 5.1544 | -0.1458 | 0.014159 | 0.15485 | 0.08665 | 0.00039951 | 0.017256 | 0.015637 | -0.065399 | 0.27874 | 0.16383 | -0.0034585 | 0.073732 | 0.086887 | 0.0012967 | -0.045776 | 5.1453 | -116.8854 | 5.0766 | -4.7241 | -114.3219 | -114.0656 | 0.009155 | 5.1498 | -116.6977 | -1.2726 | 5.2231 | -4.7974 | 1.2222 | -0.15106 | -0.009155 | -0.07782 | -0.068665 | 112.0102 | 112.1613 | 2.6688 | 2.4124 | 2.8198 | 2.5635 | 0.036621 | -0.004578 | -6.427 | -6.4224 | 121.9574 | 121.9208 | 117.9565 | 111.937 | 117.9199 | 111.9003
2383.1719 | 122.1039 | -116.748 | 5.2185 | -0.092392 | 0.011382 | 0.15314 | 0.085795 | 0.0025357 | 0.015334 | 0.013715 | -0.15085 | 0.27831 | 0.16682 | 0.0014547 | 0.070741 | 0.088596 | 0.0017239 | 0.013733 | 5.159 | -116.9724 | 4.9942 | -4.8019 | -114.3539 | -114.3723 | 0.059509 | 5.1682 | -116.7526 | -1.4145 | 5.159 | -4.8203 | 1.1856 | -0.2243 | -0.059509 | -0.2243 | -0.1648 | 111.9461 | 112.1704 | 2.3758 | 2.3941 | 2.6001 | 2.6184 | -0.004578 | -0.050354 | -6.633 | -6.5826 | 121.907 | 121.9116 | 117.9337 | 111.9278 | 117.9382 | 111.9324
2384.1719 | 122.049 | -116.7664 | 5.2002 | -0.069504 | 0.013091 | 0.15271 | 0.086863 | 0.0012539 | 0.015761 | 0.017133 | -0.16306 | 0.28002 | 0.16383 | -0.0011087 | 0.072023 | 0.08646 | 0.00086952 | -0.059509 | 5.1498 | -116.9769 | 4.9301 | -4.4815 | -114.4867 | -114.2761 | -0.11444 | 5.2094 | -116.7572 | -1.4832 | 5.0308 | -4.8386 | 1.1353 | -0.21057 | -0.050354 | -0.27008 | -0.21973 | 112.2849 | 112.4954 | 2.4902 | 2.2797 | 2.7008 | 2.4902 | 0.009155 | 0.009156 | -6.6833 | -6.6925 | 121.7972 | 121.788 | 117.9016 | 111.9278 | 117.8925 | 111.9186
2385.1719 | 122.0627 | -116.7847 | 5.2139 | -0.17632 | 0.013304 | 0.15207 | 0.086222 | -0.0017367 | 0.014266 | 0.015424 | -0.12949 | 0.2828 | 0.16682 | 0.0012411 | 0.072877 | 0.086887 | 0.0015103 | -0.009155 | 5.2231 | -116.954 | 4.9118 | -4.8111 | -114.418 | -114.2487 | -0.013733 | 5.1773 | -116.7892 | -1.4694 | 5.0858 | -4.8386 | 1.1398 | -0.16937 | 0.009155 | -0.30212 | -0.31128 | 111.9736 | 112.1429 | 2.536 | 2.3666 | 2.7054 | 2.536 | -0.004578 | -0.036621 | -6.6833 | -6.6467 | 121.8704 | 121.875 | 117.9245 | 111.9461 | 117.9291 | 111.9507
2386.1719 | 122.0764 | -116.7801 | 5.1865 | -0.084763 | 0.011168 | 0.15058 | 0.082377 | -0.0038729 | 0.014479 | 0.016919 | -0.11728 | 0.28109 | 0.16554 | -0.0028177 | 0.07245 | 0.085391 | -0.0027621 | -0.022888 | 5.1682 | -116.9815 | 4.866 | -4.8294 | -114.418 | -114.299 | -0.004578 | 5.2094 | -116.7526 | -1.4832 | 5.0858 | -4.7928 | 1.1444 | -0.20142 | -0.01831 | -0.32043 | -0.30212 | 111.9507 | 112.1521 | 2.4811 | 2.3621 | 2.6825 | 2.5635 | 0.027465 | 0.022889 | -6.6696 | -6.6925 | 121.8658 | 121.8384 | 117.9245 | 111.9873 | 117.897 | 111.9598
2387.1719 | 122.081 | -116.7664 | 5.1865 | -0.43571 | 0.013304 | 0.1525 | 0.089213 | 0.0031765 | 0.013839 | 0.015424 | -0.15695 | 0.28087 | 0.16404 | -0.0021769 | 0.072877 | 0.086032 | -0.0029757 | -0.13733 | 5.1819 | -116.9312 | 4.9301 | -4.8889 | -114.4501 | -114.2899 | -0.10529 | 5.159 | -116.7709 | -1.4603 | 5.072 | -4.7516 | 1.1444 | -0.16479 | -0.004577 | -0.25635 | -0.25177 | 111.8774 | 112.0422 | 2.4765 | 2.3163 | 2.6413 | 2.4811 | -0.004578 | -0.027465 | -6.6467 | -6.6193 | 121.8384 | 121.843 | 117.9108 | 112.0148 | 117.9153 | 112.0193
2388.1719 | 122.0856 | -116.6931 | 5.2322 | -0.10002 | 0.01245 | 0.15186 | 0.087077 | 0.0029629 | 0.013411 | 0.018201 | -0.16916 | 0.27938 | 0.16725 | -0.0040995 | 0.07245 | 0.084751 | 0.0023647 | -0.004578 | 5.1865 | -116.9266 | 5.0354 | -4.8248 | -114.4455 | -114.2303 | 0.19226 | 5.2277 | -116.6199 | -1.3916 | 5.1682 | -4.7424 | 1.2039 | -0.23346 | -0.045777 | -0.19684 | -0.15106 | 111.8683 | 112.1017 | 2.4628 | 2.2476 | 2.6962 | 2.4811 | 0.073242 | -0.004578 | -6.6238 | -6.6193 | 121.8613 | 121.788 | 117.897 | 111.9507 | 117.8238 | 111.8774
2389.1875 | 122.1634 | -116.6931 | 5.1865 | -0.16869 | 0.011809 | 0.15378 | 0.087718 | 0.0018949 | 0.015761 | 0.018628 | -0.14169 | 0.28045 | 0.16212 | -0.0043131 | 0.070955 | 0.087528 | 0.0036465 | -0.059509 | 5.1498 | -116.9357 | 5.0537 | -4.7607 | -114.2807 | -114.2166 | -0.009155 | 5.1636 | -116.716 | -1.3504 | 5.2368 | -4.6875 | 1.2726 | -0.24262 | -0.036621 | -0.13275 | -0.09613 | 111.9324 | 112.175 | 2.4765 | 2.4124 | 2.7191 | 2.655 | -0.022888 | -0.022888 | -6.5369 | -6.514 | 121.9299 | 121.9528 | 117.9657 | 112.0056 | 117.9886 | 112.0285
2390.1719 | 122.081 | -116.7664 | 5.2002 | -0.084763 | 0.01245 | 0.15143 | 0.089213 | -0.00045509 | 0.015548 | 0.014996 | -0.13254 | 0.28066 | 0.16596 | 0.0003867 | 0.071809 | 0.087528 | 0.0010831 | -0.022888 | 5.1773 | -116.922 | 4.9622 | -4.8843 | -114.4318 | -114.2807 | 0.013733 | 5.1956 | -116.748 | -1.4648 | 5.1315 | -4.6967 | 1.1581 | -0.15564 | -0.022888 | -0.23804 | -0.21515 | 111.882 | 112.0377 | 2.4857 | 2.3346 | 2.6413 | 2.4902 | 0.01831 | -0.004577 | -6.665 | -6.6605 | 121.8979 | 121.8796 | 117.9245 | 112.0697 | 117.9062 | 112.0514
2391.1719 | 122.1085 | -116.7801 | 5.2185 | -0.046616 | 0.012236 | 0.1525 | 0.086863 | 0.00039951 | 0.016616 | 0.013715 | -0.13559 | 0.28237 | 0.16404 | 0.0042319 | 0.071596 | 0.087955 | -0.0010531 | -0.032043 | 5.1865 | -116.9495 | 4.9667 | -4.8752 | -114.3997 | -114.2944 | -0.018311 | 5.1956 | -116.7709 | -1.506 | 5.1041 | -4.6646 | 1.1444 | -0.16937 | -0.032044 | -0.25177 | -0.21973 | 111.9049 | 112.0743 | 2.4857 | 2.3804 | 2.655 | 2.5497 | 0.009155 | -0.022888 | -6.7245 | -6.7017 | 121.8842 | 121.875 | 117.9245 | 112.1155 | 117.9153 | 112.1063
2392.1719 | 122.0764 | -116.7389 | 5.2231 | -0.13054 | 0.013304 | 0.15186 | 0.087077 | 0.0010403 | 0.015548 | 0.016492 | -0.19662 | 0.28258 | 0.16383 | 0.0003867 | 0.071809 | 0.083896 | -0.00019868 | -0.10529 | 5.1819 | -116.9769 | 4.9347 | -4.8935 | -114.4684 | -114.3219 | -0.12817 | 5.2002 | -116.7343 | -1.4511 | 5.0079 | -4.6005 | 1.1902 | -0.23804 | -0.041198 | -0.28839 | -0.24719 | 111.8454 | 112.0834 | 2.417 | 2.2705 | 2.655 | 2.5085 | 0.004578 | -0.022888 | -6.6742 | -6.6513 | 121.7468 | 121.7422 | 117.9291 | 112.1384 | 117.9245 | 112.1338
2393.1719 | 122.081 | -116.7801 | 5.2048 | -0.092392 | 0.013518 | 0.15293 | 0.084513 | -0.0013095 | 0.014907 | 0.016492 | -0.095917 | 0.27895 | 0.16212 | -0.0023905 | 0.072237 | 0.085605 | 0.00086952 | 0.013733 | 5.159 | -116.9174 | 4.9576 | -4.8706 | -114.4226 | -114.2715 | 0.032043 | 5.1773 | -116.7709 | -1.5198 | 5.0903 | -4.5776 | 1.1307 | -0.13733 | -0.045776 | -0.24719 | -0.20142 | 111.9095 | 112.0468 | 2.5085 | 2.3575 | 2.6459 | 2.4948 | 0.009155 | -0.027466 | -6.7245 | -6.6971 | 121.8704 | 121.8613 | 117.9108 | 112.2025 | 117.9016 | 112.1933
2394.1719 | 122.0718 | -116.7618 | 5.2505 | -0.069504 | 0.011595 | 0.15186 | 0.085154 | -0.0017367 | 0.011702 | 0.019055 | -0.12338 | 0.28002 | 0.16425 | 0.0065817 | 0.073091 | 0.083042 | -0.0010531 | -0.050354 | 5.1636 | -116.9266 | 4.9347 | -4.8523 | -114.4363 | -114.3082 | -0.009155 | 5.1865 | -116.7801 | -1.3779 | 5.1636 | -4.7333 | 1.0757 | -0.16479 | -0.086975 | -0.31586 | -0.22888 | 111.9095 | 112.0743 | 2.4536 | 2.3254 | 2.6184 | 2.4902 | -0.01831 | -0.064087 | -6.6284 | -6.5643 | 121.9254 | 121.9437 | 117.8375 | 112.0285 | 117.8558 | 112.0468
2395.1719 | 122.0718 | -116.7389 | 5.2185 | -0.15343 | 0.012236 | 0.15036 | 0.085582 | -0.0045139 | 0.015761 | 0.014569 | -0.12949 | 0.2813 | 0.16319 | -0.0030313 | 0.072237 | 0.086673 | 0.0015103 | 0.14191 | 5.2231 | -116.8762 | 4.9393 | -4.8752 | -114.3402 | -114.2807 | 0 | 5.2002 | -116.7618 | -1.506 | 5.1041 | -4.5868 | 1.1444 | -0.13733 | 0.004577 | -0.27924 | -0.28381 | 111.8637 | 112.001 | 2.4582 | 2.3987 | 2.5955 | 2.536 | -0.022888 | -0.018311 | -6.7245 | -6.7062 | 121.843 | 121.8658 | 117.8833 | 112.1521 | 117.9062 | 112.175
2396.1719 | 122.1313 | -116.7435 | 5.2277 | -0.069504 | 0.011809 | 0.15186 | 0.088786 | 0.0010403 | 0.015334 | 0.017133 | -0.15695 | 0.28045 | 0.16297 | 0.0012411 | 0.07245 | 0.086032 | -0.0010531 | 0.004578 | 5.1819 | -116.9403 | 4.9896 | -4.8523 | -114.418 | -114.2761 | 0.004578 | 5.2139 | -116.748 | -1.4969 | 5.1178 | -4.5273 | 1.1719 | -0.19684 | -0.045776 | -0.23804 | -0.19226 | 111.8912 | 112.088 | 2.4673 | 2.3254 | 2.6642 | 2.5223 | -0.004578 | -0.013733 | -6.7245 | -6.7108 | 121.8613 | 121.8658 | 117.9153 | 112.2162 | 117.9199 | 112.2208
2397.1719 | 122.1085 | -116.6977 | 5.2414 | -0.13817 | 0.013091 | 0.15378 | 0.088145 | -2.7692e-05 | 0.014479 | 0.018201 | -0.17526 | 0.28216 | 0.16233 | -0.0017495 | 0.072664 | 0.088169 | 0.0023647 | 0.15564 | 5.1773 | -116.9357 | 5.0262 | -4.9072 | -114.6194 | -114.3402 | 0.12817 | 5.2139 | -116.6016 | -1.4374 | 5.1727 | -4.5227 | 1.1993 | -0.23804 | -0.064087 | -0.21515 | -0.15106 | 111.7905 | 112.0285 | 2.3575 | 2.0782 | 2.5955 | 2.3163 | 0.096131 | -0.027466 | -6.6788 | -6.6513 | 121.8704 | 121.7743 | 117.897 | 112.175 | 117.8009 | 112.0789
2398.1719 | 122.1542 | -116.6931 | 5.2002 | -0.13054 | 0.011168 | 0.15399 | 0.088786 | -0.0010959 | 0.014907 | 0.017987 | -0.13864 | 0.27959 | 0.16297 | 0.0010275 | 0.071809 | 0.085819 | 0.0017239 | -0.10529 | 5.1544 | -116.9037 | 5.0217 | -4.8065 | -114.3539 | -114.2349 | -0.059509 | 5.1773 | -116.7435 | -1.3916 | 5.1819 | -4.4815 | 1.2451 | -0.21057 | -0.045776 | -0.17853 | -0.13275 | 111.8866 | 112.0972 | 2.4582 | 2.3392 | 2.6688 | 2.5497 | -0.050354 | -0.022888 | -6.5918 | -6.5689 | 121.875 | 121.9254 | 117.9382 | 112.2116 | 117.9886 | 112.262
2399.1719 | 122.1634 | -116.7389 | 5.2231 | -0.13054 | 0.011809 | 0.15271 | 0.085368 | -0.0019503 | 0.014266 | 0.016064 | -0.089813 | 0.28151 | 0.1649 | -0.0055947 | 0.071382 | 0.086032 | 0.0021511 | 0.032043 | 5.1773 | -116.8625 | 5.0812 | -4.7287 | -114.2624 | -114.1708 | 0.10986 | 5.1819 | -116.6702 | -1.4282 | 5.2185 | -4.4495 | 1.2314 | -0.1236 | -0.045776 | -0.14191 | -0.09613 | 112.0102 | 112.1338 | 2.5681 | 2.4765 | 2.6917 | 2.6001 | 0.068665 | -0.041198 | -6.6513 | -6.6101 | 121.9574 | 121.8887 | 117.9703 | 112.2894 | 117.9016 | 112.2208
2400.1719 | 122.0947 | -116.6931 | 5.2963 | -0.10002 | 0.01245 | 0.15186 | 0.0843 | -0.00045509 | 0.015548 | 0.016919 | -0.080658 | 0.27959 | 0.1634 | 0.0006003 | 0.072877 | 0.084751 | 0.0032193 | 0.036621 | 5.2048 | -116.8716 | 5.0766 | -4.8019 | -114.267 | -114.1342 | 0.073242 | 5.2277 | -116.6656 | -1.4969 | 5.1727 | -4.4861 | 1.1902 | -0.17853 | -0.091553 | -0.21973 | -0.12817 | 111.8912 | 112.0697 | 2.5589 | 2.4261 | 2.7374 | 2.6047 | 0.027466 | -0.068665 | -6.7932 | -6.7245 | 121.8658 | 121.8384 | 117.8833 | 112.207 | 117.8558 | 112.1796
2401.1875 | 122.1451 | -116.6702 | 5.2231 | -0.15343 | 0.013091 | 0.15357 | 0.086863 | 0.0027493 | 0.013411 | 0.014569 | -0.13864 | 0.27938 | 0.16554 | 0.0010275 | 0.069246 | 0.085605 | 0.00065592 | -0.032043 | 5.1727 | -116.8167 | 5.1315 | -4.7241 | -114.2303 | -114.6698 | -0.1831 | 5.1682 | -116.7252 | -1.4008 | 5.2551 | -4.4312 | 1.2589 | -0.14648 | -0.050354 | -0.091552 | -0.041198 | 111.9461 | 112.0926 | 2.0004 | 2.4399 | 2.1469 | 2.5864 | -0.054932 | -0.054931 | -6.6238 | -6.5689 | 121.9254 | 121.9803 | 117.9291 | 112.2391 | 117.984 | 112.294
2402.1719 | 122.0764 | -116.7114 | 5.3009 | -0.061874 | 0.014159 | 0.15293 | 0.083232 | 0.00082671 | 0.016188 | 0.015424 | -0.12643 | 0.28002 | 0.16255 | 0.0012411 | 0.072664 | 0.085819 | -0.0029757 | -0.009155 | 5.2277 | -116.8442 | 5.0674 | -4.7836 | -114.299 | -113.1638 | 0.032043 | 5.246 | -116.6794 | -1.5244 | 5.1727 | -4.5273 | 1.1856 | -0.13275 | -0.073242 | -0.23346 | -0.16022 | 111.9278 | 112.0605 | 3.5477 | 2.4124 | 3.6804 | 2.5452 | 0.032044 | -0.054931 | -6.8253 | -6.7703 | 121.8842 | 121.8521 | 117.897 | 112.1841 | 117.865 | 112.1521
2403.1719 | 122.1176 | -116.6061 | 5.3467 | -0.10765 | 0.010314 | 0.15293 | 0.086863 | 0.0016813 | 0.016402 | 0.015851 | -0.14474 | 0.28045 | 0.16468 | -0.0015359 | 0.071809 | 0.086673 | -0.00083948 | -0.022888 | 5.278 | -116.8533 | 5.04 | -4.7882 | -114.299 | -114.2029 | 0.009155 | 5.2643 | -116.6748 | -1.5335 | 5.1682 | -4.5319 | 1.2909 | -0.24719 | -0.068665 | -0.3067 | -0.23804 | 111.8179 | 112.0651 | 2.4033 | 2.3071 | 2.6505 | 2.5543 | -0.068665 | -0.082398 | -6.8802 | -6.7978 | 121.7743 | 121.843 | 117.897 | 112.0743 | 117.9657 | 112.1429
2404.1719 | 122.0673 | -116.5009 | 5.4611 | 0.044937 | 0.0098864 | 0.15186 | 0.089427 | 0.00082671 | 0.014693 | 0.01521 | -0.18442 | 0.28045 | 0.1634 | 0.0038047 | 0.070955 | 0.0871 | -0.00083948 | -0.059509 | 5.4794 | -116.6428 | 5.2597 | -4.7562 | -114.1342 | -114.0793 | 0.009155 | 5.4291 | -116.4963 | -1.474 | 5.365 | -4.4952 | 1.3596 | -0.14191 | 0.01831 | -0.20142 | -0.21973 | 111.7447 | 111.8866 | 2.4216 | 2.3666 | 2.5635 | 2.5085 | 0.004577 | -0.032044 | -6.9351 | -6.9031 | 121.8658 | 121.8613 | 117.8604 | 112.0056 | 117.8558 | 112.001
2405.1719 | 122.1313 | -116.4551 | 5.4382 | -0.046616 | 0.013945 | 0.15164 | 0.08195 | -0.0015231 | 0.014052 | 0.013928 | -0.059296 | 0.27788 | 0.1634 | -0.0073037 | 0.071382 | 0.085391 | 0.0015103 | 0.009155 | 5.4016 | -116.6336 | 5.3192 | -4.6417 | -114.0839 | -113.9328 | 0.1236 | 5.4291 | -116.4551 | -1.4236 | 5.452 | -4.4083 | 1.4557 | -0.17853 | -0.036621 | -0.11902 | -0.082397 | 111.8134 | 111.9919 | 2.5223 | 2.3712 | 2.7008 | 2.5497 | 0 | -0.009155 | -6.8619 | -6.8527 | 121.907 | 121.907 | 117.9108 | 112.0468 | 117.9108 | 112.0468
2406.1875 | 122.1725 | -116.4001 | 5.4977 | -0.1458 | 0.013304 | 0.15229 | 0.08665 | -2.7692e-05 | 0.015975 | 0.015424 | -0.14169 | 0.27895 | 0.16126 | -0.0023905 | 0.073091 | 0.086032 | 0.0021511 | -0.082397 | 5.4337 | -116.5741 | 5.4428 | -4.6555 | -114.093 | -113.8779 | 0.022888 | 5.4245 | -116.4139 | -1.387 | 5.5206 | -4.4174 | 1.4877 | -0.17395 | -0.064087 | -0.054932 | 0.009155 | 111.7447 | 111.9186 | 2.5223 | 2.3071 | 2.6962 | 2.4811 | -0.013733 | -0.073242 | -6.8848 | -6.8115 | 121.9208 | 121.9345 | 117.8879 | 111.9827 | 117.9016 | 111.9965
2407.1719 | 122.1085 | -116.4093 | 5.5252 | -0.11528 | 0.011809 | 0.15229 | 0.089213 | -0.0015231 | 0.011916 | 0.016492 | -0.18747 | 0.27852 | 0.16554 | 0.0001731 | 0.069459 | 0.083896 | 0.0021511 | 0.009155 | 5.4886 | -116.6107 | 5.3513 | -4.6875 | -114.0747 | -114.1525 | 0.045776 | 5.5023 | -116.4322 | -1.4465 | 5.4611 | -4.4586 | 1.4557 | -0.20142 | -0.036622 | -0.17395 | -0.13733 | 111.7218 | 111.9232 | 2.2568 | 2.3346 | 2.4582 | 2.536 | -0.022888 | -0.022889 | -6.9717 | -6.9489 | 121.8704 | 121.8933 | 117.865 | 111.9507 | 117.8879 | 111.9736
2408.1719 | 122.0627 | -116.4505 | 5.5161 | -0.054245 | 0.012023 | 0.15314 | 0.085795 | 0.0012539 | 0.01512 | 0.015637 | -0.10812 | 0.27981 | 0.16404 | 0.0001731 | 0.074373 | 0.084751 | 0.0034329 | 0.022888 | 5.4611 | -116.6107 | 5.3101 | -4.6921 | -114.0976 | -113.942 | 0.041199 | 5.4932 | -116.4642 | -1.5244 | 5.4199 | -4.4998 | 1.3733 | -0.16022 | -0.054931 | -0.20599 | -0.15106 | 111.7584 | 111.9186 | 2.5085 | 2.3529 | 2.6688 | 2.5131 | -0.013733 | -0.022888 | -7.0404 | -7.0175 | 121.8704 | 121.8842 | 117.8238 | 111.9507 | 117.8375 | 111.9644
2409.1719 | 122.1542 | -116.3635 | 5.5161 | 0.067825 | 0.012236 | 0.15357 | 0.084941 | -0.0015231 | 0.016829 | 0.01286 | -0.071503 | 0.27895 | 0.1634 | 0.0012411 | 0.072023 | 0.086887 | 1.4916e-05 | 0 | 5.452 | -116.5787 | 5.4291 | -4.5776 | -114.006 | -113.8412 | 0.1236 | 5.4886 | -116.3452 | -1.4053 | 5.5527 | -4.39 | 1.5015 | -0.21515 | -0.064087 | -0.086975 | -0.022888 | 111.7859 | 112.001 | 2.5223 | 2.3575 | 2.7374 | 2.5726 | 0.01831 | -0.027466 | -6.9214 | -6.8939 | 121.9162 | 121.8979 | 117.865 | 111.9736 | 117.8467 | 111.9553
2410.1719 | 122.0993 | -116.3727 | 5.5618 | -0.077133 | 0.010314 | 0.15207 | 0.086863 | 0.0016813 | 0.01213 | 0.018414 | -0.16611 | 0.27874 | 0.16383 | 0.0057273 | 0.072237 | 0.087314 | -0.00083948 | 0.004578 | 5.5023 | -116.6107 | 5.3284 | -4.7653 | -114.0656 | -113.9374 | 0.045776 | 5.5069 | -116.3956 | -1.4648 | 5.4749 | -4.4724 | 1.4465 | -0.23804 | -0.05951 | -0.23346 | -0.17395 | 111.6074 | 111.8454 | 2.4353 | 2.3071 | 2.6733 | 2.5452 | -0.022888 | -0.054932 | -7.0267 | -6.9717 | 121.8475 | 121.8704 | 117.8192 | 111.9003 | 117.8421 | 111.9232
2411.1719 | 122.0215 | -116.3544 | 5.5756 | -0.077133 | 0.012023 | 0.15357 | 0.08665 | -0.0010959 | 0.016616 | 0.018628 | -0.22409 | 0.28151 | 0.16361 | 0.0059409 | 0.071168 | 0.087741 | -0.00083948 | -0.059509 | 5.5161 | -116.597 | 5.2963 | -4.8431 | -114.0701 | -113.9694 | -0.11902 | 5.5801 | -116.4093 | -1.5198 | 5.4199 | -4.5776 | 1.4236 | -0.24262 | -0.05951 | -0.27924 | -0.21973 | 111.5112 | 111.7538 | 2.3849 | 2.2842 | 2.6276 | 2.5269 | -0.054932 | 0.004577 | -7.0953 | -7.0999 | 121.7743 | 121.8292 | 117.778 | 111.7767 | 117.8329 | 111.8317
2412.1719 | 122.0444 | -116.3315 | 5.6213 | -0.092392 | 0.011382 | 0.15207 | 0.083873 | -0.0034457 | 0.015334 | 0.014783 | -0.13864 | 0.28002 | 0.16276 | -0.0051675 | 0.072877 | 0.083042 | 0.0032193 | -0.027466 | 5.5618 | -116.5375 | 5.3329 | -4.7607 | -114.061 | -113.942 | 0 | 5.5664 | -116.3773 | -1.4877 | 5.3925 | -4.4495 | 1.5244 | -0.20599 | -0.059509 | -0.28839 | -0.22888 | 111.5707 | 111.7767 | 2.3895 | 2.2705 | 2.5955 | 2.4765 | -0.045776 | -0.054932 | -7.1091 | -7.0541 | 121.7239 | 121.7697 | 117.8558 | 111.882 | 117.9016 | 111.9278
2413.1719 | 122.0901 | -116.3544 | 5.5847 | -0.13054 | 0.011382 | 0.15164 | 0.086222 | 0.0014675 | 0.014693 | 0.016064 | -0.15085 | 0.2781 | 0.16212 | 0.0031637 | 0.072237 | 0.084323 | -0.0010531 | -0.013733 | 5.5389 | -116.5741 | 5.3787 | -4.7516 | -114.006 | -113.9282 | 0.036621 | 5.5573 | -116.3635 | -1.4786 | 5.5435 | -4.4907 | 1.5427 | -0.21973 | -0.045777 | -0.20599 | -0.16022 | 111.6028 | 111.8225 | 2.4261 | 2.3483 | 2.6459 | 2.5681 | -0.009155 | -0.027466 | -7.0633 | -7.0358 | 121.8979 | 121.907 | 117.897 | 111.8637 | 117.9062 | 111.8729
2414.1719 | 122.2 | -116.3086 | 5.5573 | -0.12291 | 0.012877 | 0.15271 | 0.084086 | -0.0025913 | 0.014479 | 0.012219 | -0.065399 | 0.27895 | 0.16361 | -0.0036721 | 0.07245 | 0.08411 | 0.0012967 | 0.013733 | 5.5252 | -116.4688 | 5.4794 | -4.5868 | -113.8641 | -113.7955 | 0.11902 | 5.5389 | -116.3452 | -1.3916 | 5.6488 | -4.3716 | 1.6434 | -0.16022 | -0.032043 | -0.07782 | -0.045777 | 111.7218 | 111.882 | 2.5131 | 2.4445 | 2.6733 | 2.6047 | -0.036621 | -0.018311 | -6.9489 | -6.9305 | 121.9574 | 121.994 | 117.952 | 111.937 | 117.9886 | 111.9736
2415.1719 | 122.1222 | -116.3269 | 5.5984 | -0.12291 | 0.013091 | 0.15058 | 0.084941 | -0.0051547 | 0.017043 | 0.014569 | -0.080658 | 0.27724 | 0.16233 | -0.0049539 | 0.069246 | 0.084323 | -0.0016941 | 0.045776 | 5.5389 | -116.4688 | 5.4794 | -4.5639 | -114.0244 | -113.8 | 0.10986 | 5.5344 | -116.2857 | -1.4465 | 5.5984 | -4.4495 | 1.5656 | -0.14191 | -0.05951 | -0.11902 | -0.059509 | 111.763 | 111.9049 | 2.5269 | 2.3026 | 2.6688 | 2.4445 | 0.041198 | -0.064087 | -7.045 | -6.9809 | 121.9254 | 121.8842 | 117.8925 | 111.8774 | 117.8513 | 111.8362
2416.1719 | 122.0901 | -116.3589 | 5.6168 | -0.13817 | 0.011595 | 0.15122 | 0.090068 | -0.0015231 | 0.015975 | 0.013501 | -0.16306 | 0.27831 | 0.16297 | 0.0023093 | 0.070741 | 0.085819 | -0.0023349 | -0.032043 | 5.6122 | -116.5192 | 5.3879 | -4.7104 | -113.9877 | -113.8916 | -0.004578 | 5.603 | -116.3498 | -1.5198 | 5.5161 | -4.5319 | 1.5152 | -0.16022 | -0.004577 | -0.22888 | -0.22431 | 111.6486 | 111.8088 | 2.4673 | 2.3712 | 2.6276 | 2.5314 | 0.009156 | -0.013733 | -7.1365 | -7.1228 | 121.875 | 121.8658 | 117.8741 | 111.8271 | 117.865 | 111.8179
2417.1719 | 122.2 | -116.2766 | 5.603 | -0.10765 | 0.013091 | 0.15164 | 0.086436 | 0.0010403 | 0.016188 | 0.013715 | -0.077606 | 0.27852 | 0.16276 | -0.0077311 | 0.071168 | 0.082828 | -0.00083948 | 0.022888 | 5.5573 | -116.4505 | 5.5023 | -4.5319 | -113.8687 | -113.7497 | 0.13275 | 5.5664 | -116.2903 | -1.3824 | 5.6671 | -4.4083 | 1.6067 | -0.17395 | -0.045776 | -0.10071 | -0.054932 | 111.7447 | 111.9186 | 2.5269 | 2.4078 | 2.7008 | 2.5818 | -0.013733 | -0.036621 | -6.9855 | -6.9489 | 121.9437 | 121.9574 | 117.8833 | 111.8683 | 117.897 | 111.882
2418.1719 | 122.2 | -116.2582 | 5.5984 | -0.11528 | 0.012023 | 0.1525 | 0.085795 | -0.0034457 | 0.01512 | 0.015851 | -0.071503 | 0.27788 | 0.16148 | -0.0034585 | 0.071168 | 0.085605 | 0.0027921 | 0.009155 | 5.5893 | -116.4688 | 5.5206 | -4.5227 | -113.8779 | -113.7268 | 0.12817 | 5.571 | -116.2674 | -1.3824 | 5.6763 | -4.3854 | 1.6479 | -0.21057 | -0.009156 | -0.07782 | -0.068664 | 111.7355 | 111.9461 | 2.5314 | 2.3804 | 2.742 | 2.5909 | -0.009155 | -0.027466 | -6.9809 | -6.9534 | 121.9345 | 121.9437 | 117.9062 | 111.8729 | 117.9153 | 111.882
2419.1719 | 122.0901 | -116.2628 | 5.6763 | -0.046616 | 0.010527 | 0.1525 | 0.087291 | 0.00061311 | 0.012557 | 0.017346 | -0.16 | 0.27895 | 0.1634 | 0.0023093 | 0.069032 | 0.087314 | -0.0019077 | -0.027466 | 5.6305 | -116.4871 | 5.4382 | -4.7058 | -113.9694 | -113.8687 | -0.009155 | 5.6671 | -116.2674 | -1.5106 | 5.5893 | -4.5227 | 1.5976 | -0.22431 | -0.045777 | -0.23804 | -0.19226 | 111.557 | 111.7813 | 2.3941 | 2.2934 | 2.6184 | 2.5177 | -0.004578 | -0.009156 | -7.1869 | -7.1777 | 121.8521 | 121.8567 | 117.8604 | 111.7401 | 117.865 | 111.7447
2420.1719 | 122.113 | -116.217 | 5.6717 | -0.077133 | 0.012236 | 0.15335 | 0.084727 | 0.00061311 | 0.015761 | 0.015424 | -0.077606 | 0.27874 | 0.16319 | -0.0030313 | 0.072877 | 0.087528 | -0.00083948 | 0.059509 | 5.5801 | -116.4368 | 5.5664 | -4.5685 | -113.8596 | -113.7222 | 0.12817 | 5.6259 | -116.2445 | -1.4145 | 5.6808 | -4.422 | 1.6342 | -0.21973 | -0.091553 | -0.10529 | -0.013733 | 111.6486 | 111.8683 | 2.4948 | 2.3575 | 2.7145 | 2.5772 | -0.027466 | -0.045776 | -7.0862 | -7.0404 | 121.8979 | 121.9254 | 117.8513 | 111.795 | 117.8787 | 111.8225
2421.1719 | 122.1451 | -116.2079 | 5.6992 | -0.11528 | 0.011809 | 0.1525 | 0.083659 | 0.00082671 | 0.015334 | 0.016705 | -0.10507 | 0.28087 | 0.16447 | -0.0051675 | 0.070314 | 0.08411 | -0.0014805 | -0.027466 | 5.6396 | -116.3635 | 5.5664 | -4.6051 | -113.8412 | -113.7314 | 0.054932 | 5.6625 | -116.2125 | -1.4008 | 5.7129 | -4.4128 | 1.6846 | -0.15564 | -0.05951 | -0.13275 | -0.073242 | 111.6028 | 111.7584 | 2.4765 | 2.3666 | 2.6321 | 2.5223 | -0.004577 | -0.036621 | -7.0999 | -7.0633 | 121.9208 | 121.9254 | 117.8925 | 111.795 | 117.897 | 111.7996
2422.1719 | 122.1268 | -116.2079 | 5.722 | -0.1458 | 0.012877 | 0.15058 | 0.086436 | -0.0051547 | 0.012984 | 0.017987 | -0.11728 | 0.27938 | 0.16511 | 0.0018821 | 0.072237 | 0.085391 | 0.0010831 | 0.059509 | 5.6671 | -116.3818 | 5.603 | -4.5685 | -112.9715 | -113.7131 | 0.10986 | 5.6946 | -116.2125 | -1.4648 | 5.6992 | -4.4632 | 1.6571 | -0.17395 | -0.054932 | -0.11902 | -0.064087 | 111.6394 | 111.8134 | 2.4948 | 3.2364 | 2.6688 | 3.4103 | -0.004577 | -0.027466 | -7.1869 | -7.1594 | 121.907 | 121.9116 | 117.865 | 111.7447 | 117.8696 | 111.7493
2423.1719 | 122.2092 | -116.185 | 5.6854 | -0.10002 | 0.0086048 | 0.15058 | 0.0843 | -0.0047275 | 0.015975 | 0.015424 | -0.092865 | 0.28023 | 0.16425 | -0.0060221 | 0.069246 | 0.082401 | 0.0021511 | 0.10986 | 5.6625 | -116.3589 | 5.6168 | -4.5319 | -113.53 | -113.6856 | 0.14191 | 5.6717 | -116.1804 | -1.355 | 5.5527 | -4.4174 | 1.7349 | -0.17395 | -0.022888 | -0.068665 | -0.045777 | 111.6531 | 111.8271 | 2.4994 | 2.655 | 2.6733 | 2.829 | 0.004578 | -0.013733 | -7.0404 | -7.0267 | 121.7377 | 121.7331 | 117.9199 | 111.7676 | 117.9153 | 111.763
2424.1719 | 122.1222 | -116.1896 | 5.7083 | -0.084763 | 0.012664 | 0.15399 | 0.08665 | 0.0018949 | 0.015975 | 0.014996 | -0.12643 | 0.27724 | 0.1602 | 0.0023093 | 0.069887 | 0.083683 | 0.0017239 | -0.050354 | 5.6396 | -116.3773 | 5.6213 | -4.5639 | -113.9694 | -113.6993 | 0.13275 | 5.6854 | -116.2033 | -1.4328 | 5.7266 | -4.4632 | 1.68 | -0.18768 | -0.068665 | -0.086975 | -0.01831 | 111.6257 | 111.8134 | 2.4902 | 2.2202 | 2.6779 | 2.4078 | -0.013733 | -0.022888 | -7.1411 | -7.1182 | 121.9162 | 121.9299 | 117.8696 | 111.7264 | 117.8833 | 111.7401
2425.1719 | 122.0444 | -116.2262 | 5.7312 | -0.077133 | 0.011809 | 0.15207 | 0.088145 | -0.00045509 | 0.013411 | 0.015637 | -0.16611 | 0.27788 | 0.16468 | -0.0032449 | 0.068819 | 0.086246 | 0.0023647 | -0.018311 | 5.6946 | -116.4322 | 5.484 | -4.7607 | -113.8779 | -113.768 | 0.013733 | 5.7175 | -116.2308 | -1.5427 | 5.6259 | -4.4998 | 1.593 | -0.20599 | -0.036621 | -0.24719 | -0.21057 | 111.4655 | 111.6714 | 2.4582 | 2.3483 | 2.6642 | 2.5543 | -0.004578 | -0.013733 | -7.2739 | -7.2601 | 121.8521 | 121.8567 | 117.8192 | 111.7264 | 117.8238 | 111.731
2426.1719 | 122.1405 | -116.1621 | 5.7083 | -0.13054 | 0.010955 | 0.15464 | 0.085368 | 0.0027493 | 0.013198 | 0.016492 | -0.12338 | 0.27852 | 0.16468 | -0.0026041 | 0.07245 | 0.084964 | -0.0010531 | -0.004578 | 5.6763 | -116.3589 | 5.5664 | -4.6509 | -113.8184 | -113.7543 | 0.041199 | 5.6854 | -116.2125 | -1.4282 | 5.7632 | -4.4083 | 1.6937 | -0.19684 | -0.032043 | -0.14191 | -0.10986 | 111.5112 | 111.7081 | 2.4078 | 2.3438 | 2.6047 | 2.5406 | -0.050354 | -0.022888 | -7.1365 | -7.1136 | 121.9254 | 121.9757 | 117.8558 | 111.7538 | 117.9062 | 111.8042
2427.1719 | 122.0901 | -116.1713 | 5.7541 | -0.092392 | 0.010955 | 0.15335 | 0.088786 | 0.0023221 | 0.015975 | 0.018201 | -0.23935 | 0.27895 | 0.1619 | 0.0006003 | 0.0701 | 0.087955 | 0.0019375 | 0.013733 | 5.69 | -116.3956 | 5.5527 | -4.6921 | -113.8229 | -114.0656 | 0.20142 | 5.7312 | -116.0751 | -1.4374 | 5.6717 | -4.5731 | 1.6388 | -0.2243 | -0.064087 | -0.20142 | -0.13733 | 111.4792 | 111.7035 | 2.1057 | 2.3483 | 2.33 | 2.5726 | 0.096131 | -0.022888 | -7.1915 | -7.1686 | 121.843 | 121.7468 | 117.8101 | 111.5982 | 117.7139 | 111.5021
2428.1719 | 122.1313 | -116.1987 | 5.7037 | -0.061874 | 0.011168 | 0.15036 | 0.085368 | -0.00066869 | 0.015975 | 0.013715 | -0.098968 | 0.27852 | 0.16105 | -0.0070901 | 0.070314 | 0.084964 | -0.0036167 | 0.027466 | 5.6488 | -116.3315 | 5.6351 | -4.5731 | -113.7268 | -113.6993 | 0.13733 | 5.6671 | -116.185 | -1.4191 | 5.7632 | -4.4357 | 1.6937 | -0.13275 | -0.054931 | -0.068664 | -0.013733 | 111.6257 | 111.7584 | 2.4994 | 2.4719 | 2.6321 | 2.6047 | 0.013732 | -0.036621 | -7.1228 | -7.0862 | 121.962 | 121.9482 | 117.8925 | 111.763 | 117.8787 | 111.7493
2429.1719 | 122.1451 | -118.4875 | 3.4241 | -0.10765 | 0.0096728 | 0.15207 | 0.084513 | 0.00039951 | 0.017256 | 0.016705 | -0.12033 | 0.28258 | 0.16425 | -0.0019631 | 0.071168 | 0.085605 | -0.0016941 | -0.050354 | 3.3691 | -118.689 | 3.1311 | -4.866 | -116.217 | -116.1575 | 0 | 3.3051 | -118.5379 | -1.474 | 3.3829 | -4.5685 | -0.27466 | -0.20142 | -0.054931 | -0.29297 | -0.23804 | 113.6215 | 113.8229 | 2.33 | 2.2705 | 2.5314 | 2.4719 | -0.050354 | -0.11902 | -4.8981 | -4.7791 | 121.8704 | 121.9208 | 118.2129 | 113.9191 | 118.2632 | 113.9694
2430.1719 | 122.1725 | -118.5242 | 3.3737 | -0.12291 | 0.01245 | 0.15549 | 0.085582 | -0.0032321 | 0.016188 | 0.01521 | -0.24545 | 0.28173 | 0.1634 | -0.0032449 | 0.07245 | 0.088382 | -0.0042575 | 0.091553 | 3.3005 | -118.7073 | 3.2547 | -4.6875 | -115.7593 | -116.3406 | 0.21973 | 3.3279 | -118.5425 | -1.4557 | 3.392 | -4.5273 | -0.29297 | -0.1831 | -0.073242 | -0.11902 | -0.045776 | 113.8367 | 114.0198 | 2.1835 | 2.7649 | 2.3666 | 2.948 | -0.01831 | -0.045776 | -4.8294 | -4.7836 | 121.9162 | 121.9345 | 118.2312 | 113.9969 | 118.2495 | 114.0152
2431.1719 | 122.1817 | -118.5699 | 3.3371 | -0.11528 | 0.011382 | 0.15229 | 0.084513 | 0.0014675 | 0.015548 | 0.017773 | -0.10202 | 0.28515 | 0.16554 | 0.0016685 | 0.070955 | 0.087741 | -0.0021213 | 0 | 3.3051 | -118.721 | 3.2272 | -4.66 | -116.1392 | -116.0202 | 0.10986 | 3.3188 | -118.5196 | -1.4191 | 3.3875 | -4.4769 | -0.27008 | -0.15106 | -0.032043 | -0.10986 | -0.07782 | 113.9099 | 114.061 | 2.5497 | 2.4307 | 2.7008 | 2.5818 | 0.050354 | -0.01831 | -4.7562 | -4.7379 | 121.9574 | 121.907 | 118.2999 | 114.093 | 118.2495 | 114.0427
2432.1719 | 122.1725 | -118.5654 | 3.3188 | -0.12291 | 0.013945 | 0.15378 | 0.088359 | -0.0030185 | 0.016616 | 0.015637 | -0.12033 | 0.28216 | 0.16383 | 0.0008139 | 0.070741 | 0.086032 | -0.00019868 | -0.013733 | 3.3096 | -118.7393 | 3.1952 | -4.6646 | -116.1392 | -116.1026 | 0.10986 | 3.3234 | -118.5608 | -1.4328 | 3.3737 | -4.4907 | -0.28839 | -0.17395 | -0.009156 | -0.1236 | -0.11444 | 113.9008 | 114.0747 | 2.4628 | 2.4261 | 2.6367 | 2.6001 | 0.004578 | 0.004577 | -4.7516 | -4.7562 | 121.9391 | 121.9345 | 118.277 | 114.0747 | 118.2724 | 114.0701
2433.1719 | 121.9757 | -118.5745 | 3.4012 | -0.084763 | 0.010741 | 0.15229 | 0.085795 | -0.0045139 | 0.012984 | 0.016919 | -0.17526 | 0.28472 | 0.16468 | 0.0050863 | 0.069246 | 0.086246 | -0.0019077 | -0.045776 | 3.3279 | -118.7851 | 3.0762 | -4.8203 | -116.272 | -116.2262 | 0.009155 | 3.3234 | -118.6295 | -1.5335 | 3.4882 | -4.4952 | -0.36163 | -0.21057 | -0.073242 | -0.32501 | -0.25177 | 113.7543 | 113.9648 | 2.3483 | 2.3026 | 2.5589 | 2.5131 | -0.054932 | -0.07782 | -4.9347 | -4.8569 | 122.0627 | 122.1176 | 118.2129 | 114.0793 | 118.2678 | 114.1342
2434.1719 | 122.1085 | -118.5242 | 3.392 | -0.13817 | 0.012023 | 0.1557 | 0.084513 | 0.00039951 | 0.016616 | 0.014996 | -0.15695 | 0.2828 | 0.1634 | 0.0057273 | 0.071382 | 0.086887 | 0.00022852 | -0.07782 | 3.3371 | -118.6981 | 3.2135 | -4.6829 | -116.1392 | -116.0568 | -0.036621 | 3.3646 | -118.5242 | -1.4923 | 3.3508 | -4.5227 | -0.29755 | -0.17395 | -0.054932 | -0.17853 | -0.1236 | 113.8412 | 114.0152 | 2.4673 | 2.3849 | 2.6413 | 2.5589 | 0 | -0.027466 | -4.8843 | -4.8569 | 121.875 | 121.875 | 118.2266 | 114.0015 | 118.2266 | 114.0015
2435.1719 | 122.113 | -118.5196 | 3.4058 | -0.069504 | 0.013091 | 0.15357 | 0.085368 | -0.00066869 | 0.016188 | 0.016278 | -0.14474 | 0.28451 | 0.16575 | -0.0026041 | 0.070314 | 0.085178 | -0.0014805 | 0.036621 | 3.3188 | -118.6844 | 3.2593 | -4.6371 | -116.0934 | -116.0568 | 0.10071 | 3.3646 | -118.5104 | -1.4877 | 3.3737 | -4.5227 | -0.28839 | -0.16479 | -0.086975 | -0.14649 | -0.05951 | 113.8824 | 114.0472 | 2.4628 | 2.4261 | 2.6276 | 2.5909 | 0.009155 | -0.041199 | -4.8935 | -4.8523 | 121.8933 | 121.8842 | 118.2312 | 113.9969 | 118.222 | 113.9877
2436.1719 | 122.0764 | -118.5791 | 3.3463 | -0.069504 | 0.010741 | 0.15314 | 0.087504 | 0.0025357 | 0.013411 | 0.014783 | -0.10202 | 0.28386 | 0.16212 | -0.0004679 | 0.071809 | 0.086673 | 0.0017239 | 0.004578 | 3.3096 | -118.7119 | 3.273 | -4.6234 | -116.1072 | -116.0294 | 0.13275 | 3.3371 | -118.5333 | -1.4282 | 3.3096 | -4.3854 | -0.14648 | -0.13275 | -0.036621 | -0.073242 | -0.036621 | 113.9557 | 114.0884 | 2.5497 | 2.4719 | 2.6825 | 2.6047 | 0.045777 | -0.009155 | -4.7745 | -4.7653 | 121.8887 | 121.843 | 118.4326 | 114.1937 | 118.3868 | 114.1479
2437.1719 | 122.1588 | -118.4921 | 3.3829 | -0.084763 | 0.0066822 | 0.15399 | 0.088359 | -0.0028049 | 0.014479 | 0.017133 | -0.12338 | 0.28408 | 0.16383 | -0.0019631 | 0.070955 | 0.087314 | 1.4916e-05 | 0.013733 | 3.3279 | -118.6844 | 3.2684 | -4.541 | -116.066 | -116.011 | 0.16022 | 3.3508 | -118.5059 | -1.4374 | 3.4287 | -4.66 | -0.25635 | -0.19226 | -0.054932 | -0.11444 | -0.059509 | 113.9511 | 114.1434 | 2.4811 | 2.4261 | 2.6733 | 2.6184 | -0.013733 | -0.032044 | -4.8203 | -4.7882 | 121.9208 | 121.9345 | 118.2358 | 113.8321 | 118.2495 | 113.8458
2438.1719 | 122.0947 | -118.5196 | 3.4103 | -0.10765 | 0.011168 | 0.15357 | 0.087931 | -0.0010959 | 0.014479 | 0.018201 | -0.15695 | 0.28451 | 0.16618 | -0.0002543 | 0.071168 | 0.087741 | -0.00041228 | 0.059509 | 3.3371 | -118.7119 | 3.2455 | -4.6188 | -116.0934 | -116.1713 | 0.086975 | 3.3371 | -118.5196 | -1.4786 | 3.3508 | -4.6371 | -0.32501 | -0.19226 | -0.073242 | -0.16479 | -0.091553 | 113.9008 | 114.093 | 2.3483 | 2.4261 | 2.5406 | 2.6184 | 0 | -0.073242 | -4.8889 | -4.8157 | 121.8704 | 121.8704 | 118.1946 | 113.8824 | 118.1946 | 113.8824
2439.1719 | 122.1405 | -118.5608 | 3.3829 | -0.1458 | 0.0098864 | 0.15506 | 0.085368 | -0.0034457 | 0.013839 | 0.016492 | -0.13559 | 0.28557 | 0.16319 | -0.0045267 | 0.071809 | 0.087314 | 1.4916e-05 | -0.07782 | 3.3829 | -118.6798 | 3.241 | -4.5181 | -116.153 | -116.0614 | -0.036621 | 3.4332 | -118.4692 | -1.506 | 3.392 | -4.66 | -0.27008 | -0.11902 | 0 | -0.14191 | -0.14191 | 114.0427 | 114.1617 | 2.4994 | 2.4078 | 2.6184 | 2.5269 | 0.091553 | 0.050354 | -4.8889 | -4.9393 | 121.9528 | 121.8613 | 118.2907 | 113.9008 | 118.1992 | 113.8092
2440.1875 | 122.0856 | -118.5013 | 3.4058 | -0.023727 | 0.011168 | 0.15442 | 0.085154 | -0.00066869 | 0.01277 | 0.017773 | -0.22104 | 0.28515 | 0.16297 | 0.0012411 | 0.069246 | 0.086887 | 0.0019375 | -0.027466 | 3.3875 | -118.721 | 3.1631 | -4.7928 | -116.1804 | -116.0797 | -0.027466 | 3.4012 | -118.5333 | -1.5244 | 3.3051 | -4.6646 | -0.33417 | -0.21973 | -0.018311 | -0.24262 | -0.2243 | 113.7085 | 113.9282 | 2.4216 | 2.3209 | 2.6413 | 2.5406 | -0.032043 | -0.004578 | -4.9301 | -4.9255 | 121.8063 | 121.8384 | 118.1671 | 113.8367 | 118.1992 | 113.8687
2441.1719 | 122.0764 | -118.5104 | 3.3875 | -0.084763 | 0.011809 | 0.15271 | 0.086009 | -0.0019503 | 0.014052 | 0.015851 | -0.044037 | 0.28472 | 0.16148 | -0.0034585 | 0.071809 | 0.088169 | 0.0021511 | -0.10071 | 3.3371 | -118.7027 | 3.2227 | -4.7195 | -116.1255 | -116.0568 | 0 | 3.3188 | -118.5974 | -1.4786 | 3.3783 | -4.6005 | -0.27466 | -0.19226 | -0.050354 | -0.16479 | -0.11444 | 113.7909 | 113.9832 | 2.4536 | 2.3849 | 2.6459 | 2.5772 | -0.086975 | -0.068664 | -4.866 | -4.7974 | 121.8887 | 121.9757 | 118.2358 | 113.9099 | 118.3228 | 113.9969
2442.1719 | 122.1497 | -118.5059 | 3.3737 | -0.18394 | 0.01245 | 0.15634 | 0.085795 | -0.00045509 | 0.015548 | 0.017773 | -0.18137 | 0.28408 | 0.16255 | -0.0053811 | 0.070528 | 0.087314 | 0.0012967 | 0.022888 | 3.2822 | -118.6981 | 3.2867 | -4.6417 | -116.1392 | -116.1118 | 0.018311 | 3.3691 | -118.4875 | -1.4099 | 3.4012 | -4.5227 | -0.26093 | -0.19226 | -0.091552 | -0.086975 | 0.004577 | 113.8641 | 114.0564 | 2.3941 | 2.3666 | 2.5864 | 2.5589 | 0.01831 | -0.004577 | -4.7836 | -4.7791 | 121.907 | 121.8887 | 118.2449 | 113.9832 | 118.2266 | 113.9648
2443.1719 | 122.0535 | -118.5104 | 3.4332 | -0.1458 | 0.0098864 | 0.15357 | 0.083018 | -0.0077183 | 0.013411 | 0.014783 | -0.14169 | 0.28301 | 0.16383 | -0.0006815 | 0.069246 | 0.086887 | -0.0063937 | 0.027466 | 3.3646 | -118.6844 | 3.2318 | -4.7012 | -116.1072 | -116.1026 | 0.064087 | 3.4012 | -118.5013 | -1.5381 | 3.3142 | -4.5868 | -0.32959 | -0.17395 | -0.068665 | -0.20142 | -0.13275 | 113.8092 | 113.9832 | 2.4078 | 2.4033 | 2.5818 | 2.5772 | 0.009155 | -0.032044 | -4.9713 | -4.9393 | 121.8246 | 121.8155 | 118.1808 | 113.9236 | 118.1717 | 113.9145
2444.1719 | 122.1725 | -118.4784 | 3.4103 | -0.084763 | 0.010741 | 0.15293 | 0.084941 | -2.7692e-05 | 0.014052 | 0.017133 | -0.11728 | 0.28365 | 0.16468 | 0.0046591 | 0.070955 | 0.08646 | -0.0023349 | -0.041199 | 3.3508 | -118.689 | 3.2364 | -4.6555 | -116.0797 | -116.0614 | 0.050354 | 3.392 | -118.4921 | -1.4282 | 3.4195 | -4.4724 | -0.25177 | -0.21057 | -0.059509 | -0.17395 | -0.11444 | 113.8229 | 114.0335 | 2.417 | 2.3987 | 2.6276 | 2.6093 | -0.013732 | -0.01831 | -4.8386 | -4.8203 | 121.8979 | 121.9116 | 118.2266 | 114.006 | 118.2404 | 114.0198
2445.1875 | 122.0856 | -118.5242 | 3.4195 | -0.084763 | 0.011168 | 0.15293 | 0.085795 | -0.0032321 | 0.012984 | 0.015424 | -0.20273 | 0.28536 | 0.16233 | 0.0031637 | 0.069459 | 0.087528 | 0.00044212 | -0.013733 | 3.3783 | -118.7347 | 3.1769 | -4.66 | -116.1255 | -116.2033 | 0 | 3.3646 | -118.515 | -1.5106 | 3.3234 | -4.5959 | -0.31586 | -0.21057 | -0.041199 | -0.24261 | -0.20142 | 113.8641 | 114.0747 | 2.3209 | 2.3987 | 2.5314 | 2.6093 | 0.009155 | -0.054932 | -4.9301 | -4.8752 | 121.8475 | 121.8384 | 118.2083 | 113.9282 | 118.1992 | 113.9191
2446.1719 | 122.113 | -118.4967 | 3.392 | -0.17632 | 0.012664 | 0.15506 | 0.086436 | -0.00088229 | 0.018325 | 0.016278 | -0.15085 | 0.28408 | 0.16169 | -0.0034585 | 0.071168 | 0.087314 | -0.0016941 | -0.050354 | 3.3096 | -118.6798 | 3.2593 | -4.541 | -116.1392 | -116.0934 | -0.013733 | 3.36 | -118.5242 | -1.4557 | 3.392 | -4.5136 | -0.2655 | -0.18311 | -0.082398 | -0.13275 | -0.050354 | 113.9557 | 114.1388 | 2.4033 | 2.3575 | 2.5864 | 2.5406 | -0.027466 | -0.032044 | -4.8477 | -4.8157 | 121.8887 | 121.9162 | 118.2312 | 113.9832 | 118.2587 | 114.0106
2447.1719 | 122.1313 | -118.4875 | 3.4103 | -0.11528 | 0.01245 | 0.15399 | 0.087718 | -0.00066869 | 0.016829 | 0.017133 | -0.11728 | 0.28386 | 0.16148 | -0.0051675 | 0.069459 | 0.0871 | -0.00019868 | -0.086975 | 3.3371 | -118.6707 | 3.2364 | -4.6326 | -116.1118 | -116.0843 | -0.045776 | 3.3737 | -118.483 | -1.474 | 3.3737 | -4.5136 | -0.27924 | -0.1831 | -0.073242 | -0.17395 | -0.10071 | 113.855 | 114.0381 | 2.4033 | 2.3758 | 2.5864 | 2.5589 | 0.004578 | -0.036621 | -4.8843 | -4.8477 | 121.8613 | 121.8567 | 118.2083 | 113.974 | 118.2037 | 113.9694
2448.1719 | 122.049 | -118.5425 | 3.4378 | -0.077133 | 0.010741 | 0.15314 | 0.08665 | -0.0015231 | 0.015334 | 0.013928 | -0.19357 | 0.28579 | 0.1634 | 0.0023093 | 0.07245 | 0.088596 | -0.0010531 | -0.036621 | 3.3783 | -118.7393 | 3.1723 | -4.6967 | -116.1346 | -116.1896 | 0.004578 | 3.4058 | -118.4647 | -1.5656 | 3.36 | -4.6143 | -0.27924 | -0.19684 | -0.059509 | -0.2655 | -0.20599 | 113.8458 | 114.0427 | 2.3529 | 2.4078 | 2.5497 | 2.6047 | 0.077819 | -0.032043 | -5.0034 | -4.9713 | 121.9025 | 121.8246 | 118.2632 | 113.9282 | 118.1854 | 113.8504
2449.1719 | 122.1222 | -118.5013 | 3.4058 | -0.15343 | 0.012877 | 0.15442 | 0.086436 | -0.00088229 | 0.015548 | 0.014996 | -0.12338 | 0.28429 | 0.16447 | 0.0044455 | 0.072237 | 0.0871 | 1.4916e-05 | -0.073242 | 3.3508 | -118.6798 | 3.2043 | -4.66 | -115.9927 | -116.1026 | -0.032043 | 3.3875 | -118.4967 | -1.4832 | 3.3737 | -4.5181 | -0.27924 | -0.17853 | -0.054932 | -0.20142 | -0.14648 | 113.8412 | 114.0198 | 2.3987 | 2.5085 | 2.5772 | 2.6871 | 0.004578 | -0.018311 | -4.8889 | -4.8706 | 121.875 | 121.8704 | 118.222 | 113.9832 | 118.2175 | 113.9786
2450.1719 | 122.1313 | -118.5104 | 3.3966 | -0.16869 | 0.012023 | 0.15421 | 0.085582 | 0.00061311 | 0.016402 | 0.014783 | -0.16306 | 0.286 | 0.16255 | -0.0073037 | 0.071596 | 0.088382 | 0.0034329 | -0.07782 | 3.3371 | -118.6935 | 3.273 | -4.5502 | -116.0843 | -116.1255 | 0.009155 | 3.3783 | -118.4875 | -1.3916 | 3.3875 | -4.4678 | -0.16937 | -0.1831 | -0.059509 | -0.1236 | -0.064087 | 113.9603 | 114.1434 | 2.3849 | 2.4261 | 2.5681 | 2.6093 | 0.022888 | -0.01831 | -4.7882 | -4.7699 | 121.8979 | 121.875 | 118.3411 | 114.0427 | 118.3182 | 114.0198
2451.1719 | 122.0947 | -118.4921 | 3.3783 | -0.16869 | 0.01245 | 0.15335 | 0.084941 | -0.0013095 | 0.014479 | 0.016919 | -0.12949 | 0.28386 | 0.16383 | 0.0008139 | 0.071809 | 0.086246 | 0.0036465 | -0.073242 | 3.36 | -118.689 | 3.2547 | -4.5364 | -116.1026 | -116.0889 | -0.032043 | 3.3646 | -118.4921 | -1.4603 | 3.36 | -4.5364 | -0.28381 | -0.19684 | -0.018311 | -0.1236 | -0.10528 | 113.9557 | 114.1525 | 2.4033 | 2.3895 | 2.6001 | 2.5864 | 0 | -0.013733 | -4.8386 | -4.8248 | 121.8521 | 121.8521 | 118.2083 | 113.9557 | 118.2083 | 113.9557
2452.1719 | 122.1405 | -118.5104 | 3.3646 | -0.15343 | 0.010955 | 0.15464 | 0.088786 | 0.0025357 | 0.014479 | 0.018842 | -0.17221 | 0.28408 | 0.16169 | -0.0017495 | 0.070955 | 0.087741 | 0.0019375 | -0.032043 | 3.3188 | -118.7576 | 3.2227 | -4.6555 | -116.0889 | -116.1255 | 0.022888 | 3.3463 | -118.483 | -1.3962 | 3.392 | -4.509 | -0.26093 | -0.24719 | -0.045776 | -0.14191 | -0.096131 | 113.855 | 114.1022 | 2.3849 | 2.4216 | 2.6321 | 2.6688 | 0.027466 | -0.018311 | -4.7607 | -4.7424 | 121.9025 | 121.875 | 118.2495 | 114.0015 | 118.222 | 113.974
2453.1719 | 122.0398 | -118.689 | 3.2822 | -0.031357 | 0.0096728 | 0.15421 | 0.088786 | -0.0028049 | 0.014266 | 0.019055 | -0.16916 | 0.28643 | 0.16425 | 0.0012411 | 0.072877 | 0.085819 | 0.00022852 | -0.027466 | 3.2272 | -118.84 | 3.0121 | -4.6692 | -116.2262 | -116.3132 | 0.004578 | 3.2684 | -118.6798 | -1.5106 | 3.1631 | -4.5914 | -0.4715 | -0.15106 | -0.054932 | -0.27008 | -0.21515 | 114.0198 | 114.1708 | 2.3758 | 2.4628 | 2.5269 | 2.6138 | 0.009155 | -0.013733 | -4.7928 | -4.7791 | 121.8521 | 121.843 | 118.2175 | 114.0976 | 118.2083 | 114.0884
2454.1719 | 122.113 | -118.6295 | 3.1998 | -0.19157 | 0.014373 | 0.1572 | 0.08665 | 0.00082671 | 0.016402 | 0.016492 | -0.19357 | 0.28386 | 0.16554 | -0.0047403 | 0.069032 | 0.091159 | -0.0023349 | -0.041199 | 3.1952 | -118.8126 | 3.1311 | -4.6005 | -116.2582 | -116.2308 | 0.009155 | 3.2135 | -118.6249 | -1.506 | 3.2959 | -4.3854 | -0.3067 | -0.1831 | -0.004578 | -0.068664 | -0.064086 | 114.0289 | 114.212 | 2.3987 | 2.3712 | 2.5818 | 2.5543 | 0.004578 | 0.013733 | -4.7058 | -4.7195 | 121.9254 | 121.9208 | 118.3228 | 114.2441 | 118.3182 | 114.2395
2455.1719 | 122.0398 | -118.6661 | 3.2867 | -0.061874 | 0.012236 | 0.15528 | 0.084086 | 0.0010403 | 0.014693 | 0.014355 | -0.12338 | 0.28451 | 0.1619 | -0.0030313 | 0.068391 | 0.088169 | -0.0021213 | 0.018311 | 3.1998 | -118.8263 | 3.067 | -4.7058 | -116.2766 | -116.2582 | 0.041199 | 3.2318 | -118.6386 | -1.474 | 3.1815 | -4.5731 | -0.44403 | -0.16022 | -0.086975 | -0.21973 | -0.13275 | 113.9603 | 114.1205 | 2.4078 | 2.3895 | 2.5681 | 2.5497 | 0.027466 | -0.054931 | -4.7607 | -4.7058 | 121.8475 | 121.8201 | 118.222 | 114.093 | 118.1946 | 114.0656
2456.1719 | 122.1268 | -118.6249 | 3.2135 | -0.12291 | 0.0092456 | 0.15399 | 0.08665 | 0.0018949 | 0.01747 | 0.016278 | -0.16 | 0.28429 | 0.16233 | 0.0023093 | 0.069887 | 0.087314 | 0.0017239 | -0.036621 | 3.1631 | -118.8217 | 3.067 | -4.4312 | -116.2445 | -116.3086 | 0.027466 | 3.1403 | -118.6569 | -1.3458 | 3.2364 | -4.4724 | -0.37079 | -0.19684 | -0.050354 | -0.14648 | -0.09613 | 114.1937 | 114.3906 | 2.3163 | 2.3804 | 2.5131 | 2.5772 | -0.032043 | -0.073242 | -4.5593 | -4.4861 | 121.8613 | 121.8933 | 118.2541 | 114.1525 | 118.2861 | 114.1846
2457.1719 | 122.1771 | -117.0731 | 4.66 | -0.16869 | 0.011595 | 0.15592 | 0.083873 | 0.00018591 | 0.016616 | 0.016919 | -0.1539 | 0.27617 | 0.1602 | -0.0053811 | 0.072877 | 0.086887 | 0.0012967 | -0.054932 | 4.6097 | -117.3615 | 4.5456 | -4.2892 | -114.8575 | -114.7659 | 0.009155 | 4.6234 | -117.1646 | -1.387 | 4.6875 | -3.6072 | 0.90179 | -0.28839 | -0.050354 | -0.11444 | -0.064087 | 112.7838 | 113.0722 | 2.3071 | 2.2156 | 2.5955 | 2.504 | -0.091553 | -0.036621 | -6.0471 | -6.0104 | 121.7606 | 121.8521 | 117.9749 | 113.4659 | 118.0664 | 113.5574
2458.1875 | 122.1817 | -117.2104 | 4.5685 | -0.084763 | 0.012023 | 0.15378 | 0.087077 | -0.0017367 | 0.011061 | 0.016492 | -0.13254 | 0.27746 | 0.15656 | 0.0010275 | 0.071596 | 0.089237 | -0.0016941 | 0.004578 | 4.5044 | -117.4484 | 4.335 | -4.0329 | -114.9353 | -115.0909 | 0.16937 | 4.4769 | -117.3157 | -1.4694 | 4.6417 | -3.4515 | 0.82855 | -0.23804 | -0.064086 | -0.23346 | -0.16937 | 113.1775 | 113.4155 | 2.1194 | 2.2751 | 2.3575 | 2.5131 | -0.10529 | -0.091552 | -6.0379 | -5.9464 | 121.8521 | 121.9574 | 118.0389 | 113.7588 | 118.1442 | 113.8641
2459.1719 | 122.1588 | -117.54 | 4.3076 | -0.32127 | 0.01245 | 0.15399 | 0.090281 | -0.0010959 | 0.012984 | 0.017773 | -0.16611 | 0.27874 | 0.15806 | 0.0025229 | 0.071596 | 0.087741 | -0.0019077 | 0.013733 | 4.184 | -117.7689 | 4.1702 | -4.1977 | -115.1642 | -115.2191 | 0.009155 | 4.2709 | -117.54 | -1.4694 | 4.3259 | -3.4561 | 0.56305 | -0.22888 | -0.1236 | -0.13733 | -0.013733 | 113.3423 | 113.5712 | 2.3209 | 2.3758 | 2.5497 | 2.6047 | 0 | -0.036621 | -5.777 | -5.7404 | 121.8658 | 121.8658 | 118.103 | 114.0839 | 118.103 | 114.0839
2460.1719 | 122.0764 | -117.7597 | 4.1382 | -0.13054 | 0.013091 | 0.15464 | 0.087718 | -0.0032321 | 0.014052 | 0.013715 | -0.1539 | 0.27895 | 0.16084 | 0.0016685 | 0.068819 | 0.087955 | 0.00044212 | 0.027466 | 4.0787 | -117.9245 | 3.9871 | -4.248 | -115.3519 | -115.2695 | 0.07782 | 4.1061 | -117.7322 | -1.5839 | 4.0924 | -3.511 | 0.36621 | -0.16479 | -0.05951 | -0.15106 | -0.091552 | 113.5117 | 113.6765 | 2.4902 | 2.4078 | 2.655 | 2.5726 | 0.027466 | -0.032044 | -5.722 | -5.69 | 121.8521 | 121.8246 | 118.1259 | 114.2487 | 118.0985 | 114.2212
2461.1719 | 122.1085 | -117.8604 | 4.01 | -0.11528 | 0.01245 | 0.15464 | 0.087504 | -2.7692e-05 | 0.014693 | 0.014996 | -0.15695 | 0.27916 | 0.15913 | 0.0046591 | 0.072877 | 0.087955 | -0.0023349 | -0.050354 | 3.9459 | -118.0481 | 3.8589 | -4.1977 | -115.5304 | -115.242 | 0.013733 | 3.9597 | -117.8558 | -1.5152 | 4.0009 | -3.4515 | 0.29755 | -0.18768 | -0.064087 | -0.15106 | -0.086975 | 113.6627 | 113.8504 | 2.6184 | 2.33 | 2.8061 | 2.5177 | 0.004578 | -0.050354 | -5.5252 | -5.4749 | 121.8613 | 121.8567 | 118.158 | 114.4089 | 118.1534 | 114.4043
2462.1875 | 122.0764 | -117.9428 | 3.9871 | -0.077133 | 0.011595 | 0.15485 | 0.08665 | -0.0017367 | 0.014479 | 0.01521 | -0.21188 | 0.27981 | 0.15913 | 0.0065817 | 0.069673 | 0.088169 | 0.0019375 | 0.018311 | 3.9322 | -118.1488 | 3.7445 | -4.2297 | -115.6082 | -115.6631 | 0.013733 | 4.0054 | -117.9153 | -1.5976 | 3.9001 | -3.5385 | 0.23346 | -0.20599 | -0.054932 | -0.24261 | -0.18768 | 113.7131 | 113.9191 | 2.2797 | 2.3346 | 2.4857 | 2.5406 | 0.027466 | 0.01831 | -5.5847 | -5.603 | 121.843 | 121.8155 | 118.1763 | 114.4043 | 118.1488 | 114.3768
2463.1719 | 122.113 | -117.9245 | 3.9322 | -0.17632 | 0.011168 | 0.15335 | 0.087077 | -0.00024129 | 0.013625 | 0.017773 | -0.13254 | 0.28002 | 0.16084 | -0.0002543 | 0.074586 | 0.0871 | -0.0023349 | -0.054932 | 3.8864 | -118.0801 | 3.8132 | -4.1153 | -115.4572 | -115.5396 | -0.027466 | 3.923 | -117.9016 | -1.5656 | 3.9139 | -3.479 | 0.26093 | -0.15564 | -0.045776 | -0.11902 | -0.073243 | 113.8092 | 113.9648 | 2.3849 | 2.4673 | 2.5406 | 2.623 | 0.022889 | -0.009155 | -5.4977 | -5.4886 | 121.8384 | 121.8155 | 118.1854 | 114.4455 | 118.1625 | 114.4226
2464.1719 | 122.1542 | -117.952 | 3.9276 | -0.12291 | 0.011382 | 0.15506 | 0.089213 | -0.00088229 | 0.016188 | 0.017133 | -0.20273 | 0.28087 | 0.15934 | 0.0050863 | 0.069032 | 0.088809 | -0.0029757 | -0.091553 | 3.8956 | -118.1396 | 3.772 | -4.129 | -115.4938 | -115.5075 | -0.022888 | 3.9001 | -117.9337 | -1.5244 | 3.9413 | -3.4698 | 0.26093 | -0.18768 | -0.032043 | -0.15564 | -0.1236 | 113.8229 | 114.0106 | 2.4445 | 2.4582 | 2.6321 | 2.6459 | 0.01831 | -0.027466 | -5.452 | -5.4245 | 121.8933 | 121.875 | 118.2129 | 114.4821 | 118.1946 | 114.4638
2465.1875 | 122.1359 | -118.1488 | 3.7216 | -0.069504 | 0.011595 | 0.1525 | 0.083873 | -0.0021641 | 0.014907 | 0.014783 | -0.095917 | 0.27938 | 0.15891 | 0.0006003 | 0.071382 | 0.087314 | 1.4916e-05 | 0.045776 | 3.6621 | -118.3136 | 3.6713 | -4.1794 | -115.9287 | -115.6357 | 0.073242 | 3.6621 | -118.1625 | -1.506 | 3.7674 | -3.4241 | 0.11444 | -0.16479 | -0.05951 | -0.050354 | 0.009156 | 113.9694 | 114.1342 | 2.5131 | 2.2202 | 2.6779 | 2.3849 | -0.013733 | -0.05951 | -5.2277 | -5.1682 | 121.9162 | 121.9299 | 118.2632 | 114.7247 | 118.277 | 114.7385
2466.1719 | 122.1176 | -118.1808 | 3.6987 | -0.18394 | 0.012877 | 0.15699 | 0.086436 | -0.0021641 | 0.016188 | 0.01521 | -0.19052 | 0.28066 | 0.16126 | -0.0015359 | 0.070528 | 0.086673 | 0.00065592 | -0.073242 | 3.5889 | -118.3868 | 3.5889 | -4.2389 | -115.7455 | -115.7639 | 0.022888 | 3.6713 | -118.1717 | -1.506 | 3.7308 | -3.4424 | 0.10071 | -0.20599 | -0.10986 | -0.10986 | 0 | 113.942 | 114.1479 | 2.417 | 2.4353 | 2.623 | 2.6413 | 0.009155 | -0.027465 | -5.2048 | -5.1773 | 121.9116 | 121.9025 | 118.2816 | 114.7385 | 118.2724 | 114.7293
2467.1719 | 122.1451 | -118.1717 | 3.6942 | -0.12291 | 0.014586 | 0.15378 | 0.086222 | -0.00045509 | 0.011702 | 0.016278 | -0.092865 | 0.27895 | 0.16041 | 0.0020957 | 0.071596 | 0.087314 | 0.0025785 | 0.036621 | 3.6484 | -118.3319 | 3.6575 | -4.184 | -115.9012 | -115.7272 | 0.07782 | 3.6758 | -118.1534 | -1.4603 | 3.7445 | -3.4058 | 0.091553 | -0.16022 | -0.045777 | -0.036621 | 0.009156 | 113.9877 | 114.1479 | 2.4445 | 2.2705 | 2.6047 | 2.4307 | 0.018311 | -0.018311 | -5.1544 | -5.1361 | 121.9162 | 121.8979 | 118.2632 | 114.7659 | 118.2449 | 114.7476
2468.1719 | 122.1405 | -118.2861 | 3.5889 | -0.10765 | 0.0094592 | 0.15656 | 0.088145 | -0.0010959 | 0.012984 | 0.017346 | -0.20883 | 0.28194 | 0.15977 | 0.0029501 | 0.072023 | 0.089023 | -0.00062588 | -0.086975 | 3.566 | -118.4921 | 3.4149 | -4.2252 | -115.6723 | -115.947 | -0.045776 | 3.5797 | -118.2861 | -1.4603 | 3.5843 | -3.4744 | -0.032043 | -0.20599 | -0.022888 | -0.17395 | -0.15106 | 114.061 | 114.267 | 2.3392 | 2.6138 | 2.5452 | 2.8198 | 0 | -0.009155 | -5.0491 | -5.04 | 121.8704 | 121.8704 | 118.2541 | 114.8117 | 118.2541 | 114.8117
2469.1719 | 122.0993 | -118.3365 | 3.5797 | -0.1458 | 0.010314 | 0.15464 | 0.089 | -0.00024129 | 0.015761 | 0.016705 | -0.22104 | 0.28109 | 0.15721 | -0.0023905 | 0.070528 | 0.088596 | 0.00086952 | -0.091553 | 3.5522 | -118.5333 | 3.3554 | -4.2801 | -115.9332 | -116.0248 | -0.086975 | 3.5751 | -118.3228 | -1.506 | 3.5156 | -3.5019 | -0.036621 | -0.19684 | -0.027466 | -0.2243 | -0.19684 | 114.0564 | 114.2532 | 2.3117 | 2.4033 | 2.5085 | 2.6001 | 0.013733 | -0.004578 | -5.0858 | -5.0812 | 121.8521 | 121.8384 | 118.2999 | 114.8346 | 118.2861 | 114.8209
2470.1719 | 122.1085 | -118.364 | 3.5248 | -0.13817 | 0.012664 | 0.1572 | 0.087077 | 0.0016813 | 0.017897 | 0.016705 | -0.22104 | 0.28023 | 0.15699 | -0.0028177 | 0.069246 | 0.088596 | 0.00044212 | -0.045776 | 3.4561 | -118.4967 | 3.4012 | -4.1519 | -115.9287 | -115.947 | -0.09613 | 3.4561 | -118.4097 | -1.4511 | 3.5339 | -3.4332 | -0.013733 | -0.13275 | -0.068664 | -0.1236 | -0.054932 | 114.212 | 114.3448 | 2.417 | 2.4353 | 2.5497 | 2.5681 | -0.045776 | -0.068664 | -4.9759 | -4.9072 | 121.8979 | 121.9437 | 118.3502 | 114.9307 | 118.396 | 114.9765
2471.1719 | 122.0352 | -118.3411 | 3.5294 | -0.13817 | 0.014159 | 0.15677 | 0.089213 | 0.00018591 | 0.015761 | 0.013928 | -0.21188 | 0.28151 | 0.15913 | 0.0003867 | 0.070741 | 0.087314 | 0.0017239 | -0.24261 | 3.5248 | -118.5699 | 3.3005 | -4.2618 | -115.9882 | -116.0568 | -0.11444 | 3.5385 | -118.3868 | -1.5518 | 3.4195 | -3.5568 | -0.12817 | -0.22888 | -0.004578 | -0.22888 | -0.2243 | 114.0793 | 114.3082 | 2.2842 | 2.3529 | 2.5131 | 2.5818 | -0.045777 | 0.009155 | -5.0812 | -5.0903 | 121.7606 | 121.8063 | 118.2129 | 114.7842 | 118.2587 | 114.83
2472.1719 | 121.994 | -118.1168 | 3.7903 | -0.12291 | 0.0088184 | 0.1557 | 0.089 | 0.0029629 | 0.014266 | 0.019482 | -0.34921 | 0.27938 | 0.15721 | 0.0070089 | 0.069459 | 0.08945 | -0.00083948 | -0.09613 | 3.7537 | -118.3228 | 3.511 | -4.2389 | -115.8783 | -115.8646 | -0.013733 | 3.7857 | -118.1396 | -1.5793 | 3.6667 | -3.5477 | 0.064087 | -0.20599 | -0.036621 | -0.27924 | -0.24261 | 113.8779 | 114.0839 | 2.2522 | 2.2385 | 2.4582 | 2.4445 | -0.022888 | -0.004577 | -5.3696 | -5.365 | 121.7834 | 121.8063 | 118.1808 | 114.5691 | 118.2037 | 114.592
2473.1719 | 122.1542 | -118.1259 | 3.7033 | -0.1458 | 0.012664 | 0.15293 | 0.085582 | -0.0028049 | 0.016402 | 0.015637 | -0.15695 | 0.27938 | 0.15656 | 0.0012411 | 0.071382 | 0.087314 | 0.0017239 | -0.004578 | 3.6942 | -118.2907 | 3.6438 | -4.0421 | -115.6952 | -115.7364 | 0.004578 | 3.7125 | -118.1442 | -1.442 | 3.7857 | -3.4653 | 0.11902 | -0.16479 | -0.009155 | -0.059509 | -0.050354 | 114.0839 | 114.2487 | 2.3895 | 2.4307 | 2.5543 | 2.5955 | -0.01831 | 0.009155 | -5.1453 | -5.1544 | 121.9116 | 121.9299 | 118.2449 | 114.6606 | 118.2632 | 114.679
2474.1719 | 122.0764 | -118.2495 | 3.6484 | -0.11528 | 0.012664 | 0.1572 | 0.086863 | -0.0030185 | 0.014693 | 0.013715 | -0.19968 | 0.28173 | 0.16084 | 0.0038047 | 0.069246 | 0.089023 | 0.00065592 | -0.11444 | 3.6209 | -118.4601 | 3.4195 | -4.1748 | -115.8966 | -115.9607 | -0.10071 | 3.6392 | -118.2724 | -1.5335 | 3.5843 | -3.4744 | 0.013733 | -0.21057 | -0.027465 | -0.22888 | -0.20142 | 114.0747 | 114.2853 | 2.2888 | 2.3529 | 2.4994 | 2.5635 | -0.022888 | -0.009155 | -5.1819 | -5.1727 | 121.8338 | 121.8567 | 118.2632 | 114.7751 | 118.2861 | 114.798
2475.1719 | 122.1085 | -118.2999 | 3.5568 | -0.18394 | 0.012236 | 0.15421 | 0.087077 | -0.0019503 | 0.013411 | 0.015424 | -0.12949 | 0.27981 | 0.15785 | -0.0011087 | 0.071809 | 0.087955 | 0.00044212 | -0.068665 | 3.5339 | -118.4418 | 3.4424 | -4.0329 | -115.8783 | -115.892 | -0.004578 | 3.5522 | -118.3136 | -1.4969 | 3.5797 | -3.4332 | 0.004578 | -0.14191 | -0.022888 | -0.11444 | -0.091553 | 114.267 | 114.4089 | 2.4078 | 2.4216 | 2.5497 | 2.5635 | -0.013733 | -0.004578 | -5.0537 | -5.0491 | 121.8796 | 121.8933 | 118.3044 | 114.8666 | 118.3182 | 114.8804
2476.1719 | 121.994 | -118.4006 | 3.5339 | -0.13054 | 0.010527 | 0.15549 | 0.089213 | 0.0053129 | 0.014907 | 0.015851 | -0.22104 | 0.28045 | 0.15934 | 0.0057273 | 0.070955 | 0.088596 | -0.0019077 | -0.045776 | 3.4973 | -118.6066 | 3.2959 | -4.1611 | -116.0614 | -116.0797 | -0.050354 | 3.5339 | -118.4006 | -1.6159 | 3.4012 | -3.5706 | -0.16022 | -0.20599 | -0.036622 | -0.23804 | -0.20142 | 114.2395 | 114.4455 | 2.3209 | 2.3392 | 2.5269 | 2.5452 | 0 | 0 | -5.1498 | -5.1498 | 121.8018 | 121.8018 | 118.2404 | 114.83 | 118.2404 | 114.83
2477.1719 | 122.0123 | -118.4189 | 3.5294 | -0.16106 | 0.011382 | 0.15442 | 0.088786 | 0.0023221 | 0.013411 | 0.016492 | -0.20883 | 0.28109 | 0.15977 | 0.0010275 | 0.070314 | 0.089023 | 0.00022852 | -0.10986 | 3.4882 | -118.6066 | 3.2501 | -4.2435 | -116.1484 | -116.1484 | -0.009155 | 3.5019 | -118.3731 | -1.6617 | 3.4149 | -3.6392 | -0.14191 | -0.18768 | -0.041199 | -0.27924 | -0.23804 | 114.1754 | 114.3631 | 2.2705 | 2.2705 | 2.4582 | 2.4582 | 0.045776 | -0.027466 | -5.191 | -5.1636 | 121.8338 | 121.788 | 118.277 | 114.7797 | 118.2312 | 114.7339
2478.1875 | 122.0673 | -118.4052 | 3.479 | -0.19157 | 0.013945 | 0.15378 | 0.083445 | -0.0025913 | 0.016188 | 0.012646 | -0.095917 | 0.27895 | 0.15763 | 0.0010275 | 0.072877 | 0.086887 | 0.0021511 | -0.059509 | 3.4424 | -118.5928 | 3.3005 | -4.0558 | -116.0477 | -116.0751 | -0.045776 | 3.4378 | -118.4052 | -1.5656 | 3.447 | -3.5339 | -0.1236 | -0.18768 | -0.036621 | -0.17853 | -0.14191 | 114.3494 | 114.537 | 2.33 | 2.3575 | 2.5177 | 2.5452 | 0 | -0.041199 | -5.0446 | -5.0034 | 121.8521 | 121.8521 | 118.2816 | 114.8712 | 118.2816 | 114.8712
2479.1719 | 122.049 | -118.3914 | 3.5065 | -0.031357 | 0.011168 | 0.15378 | 0.086863 | -0.0032321 | 0.013625 | 0.016278 | -0.21494 | 0.27981 | 0.15934 | 0.0023093 | 0.071596 | 0.089237 | 0.00086952 | -0.027466 | 3.5065 | -118.5608 | 3.3371 | -4.3442 | -116.0156 | -116.185 | -0.032043 | 3.5156 | -118.3777 | -1.6251 | 3.4287 | -3.5156 | -0.11444 | -0.16937 | 0 | -0.16937 | -0.16937 | 114.0472 | 114.2166 | 2.2064 | 2.3758 | 2.3758 | 2.5452 | 0.013732 | 0.009155 | -5.1315 | -5.1407 | 121.8201 | 121.8063 | 118.277 | 114.8758 | 118.2632 | 114.8621
2480.1875 | 122.081 | -118.3273 | 3.5522 | -0.19157 | 0.012023 | 0.15699 | 0.084727 | 0.0038175 | 0.017256 | 0.015851 | -0.12338 | 0.28045 | 0.1602 | -0.0028177 | 0.070741 | 0.088382 | 0.00086952 | -0.041199 | 3.4698 | -118.4921 | 3.4103 | -4.3213 | -115.9424 | -115.947 | 0.018311 | 3.5202 | -118.3273 | -1.5335 | 3.5522 | -3.3875 | -0.022888 | -0.16479 | -0.082397 | -0.14191 | -0.05951 | 114.006 | 114.1708 | 2.3804 | 2.3849 | 2.5452 | 2.5497 | 0 | -0.032043 | -5.0858 | -5.0537 | 121.8796 | 121.8796 | 118.3044 | 114.9399 | 118.3044 | 114.9399
2481.1719 | 122.0169 | -118.3777 | 3.5339 | -0.13054 | 0.013091 | 0.15314 | 0.088786 | 0.0025357 | 0.015334 | 0.014355 | -0.19052 | 0.27831 | 0.16041 | 0.0016685 | 0.071168 | 0.088809 | 0.00044212 | -0.050354 | 3.5431 | -118.5562 | 3.3646 | -6.8665 | -116.0248 | -116.098 | -0.013733 | 3.5431 | -118.3502 | -1.6479 | 3.4241 | -3.5065 | -0.091553 | -0.17853 | 0.009155 | -0.16937 | -0.17853 | 111.5112 | 111.6898 | 2.2797 | 2.3529 | 2.4582 | 2.5314 | 0.027466 | 0.009155 | -5.1819 | -5.191 | 121.8018 | 121.7743 | 118.2861 | 114.8712 | 118.2587 | 114.8438
2482.1719 | 122.0261 | -118.3868 | 3.5431 | -0.13054 | 0.010314 | 0.15549 | 0.086436 | 0.0046719 | 0.015548 | 0.016278 | -0.19662 | 0.27895 | 0.16062 | -0.0004679 | 0.071382 | 0.087528 | -0.00062588 | -0.054932 | 3.4836 | -118.5791 | 3.2822 | -6.6559 | -116.0614 | -116.1163 | -0.064087 | 3.5294 | -118.3777 | -1.6571 | 3.392 | -3.5202 | -0.13733 | -0.19226 | -0.059509 | -0.26092 | -0.20142 | 111.731 | 111.9232 | 2.2705 | 2.3254 | 2.4628 | 2.5177 | 0.009155 | -0.013733 | -5.2002 | -5.1865 | 121.7789 | 121.7697 | 118.2495 | 114.8666 | 118.2404 | 114.8575
2483.1719 | 122.0261 | -118.4601 | 3.4607 | -0.1458 | 0.012023 | 0.15271 | 0.088786 | 0.00082671 | 0.01277 | 0.014996 | -0.22409 | 0.27852 | 0.1587 | -0.0036721 | 0.071809 | 0.087528 | 0.00022852 | -0.091553 | 3.4195 | -118.6844 | 3.1586 | -6.5826 | -116.2033 | -116.3361 | -0.004578 | 3.4607 | -118.4692 | -1.68 | 3.3371 | -3.5019 | -0.19226 | -0.2243 | -0.041198 | -0.30212 | -0.26093 | 111.8774 | 112.1017 | 2.124 | 2.2568 | 2.3483 | 2.4811 | -0.009155 | 0 | -5.1407 | -5.1407 | 121.7972 | 121.8063 | 118.2678 | 114.9582 | 118.277 | 114.9673
2484.1719 | 122.0673 | -118.5562 | 3.3646 | -0.084763 | 0.013732 | 0.15229 | 0.085582 | -0.0010959 | 0.014907 | 0.014142 | -0.1539 | 0.27981 | 0.15827 | -0.0015359 | 0.069887 | 0.087955 | 0.00086952 | 0.022888 | 3.3279 | -118.7347 | 3.1586 | -6.6879 | -116.1392 | -116.3361 | 0.050354 | 3.3371 | -118.5699 | -1.6022 | 3.3051 | -3.447 | -0.23804 | -0.17853 | -0.036621 | -0.20599 | -0.16937 | 111.8683 | 112.0468 | 2.2202 | 2.417 | 2.3987 | 2.5955 | -0.013733 | -0.027466 | -4.9667 | -4.9393 | 121.8613 | 121.875 | 118.3182 | 115.1093 | 118.3319 | 115.123
2485.1875 | 122.1039 | -118.4052 | 3.4973 | -0.15343 | 0.011595 | 0.15271 | 0.08195 | -0.0021641 | 0.014693 | 0.014355 | -0.086761 | 0.27938 | 0.16084 | -0.0021769 | 0.070955 | 0.086887 | 0.00065592 | -0.054932 | 3.4195 | -118.5516 | 3.3325 | -6.9351 | -115.9927 | -116.1072 | 0 | 3.4515 | -118.4052 | -1.5839 | 3.4744 | -3.3737 | -0.1236 | -0.14648 | -0.077819 | -0.16479 | -0.086975 | 111.47 | 111.6165 | 2.298 | 2.4124 | 2.4445 | 2.5589 | 0 | -0.045776 | -5.0812 | -5.0354 | 121.8796 | 121.8796 | 118.2816 | 115.0314 | 118.2816 | 115.0314
2486.1719 | 122.0261 | -118.4647 | 3.4973 | 0.029678 | 0.011809 | 0.15442 | 0.088572 | 0.00018591 | 0.016402 | 0.013074 | -0.19357 | 0.27938 | 0.16062 | 0.0018821 | 0.069887 | 0.088809 | 0.0023647 | -0.13733 | 3.5248 | -118.5791 | 3.2639 | -6.8115 | -116.217 | -116.1987 | -0.009155 | 3.4515 | -118.428 | -1.6296 | 3.3737 | -3.4515 | -0.15564 | -0.11444 | 0.027466 | -0.23346 | -0.26092 | 111.6531 | 111.7676 | 2.2659 | 2.2476 | 2.3804 | 2.3621 | 0.036621 | -0.045776 | -5.127 | -5.0812 | 121.8384 | 121.8018 | 118.309 | 115.0131 | 118.2724 | 114.9765
2487.1719 | 122.0535 | -118.5059 | 3.4058 | -0.077133 | 0.010741 | 0.15207 | 0.087077 | -2.7692e-05 | 0.015975 | 0.015637 | -0.1478 | 0.28151 | 0.15891 | -0.0011087 | 0.069673 | 0.086246 | 0.0012967 | -0.054932 | 3.4195 | -118.689 | 3.1448 | -6.6513 | -116.2125 | -116.2766 | -0.013733 | 3.4058 | -118.5059 | -1.6754 | 3.3005 | -3.4332 | -0.24719 | -0.18311 | 0.013733 | -0.26093 | -0.27466 | 111.8546 | 112.0377 | 2.2293 | 2.2934 | 2.4124 | 2.4765 | 0 | 0 | -5.0812 | -5.0812 | 121.8063 | 121.8063 | 118.2587 | 115.0726 | 118.2587 | 115.0726
2488.1719 | 122.0123 | -118.5287 | 3.4012 | -0.13817 | 0.011168 | 0.15293 | 0.087077 | -0.00066869 | 0.013625 | 0.01286 | -0.20273 | 0.27724 | 0.16169 | -0.0026041 | 0.069246 | 0.087955 | 0.0032193 | -0.082397 | 3.3417 | -118.7256 | 3.1494 | -6.9443 | -116.2033 | -116.3818 | -0.082397 | 3.3691 | -118.5242 | -1.6617 | 3.273 | -3.4836 | -0.27008 | -0.19684 | -0.059509 | -0.25177 | -0.19226 | 111.5845 | 111.7813 | 2.1469 | 2.3254 | 2.3438 | 2.5223 | 0.004578 | -0.032043 | -5.0629 | -5.0308 | 121.8018 | 121.7972 | 118.2587 | 115.0452 | 118.2541 | 115.0406
2489.1719 | 122.0078 | -118.4967 | 3.4836 | -0.038986 | 0.012023 | 0.15293 | 0.083873 | -0.0043003 | 0.011275 | 0.012006 | -0.098968 | 0.28109 | 0.1587 | -0.0049539 | 0.072877 | 0.088809 | 0.00065592 | -0.045776 | 3.4103 | -118.6844 | 3.154 | -6.5506 | -116.153 | -116.2949 | -0.004578 | 3.4332 | -118.483 | -1.7166 | 3.3325 | -3.3737 | -0.21057 | -0.18768 | -0.073243 | -0.32959 | -0.25635 | 111.9461 | 112.1338 | 2.2018 | 2.3438 | 2.3895 | 2.5314 | 0.013733 | -0.050354 | -5.2002 | -5.1498 | 121.8292 | 121.8155 | 118.2861 | 115.123 | 118.2724 | 115.1093
2490.1875 | 122.1359 | -118.515 | 3.3646 | -0.16106 | 0.0098864 | 0.15442 | 0.085795 | -0.00045509 | 0.015548 | 0.014355 | -0.14474 | 0.2781 | 0.15998 | -0.0064493 | 0.070955 | 0.090091 | -0.0016941 | -0.036621 | 3.2867 | -118.6707 | 3.2364 | -6.3675 | -116.1713 | -116.5054 | 0.11444 | 3.2684 | -118.5654 | -1.5518 | 3.3554 | -3.2959 | -0.20142 | -0.15564 | -0.07782 | -0.12817 | -0.050354 | 112.1475 | 112.3032 | 2.0096 | 2.3438 | 2.1652 | 2.4994 | -0.050354 | -0.09613 | -4.9164 | -4.8203 | 121.8704 | 121.9208 | 118.3136 | 115.2191 | 118.364 | 115.2695
2491.1719 | 122.0215 | -118.5104 | 3.4378 | -0.031357 | 0.012023 | 0.15186 | 0.083659 | -0.00024129 | 0.017043 | 0.011792 | -0.12338 | 0.27746 | 0.15913 | 0.0046591 | 0.071596 | 0.086246 | -0.0016941 | -0.022888 | 3.392 | -118.6798 | 3.1403 | -6.3675 | -116.1758 | -116.2857 | 0.21057 | 3.4149 | -118.4326 | -1.6479 | 3.2639 | -3.4012 | -0.25177 | -0.16937 | -0.045776 | -0.29755 | -0.25177 | 112.1429 | 112.3123 | 2.2247 | 2.3346 | 2.3941 | 2.504 | 0.07782 | -0.022888 | -5.0858 | -5.0629 | 121.7743 | 121.6965 | 118.2587 | 115.1093 | 118.1808 | 115.0314
2492.1719 | 122.0398 | -118.4692 | 3.4653 | -0.19157 | 0.011595 | 0.15143 | 0.086222 | -0.00066869 | 0.014693 | 0.014996 | -0.16916 | 0.27981 | 0.1619 | -0.0038857 | 0.071596 | 0.088596 | 0.0027921 | -0.027466 | 3.4882 | -118.6386 | 3.186 | -6.4316 | -115.9561 | -116.2766 | -0.022888 | 3.4561 | -118.4555 | -1.593 | 3.3417 | -3.3737 | -0.21057 | -0.16937 | 0.022888 | -0.27924 | -0.30212 | 112.0377 | 112.207 | 2.1927 | 2.5131 | 2.3621 | 2.6825 | 0.013733 | -0.009155 | -5.0583 | -5.0491 | 121.8109 | 121.7972 | 118.2587 | 115.0955 | 118.2449 | 115.0818
2493.1719 | 122.0215 | -118.3182 | 3.6163 | -0.16869 | 0.011168 | 0.15421 | 0.087718 | 0.0010403 | 0.01277 | 0.012646 | -0.17831 | 0.27831 | 0.16169 | 0.0046591 | 0.070314 | 0.086673 | -0.0010531 | -0.022888 | 3.5751 | -118.4921 | 3.3737 | -6.3354 | -115.9607 | -116.0751 | -0.064087 | 3.6163 | -118.309 | -1.6022 | 3.4698 | -3.3966 | -0.091553 | -0.17395 | -0.041199 | -0.24261 | -0.20142 | 111.9827 | 112.1567 | 2.243 | 2.3575 | 2.417 | 2.5314 | 0.009155 | 0 | -5.2185 | -5.2185 | 121.788 | 121.7789 | 118.2266 | 114.9216 | 118.2175 | 114.9124
2494.1719 | 122.0444 | -118.3868 | 3.5385 | -0.038986 | 0.010527 | 0.15293 | 0.086222 | -0.0015231 | 0.015334 | 0.016919 | -0.13559 | 0.27959 | 0.15956 | 0.0059409 | 0.071596 | 0.087741 | -0.0012667 | -0.036621 | 3.5385 | -118.5654 | 3.2593 | -6.3217 | -115.892 | -116.185 | -0.013733 | 3.5019 | -118.4052 | -1.6113 | 3.4012 | -3.3691 | -0.16479 | -0.17853 | 0 | -0.27924 | -0.27924 | 112.0651 | 112.2437 | 2.2018 | 2.4948 | 2.3804 | 2.6733 | -0.01831 | -0.036621 | -5.1498 | -5.1132 | 121.788 | 121.8063 | 118.222 | 115.0177 | 118.2404 | 115.036
2495.1719 | 122.0718 | -118.4601 | 3.4332 | -0.084763 | 0.01245 | 0.15464 | 0.086436 | -0.0034457 | 0.014907 | 0.012006 | -0.16611 | 0.28045 | 0.16084 | 0.0048727 | 0.072664 | 0.087741 | 0.0042875 | -0.027466 | 3.4058 | -118.6935 | 3.1815 | -6.3629 | -116.0797 | -116.2537 | 0.027466 | 3.4103 | -118.5013 | -1.6251 | 3.3417 | -3.3279 | -0.20599 | -0.23346 | -0.027466 | -0.25177 | -0.2243 | 112.0972 | 112.3306 | 2.2064 | 2.3804 | 2.4399 | 2.6138 | -0.041199 | -0.022889 | -5.0583 | -5.0354 | 121.8018 | 121.843 | 118.2541 | 115.1321 | 118.2953 | 115.1733
2496.1719 | 122.0581 | -118.4235 | 3.4149 | -0.092392 | 0.0096728 | 0.15464 | 0.084086 | 0.0016813 | 0.014052 | 0.01286 | -0.047089 | 0.27831 | 0.16041 | -0.0006815 | 0.0701 | 0.087314 | -0.00019868 | 0.045776 | 3.3829 | -118.5974 | 3.3279 | -6.0974 | -116.0202 | -116.1026 | 0.16022 | 3.4378 | -118.4143 | -1.7441 | 3.5431 | -3.1723 | -0.09613 | -0.17395 | -0.032043 | -0.086975 | -0.054932 | 112.326 | 112.5 | 2.3209 | 2.4033 | 2.4948 | 2.5772 | 0.009155 | 0.022888 | -5.159 | -5.1819 | 121.9666 | 121.9574 | 118.3273 | 115.2512 | 118.3182 | 115.242
2497.1719 | 122.1313 | -118.4326 | 3.4149 | -0.061874 | 0.011595 | 0.15399 | 0.086222 | -0.00066869 | 0.01512 | 0.016705 | -0.065399 | 0.28087 | 0.15977 | 0.0023093 | 0.068819 | 0.087741 | 0.0019375 | 0.12817 | 3.3417 | -118.6707 | 3.3325 | -6.3629 | -116.2628 | -116.1209 | 0.1236 | 3.3737 | -118.4692 | -1.5335 | 3.4195 | -3.1906 | -0.13733 | -0.23804 | -0.073242 | -0.082397 | -0.009155 | 112.0697 | 112.3077 | 2.3117 | 2.1698 | 2.5497 | 2.4078 | -0.036621 | -0.041199 | -4.9484 | -4.9072 | 121.8521 | 121.8887 | 118.2953 | 115.242 | 118.3319 | 115.2786
2498.1875 | 122.0398 | -118.3868 | 3.5431 | -0.10002 | 0.012664 | 0.15207 | 0.090495 | 0.0050991 | 0.014907 | 0.017773 | -0.19662 | 0.28216 | 0.15785 | 0.0029501 | 0.070314 | 0.088809 | 0.00065592 | -0.16479 | 3.6713 | -118.5699 | 3.2822 | -6.3675 | -115.4892 | -116.1621 | -0.10986 | 3.5065 | -118.396 | -1.7807 | 3.3508 | -3.3691 | -0.13733 | -0.1831 | 0.12817 | -0.26092 | -0.3891 | 112.0193 | 112.2025 | 2.2247 | 2.8976 | 2.4078 | 3.0807 | -0.009155 | -0.036621 | -5.3238 | -5.2872 | 121.7377 | 121.7468 | 118.2495 | 115.0177 | 118.2587 | 115.0269
2499.1719 | 122.1497 | -118.0664 | 3.7857 | -0.038986 | 0.012236 | 0.1525 | 0.085582 | -0.00066869 | 0.013839 | 0.015424 | -0.12949 | 0.28151 | 0.16383 | -0.0002543 | 0.069459 | 0.084323 | -0.0025485 | -0.013733 | 3.4287 | -118.6523 | 3.1769 | -6.4728 | -116.0889 | -116.2125 | 0.013733 | 3.4515 | -118.4463 | -1.8082 | 3.2684 | -3.3051 | 0.39368 | -0.58594 | -0.35706 | -0.60883 | -0.25177 | 111.5936 | 112.1796 | 1.8539 | 1.9775 | 2.4399 | 2.5635 | -0.37994 | -0.33417 | -5.5939 | -5.2597 | 121.3348 | 121.7148 | 118.4601 | 114.7614 | 118.84 | 115.1413
2500.1719 | 122.081 | -118.309 | 3.5934 | -0.077133 | 0.013518 | 0.15164 | 0.087504 | -0.0010959 | 0.014479 | 0.01521 | -0.16306 | 0.27788 | 0.15956 | -0.0017495 | 0.069032 | 0.088809 | -0.0036167 | 0.018311 | 3.5797 | -118.5333 | 3.3646 | -6.3171 | -115.9424 | -116.0568 | 0.027466 | 3.5522 | -118.2953 | -2.4033 | 3.5294 | -3.2913 | -0.054932 | -0.2243 | -0.013733 | -0.22888 | -0.21515 | 111.9919 | 112.2162 | 2.2522 | 2.3666 | 2.4765 | 2.5909 | 0.013733 | -0.041199 | -5.9967 | -5.9555 | 121.8384 | 121.8246 | 118.2541 | 115.0177 | 118.2404 | 115.004
2501.1719 | 122.0993 | -118.2907 | 3.598 | -0.084763 | 0.013732 | 0.15335 | 0.085368 | -0.0013095 | 0.014052 | 0.013074 | -0.059296 | 0.27767 | 0.15742 | -0.0008951 | 0.070528 | 0.08646 | -0.0025485 | 0.09613 | 3.5294 | -118.4463 | 3.4698 | -6.1203 | -115.6998 | -116.0019 | 0.1236 | 3.5843 | -118.2953 | -2.417 | 3.5751 | -3.1815 | -0.018311 | -0.15564 | -0.068664 | -0.12817 | -0.059509 | 112.1704 | 112.326 | 2.2888 | 2.5909 | 2.4445 | 2.7466 | -0.004578 | -0.013732 | -6.015 | -6.0013 | 121.8658 | 121.8704 | 118.2724 | 115.1093 | 118.277 | 115.1138
2502.1719 | 121.9986 | -26.1658 | 95.5994 | -0.11528 | 0.016509 | 0.095248 | 0.086863 | 0.0010403 | 0.011275 | 0.008374 | 0.023102 | 0.02004 | 0.013864 | -0.0066629 | 0.013704 | -0.0062527 | -0.0042575 | 0.068665 | 95.4849 | -26.3672 | 95.4117 | -1.4923 | -24.5636 | -25.1404 | 0.068665 | 0.18768 | -0.17395 | -0.068665 | 0.13275 | -0.032043 | 0.022888 | -0.20142 | -0.11444 | -0.18768 | -0.073242 | 24.6735 | 24.8749 | 1.0254 | 1.6022 | 1.2268 | 1.8036 | 25.9918 | -95.4117 | -95.668 | -0.25635 | 26.2985 | 0.3067 | 26.1887 | 26.1337 | 0.19684 | 0.14191
2503.1719 | 122.1451 | -26.239 | 95.5261 | -0.16869 | 0.017791 | 0.096744 | 0.089 | -0.0013095 | 0.013625 | 0.0077332 | 0.056671 | 0.020467 | 0.014291 | -0.0055947 | 0.012208 | -0.0073209 | -0.0040439 | -0.013733 | 95.4803 | -26.4267 | 95.4208 | -1.5015 | -24.527 | -25.2274 | 0.045776 | 0.15106 | -0.16479 | 0.045776 | 0.19226 | 0.050354 | 0.11902 | -0.18768 | -0.045776 | -0.10529 | -0.05951 | 24.7375 | 24.9252 | 1.0117 | 1.712 | 1.1993 | 1.8997 | 26.0742 | -95.3751 | -95.4803 | -0.10529 | 26.4313 | 0.35706 | 26.358 | 26.2894 | 0.28381 | 0.21515
2504.1719 | 122.1497 | -26.3168 | 95.4712 | -0.13817 | 0.017791 | 0.09653 | 0.091563 | 0.00018591 | 0.0099934 | 0.0051696 | 0.038361 | 0.021749 | 0.014505 | -0.0038857 | 0.013704 | -0.0064663 | -0.0031895 | 0.19684 | 95.3979 | -26.564 | 95.3613 | -1.5244 | -25.0763 | -25.3326 | 0.022888 | 0.15106 | -0.15564 | 0.036621 | 0.16937 | 0.059509 | 0.11444 | -0.24719 | -0.073242 | -0.10986 | -0.036621 | 24.7925 | 25.0397 | 0.98419 | 1.2405 | 1.2314 | 1.4877 | 26.1612 | -95.3201 | -95.4346 | -0.11444 | 26.4862 | 0.32501 | 26.4313 | 26.3763 | 0.27008 | 0.21515
2505.1719 | 122.1497 | -26.3763 | 95.4666 | -0.21446 | 0.019072 | 0.097384 | 0.090281 | -0.0021641 | 0.011702 | 0.0090148 | 0.056671 | 0.02004 | 0.013223 | -0.0008951 | 0.013277 | -0.0060391 | -0.0061801 | -0.013733 | 95.3979 | -26.5457 | 95.3842 | -1.4694 | -24.7421 | -25.296 | 0.14648 | 0.14191 | -0.17395 | 0.041199 | 0.19226 | 0.054932 | 0.1236 | -0.16937 | -0.068665 | -0.082398 | -0.013733 | 24.9069 | 25.0763 | 1.0803 | 1.6342 | 1.2497 | 1.8036 | 26.2024 | -95.3247 | -95.4254 | -0.10071 | 26.5686 | 0.36621 | 26.4999 | 26.4313 | 0.29755 | 0.22888
2506.1719 | 122.168 | -26.4633 | 95.3934 | -0.13054 | 0.018004 | 0.09888 | 0.090281 | -0.00045509 | 0.011916 | 0.0045288 | 0.078034 | 0.02239 | 0.014505 | 0.0001731 | 0.011781 | -0.0049709 | -0.0036167 | 0.050354 | 95.3384 | -26.6006 | 95.4117 | -1.506 | -25.296 | -25.3784 | 0.082397 | 0.1236 | -0.16022 | 0.059509 | 0.17853 | 0.068665 | 0.1236 | -0.13733 | -0.054932 | 0.01831 | 0.073242 | 24.9573 | 25.0946 | 1.0849 | 1.1673 | 1.2222 | 1.3046 | 26.3031 | -95.2698 | -95.3339 | -0.064087 | 26.6418 | 0.33874 | 26.5869 | 26.532 | 0.28381 | 0.22888
2507.1719 | 122.1451 | -26.5045 | 95.3659 | -0.069504 | 0.017791 | 0.09653 | 0.09199 | -0.0015231 | 0.0091388 | 0.007306 | 0.029205 | 0.019826 | 0.012155 | 0.0033773 | 0.013277 | -0.0053983 | -0.0040439 | 0.018311 | 95.3247 | -26.6464 | 95.3018 | -1.5381 | -24.8886 | -25.2594 | 0.1236 | 0.16022 | -0.17395 | 0.018311 | 0.16022 | 0.041199 | 0.091553 | -0.14191 | -0.041199 | -0.064087 | -0.022888 | 24.9664 | 25.1083 | 1.2451 | 1.6159 | 1.387 | 1.7578 | 26.3306 | -95.2057 | -95.3476 | -0.14191 | 26.6647 | 0.33417 | 26.5961 | 26.5457 | 0.2655 | 0.21515
2508.1719 | 122.0581 | -26.619 | 95.3339 | -0.031357 | 0.018218 | 0.09653 | 0.094126 | -0.00045509 | 0.011275 | 0.0077332 | -0.028778 | 0.023031 | 0.012582 | -0.0083719 | 0.016908 | -0.0026211 | -0.0016941 | -0.1236 | 95.3201 | -26.7426 | 95.1233 | -1.6891 | -25.1083 | -25.6302 | 0.004578 | 0.19684 | -0.19684 | -0.082397 | 0.12817 | -0.073242 | 0.013733 | -0.1236 | -0.013733 | -0.21057 | -0.19684 | 24.9298 | 25.0534 | 0.98877 | 1.5106 | 1.1124 | 1.6342 | 26.4221 | -95.137 | -95.4163 | -0.27924 | 26.7471 | 0.32501 | 26.6327 | 26.5457 | 0.21057 | 0.1236
2509.1719 | 122.2046 | -26.5686 | 95.3064 | -0.1458 | 0.01715 | 0.098025 | 0.093058 | -0.0017367 | 0.012557 | 0.0077332 | 0.056671 | 0.021108 | 0.013437 | -0.0028177 | 0.014772 | -0.0049709 | -0.0046847 | 0.036621 | 95.2744 | -26.7609 | 95.2103 | -1.5656 | -24.8383 | -25.5203 | 0.068665 | 0.16022 | -0.14191 | 0.050354 | 0.20599 | 0.082397 | 0.15106 | -0.19226 | -0.032043 | -0.09613 | -0.064087 | 25.0031 | 25.1953 | 1.0483 | 1.7303 | 1.2405 | 1.9226 | 26.4267 | -95.1462 | -95.256 | -0.10986 | 26.7746 | 0.3479 | 26.7197 | 26.651 | 0.29297 | 0.2243
2510.1719 | 122.1039 | -26.7197 | 95.2286 | -0.10002 | 0.018218 | 0.096744 | 0.091777 | -0.0010959 | 0.011061 | 0.0064514 | -0.010467 | 0.020254 | 0.013864 | -0.0013223 | 0.016267 | -0.0062527 | -0.0019077 | 0.13275 | 95.1736 | -26.9073 | 95.0272 | -1.6891 | -25.3098 | -25.708 | 0.018311 | 0.18768 | -0.1831 | -0.059509 | 0.064087 | -0.064087 | 0.018311 | -0.18768 | -0.054932 | -0.20142 | -0.14648 | 25.0305 | 25.2182 | 1.0117 | 1.4099 | 1.1993 | 1.5976 | 26.5366 | -95.0409 | -95.2881 | -0.24719 | 26.7838 | 0.24719 | 26.738 | 26.6556 | 0.20142 | 0.11902
2511.1719 | 122.2 | -26.7288 | 95.1599 | -0.16106 | 0.019286 | 0.098239 | 0.089854 | -0.00066869 | 0.011275 | 0.0088012 | 0.013947 | 0.022603 | 0.014505 | -0.0068765 | 0.014345 | -0.0053983 | -0.0014805 | -0.064087 | 95.1965 | -26.8341 | 95.1096 | -1.6251 | -25.0122 | -25.5798 | -0.022888 | 0.15564 | -0.15106 | 0.004578 | 0.15106 | -0.041199 | -0.10529 | -0.10528 | 0.036621 | -0.050354 | -0.086975 | 25.1038 | 25.209 | 1.149 | 1.7166 | 1.2543 | 1.8219 | 26.5778 | -95.0043 | -95.1553 | -0.15106 | 26.8799 | 0.30212 | 26.6235 | 26.6876 | 0.045776 | 0.10986
2512.1719 | 122.1771 | -26.7563 | 95.1736 | 0.060196 | 0.013945 | 0.098666 | 0.092204 | 0.00039951 | 0.010634 | 0.0075196 | 0.010895 | 0.02004 | 0.011301 | -0.0051675 | 0.013063 | -0.0071071 | -0.0016941 | -0.054932 | 95.0592 | -26.9348 | 95.0272 | -1.6479 | -25.0671 | -25.6805 | 0.022888 | 0.15106 | -0.15564 | 0.032043 | 0.17395 | 0.045776 | 0.10986 | -0.17853 | -0.11444 | -0.14648 | -0.032043 | 25.1083 | 25.2869 | 1.0757 | 1.6891 | 1.2543 | 1.8677 | 26.6006 | -95.0226 | -95.1416 | -0.11902 | 26.9302 | 0.32959 | 26.8662 | 26.8021 | 0.2655 | 0.20142
2513.1875 | 122.1222 | -26.7929 | 95.1324 | -0.10765 | 0.018004 | 0.09888 | 0.096049 | 0.0016813 | 0.01213 | 0.0092286 | -0.053192 | 0.022603 | 0.013223 | 0.0061545 | 0.014558 | -0.0045437 | -0.0072483 | -0.004578 | 95.1233 | -26.9485 | 95.0043 | -1.6754 | -25.0488 | -25.8087 | 0.045776 | 0.15564 | -0.17395 | 0 | 0.11444 | -0.041199 | 0.054932 | -0.15564 | -0.009155 | -0.12817 | -0.11902 | 25.1175 | 25.2731 | 0.98419 | 1.7441 | 1.1398 | 1.8997 | 26.619 | -94.9768 | -95.1324 | -0.15564 | 26.9073 | 0.28839 | 26.8478 | 26.7517 | 0.22888 | 0.13275
2514.1719 | 122.2458 | -26.6556 | 95.2194 | -0.077133 | 0.019072 | 0.096957 | 0.090709 | -0.0010959 | 0.01213 | 0.0053832 | 0.013947 | 0.020254 | 0.012369 | -0.0053811 | 0.016267 | -0.0043301 | -0.0016941 | 0.032043 | 95.2286 | -26.7883 | 95.2011 | -1.5518 | -24.8474 | -25.6622 | 0.13733 | 0.15564 | -0.14191 | 0.086975 | 0.2243 | 0.07782 | 0.1236 | -0.13275 | 0.009156 | -0.01831 | -0.027466 | 25.1038 | 25.2365 | 0.99335 | 1.8082 | 1.1261 | 1.9409 | 26.5137 | -95.0638 | -95.1324 | -0.068665 | 26.8799 | 0.36621 | 26.7792 | 26.7334 | 0.2655 | 0.21973
2515.1719 | 122.113 | -26.651 | 95.2927 | -0.077133 | 0.018218 | 0.09888 | 0.094126 | -0.0010959 | 0.011061 | 0.0053832 | -0.056244 | 0.020681 | 0.011087 | -0.0006815 | 0.014131 | -0.0051845 | -0.0059665 | -0.032043 | 95.2927 | -26.7975 | 95.1416 | -1.8265 | -25.1862 | -25.6668 | -0.032043 | 0.2243 | -0.15106 | -0.059509 | 0.068665 | -0.068665 | 0.004578 | -0.14648 | 0 | -0.15106 | -0.15106 | 24.8245 | 24.971 | 0.98419 | 1.4648 | 1.1307 | 1.6113 | 26.4999 | -95.0684 | -95.3522 | -0.28381 | 26.7197 | 0.21973 | 26.6556 | 26.5823 | 0.15564 | 0.082397
2516.1719 | 122.2183 | -26.5915 | 95.2972 | -0.13817 | 0.018431 | 0.097384 | 0.090495 | -0.0021641 | 0.0099934 | 0.0051696 | 0.074982 | 0.020254 | 0.01066 | -0.0032449 | 0.015199 | -0.0060391 | -0.0027621 | -0.004578 | 95.3018 | -26.7288 | 95.2194 | -1.5747 | -24.9161 | -25.5386 | 0.09613 | 0.16479 | -0.13733 | 0.068665 | 0.2243 | 0.082397 | 0.13275 | -0.13733 | 0.004578 | -0.07782 | -0.082398 | 25.0168 | 25.1541 | 1.0529 | 1.6754 | 1.1902 | 1.8127 | 26.4542 | -95.1324 | -95.2286 | -0.09613 | 26.8158 | 0.36163 | 26.7242 | 26.6739 | 0.27008 | 0.21973
2517.1719 | 122.1863 | -26.6281 | 95.2881 | -0.13817 | 0.018431 | 0.098666 | 0.09199 | -0.00024129 | 0.014907 | 0.0079468 | 0.029205 | 0.021322 | 0.013864 | -0.0064493 | 0.015199 | -0.0043301 | -0.00041228 | -0.10071 | 95.2652 | -26.7654 | 95.1691 | -1.6479 | -25.0259 | -25.4974 | -0.032043 | 0.16022 | -0.16479 | 0 | 0.1236 | 0.018311 | 0.07782 | -0.13733 | -0.022888 | -0.11902 | -0.096131 | 24.9802 | 25.1175 | 1.1307 | 1.6022 | 1.268 | 1.7395 | 26.4633 | -95.1279 | -95.2881 | -0.16022 | 26.7517 | 0.28839 | 26.7059 | 26.6464 | 0.24261 | 0.18311
2518.1875 | 122.0901 | -26.6922 | 95.2789 | -0.15343 | 0.019286 | 0.096957 | 0.095835 | 0.00039951 | 0.011702 | 0.010937 | -0.098968 | 0.020681 | 0.012582 | -0.0017495 | 0.013277 | -0.0039029 | -0.0019077 | -0.068665 | 95.256 | -26.8524 | 95.105 | -1.7624 | -25.1038 | -25.7126 | -0.10071 | 0.16937 | -0.19684 | -0.073242 | 0.013733 | -0.13275 | -0.004578 | -0.16022 | -0.022889 | -0.17395 | -0.15106 | 24.9298 | 25.09 | 0.97961 | 1.5884 | 1.1398 | 1.7487 | 26.4954 | -95.1096 | -95.3522 | -0.24261 | 26.7059 | 0.21057 | 26.6876 | 26.5594 | 0.19226 | 0.064087
2519.1719 | 122.1771 | -26.6602 | 95.256 | -0.13817 | 0.018004 | 0.09653 | 0.091349 | -0.0015231 | 0.013839 | 0.0053832 | 0.016998 | 0.019613 | 0.011728 | -0.0019631 | 0.015199 | -0.0051845 | 0.0012967 | -0.068665 | 95.1782 | -26.8936 | 95.0638 | -1.6663 | -24.646 | -25.7126 | -0.022888 | 0.086975 | -0.15106 | 0.009155 | 0.15564 | 0.027466 | 0.10986 | -0.23346 | -0.077819 | -0.19226 | -0.11444 | 24.9939 | 25.2274 | 0.94757 | 2.0142 | 1.181 | 2.2476 | 26.5091 | -95.1691 | -95.2469 | -0.07782 | 26.8158 | 0.3067 | 26.77 | 26.6876 | 0.26093 | 0.17853
2520.1875 | 122.1771 | -26.6922 | 95.224 | -0.12291 | 0.015654 | 0.098453 | 0.091349 | -0.0010959 | 0.01277 | 0.0068786 | 0.029205 | 0.021108 | 0.014719 | 0.0003867 | 0.013917 | -0.0026211 | -0.0025485 | -0.086975 | 95.2194 | -26.8524 | 95.1187 | -1.6571 | -25.0717 | -25.676 | -0.041199 | 0.15564 | -0.14648 | -0.004578 | 0.14648 | 0.009155 | 0.09613 | -0.16022 | -0.004578 | -0.10529 | -0.10071 | 25.0351 | 25.1953 | 1.0162 | 1.6205 | 1.1765 | 1.7807 | 26.5457 | -95.0684 | -95.2286 | -0.16022 | 26.8387 | 0.29297 | 26.7883 | 26.7014 | 0.24261 | 0.15564
2521.1719 | 122.2092 | -26.6876 | 95.224 | -0.15343 | 0.017577 | 0.099734 | 0.091136 | 0.0014675 | 0.011489 | 0.0070922 | 0.029205 | 0.020467 | 0.011728 | -0.0043131 | 0.013704 | -0.0053983 | -0.0025485 | -0.018311 | 95.2011 | -26.9073 | 95.1782 | -1.6434 | -25.0626 | -26.3443 | 0.032043 | 0.16479 | -0.13733 | 0.036621 | 0.1831 | 0.068665 | 0.1236 | -0.21973 | -0.022888 | -0.045776 | -0.022888 | 25.0443 | 25.264 | 0.34332 | 1.6251 | 0.56305 | 1.8448 | 26.5503 | -95.0592 | -95.1874 | -0.12817 | 26.8707 | 0.32043 | 26.8112 | 26.7563 | 0.26093 | 0.20599
2522.1719 | 122.2504 | -26.5915 | 95.2881 | -0.15343 | 0.016082 | 0.098453 | 0.092631 | -0.0010959 | 0.011489 | 0.005597 | 0.035309 | 0.02004 | 0.014078 | -0.0008951 | 0.014131 | -0.0060391 | -0.0016941 | -0.041199 | 95.137 | -26.9211 | 94.9905 | -1.5976 | -25.209 | -23.8403 | 0.09613 | 0.15106 | -0.13275 | 0.059509 | 0.21973 | 0.091553 | 0.14648 | -0.32959 | -0.15106 | -0.29755 | -0.14648 | 24.9939 | 25.3235 | 2.7512 | 1.3824 | 3.0807 | 1.712 | 26.4587 | -95.137 | -95.2286 | -0.091553 | 26.8112 | 0.35248 | 26.738 | 26.683 | 0.27924 | 0.2243
2523.1875 | 122.0947 | -26.7471 | 95.256 | -0.1458 | 0.015868 | 0.09888 | 0.094554 | -0.00045509 | 0.011061 | 0.0094422 | -0.10812 | 0.022176 | 0.013223 | 0.0078635 | 0.015413 | -0.0062527 | -0.0044711 | -0.15564 | 95.1233 | -27.0172 | 94.8624 | -1.8036 | -25.1678 | -27.356 | -0.1236 | 0.17853 | -0.19684 | -0.054932 | 0.032043 | -0.10986 | 0.009155 | -0.27008 | -0.13275 | -0.39368 | -0.26093 | 24.9435 | 25.2136 | -0.60883 | 1.5793 | -0.33874 | 1.8494 | 26.5503 | -95.0775 | -95.311 | -0.23346 | 26.7792 | 0.22888 | 26.7563 | 26.6373 | 0.20599 | 0.086975
2524.1719 | 122.2046 | -26.7471 | 95.1828 | -0.16106 | 0.019927 | 0.098453 | 0.090495 | -0.0040867 | 0.010207 | 0.005597 | 0.026154 | 0.021535 | 0.012796 | -0.0028177 | 0.014558 | -0.0068935 | -0.00019868 | -0.14648 | 95.0775 | -26.9897 | 95.137 | -1.6754 | -25.09 | -26.4359 | -0.054932 | 0.14648 | -0.13733 | 0.009155 | 0.14648 | 0.018311 | 0.086975 | -0.24262 | -0.10528 | -0.045776 | 0.059509 | 25.0717 | 25.3143 | 0.31128 | 1.6571 | 0.55389 | 1.8997 | 26.6098 | -95.0363 | -95.1736 | -0.13733 | 26.8936 | 0.28381 | 26.8341 | 26.7654 | 0.2243 | 0.15564
2525.1719 | 122.1725 | -26.9989 | 94.9493 | -0.15343 | 0.017577 | 0.098025 | 0.093699 | -0.0023777 | 0.012343 | 0.0070922 | 0.016998 | 0.020467 | 0.010446 | -0.0049539 | 0.01349 | -0.0062527 | -0.00019868 | 0.022888 | 94.9448 | -27.1408 | 94.8715 | -1.6708 | -25.0214 | -26.12 | 0.082397 | 0.16937 | -0.18768 | -0.018311 | 0.10986 | 0 | 0.064087 | -0.14191 | -0.004578 | -0.07782 | -0.073242 | 25.3281 | 25.47 | 0.87891 | 1.9775 | 1.0208 | 2.1194 | 26.8112 | -94.78 | -94.9677 | -0.18768 | 27.1088 | 0.29755 | 27.063 | 26.9989 | 0.25177 | 0.18768
2526.1719 | 122.2321 | -27.0721 | 94.8532 | -0.16106 | 0.018859 | 0.098239 | 0.092418 | -0.0025913 | 0.01213 | 0.013501 | 0.010895 | 0.020681 | 0.012155 | -0.0019631 | 0.015199 | -0.0064663 | 0.00086952 | -0.045776 | 94.8212 | -27.2552 | 94.7388 | -1.6937 | -25.4425 | -26.0147 | 0.036621 | 0.17395 | -0.15564 | 0.045776 | 0.19684 | 0.068665 | 0.1236 | -0.1831 | -0.032043 | -0.11444 | -0.082397 | 25.3784 | 25.5615 | 1.0574 | 1.6296 | 1.2405 | 1.8127 | 26.9165 | -94.6793 | -94.8074 | -0.12817 | 27.269 | 0.35248 | 27.1957 | 27.1408 | 0.27924 | 0.22431
2527.1719 | 122.2504 | -27.1133 | 94.8257 | -0.16106 | 0.016509 | 0.098025 | 0.093699 | 0.00018591 | 0.01213 | 0.0077332 | 0.0078429 | 0.020894 | 0.011728 | 0.0042319 | 0.014131 | -0.0036893 | -0.0019077 | -0.041199 | 94.9539 | -27.1225 | 94.8624 | -1.5656 | -25.296 | -25.9918 | -0.018311 | 0.14191 | -0.14648 | 0 | 0.14191 | 0.013733 | -0.15564 | -0.009156 | 0.12817 | 0.036621 | -0.091552 | 25.5478 | 25.5569 | 1.1215 | 1.8173 | 1.1307 | 1.8265 | 26.9669 | -94.6838 | -94.8257 | -0.14191 | 27.2552 | 0.28839 | 26.9577 | 27.1271 | -0.009156 | 0.16022
2528.1719 | 122.1451 | -27.1729 | 94.8212 | -0.12291 | 0.018431 | 0.098453 | 0.093058 | -2.7692e-05 | 0.013625 | 0.0075196 | -0.03183 | 0.022176 | 0.010873 | -0.0040995 | 0.014772 | -0.0051845 | -0.0040439 | 0.009155 | 94.7983 | -27.3148 | 94.6609 | -1.7441 | -25.5661 | -26.1658 | 0.045776 | 0.19226 | -0.17853 | -0.059509 | 0.091553 | -0.045776 | 0.018311 | -0.14191 | -0.022888 | -0.16022 | -0.13733 | 25.4288 | 25.5707 | 1.0071 | 1.6068 | 1.149 | 1.7487 | 26.9943 | -94.6289 | -94.8807 | -0.25177 | 27.2644 | 0.27008 | 27.1912 | 27.1271 | 0.19684 | 0.13275
2529.1719 | 122.2366 | -27.0264 | 94.9127 | -0.13817 | 0.018645 | 0.10038 | 0.093058 | -0.0019503 | 0.014907 | 0.0070922 | 0.02005 | 0.019826 | 0.013437 | -0.0011087 | 0.016695 | -0.0034757 | 0.0010831 | -0.20142 | 94.8944 | -27.2507 | 94.7433 | -1.7853 | -25.3876 | -26.033 | 0.018311 | 0.13275 | -0.13275 | 0.004578 | 0.15564 | 0.032043 | 0.09613 | -0.2243 | -0.018311 | -0.16937 | -0.15106 | 25.2411 | 25.4654 | 0.99335 | 1.6388 | 1.2177 | 1.8631 | 26.8936 | -94.78 | -94.9081 | -0.12817 | 27.182 | 0.28839 | 27.1225 | 27.0584 | 0.22888 | 0.16479
2530.1719 | 122.2458 | -27.0218 | 94.9219 | -0.13817 | 0.016509 | 0.09653 | 0.09434 | -0.0017367 | 0.011489 | 0.0079468 | 0.029205 | 0.021749 | 0.012582 | -0.0036721 | 0.01349 | -0.0083889 | -0.0038303 | -0.018311 | 94.8715 | -27.1957 | 94.8029 | -1.5564 | -24.5132 | -26.033 | 0.064087 | 0.16479 | -0.14648 | 0.068665 | 0.20599 | 0.073242 | 0.14191 | -0.17395 | -0.050354 | -0.11902 | -0.068665 | 25.4654 | 25.6393 | 0.98877 | 2.5085 | 1.1627 | 2.6825 | 26.8753 | -94.7571 | -94.8532 | -0.09613 | 27.2278 | 0.35248 | 27.1637 | 27.095 | 0.28839 | 0.21973
2531.1719 | 122.2504 | -27.0813 | 94.8624 | -0.11528 | 0.01715 | 0.098666 | 0.093058 | -0.0032321 | 0.01277 | 0.0064514 | 0.02005 | 0.020467 | 0.014291 | -0.0068765 | 0.013917 | -0.0039029 | -0.0021213 | -0.027466 | 94.8532 | -27.2415 | 94.7525 | -1.7395 | -25.3418 | -26.0605 | 0.054932 | 0.16022 | -0.13275 | 0.064087 | 0.19684 | 0.068665 | 0.13275 | -0.16022 | -0.009156 | -0.10986 | -0.10071 | 25.3418 | 25.502 | 1.0208 | 1.7395 | 1.181 | 1.8997 | 26.9485 | -94.7021 | -94.7983 | -0.09613 | 27.2781 | 0.32959 | 27.2141 | 27.15 | 0.2655 | 0.20142
2532.1719 | 122.2137 | -27.0996 | 94.8761 | -0.10765 | 0.015654 | 0.097598 | 0.092418 | 0.0018949 | 0.0091388 | 0.0092286 | -0.0043641 | 0.019826 | 0.011942 | -0.0034585 | 0.015626 | -0.0056119 | -0.0021213 | -0.018311 | 94.9677 | -27.1454 | 94.812 | -1.7258 | -25.3326 | -26.0147 | -0.045776 | 0.15106 | -0.12817 | 0.004578 | 0.13275 | 0.009155 | 0.086975 | -0.045777 | 0.091552 | -0.064087 | -0.15564 | 25.3738 | 25.4196 | 1.0849 | 1.767 | 1.1307 | 1.8127 | 26.9714 | -94.725 | -94.8715 | -0.14648 | 27.2324 | 0.26093 | 27.1866 | 27.1088 | 0.21515 | 0.13733
2533.1719 | 122.2229 | -27.095 | 94.8898 | -0.12291 | 0.0195 | 0.098453 | 0.093913 | 0.00018591 | 0.012984 | 0.0070922 | -0.019623 | 0.021322 | 0.011514 | 0.0018821 | 0.012849 | -0.0032619 | -0.0036167 | 0.059509 | 94.8212 | -27.3102 | 94.8257 | -1.7303 | -26.5411 | -26.0468 | 0.091553 | 0.16022 | -0.16479 | 0.004578 | 0.13275 | 0.009155 | 0.082397 | -0.21515 | -0.068665 | -0.064087 | 0.004578 | 25.3647 | 25.5798 | 1.0483 | 0.55389 | 1.2634 | 0.76904 | 26.9302 | -94.7296 | -94.8853 | -0.15564 | 27.2278 | 0.29755 | 27.1774 | 27.1042 | 0.24719 | 0.17395
2534.1719 | 122.2916 | -27.1362 | 94.812 | -0.10765 | 0.017577 | 0.097384 | 0.093913 | -0.0034457 | 0.01277 | 0.005597 | 0.016998 | 0.020681 | 0.011728 | -0.0017495 | 0.014558 | -0.0043301 | -0.0066073 | -0.013733 | 94.812 | -27.2644 | 94.78 | -1.6251 | -25.3464 | -26.0147 | 0.1236 | 0.16022 | -0.13275 | 0.073242 | 0.21973 | 0.09613 | 0.14191 | -0.12817 | 0 | -0.032044 | -0.032044 | 25.5112 | 25.6393 | 1.1215 | 1.7899 | 1.2497 | 1.918 | 27.0035 | -94.6518 | -94.7388 | -0.086975 | 27.356 | 0.35248 | 27.2781 | 27.2324 | 0.27466 | 0.22888
2535.1719 | 122.2183 | -27.1362 | 94.8669 | -0.18394 | 0.018218 | 0.099307 | 0.093272 | -0.0010959 | 0.012557 | 0.005597 | 0.0017395 | 0.023458 | 0.012155 | -4.0504e-05 | 0.014772 | -0.0049709 | -0.0048983 | 0.041199 | 94.8074 | -27.2964 | 94.7525 | -1.6891 | -25.415 | -25.8774 | 0.11902 | 0.16937 | -0.13733 | 0.018311 | 0.15106 | 0.036621 | 0.09613 | -0.16022 | -0.059509 | -0.11444 | -0.054932 | 25.4471 | 25.6073 | 1.2589 | 1.7212 | 1.4191 | 1.8814 | 26.9989 | -94.6976 | -94.8486 | -0.15106 | 27.2873 | 0.28839 | 27.2324 | 27.1729 | 0.23346 | 0.17395
2536.1719 | 122.2595 | -27.0035 | 94.9585 | -0.19157 | 0.0195 | 0.097171 | 0.093699 | -0.00024129 | 0.010421 | 0.0060242 | 0.013947 | 0.021749 | 0.010446 | -0.0023905 | 0.014772 | -0.0056119 | 0.00044212 | -0.054932 | 94.9448 | -27.1591 | 94.812 | -1.7349 | -25.0763 | -25.6714 | 0.032043 | 0.15564 | -0.1236 | 0.054932 | 0.19226 | 0.059509 | 0.10986 | -0.15564 | -0.013733 | -0.14648 | -0.13275 | 25.2686 | 25.4242 | 1.3321 | 1.9272 | 1.4877 | 2.0828 | 26.8799 | -94.8029 | -94.9036 | -0.10071 | 27.1957 | 0.31586 | 27.1133 | 27.063 | 0.23346 | 0.1831
2537.1719 | 122.2458 | -26.9943 | 94.9722 | -0.11528 | 0.019072 | 0.099307 | 0.093272 | -0.0021641 | 0.012343 | 0.0062378 | -0.0013123 | 0.021535 | 0.014505 | 0.0012411 | 0.014558 | -0.0051845 | -0.0044711 | -0.10071 | 94.7845 | -27.3056 | 94.7479 | -1.7303 | -25.4196 | -26.9852 | 0.018311 | 0.16479 | -0.13733 | 0.036621 | 0.17853 | 0.045776 | 0.11902 | -0.31128 | -0.18768 | -0.2243 | -0.036621 | 25.264 | 25.5753 | 0.009156 | 1.5747 | 0.32043 | 1.886 | 26.857 | -94.8074 | -94.9356 | -0.12817 | 27.1729 | 0.31586 | 27.1133 | 27.0401 | 0.25635 | 0.18311
2538.1719 | 122.1497 | -27.0767 | 94.9631 | -0.10002 | 0.018218 | 0.09888 | 0.093913 | -0.0021641 | 0.011275 | 0.0064514 | -0.062347 | 0.022603 | 0.01066 | 0.0040183 | 0.01349 | -0.0060391 | -0.0031895 | -0.004578 | 94.9631 | -27.2415 | 94.7296 | -1.8677 | -25.502 | -26.152 | 0.19684 | 0.2243 | -0.13733 | -0.064087 | 0.082397 | -0.045776 | 0.009155 | -0.1648 | 0 | -0.23346 | -0.23346 | 25.209 | 25.3738 | 0.92468 | 1.5747 | 1.0895 | 1.7395 | 26.9394 | -94.7388 | -95.0272 | -0.28839 | 27.1591 | 0.21973 | 27.0859 | 27.0309 | 0.14648 | 0.091553
2539.1719 | 122.3007 | -26.9852 | 94.9539 | -0.10002 | 0.019927 | 0.099521 | 0.091563 | -0.0025913 | 0.010848 | 0.0068786 | 0.0078429 | 0.020681 | 0.014078 | 0.0006003 | 0.014345 | -0.0049709 | -0.0021213 | 0.045776 | 94.8944 | -27.2049 | 94.9722 | -1.7395 | -25.2045 | -27.0264 | 0.07782 | 0.16022 | -0.10986 | 0.064087 | 0.20599 | 0.086975 | 0.14648 | -0.21973 | -0.059509 | 0.018311 | 0.07782 | 25.2457 | 25.4654 | -0.041199 | 1.7807 | 0.17853 | 2.0004 | 26.8753 | -94.7937 | -94.8898 | -0.09613 | 27.1912 | 0.31586 | 27.1317 | 27.0721 | 0.25635 | 0.19684
2540.1719 | 122.2458 | -27.0767 | 94.9219 | -0.10765 | 0.01715 | 0.09888 | 0.093058 | -0.0049411 | 0.01213 | 0.0077332 | -0.0013123 | 0.022176 | 0.012582 | 0.0046591 | 0.013704 | -0.0066799 | -0.0042575 | -0.091553 | 94.9265 | -27.2232 | 94.7067 | -1.7853 | -25.2823 | -26.0696 | -0.009155 | 0.10071 | -0.16022 | 0.013733 | 0.16022 | 0.013733 | 0.10071 | -0.14649 | 0.004578 | -0.21515 | -0.21973 | 25.2914 | 25.4379 | 1.0071 | 1.7944 | 1.1536 | 1.9409 | 26.9165 | -94.8212 | -94.9081 | -0.086975 | 27.2369 | 0.32043 | 27.1774 | 27.0905 | 0.26093 | 0.17395
2541.1719 | 122.2687 | -26.9485 | 94.9814 | -0.15343 | 0.016936 | 0.099307 | 0.092845 | -0.0010959 | 0.01213 | 0.0070922 | 0.016998 | 0.021108 | 0.011942 | -0.0043131 | 0.015413 | -0.0041165 | -0.0038303 | -0.036621 | 94.9677 | -27.1317 | 94.8898 | -1.6479 | -25.3738 | -25.8362 | 0.045776 | 0.16479 | -0.13733 | 0.041199 | 0.20142 | 0.07782 | 0.13275 | -0.18311 | -0.013733 | -0.091552 | -0.077819 | 25.3006 | 25.4837 | 1.1124 | 1.5747 | 1.2955 | 1.7578 | 26.8112 | -94.8166 | -94.9402 | -0.1236 | 27.15 | 0.33875 | 27.0813 | 27.0264 | 0.27008 | 0.21515
2542.1719 | 122.2595 | -27.0493 | 94.9219 | -0.10765 | 0.020354 | 0.10016 | 0.094981 | 0.0012539 | 0.012343 | 0.0094422 | -0.016571 | 0.019826 | 0.011514 | -0.0053811 | 0.011995 | -0.0058255 | -0.0021213 | -0.10986 | 94.9081 | -27.1957 | 94.8395 | -1.6479 | -25.4654 | -26.001 | 0.036621 | 0.14648 | -0.14191 | 0.064087 | 0.18768 | 0.064087 | 0.13733 | -0.14649 | -0.013733 | -0.082397 | -0.068664 | 25.4013 | 25.5478 | 1.0483 | 1.5839 | 1.1948 | 1.7303 | 26.9073 | -94.7754 | -94.8578 | -0.082397 | 27.2369 | 0.32959 | 27.1866 | 27.1133 | 0.27924 | 0.20599
2543.1719 | 122.2458 | -27.095 | 94.9173 | -0.092392 | 0.018004 | 0.098666 | 0.093699 | -0.0010959 | 0.011916 | 0.0081604 | -0.0043641 | 0.020467 | 0.012369 | 0.0023093 | 0.01349 | -0.0049709 | -0.0031895 | 0.032043 | 94.8807 | -27.2141 | 94.8257 | -1.7166 | -25.3326 | -25.9872 | 0.11444 | 0.14648 | -0.13733 | 0.045776 | 0.16479 | 0.041199 | 0.10986 | -0.11902 | -0.036621 | -0.091552 | -0.054931 | 25.3784 | 25.4974 | 1.1078 | 1.7624 | 1.2268 | 1.8814 | 26.9577 | -94.7708 | -94.8715 | -0.10071 | 27.2598 | 0.30212 | 27.2049 | 27.1362 | 0.24719 | 0.17853
2544.1719 | 122.2595 | -26.7883 | 95.1736 | -0.15343 | 0.017791 | 0.097384 | 0.091563 | 0.0010403 | 0.011916 | 0.007306 | 0.016998 | 0.021749 | 0.011728 | -0.0034585 | 0.014772 | -0.0058255 | -0.0027621 | -0.086975 | 95.1096 | -27.0264 | 95.0317 | -1.3458 | -25.1083 | -25.7996 | 0.013733 | 0.13275 | -0.14191 | 0.036621 | 0.16022 | 0.045776 | 0.10071 | -0.23804 | -0.064087 | -0.14191 | -0.07782 | 25.4425 | 25.6805 | 0.98877 | 1.68 | 1.2268 | 1.918 | 26.6464 | -95.0409 | -95.137 | -0.09613 | 26.9485 | 0.30212 | 26.889 | 26.8341 | 0.24262 | 0.18768
2545.1875 | 122.2046 | -27.0081 | 95.0317 | -0.077133 | 0.016509 | 0.097812 | 0.096263 | 0.0025357 | 0.013198 | 0.008374 | -0.074554 | 0.02239 | 0.011942 | -0.0015359 | 0.01349 | -0.0034757 | -0.0081027 | 0.022888 | 94.9997 | -27.182 | 94.8532 | -1.8356 | -25.5203 | -26.0742 | 0.050354 | 0.1831 | -0.17853 | -0.013733 | 0.10529 | -0.032043 | 0.064087 | -0.17395 | -0.032043 | -0.17853 | -0.14648 | 25.1724 | 25.3464 | 0.93384 | 1.4877 | 1.1078 | 1.6617 | 26.8295 | -94.8486 | -95.0455 | -0.19684 | 27.1133 | 0.28381 | 27.0721 | 26.976 | 0.24261 | 0.14648
2546.1719 | 122.2275 | -26.9852 | 95.0134 | -0.16106 | 0.016509 | 0.098453 | 0.093913 | 0.0031765 | 0.012557 | 0.01051 | 0.0078429 | 0.020681 | 0.011942 | -0.0060221 | 0.013917 | -0.0060391 | -0.0046847 | -0.082397 | 95.018 | -27.1133 | 94.9081 | -1.7899 | -25.2914 | -25.7217 | -0.036621 | 0.15106 | -0.14648 | -0.013733 | 0.1236 | 0 | 0.068665 | -0.12817 | 0.004577 | -0.10529 | -0.10986 | 25.1953 | 25.3235 | 1.2634 | 1.6937 | 1.3916 | 1.8219 | 26.8387 | -94.8624 | -95.0272 | -0.16479 | 27.1088 | 0.27008 | 27.0538 | 26.9852 | 0.21515 | 0.14648
2547.1719 | 122.3145 | -26.9348 | 95.0272 | -0.11528 | 0.018431 | 0.099948 | 0.09199 | 0.0016813 | 0.010848 | 0.0077332 | -0.0074159 | 0.022603 | 0.011514 | -0.0049539 | 0.016054 | -0.0060391 | -0.0031895 | 0.050354 | 95.0226 | -27.0721 | 94.9905 | -1.7075 | -25.2182 | -25.8865 | 0.091553 | 0.15106 | -0.12817 | 0.073242 | 0.21515 | 0.09613 | 0.14191 | -0.13733 | -0.004578 | -0.036621 | -0.032043 | 25.2274 | 25.3647 | 1.0483 | 1.7166 | 1.1856 | 1.8539 | 26.8066 | -94.8761 | -94.9539 | -0.07782 | 27.15 | 0.34332 | 27.0767 | 27.0309 | 0.27008 | 0.2243
2548.1719 | 122.2961 | -27.0447 | 94.899 | -0.11528 | 0.018859 | 0.096316 | 0.092631 | -0.0025913 | 0.010634 | 0.0070922 | 0.013947 | 0.022603 | 0.012155 | -0.0008951 | 0.014772 | -0.0056119 | -0.0016941 | 0.21515 | 94.9265 | -27.2141 | 94.8853 | -1.6754 | -25.2731 | -26.2436 | 0.13733 | 0.14648 | -0.14648 | 0.041199 | 0.17853 | 0.064087 | 0.1236 | -0.16937 | 0.027466 | -0.013733 | -0.041199 | 25.3693 | 25.5386 | 0.80109 | 1.7715 | 0.97046 | 1.9409 | 26.8982 | -94.7525 | -94.8578 | -0.10529 | 27.2232 | 0.32501 | 27.1683 | 27.1088 | 0.27008 | 0.21057
2549.1719 | 122.2687 | -27.0676 | 94.9036 | -0.10765 | 0.0195 | 0.10059 | 0.095835 | 0.0012539 | 0.011061 | 0.0060242 | 0.02005 | 0.021322 | 0.011728 | -0.0051675 | 0.016267 | -0.0066799 | -0.0012667 | -0.054932 | 94.9036 | -27.2232 | 94.7891 | -1.7853 | -25.3738 | -25.9689 | 0.032043 | 0.16022 | -0.14648 | 0.041199 | 0.18768 | 0.064087 | 0.13275 | -0.15564 | 0 | -0.11444 | -0.11444 | 25.2823 | 25.4379 | 1.0986 | 1.6937 | 1.2543 | 1.8494 | 26.9211 | -94.7433 | -94.8624 | -0.11902 | 27.2552 | 0.33417 | 27.2003 | 27.1317 | 0.27924 | 0.21057
2550.1719 | 122.3053 | -27.0493 | 94.899 | -0.13817 | 0.015227 | 0.09888 | 0.095195 | 0.00018591 | 0.013411 | 0.0053832 | 0.0078429 | 0.020894 | 0.012155 | -0.0040995 | 0.014772 | -0.0041165 | -0.0044711 | -0.013733 | 94.8944 | -27.1866 | 94.8257 | -1.7212 | -25.3555 | -25.9598 | 0.091553 | 0.16022 | -0.10986 | 0.073242 | 0.21057 | 0.082397 | 0.14648 | -0.13733 | -0.004578 | -0.073242 | -0.068664 | 25.3281 | 25.4654 | 1.0895 | 1.6937 | 1.2268 | 1.8311 | 26.9394 | -94.7388 | -94.8257 | -0.086975 | 27.2598 | 0.32043 | 27.1957 | 27.1317 | 0.25635 | 0.19226
2551.1719 | 122.2778 | -26.9257 | 95.0592 | -0.13817 | 0.018431 | 0.098025 | 0.092845 | -0.00024129 | 0.010634 | 0.010937 | 0.023102 | 0.019826 | 0.012796 | -0.0040995 | 0.011995 | -0.0047573 | -0.0057529 | -0.018311 | 95.0592 | -27.0767 | 94.9677 | -1.7166 | -25.1999 | -25.8728 | 0.073242 | 0.16022 | -0.13733 | 0.054932 | 0.19684 | 0.059509 | 0.13733 | -0.15106 | 0 | -0.091553 | -0.091553 | 25.209 | 25.3601 | 1.0529 | 1.7258 | 1.2039 | 1.8768 | 26.7883 | -94.899 | -95.0043 | -0.10528 | 27.1225 | 0.33417 | 27.063 | 26.9852 | 0.27466 | 0.19684
2552.1719 | 122.2687 | -26.619 | 95.3522 | -0.038986 | 0.018004 | 0.10059 | 0.094767 | -0.0028049 | 0.011275 | 0.0068786 | -0.047089 | 0.02239 | 0.01365 | -0.0040995 | 0.013917 | -0.0051845 | -0.0053257 | -0.004578 | 95.3659 | -26.7563 | 95.3018 | -1.2131 | -24.9207 | -25.9644 | 0.032043 | 0.13733 | -0.14191 | 0.082397 | 0.17395 | 0.045776 | 0.13275 | -0.13733 | 0.013733 | -0.050354 | -0.064087 | 25.4059 | 25.5432 | 0.6546 | 1.6983 | 0.79193 | 1.8356 | 26.4771 | -95.2148 | -95.2698 | -0.054932 | 26.7929 | 0.31586 | 26.7517 | 26.6647 | 0.27466 | 0.18768
2553.1719 | 122.2458 | -26.9806 | 94.9631 | -0.10002 | 0.018218 | 0.09888 | 0.090709 | -0.0019503 | 0.011489 | 0.0075196 | -0.040985 | 0.020681 | 0.014505 | -0.0026041 | 0.015626 | -0.0058255 | -0.0036167 | -0.041199 | 94.9493 | -27.1454 | 94.8807 | -1.6846 | -25.3555 | -25.9232 | -0.004578 | 0.13733 | -0.13733 | 0.054932 | 0.15106 | 0.032043 | 0.10529 | -0.16479 | -0.013733 | -0.082398 | -0.068665 | 25.296 | 25.4608 | 1.0574 | 1.6251 | 1.2222 | 1.7899 | 26.8433 | -94.8257 | -94.9081 | -0.082397 | 27.1317 | 0.28839 | 27.0859 | 27.0126 | 0.24262 | 0.16937
2554.1719 | 122.2504 | -26.7059 | 95.2652 | -0.18394 | 0.017363 | 0.098239 | 0.092204 | -0.00024129 | 0.015334 | 0.0079468 | -0.013519 | 0.021535 | 0.012582 | -0.0066629 | 0.014345 | -0.0062527 | -0.00041228 | -0.073242 | 95.2515 | -26.889 | 95.1691 | -1.2451 | -25.0534 | -25.621 | -0.064087 | 0.14648 | -0.1236 | 0.032043 | 0.14191 | 0.013733 | 0.086975 | -0.1831 | -0.013733 | -0.096131 | -0.082398 | 25.4608 | 25.6439 | 1.0849 | 1.6525 | 1.268 | 1.8356 | 26.5823 | -95.1187 | -95.2332 | -0.11444 | 26.8478 | 0.2655 | 26.7929 | 26.7197 | 0.21057 | 0.13733
2555.1719 | 122.2916 | -26.9348 | 95.018 | -0.13054 | 0.016509 | 0.099521 | 0.091777 | -0.0055819 | 0.013411 | 0.0070922 | 0.026154 | 0.021322 | 0.012369 | -0.0038857 | 0.015626 | -0.0049709 | -0.0025485 | 0 | 95.0272 | -27.0905 | 94.9585 | -1.7075 | -25.2686 | -25.8911 | 0.10071 | 0.16937 | -0.11902 | 0.064087 | 0.21057 | 0.091553 | 0.14191 | -0.15564 | 0.009156 | -0.059509 | -0.068665 | 25.2274 | 25.383 | 1.0437 | 1.6663 | 1.1993 | 1.8219 | 26.8158 | -94.8486 | -94.9539 | -0.10529 | 27.1454 | 0.32959 | 27.0767 | 27.0264 | 0.26093 | 0.21057
2556.1719 | 122.3053 | -26.9943 | 95.0272 | -0.13054 | 0.016295 | 0.097812 | 0.092204 | 0.00039951 | 0.014052 | 0.010297 | 0.0078429 | 0.019399 | 0.011942 | 0.0067953 | 0.016481 | -0.0041165 | -0.0059665 | -0.004578 | 94.9951 | -27.1454 | 94.899 | -1.7532 | -25.3189 | -25.9048 | 0.059509 | 0.15106 | -0.12817 | 0.068665 | 0.19684 | 0.086975 | 0.14648 | -0.15106 | -0.032044 | -0.12817 | -0.09613 | 25.2411 | 25.3922 | 1.0895 | 1.6754 | 1.2405 | 1.8265 | 26.8661 | -94.8761 | -94.9585 | -0.082397 | 27.1912 | 0.32501 | 27.1408 | 27.0813 | 0.27466 | 0.21515
2557.1719 | 122.2641 | -26.9073 | 95.0729 | -0.17632 | 0.017791 | 0.099093 | 0.094126 | -0.00066869 | 0.010848 | 0.008374 | -0.013519 | 0.021749 | 0.012582 | 0.0029501 | 0.016054 | -0.0056119 | -0.0010531 | -0.045776 | 95.0592 | -27.0676 | 94.9539 | -1.767 | -25.2182 | -25.9277 | -0.036621 | 0.15106 | -0.14191 | 0.013733 | 0.13733 | 0.022888 | 0.10071 | -0.16022 | -0.013733 | -0.11902 | -0.10529 | 25.1404 | 25.3006 | 0.97961 | 1.6891 | 1.1398 | 1.8494 | 26.7654 | -94.9219 | -95.0592 | -0.13733 | 27.0447 | 0.27924 | 27.0081 | 26.9302 | 0.24262 | 0.16479
2558.1875 | 122.2092 | -27.0401 | 95.018 | -0.092392 | 0.018218 | 0.098453 | 0.092204 | -0.0025913 | 0.011275 | 0.0085876 | -0.022675 | 0.021108 | 0.012796 | 0.0001731 | 0.01584 | -0.0068935 | 0.00086952 | -0.032043 | 94.9951 | -27.1591 | 94.8624 | -1.7944 | -25.3326 | -25.9735 | 0.054932 | 0.18768 | -0.15564 | -0.041199 | 0.059509 | -0.036621 | 0.013733 | -0.11902 | -0.022888 | -0.15564 | -0.13275 | 25.2457 | 25.3647 | 1.0666 | 1.7075 | 1.1856 | 1.8265 | 26.8845 | -94.8303 | -95.0592 | -0.22888 | 27.0996 | 0.21515 | 27.0538 | 27.0035 | 0.16937 | 0.11902
2559.1719 | 122.319 | -27.0493 | 94.8944 | -0.15343 | 0.015654 | 0.10144 | 0.095622 | 0.0023221 | 0.01277 | 0.0085876 | -0.059296 | 0.022817 | 0.011301 | 0.0029501 | 0.015199 | -0.0024075 | 0.0012967 | -0.059509 | 94.899 | -27.2049 | 94.8761 | -1.6937 | -25.4059 | -26.0696 | -0.004578 | 0.13733 | -0.14191 | -0.059509 | -0.036621 | 0.12817 | 0.19226 | -0.15564 | 0.004578 | -0.01831 | -0.022888 | 25.3555 | 25.5112 | 0.97961 | 1.6434 | 1.1353 | 1.799 | 26.9073 | -94.7571 | -94.9539 | -0.19684 | 27.0126 | 0.10529 | 27.2415 | 27.1774 | 0.33417 | 0.27008
2560.1719 | 122.2778 | -26.9531 | 94.9768 | -0.1458 | 0.016936 | 0.10208 | 0.091777 | -0.0019503 | 0.013198 | 0.010724 | -0.022675 | 0.020467 | 0.011942 | 0.0027365 | 0.012636 | -0.0043301 | -0.0040439 | -0.14648 | 94.9722 | -27.1271 | 94.9631 | -1.6983 | -25.1587 | -25.8133 | 0.013733 | 0.12817 | -0.12817 | 0.07782 | 0.1831 | 0.059509 | 0.1236 | -0.17395 | -0.004578 | -0.013733 | -0.009155 | 25.2548 | 25.4288 | 1.1398 | 1.7944 | 1.3138 | 1.9684 | 26.825 | -94.8486 | -94.899 | -0.050354 | 27.1362 | 0.31128 | 27.0767 | 27.0126 | 0.25177 | 0.18768
2561.1719 | 122.287 | -26.889 | 95.0821 | -0.13817 | 0.017363 | 0.098453 | 0.095622 | 0.0010403 | 0.0099934 | 0.012006 | -0.074554 | 0.020894 | 0.012155 | -0.0040995 | 0.01349 | -0.0056119 | -0.0023349 | -0.07782 | 95.0867 | -27.0996 | 94.9631 | -1.6067 | -25.264 | -25.9186 | -0.036621 | 0.16022 | -0.14191 | 0.064087 | 0.16937 | 0.041199 | 0.13733 | -0.21057 | 0.004578 | -0.11902 | -0.1236 | 25.2823 | 25.4929 | 0.97046 | 1.6251 | 1.181 | 1.8356 | 26.7471 | -94.9219 | -95.018 | -0.09613 | 27.0584 | 0.31128 | 27.0264 | 26.9302 | 0.27924 | 0.18311
2562.1719 | 122.2961 | -26.9073 | 95.0317 | -0.21446 | 0.017577 | 0.099948 | 0.091563 | 0.0012539 | 0.0099934 | 0.0090148 | -0.034882 | 0.021963 | 0.01301 | 0.0001731 | 0.014772 | -0.0064663 | -0.0042575 | -0.050354 | 95.0638 | -27.0676 | 95.0088 | -1.6205 | -25.2182 | -25.9003 | -0.004578 | 0.16479 | -0.15564 | 0.07782 | 0.18768 | 0.045776 | 0.12817 | -0.16022 | 0.032044 | -0.022888 | -0.054932 | 25.2869 | 25.4471 | 1.0071 | 1.6891 | 1.1673 | 1.8494 | 26.7517 | -94.8669 | -94.9539 | -0.086975 | 27.095 | 0.34332 | 27.0355 | 26.9531 | 0.28381 | 0.20142
2563.1719 | 122.319 | -26.976 | 95.0043 | -0.15343 | 0.017791 | 0.099307 | 0.089 | -0.00045509 | 0.014907 | 0.0098694 | 0.029205 | 0.019185 | 0.012155 | -0.0006815 | 0.017122 | -0.0047573 | -0.0025485 | -0.027466 | 95.018 | -27.1179 | 94.9219 | -1.7349 | -25.2411 | -25.8957 | 0.082397 | 0.16022 | -0.12817 | 0.068665 | 0.2243 | 0.10986 | 0.15564 | -0.14191 | 0.013733 | -0.082397 | -0.09613 | 25.2411 | 25.383 | 1.0803 | 1.7349 | 1.2222 | 1.8768 | 26.8478 | -94.8441 | -94.9356 | -0.091552 | 27.2003 | 0.35248 | 27.1317 | 27.0859 | 0.28381 | 0.23804
2564.1719 | 122.2778 | -26.9577 | 95.0272 | -0.18394 | 0.018645 | 0.099093 | 0.092418 | -2.7692e-05 | 0.013198 | 0.0062378 | -0.0043641 | 0.019399 | 0.011301 | -0.0023905 | 0.01584 | -0.0049709 | -0.0044711 | -0.036621 | 94.9905 | -27.1317 | 94.9448 | -1.7166 | -25.3326 | -25.9048 | 0.004578 | 0.14648 | -0.1236 | 0.059509 | 0.16479 | 0.045776 | 0.12817 | -0.17395 | -0.036621 | -0.082398 | -0.045777 | 25.2411 | 25.415 | 1.0529 | 1.6251 | 1.2268 | 1.799 | 26.8341 | -94.8807 | -94.9677 | -0.086975 | 27.1225 | 0.28839 | 27.0859 | 27.0035 | 0.25177 | 0.16937
2565.1719 | 122.2778 | -26.7792 | 95.2148 | -0.15343 | 0.016295 | 0.099307 | 0.091563 | 0.00082671 | 0.011702 | 0.0077332 | 0.0078429 | 0.021535 | 0.011728 | -0.0015359 | 0.014131 | -0.0056119 | -0.0046847 | 0.045776 | 95.2057 | -26.944 | 95.137 | -1.2268 | -25.1129 | -25.7355 | 0.10529 | 0.14648 | -0.1236 | 0.041199 | 0.17853 | 0.073242 | 0.11902 | -0.1648 | -0.009156 | -0.07782 | -0.068664 | 25.5524 | 25.7172 | 1.0437 | 1.6663 | 1.2085 | 1.8311 | 26.6556 | -95.0684 | -95.1736 | -0.10529 | 26.9577 | 0.30212 | 26.8982 | 26.8524 | 0.24261 | 0.19684
2566.1719 | 122.1954 | -26.7746 | 95.2927 | -0.15343 | 0.016936 | 0.10059 | 0.094554 | 0.00061311 | 0.010207 | 0.012646 | -0.11118 | 0.023458 | 0.014291 | 0.01064 | 0.015413 | -0.0039029 | -0.0012667 | -0.09613 | 95.2927 | -26.9211 | 95.0592 | -1.3641 | -25.2411 | -25.7446 | -0.13733 | 0.17853 | -0.16937 | -0.018311 | 0.15106 | -0.17853 | -0.036621 | -0.14648 | 0 | -0.23346 | -0.23346 | 25.4105 | 25.5569 | 1.03 | 1.5335 | 1.1765 | 1.68 | 26.6052 | -95.1141 | -95.311 | -0.19684 | 26.9257 | 0.32044 | 26.738 | 26.5961 | 0.13275 | -0.009155
2567.1719 | 122.1863 | -26.8936 | 95.1691 | -0.48912 | 0.020568 | 0.096744 | 0.093058 | -0.0015231 | 0.0099934 | 0.011578 | -0.08371 | 0.020894 | 0.014291 | -0.0006815 | 0.013917 | -0.0041165 | -0.0042575 | -0.050354 | 95.192 | -27.0172 | 94.9448 | -1.355 | -25.3555 | -25.9552 | -0.10986 | 0.17395 | -0.16479 | -0.032043 | 0.059509 | -0.082397 | 0.018311 | -0.1236 | 0.022889 | -0.2243 | -0.24719 | 25.5386 | 25.6622 | 0.93842 | 1.5381 | 1.062 | 1.6617 | 26.7288 | -94.9951 | -95.2011 | -0.20599 | 26.9531 | 0.2243 | 26.9119 | 26.8112 | 0.18311 | 0.082398
2568.1875 | 122.2595 | -26.8616 | 95.1553 | -0.16869 | 0.016936 | 0.099734 | 0.090495 | -0.0032321 | 0.013198 | 0.004956 | 0.035309 | 0.018972 | 0.014719 | -0.0023905 | 0.014345 | -0.0064663 | -0.0014805 | -0.068665 | 95.1004 | -27.0676 | 94.9997 | -1.2085 | -25.1907 | -25.8636 | -0.009155 | 0.15106 | -0.15564 | 0.032043 | 0.15564 | 0.045776 | 0.11902 | -0.20599 | -0.054931 | -0.15564 | -0.10071 | 25.6531 | 25.8591 | 0.99792 | 1.6708 | 1.2039 | 1.8768 | 26.7059 | -95.0043 | -95.1233 | -0.11902 | 27.0172 | 0.31128 | 26.9806 | 26.9073 | 0.27466 | 0.20142
2569.1719 | 122.1909 | -26.7929 | 95.2423 | -0.13054 | 0.018004 | 0.099521 | 0.096049 | -2.7692e-05 | 0.010421 | 0.0079468 | -0.074554 | 0.022817 | 0.01365 | 0.0038047 | 0.012849 | -0.0045437 | -0.0068209 | -0.07782 | 95.2789 | -26.9119 | 95.0684 | -1.3458 | -25.2869 | -25.7309 | -0.11902 | 0.1831 | -0.1831 | -0.027466 | 0.059509 | -0.073242 | 0.027466 | -0.11902 | 0.036621 | -0.17395 | -0.21057 | 25.4471 | 25.5661 | 1.062 | 1.506 | 1.181 | 1.6251 | 26.6098 | -95.0592 | -95.2698 | -0.21057 | 26.8524 | 0.24261 | 26.8204 | 26.7197 | 0.21057 | 0.10986
2570.1719 | 122.2916 | -26.7334 | 95.2652 | -0.092392 | 0.018645 | 0.099307 | 0.091563 | 0.00018591 | 0.012343 | 0.012006 | -0.040985 | 0.02004 | 0.011301 | 0.0038047 | 0.013917 | -0.0043301 | -0.0023349 | -0.054932 | 95.1645 | -26.9257 | 95.1874 | -1.1856 | -25.0763 | -25.882 | 0.022888 | 0.15106 | -0.13733 | 0.086975 | 0.17395 | 0.036621 | 0.14648 | -0.19226 | -0.10071 | -0.07782 | 0.022888 | 25.5478 | 25.7401 | 0.85144 | 1.6571 | 1.0437 | 1.8494 | 26.5961 | -95.1141 | -95.1782 | -0.064087 | 26.9073 | 0.31128 | 26.8799 | 26.77 | 0.28381 | 0.17395
2571.1719 | 122.2092 | -26.77 | 95.2423 | -0.077133 | 0.018218 | 0.10016 | 0.095622 | -0.0040867 | 0.011061 | 0.0077332 | -0.08371 | 0.021322 | 0.012582 | 0.0038047 | 0.014986 | -0.0036893 | -0.0012667 | -0.10071 | 95.2515 | -26.9394 | 95.0455 | -1.2817 | -25.0946 | -25.8591 | -0.10986 | 0.16022 | -0.17853 | -0.027466 | 0.073242 | -0.059509 | 0.054932 | -0.16937 | 0.009155 | -0.19684 | -0.20599 | 25.4883 | 25.6577 | 0.91095 | 1.6754 | 1.0803 | 1.8448 | 26.5915 | -95.0821 | -95.2698 | -0.18768 | 26.8433 | 0.25177 | 26.825 | 26.7105 | 0.23346 | 0.11902
2572.1719 | 122.2687 | -26.7151 | 95.3018 | -0.13054 | 0.018004 | 0.099948 | 0.092631 | -0.0028049 | 0.013198 | 0.008374 | -0.062347 | 0.022817 | 0.014078 | -0.0011087 | 0.013917 | -0.0049709 | -0.0036167 | -0.10529 | 95.3293 | -26.8845 | 95.1187 | -1.2451 | -25.1129 | -25.7217 | -0.082397 | 0.16022 | -0.16022 | 0.032043 | 0.12817 | -0.013733 | 0.10071 | -0.16937 | 0.027466 | -0.18311 | -0.21057 | 25.47 | 25.6393 | 0.99335 | 1.6022 | 1.1627 | 1.7715 | 26.5549 | -95.1416 | -95.2698 | -0.12817 | 26.8433 | 0.28839 | 26.8158 | 26.7014 | 0.26093 | 0.14648
2573.1719 | 122.2458 | -26.7151 | 95.2881 | -0.084763 | 0.01715 | 0.099093 | 0.093913 | 0.00082671 | 0.011061 | 0.005597 | -0.040985 | 0.024099 | 0.014291 | -0.0006815 | 0.015413 | -0.0034757 | -0.0021213 | -0.12817 | 95.2881 | -26.9028 | 95.0821 | -1.268 | -25.1816 | -25.6027 | -0.10529 | 0.1831 | -0.15564 | 0.004578 | 0.091553 | -0.050354 | 0.041199 | -0.18768 | 0 | -0.20599 | -0.20599 | 25.4471 | 25.6348 | 1.1124 | 1.5335 | 1.3 | 1.7212 | 26.5594 | -95.105 | -95.2835 | -0.17853 | 26.8066 | 0.24719 | 26.7563 | 26.6647 | 0.19684 | 0.10529
2574.1719 | 122.2137 | -26.8707 | 95.192 | -0.10765 | 0.017791 | 0.099307 | 0.093272 | 0.0044583 | 0.013198 | 0.0070922 | -0.059296 | 0.02239 | 0.012369 | 0.0008139 | 0.015626 | -0.0056119 | -0.0036167 | -0.059509 | 95.2286 | -27.0172 | 95.0363 | -1.3412 | -25.6485 | -25.8224 | -0.064087 | 0.17395 | -0.16022 | -0.059509 | 0.036621 | -0.09613 | 0.009155 | -0.14648 | 0.036621 | -0.15564 | -0.19226 | 25.5295 | 25.676 | 1.0483 | 1.2222 | 1.1948 | 1.3687 | 26.7105 | -95.018 | -95.2515 | -0.23346 | 26.9073 | 0.19684 | 26.8799 | 26.7746 | 0.16937 | 0.064087
2575.1719 | 122.287 | -27.0721 | 94.9081 | -0.1458 | 0.017791 | 0.10294 | 0.092204 | -0.0017367 | 0.015548 | 0.0079468 | 0.013947 | 0.022176 | 0.011301 | -0.0006815 | 0.015199 | -0.0047573 | -0.0029757 | -0.27008 | 94.8441 | -27.2598 | 94.8395 | -1.7349 | -25.2136 | -25.9415 | 0.022888 | 0.11902 | -0.13733 | 0.064087 | 0.18768 | 0.054932 | 0.15106 | -0.18768 | -0.064087 | -0.068664 | -0.004577 | 25.3372 | 25.5249 | 1.1307 | 1.8585 | 1.3184 | 2.0462 | 26.9348 | -94.7891 | -94.8441 | -0.054932 | 27.2598 | 0.32501 | 27.2232 | 27.1271 | 0.28839 | 0.19226
2576.1719 | 122.2137 | -27.1408 | 94.8761 | -0.10002 | 0.018218 | 0.099307 | 0.093058 | 0.00039951 | 0.013198 | 0.0088012 | -0.059296 | 0.021535 | 0.014078 | 0.0040183 | 0.014772 | -0.0053983 | -0.0019077 | -0.10071 | 94.9127 | -27.3148 | 94.6976 | -1.8631 | -25.7126 | -26.2024 | -0.1236 | 0.17395 | -0.16937 | -0.032043 | 0.068665 | -0.073242 | 0.032043 | -0.17395 | 0.036621 | -0.17853 | -0.21515 | 25.2777 | 25.4517 | 0.93841 | 1.4282 | 1.1124 | 1.6022 | 26.9714 | -94.7021 | -94.9081 | -0.20599 | 27.2095 | 0.23804 | 27.1729 | 27.0676 | 0.20142 | 0.096131
2577.1719 | 122.2046 | -26.7883 | 95.2423 | -0.077133 | 0.018431 | 0.1008 | 0.094126 | 0.0018949 | 0.0099934 | 0.010724 | -0.059296 | 0.023458 | 0.012369 | 0.0031637 | 0.015199 | -0.0039029 | -0.0010531 | -0.054932 | 95.2698 | -26.944 | 95.0363 | -1.3458 | -25.1862 | -25.8408 | -0.041199 | 0.15106 | -0.16022 | -0.045776 | 0.050354 | -0.091553 | 0.009155 | -0.15564 | 0.027465 | -0.20599 | -0.23346 | 25.4425 | 25.5981 | 0.94757 | 1.6022 | 1.1032 | 1.7578 | 26.6281 | -95.0912 | -95.2881 | -0.19684 | 26.8387 | 0.21057 | 26.7975 | 26.6968 | 0.16937 | 0.068664
2578.1719 | 122.1863 | -26.8982 | 95.1691 | -0.10002 | 0.019713 | 0.097384 | 0.093058 | 0.00082671 | 0.014266 | 0.0075196 | -0.056244 | 0.021535 | 0.014078 | -0.0021769 | 0.014558 | -0.0045437 | -0.00083948 | -0.032043 | 95.1599 | -27.0218 | 94.9631 | -1.3824 | -25.264 | -25.9369 | -0.022888 | 0.18768 | -0.17395 | -0.045776 | 0.045776 | -0.059509 | 0.018311 | -0.1236 | -0.009155 | -0.20599 | -0.19684 | 25.5157 | 25.6393 | 0.9613 | 1.6342 | 1.0849 | 1.7578 | 26.7242 | -94.9814 | -95.2148 | -0.23346 | 26.944 | 0.21973 | 26.9165 | 26.8387 | 0.19226 | 0.11444
2579.1719 | 122.2 | -26.825 | 95.2606 | -0.12291 | 0.015654 | 0.099093 | 0.093699 | 0.0018949 | 0.011916 | 0.0064514 | -0.016571 | 0.022603 | 0.013223 | 0.0023093 | 0.013704 | -0.0032619 | 0.00022852 | -0.064087 | 95.2286 | -26.976 | 95.0226 | -1.3687 | -24.765 | -25.8774 | -0.091553 | 0.17395 | -0.17853 | -0.068665 | 0.018311 | -0.11902 | 0 | -0.15106 | -0.032043 | -0.23804 | -0.20599 | 25.4562 | 25.6073 | 0.94757 | 2.0599 | 1.0986 | 2.211 | 26.6464 | -95.0867 | -95.3293 | -0.24261 | 26.8433 | 0.19684 | 26.825 | 26.7059 | 0.17853 | 0.059509
2580.1719 | 122.2733 | -27.0813 | 94.9081 | -0.1992 | 0.017363 | 0.098025 | 0.093913 | -0.0028049 | 0.010207 | 0.007306 | -0.025726 | 0.023244 | 0.011087 | -0.0019631 | 0.015199 | -0.0043301 | -0.00041228 | -0.09613 | 94.9173 | -27.2964 | 94.7525 | -1.8448 | -25.7858 | -26.1108 | -0.082397 | 0.16479 | -0.16022 | 0.027466 | 0.10986 | -0.009155 | 0.091553 | -0.21515 | 0.009155 | -0.15564 | -0.16479 | 25.2365 | 25.4517 | 0.97046 | 1.2955 | 1.1856 | 1.5106 | 26.9211 | -94.7433 | -94.8807 | -0.13733 | 27.1912 | 0.27008 | 27.1729 | 27.0721 | 0.25177 | 0.15106
2581.1719 | 122.1909 | -27.0859 | 94.9631 | -0.1458 | 0.016082 | 0.10123 | 0.093913 | 0.00082671 | 0.011916 | 0.01051 | -0.059296 | 0.021963 | 0.014719 | 0.0003867 | 0.015413 | -0.0041165 | -0.0021213 | -0.073242 | 94.9493 | -27.2461 | 94.7388 | -1.886 | -25.5569 | -26.12 | -0.12817 | 0.15564 | -0.16937 | -0.068665 | 0.018311 | -0.10529 | 0.018311 | -0.16022 | -0.013733 | -0.2243 | -0.21057 | 25.1999 | 25.3601 | 0.96588 | 1.5289 | 1.1261 | 1.6891 | 26.9165 | -94.8074 | -95.0317 | -0.22431 | 27.1042 | 0.18768 | 27.1042 | 26.9806 | 0.18768 | 0.064087
2582.1719 | 122.2275 | -27.0584 | 94.9951 | -0.092392 | 0.017791 | 0.1008 | 0.09434 | -0.0038729 | 0.012557 | 0.0064514 | -0.037933 | 0.022603 | 0.014719 | -0.0013223 | 0.014558 | -0.0043301 | -0.0070347 | -0.009155 | 95.0226 | -27.1545 | 94.8257 | -1.7029 | -25.4288 | -26.088 | -0.004578 | 0.20599 | -0.09613 | -0.013733 | 0.082397 | -0.041199 | 0.050354 | -0.09613 | 0.027466 | -0.16937 | -0.19684 | 25.3555 | 25.4517 | 0.97046 | 1.6296 | 1.0666 | 1.7258 | 26.9623 | -94.7891 | -95.0088 | -0.21973 | 27.1408 | 0.17853 | 27.1088 | 27.0172 | 0.14648 | 0.054931
2583.1719 | 122.3328 | -26.8433 | 95.1233 | -0.15343 | 0.018431 | 0.10272 | 0.093486 | 0.0012539 | 0.012343 | 0.010937 | 0.013947 | 0.019613 | 0.012155 | -0.0045267 | 0.015413 | -0.0032619 | -0.0042575 | -0.054932 | 95.1004 | -27.0126 | 95.0226 | -1.712 | -25.1404 | -25.8408 | -0.009155 | 0.13275 | -0.14648 | 0.064087 | 0.17853 | 0.045776 | 0.14648 | -0.16937 | -0.022888 | -0.10071 | -0.07782 | 25.1312 | 25.3006 | 1.0025 | 1.7029 | 1.1719 | 1.8723 | 26.6968 | -94.9905 | -95.0592 | -0.068664 | 27.0218 | 0.32501 | 26.9897 | 26.889 | 0.29297 | 0.19226
2584.1875 | 122.3145 | -26.9257 | 95.0867 | -0.13817 | 0.017577 | 0.10059 | 0.093486 | -0.0025913 | 0.012343 | 0.0090148 | 0.065826 | 0.021322 | 0.013864 | 0.0029501 | 0.012849 | -0.0058255 | -0.0025485 | -0.045776 | 95.0867 | -27.0584 | 94.9493 | -1.9226 | -25.2914 | -25.7858 | -0.027466 | 0.082397 | -0.14191 | 0.073242 | 0.17853 | 0.064087 | 0.13733 | -0.13275 | 0 | -0.13733 | -0.13733 | 25.0031 | 25.1358 | 1.1398 | 1.6342 | 1.2726 | 1.767 | 26.7838 | -95.0043 | -95.0134 | -0.009155 | 27.1042 | 0.32044 | 27.063 | 26.9897 | 0.27924 | 0.20599
2585.1719 | 122.2275 | -26.6739 | 95.3888 | -0.24498 | 0.01715 | 0.097384 | 0.096263 | 0.0018949 | 0.012557 | 0.0070922 | -0.0074159 | 0.022817 | 0.011942 | 0.0076499 | 0.013917 | -0.0051845 | 0.0021511 | -0.086975 | 95.4208 | -26.8066 | 95.1645 | -1.3321 | -25.0809 | -25.7126 | -0.10529 | 0.16937 | -0.17395 | -0.041199 | 0.059509 | -0.073242 | 0.032043 | -0.13275 | 0.032043 | -0.2243 | -0.25635 | 25.3418 | 25.4745 | 0.9613 | 1.593 | 1.0941 | 1.7258 | 26.4999 | -95.2194 | -95.43 | -0.21057 | 26.7334 | 0.23346 | 26.7059 | 26.6006 | 0.20599 | 0.10071
2586.1719 | 122.2275 | -26.683 | 95.4254 | -0.084763 | 0.015441 | 0.09888 | 0.09434 | -0.00024129 | 0.013625 | 0.0045288 | -0.016571 | 0.021108 | 0.014291 | 0.0020957 | 0.015626 | -0.0075345 | -0.0027621 | -0.041199 | 95.4163 | -26.8433 | 95.1691 | -1.3092 | -25.0214 | -25.6989 | -0.009155 | 0.19226 | -0.15564 | -0.064087 | 0.059509 | -0.059509 | 0.045776 | -0.16022 | -0.009155 | -0.25635 | -0.24719 | 25.3738 | 25.5341 | 0.98419 | 1.6617 | 1.1444 | 1.8219 | 26.5274 | -95.2332 | -95.4895 | -0.25635 | 26.7426 | 0.21515 | 26.7288 | 26.6235 | 0.20142 | 0.096131
2587.1719 | 122.2504 | -26.6464 | 95.4208 | -0.069504 | 0.018004 | 0.098453 | 0.095408 | 0.0010403 | 0.01213 | 0.006665 | -0.0074159 | 0.022603 | 0.014291 | 0.0010275 | 0.013704 | -0.0058255 | 0.0010831 | -0.013733 | 95.4071 | -26.825 | 95.224 | -1.2726 | -24.9298 | -25.6485 | 0.054932 | 0.15564 | -0.16479 | -0.013733 | 0.082397 | -0.045776 | 0.059509 | -0.17853 | -0.013733 | -0.19684 | -0.1831 | 25.3738 | 25.5524 | 0.99792 | 1.7166 | 1.1765 | 1.8951 | 26.4816 | -95.2652 | -95.4346 | -0.16937 | 26.7288 | 0.24719 | 26.7059 | 26.6006 | 0.2243 | 0.11902
2588.1719 | 122.2458 | -26.6052 | 95.4712 | -0.13054 | 0.019927 | 0.10016 | 0.093058 | -0.0010959 | 0.011916 | 0.005597 | -0.019623 | 0.022603 | 0.01301 | 0.0038047 | 0.014345 | -0.0032619 | -0.0012667 | -0.091553 | 95.4529 | -26.7654 | 95.2332 | -1.2634 | -25.0214 | -25.6714 | -0.11444 | 0.17395 | -0.1831 | -0.032043 | 0.068665 | -0.064087 | 0.054932 | -0.16022 | -0.01831 | -0.23804 | -0.21973 | 25.3418 | 25.502 | 0.93384 | 1.5839 | 1.0941 | 1.7441 | 26.4221 | -95.2972 | -95.5032 | -0.20599 | 26.6739 | 0.25177 | 26.6602 | 26.5411 | 0.23804 | 0.11902
2589.1875 | 122.3923 | -26.651 | 95.3888 | -0.15343 | 0.019286 | 0.10187 | 0.093486 | 0.00061311 | 0.010634 | 0.0077332 | 0.053619 | 0.021535 | 0.015146 | -0.0049539 | 0.01584 | -0.0034757 | -0.0072483 | -0.041199 | 95.4117 | -26.7517 | 95.3201 | -1.1444 | -24.9344 | -25.5524 | 0.018311 | 0.13275 | -0.10986 | 0.054932 | 0 | -0.18768 | -0.013733 | -0.10071 | 0.022888 | -0.068665 | -0.091553 | 25.5066 | 25.6073 | 1.0986 | 1.7166 | 1.1993 | 1.8173 | 26.5411 | -95.256 | -95.3339 | -0.077819 | 26.651 | 0.10986 | 26.6373 | 26.4633 | 0.09613 | -0.07782
2590.1719 | 122.287 | -26.5823 | 95.4346 | -0.16106 | 0.016509 | 0.10038 | 0.091349 | -0.0032321 | 0.012343 | 0.010297 | 0.056671 | 0.020254 | 0.012582 | -0.0017495 | 0.014772 | -0.0060391 | 0.0017239 | -0.068665 | 95.4346 | -26.7288 | 95.3247 | -1.1398 | -24.9252 | -25.5249 | 0.013733 | 0.14648 | -0.13275 | 0.045776 | 0.16022 | 0.027466 | 0.10986 | -0.14648 | 0 | -0.10986 | -0.10986 | 25.4425 | 25.589 | 1.0574 | 1.6571 | 1.2039 | 1.8036 | 26.4496 | -95.2881 | -95.3888 | -0.10071 | 26.7426 | 0.29297 | 26.6922 | 26.6098 | 0.24261 | 0.16022
2591.1719 | 122.2 | -26.4771 | 95.6223 | -0.084763 | 0.015654 | 0.10016 | 0.094767 | 0.00039951 | 0.012343 | 0.011151 | -0.05014 | 0.021322 | 0.01301 | 0.0050863 | 0.014772 | -0.0043301 | -0.0063937 | 0.018311 | 95.5536 | -26.7105 | 95.3339 | -1.2039 | -24.9023 | -25.5112 | 0.036621 | 0.16937 | -0.14648 | -0.013733 | 0.09613 | -0.022888 | 0.35706 | -0.23346 | -0.068664 | -0.28839 | -0.21973 | 25.2731 | 25.5066 | 0.96588 | 1.5747 | 1.1993 | 1.8082 | 26.3306 | -95.4529 | -95.636 | -0.18311 | 26.5732 | 0.24261 | 26.8341 | 26.4542 | 0.50354 | 0.1236
2592.1719 | 122.2412 | -26.564 | 95.5124 | -0.11528 | 0.017363 | 0.10144 | 0.094767 | 0.0021085 | 0.010634 | 0.0081604 | -0.05014 | 0.02004 | 0.012582 | 0.0031637 | 0.014558 | -0.0064663 | -0.00062588 | 0.10529 | 95.5582 | -26.6235 | 95.3568 | -1.1948 | -24.9527 | -25.5844 | 0.013733 | 0.17853 | -0.11902 | -0.022888 | 0.086975 | -0.045776 | 0.050354 | -0.059509 | 0.045777 | -0.15564 | -0.20142 | 25.3693 | 25.4288 | 0.97961 | 1.6113 | 1.0391 | 1.6708 | 26.445 | -95.3339 | -95.5353 | -0.20142 | 26.651 | 0.20599 | 26.6144 | 26.5182 | 0.16937 | 0.073243
2593.1719 | 122.1954 | -26.5732 | 95.5582 | -0.11528 | 0.018645 | 0.10102 | 0.09669 | 0.00082671 | 0.012343 | 0.0081604 | -0.059296 | 0.021749 | 0.011087 | 0.0044455 | 0.011568 | -0.0056119 | -0.0036167 | -0.086975 | 95.5353 | -26.6922 | 95.3018 | -1.3 | -25.0626 | -25.6165 | -0.12817 | 0.17853 | -0.1831 | -0.073242 | 0.027466 | -0.12817 | 0.004578 | -0.11902 | -0.022889 | -0.25635 | -0.23346 | 25.2731 | 25.3922 | 0.95673 | 1.5106 | 1.0757 | 1.6296 | 26.3901 | -95.3796 | -95.6314 | -0.25177 | 26.6006 | 0.21057 | 26.5778 | 26.445 | 0.18768 | 0.054931
2594.1719 | 122.2092 | -26.5503 | 95.549 | -0.13054 | 0.016936 | 0.10059 | 0.097117 | 0.0033901 | 0.013625 | 0.0096558 | -0.034882 | 0.023031 | 0.012582 | -0.0006815 | 0.016481 | -0.0041165 | -0.00062588 | -0.064087 | 95.5261 | -26.7014 | 95.3064 | -1.2955 | -24.9985 | -25.621 | -0.10529 | 0.16479 | -0.16479 | -0.082397 | 0.009155 | -0.11444 | 0.004578 | -0.15106 | -0.022888 | -0.24261 | -0.21973 | 25.2548 | 25.4059 | 0.92926 | 1.5518 | 1.0803 | 1.7029 | 26.3855 | -95.3842 | -95.6314 | -0.24719 | 26.5594 | 0.17395 | 26.5549 | 26.4359 | 0.16937 | 0.050354
2595.1719 | 122.2275 | -26.5228 | 95.5673 | -0.16869 | 0.016509 | 0.10016 | 0.09669 | 0.0018949 | 0.014479 | 0.0092286 | -0.03183 | 0.020894 | 0.013864 | 0.0001731 | 0.014131 | -0.0041165 | -0.0029757 | 0.013733 | 95.6543 | -26.5823 | 95.4025 | -1.2085 | -24.8108 | -25.5112 | 0.004578 | 0.17395 | -0.14191 | -0.013733 | 0.07782 | -0.050354 | 0.054932 | -0.059509 | 0.086975 | -0.16479 | -0.25177 | 25.3143 | 25.3738 | 1.0117 | 1.712 | 1.0712 | 1.7715 | 26.3809 | -95.3934 | -95.5811 | -0.18768 | 26.6006 | 0.21973 | 26.5778 | 26.4725 | 0.19684 | 0.091553
2596.1719 | 122.2458 | -26.5045 | 95.5536 | -0.13817 | 0.017363 | 0.099734 | 0.094126 | 0.0010403 | 0.011916 | 0.005597 | -0.0013123 | 0.020894 | 0.014719 | -0.0049539 | 0.015413 | -0.0058255 | -0.0031895 | -0.009155 | 95.5994 | -26.6373 | 95.3751 | -1.2222 | -24.884 | -25.5753 | 0.054932 | 0.17395 | -0.13275 | -0.004578 | 0.10529 | -0.032043 | 0.064087 | -0.13275 | 0.045776 | -0.17853 | -0.2243 | 25.2823 | 25.415 | 0.92926 | 1.6205 | 1.062 | 1.7532 | 26.3718 | -95.3796 | -95.5582 | -0.17853 | 26.6098 | 0.23804 | 26.5686 | 26.4725 | 0.19684 | 0.10071
2597.1719 | 122.2412 | -26.5549 | 95.4849 | -0.12291 | 0.018004 | 0.099734 | 0.095408 | -0.00088229 | 0.010848 | 0.0094422 | -0.013519 | 0.021749 | 0.012582 | -0.0047403 | 0.015626 | -0.0015529 | -0.0016941 | 0.16937 | 95.517 | -26.7197 | 95.3247 | -1.2131 | -25.0534 | -25.5249 | 0.045776 | 0.16022 | -0.16937 | 0.09613 | 0.16937 | 0.009155 | 0.09613 | -0.1648 | 0.032044 | -0.16022 | -0.19226 | 25.3418 | 25.5066 | 1.03 | 1.5015 | 1.1948 | 1.6663 | 26.3855 | -95.3247 | -95.3888 | -0.064087 | 26.7242 | 0.33875 | 26.651 | 26.564 | 0.2655 | 0.17853
2598.1875 | 122.3602 | -26.5045 | 95.5078 | -0.1992 | 0.017791 | 0.099734 | 0.093058 | 0.0012539 | 0.011489 | 0.011151 | 0.032257 | 0.021108 | 0.014291 | 0.0010275 | 0.013917 | -0.0086025 | -0.0023349 | -0.07782 | 95.5032 | -26.6373 | 95.4163 | -1.1261 | -24.9115 | -25.4288 | 0.009155 | 0.11902 | -0.13733 | 0.091553 | 0.16937 | 0 | 0.12817 | -0.13275 | -0.004577 | -0.091552 | -0.086975 | 25.3784 | 25.5112 | 1.0757 | 1.593 | 1.2085 | 1.7258 | 26.3672 | -95.3888 | -95.4163 | -0.027466 | 26.6739 | 0.3067 | 26.6327 | 26.5045 | 0.2655 | 0.13733
2599.1719 | 122.2275 | -26.4587 | 95.6772 | -0.092392 | 0.018431 | 0.098239 | 0.091136 | 0.00018591 | 0.011061 | 0.0064514 | 0.062775 | 0.023885 | 0.014291 | -0.0055947 | 0.013277 | -0.0056119 | -0.00062588 | -0.041199 | 95.6497 | -26.6006 | 95.3705 | -1.2543 | -24.7787 | -25.4883 | 0.004578 | 0.17395 | -0.16479 | -0.045776 | -0.036621 | -0.013733 | 0.050354 | -0.14191 | -0.027466 | -0.3067 | -0.27924 | 25.2045 | 25.3464 | 0.97046 | 1.68 | 1.1124 | 1.8219 | 26.2939 | -95.5032 | -95.723 | -0.21973 | 26.4221 | 0.12817 | 26.5091 | 26.445 | 0.21515 | 0.15106
2600.1719 | 122.2504 | -26.4267 | 95.6497 | -0.13054 | 0.018859 | 0.099307 | 0.096904 | -0.00045509 | 0.011275 | 0.0098694 | -0.028778 | 0.02239 | 0.012369 | 0.0008139 | 0.016267 | -0.0071071 | -0.0025485 | -0.3067 | 95.7001 | -26.5411 | 95.5032 | -1.1948 | -25.6851 | -25.4013 | 0.009155 | 0.14648 | -0.16479 | -0.036621 | 0.045776 | -0.068665 | 0.041199 | -0.11444 | 0.050354 | -0.14648 | -0.19684 | 25.2319 | 25.3464 | 1.0254 | 0.74158 | 1.1398 | 0.85602 | 26.2619 | -95.5032 | -95.6863 | -0.18311 | 26.4725 | 0.21057 | 26.4679 | 26.358 | 0.20599 | 0.09613
2601.1719 | 122.319 | -26.3718 | 95.6772 | -0.10002 | 0.017577 | 0.10016 | 0.093699 | -0.0028049 | 0.013198 | 0.0090148 | 0.099396 | 0.021749 | 0.012796 | -0.0034585 | 0.013063 | -0.0053983 | -0.0051121 | 0.032043 | 95.636 | -26.5366 | 95.5353 | -1.1124 | -24.7421 | -25.3601 | 0.10529 | 0.14191 | -0.13733 | 0.022888 | 0.15106 | 0.041199 | 0.10529 | -0.1648 | -0.041199 | -0.14191 | -0.10071 | 25.2594 | 25.4242 | 1.0117 | 1.6296 | 1.1765 | 1.7944 | 26.2344 | -95.5353 | -95.6543 | -0.11902 | 26.5228 | 0.28839 | 26.4771 | 26.413 | 0.24262 | 0.17853
2602.1719 | 122.2183 | -26.3535 | 95.755 | -0.10002 | 0.015654 | 0.099093 | 0.090922 | 0.0018949 | 0.011489 | 0.010297 | 0.053619 | 0.02004 | 0.014719 | -0.0066629 | 0.013917 | -0.0058255 | -0.0016941 | 0.091553 | 95.6955 | -26.5091 | 95.5673 | -0.99335 | -24.9344 | -25.3418 | 0.073242 | 0.15564 | -0.13275 | -0.050354 | 0.082397 | -0.009155 | 0.018311 | -0.15564 | -0.059509 | -0.18768 | -0.12817 | 25.3601 | 25.5157 | 1.0117 | 1.4191 | 1.1673 | 1.5747 | 26.2207 | -95.5994 | -95.8054 | -0.20599 | 26.4359 | 0.21515 | 26.3718 | 26.3443 | 0.15106 | 0.1236
2603.1719 | 122.2504 | -26.3901 | 95.7138 | -0.17632 | 0.016509 | 0.10208 | 0.094767 | 0.00082671 | 0.014266 | 0.0058106 | -0.056244 | 0.020894 | 0.012582 | 0.0020957 | 0.012849 | -0.0058255 | -0.0040439 | -0.009155 | 95.7184 | -26.5228 | 95.5078 | -1.0529 | -24.5728 | -25.3876 | 0.022888 | 0.17853 | -0.15564 | -0.032043 | 0.064087 | -0.059509 | 0.045776 | -0.13275 | 0.004578 | -0.20599 | -0.21057 | 25.3372 | 25.47 | 1.0025 | 1.8173 | 1.1353 | 1.9501 | 26.2344 | -95.5353 | -95.7458 | -0.21057 | 26.4542 | 0.21973 | 26.4359 | 26.3306 | 0.20142 | 0.096131
2604.1719 | 122.2641 | -26.3718 | 95.7275 | -0.092392 | 0.015654 | 0.098239 | 0.093913 | 0.0042447 | 0.012557 | 0.0060242 | -0.0013123 | 0.023458 | 0.014078 | 0.0080771 | 0.014772 | -0.0053983 | -0.0014805 | 0.018311 | 95.7779 | -26.5274 | 95.4987 | -1.1124 | -24.3027 | -25.4013 | -0.013733 | 0.16937 | -0.17395 | -0.082397 | 0.064087 | -0.059509 | 0.032043 | -0.15564 | 0.050354 | -0.22888 | -0.27924 | 25.2594 | 25.415 | 0.97046 | 2.0691 | 1.1261 | 2.2247 | 26.1978 | -95.5582 | -95.8099 | -0.25177 | 26.4359 | 0.23804 | 26.4038 | 26.3123 | 0.20599 | 0.11444
2605.1719 | 122.2275 | -26.4267 | 95.6863 | -0.17632 | 0.015013 | 0.099734 | 0.092631 | -0.0021641 | 0.014693 | 0.0064514 | 0.0017395 | 0.021749 | 0.016855 | -0.0036721 | 0.014558 | -0.0032619 | -0.0029757 | 0.018311 | 95.723 | -26.5411 | 95.4575 | -1.0803 | -24.7604 | -25.4517 | -0.018311 | 0.1831 | -0.17853 | -0.050354 | 0.07782 | -0.068665 | 0.013733 | -0.11444 | 0.036621 | -0.22888 | -0.2655 | 25.3464 | 25.4608 | 0.97504 | 1.6663 | 1.0895 | 1.7807 | 26.2482 | -95.5032 | -95.7367 | -0.23346 | 26.5045 | 0.25635 | 26.4404 | 26.358 | 0.19226 | 0.10986
2606.1719 | 122.3007 | -26.4313 | 95.6268 | -0.13817 | 0.01715 | 0.099734 | 0.090068 | -0.00088229 | 0.012984 | 0.0092286 | 0.059723 | 0.019613 | 0.012582 | -0.0015359 | 0.015199 | -0.0066799 | -0.0014805 | -0.086975 | 95.6314 | -26.5549 | 95.5078 | -0.94299 | -24.8016 | -25.3784 | 0.10529 | 0.15106 | -0.13275 | -0.018311 | 0.15564 | 0.041199 | 0.10071 | -0.1236 | 0.004578 | -0.11902 | -0.1236 | 25.4883 | 25.6119 | 1.0529 | 1.6296 | 1.1765 | 1.7532 | 26.2985 | -95.4758 | -95.6451 | -0.16937 | 26.5869 | 0.28839 | 26.532 | 26.4725 | 0.23346 | 0.17395
2607.1719 | 122.3602 | -26.4404 | 95.5902 | -0.1458 | 0.015441 | 0.10102 | 0.092631 | 0.0021085 | 0.01277 | 0.0075196 | 0.087189 | 0.020467 | 0.014291 | -0.0021769 | 0.013917 | -0.0062527 | -0.0055393 | 0.041199 | 95.6223 | -26.5503 | 95.4712 | -0.89722 | -23.7122 | -25.4929 | 0.13275 | 0.14191 | -0.12817 | 0.032043 | 0.21057 | 0.091553 | 0.14191 | -0.10986 | 0.032043 | -0.11902 | -0.15106 | 25.5432 | 25.6531 | 0.94757 | 2.7283 | 1.0574 | 2.8381 | 26.3123 | -95.4483 | -95.5582 | -0.10986 | 26.651 | 0.33875 | 26.5823 | 26.532 | 0.27008 | 0.21973
2608.1719 | 122.2 | -26.445 | 95.6772 | -0.11528 | 0.013945 | 0.099521 | 0.092631 | 0.00039951 | 0.01277 | 0.0064514 | 0.013947 | 0.022603 | 0.012796 | -0.0017495 | 0.013917 | -0.0043301 | 1.4916e-05 | -0.11902 | 95.6314 | -26.5869 | 95.3293 | -1.062 | -24.9573 | -25.7126 | 0.013733 | 0.19684 | -0.16479 | -0.07782 | 0.068665 | -0.059509 | -0.004578 | -0.14191 | -0.045776 | -0.3479 | -0.30212 | 25.383 | 25.5249 | 0.73242 | 1.4877 | 0.87433 | 1.6296 | 26.2802 | -95.4803 | -95.755 | -0.27466 | 26.5137 | 0.23346 | 26.4404 | 26.3855 | 0.16022 | 0.10529
2609.1719 | 122.2321 | -26.445 | 95.5994 | -0.077133 | 0.018218 | 0.099521 | 0.090922 | -0.0028049 | 0.011916 | 0.007306 | 0.074982 | 0.020467 | 0.010873 | 0.0027365 | 0.014772 | -0.0086025 | -0.0038303 | 0.036621 | 95.6085 | -26.5594 | 95.5582 | -0.91553 | -24.5773 | -25.3647 | 0.15106 | 0.13733 | -0.12817 | -0.091553 | 0.18768 | 0.15564 | 0.2243 | -0.11444 | 0.009156 | -0.041198 | -0.050354 | 25.5295 | 25.6439 | 1.0803 | 1.8677 | 1.1948 | 1.9821 | 26.3168 | -95.462 | -95.6909 | -0.22888 | 26.6327 | 0.31586 | 26.6693 | 26.6006 | 0.35248 | 0.28381
2610.1719 | 122.3419 | -26.3626 | 95.6635 | -0.10765 | 0.01715 | 0.097812 | 0.09434 | -0.00045509 | 0.014266 | 0.0085876 | 0.0078429 | 0.021108 | 0.01301 | 0.0010275 | 0.013704 | -0.0066799 | -0.0046847 | 0.073242 | 95.6635 | -26.4771 | 95.4758 | -1.1353 | -24.4995 | -25.2869 | 0.14191 | 0.1236 | -0.13275 | 0.036621 | 0.1831 | 0.073242 | 0.11902 | -0.11444 | 0 | -0.18768 | -0.18768 | 25.2274 | 25.3418 | 1.0757 | 1.8631 | 1.1902 | 1.9775 | 26.2299 | -95.5399 | -95.6268 | -0.086975 | 26.5457 | 0.31586 | 26.4816 | 26.4359 | 0.25177 | 0.20599
2611.1719 | 122.3877 | -26.3809 | 95.6635 | -0.1458 | 0.016295 | 0.10038 | 0.092418 | -0.0038729 | 0.012984 | 0.0058106 | 0.065826 | 0.019613 | 0.01301 | -0.0013223 | 0.014986 | -0.010739 | -0.0066073 | 0.018311 | 95.6955 | -26.4954 | 95.6131 | -0.91553 | -24.234 | -25.3235 | 0.14191 | 0.13733 | -0.13275 | 0.10986 | 0.068665 | -0.09613 | 0.050354 | -0.11444 | 0.032044 | -0.050354 | -0.082398 | 25.4654 | 25.5798 | 1.0574 | 2.1469 | 1.1719 | 2.2614 | 26.2482 | -95.5261 | -95.5536 | -0.027466 | 26.4496 | 0.20142 | 26.4313 | 26.2848 | 0.18311 | 0.036621
2612.1719 | 122.374 | -26.326 | 95.7275 | -0.13817 | 0.018218 | 0.099307 | 0.094767 | 0.0014675 | 0.013411 | 0.0058106 | 0.047516 | 0.022603 | 0.014291 | -0.0006815 | 0.012636 | -0.0071071 | -0.0025485 | 0.009155 | 95.7138 | -26.4771 | 95.6406 | -0.91553 | -24.5407 | -25.2731 | 0.16479 | 0.15564 | -0.13275 | 0.041199 | 0.20142 | 0.059509 | 0.13733 | -0.15106 | -0.013733 | -0.086975 | -0.073242 | 25.4105 | 25.5615 | 1.0529 | 1.7853 | 1.2039 | 1.9363 | 26.1932 | -95.5719 | -95.6863 | -0.11444 | 26.5274 | 0.33417 | 26.4633 | 26.3855 | 0.27008 | 0.19226
2613.1719 | 122.2458 | -26.3351 | 95.8099 | -0.069504 | 0.017791 | 0.098666 | 0.093913 | -0.0015231 | 0.01213 | 0.0077332 | -0.065399 | 0.022176 | 0.015573 | 0.0033773 | 0.013277 | -0.0073209 | -0.0029757 | -0.10986 | 95.8008 | -26.4954 | 95.5811 | -0.85602 | -24.5178 | -25.5341 | 0.041199 | 0.23346 | -0.10986 | -0.07782 | 0.068665 | -0.068665 | 0.009155 | -0.16022 | -0.009156 | -0.22888 | -0.21973 | 25.4791 | 25.6393 | 0.80109 | 1.8173 | 0.9613 | 1.9775 | 26.2253 | -95.5765 | -95.8878 | -0.31128 | 26.4038 | 0.17853 | 26.3443 | 26.2665 | 0.11902 | 0.041198
2614.1719 | 122.3557 | -26.2527 | 95.7779 | -0.13054 | 0.015868 | 0.09888 | 0.09199 | -0.0038729 | 0.011702 | 0.0077332 | 0.035309 | 0.021749 | 0.014505 | -0.0026041 | 0.013917 | -0.0047573 | -0.0034031 | 0 | 95.7687 | -26.3901 | 95.6726 | -0.98877 | -24.2935 | -25.1907 | 0.073242 | 0.17395 | -0.12817 | 0.045776 | 0.20142 | 0.091553 | 0.14191 | -0.13733 | -0.009155 | -0.10529 | -0.096131 | 25.264 | 25.4013 | 1.062 | 1.9592 | 1.1993 | 2.0966 | 26.1246 | -95.6039 | -95.7321 | -0.12817 | 26.4542 | 0.32959 | 26.3947 | 26.3443 | 0.27008 | 0.21973
2615.1719 | 122.3557 | -26.3031 | 95.7367 | -0.11528 | 0.0148 | 0.099307 | 0.09434 | -0.0028049 | 0.013411 | 0.0079468 | 0.10245 | 0.021535 | 0.013223 | -0.0015359 | 0.014345 | -0.0041165 | -0.0023349 | 0.022888 | 95.7504 | -26.413 | 95.6818 | -0.94299 | -24.5316 | -25.0809 | -0.013733 | 0.14191 | -0.14191 | 0.068665 | 0.1831 | 0.073242 | 0.1236 | -0.10986 | 0.013733 | -0.054931 | -0.068664 | 25.3601 | 25.47 | 1.2222 | 1.7715 | 1.3321 | 1.8814 | 26.1612 | -95.5948 | -95.668 | -0.073242 | 26.4862 | 0.32501 | 26.4267 | 26.3763 | 0.2655 | 0.21515
2616.1719 | 122.3694 | -26.2573 | 95.7642 | -0.10765 | 0.016082 | 0.099521 | 0.091349 | 0.0016813 | 0.011916 | 0.0030334 | 0.038361 | 0.021963 | 0.011728 | -0.0004679 | 0.014772 | -0.0073209 | -0.0014805 | -0.004578 | 95.7825 | -26.3947 | 95.6909 | -1.0117 | -24.527 | -25.177 | 0.086975 | 0.15106 | -0.09613 | 0.068665 | 0.21057 | 0.082397 | 0.11902 | -0.13733 | 0.018311 | -0.073242 | -0.091553 | 25.2457 | 25.383 | 1.0803 | 1.7303 | 1.2177 | 1.8677 | 26.1612 | -95.6131 | -95.6955 | -0.082397 | 26.4679 | 0.3067 | 26.3763 | 26.3397 | 0.21515 | 0.17853
2617.1719 | 122.3557 | -26.2573 | 95.8054 | -0.13817 | 0.015227 | 0.09888 | 0.093486 | -2.7692e-05 | 0.010634 | 0.004956 | 0.026154 | 0.018331 | 0.011942 | -0.0030313 | 0.015413 | -0.0081753 | -0.0042575 | -0.29297 | 95.7916 | -26.3855 | 95.6497 | -0.99792 | -24.8062 | -25.2136 | 0.045776 | 0.14648 | -0.11902 | 0.059509 | 0.1831 | 0.073242 | 0.1236 | -0.12817 | -0.013733 | -0.15564 | -0.14191 | 25.2594 | 25.3876 | 1.0437 | 1.4511 | 1.1719 | 1.5793 | 26.1383 | -95.6589 | -95.7458 | -0.086975 | 26.4404 | 0.30212 | 26.3809 | 26.3306 | 0.24261 | 0.19226
2618.1719 | 122.2366 | -26.3763 | 95.787 | -0.077133 | 0.015441 | 0.099307 | 0.096476 | -0.0023777 | 0.011702 | 0.0090148 | -0.010467 | 0.022176 | 0.01301 | 0.0044455 | 0.013917 | -0.0053983 | -0.0036167 | -0.032043 | 95.7825 | -26.4954 | 95.5307 | -1.1353 | -24.6094 | -25.177 | -0.1236 | 0.23804 | -0.11444 | -0.045776 | 0.050354 | -0.068665 | 0.004578 | -0.11902 | -0.004577 | -0.25635 | -0.25177 | 25.2411 | 25.3601 | 1.1993 | 1.767 | 1.3184 | 1.886 | 26.2619 | -95.549 | -95.8328 | -0.28381 | 26.4267 | 0.16479 | 26.3809 | 26.3077 | 0.11902 | 0.045776
2619.1719 | 122.3236 | -26.3214 | 95.7413 | -0.11528 | 0.016509 | 0.099093 | 0.09434 | 0.0033901 | 0.015761 | 0.0081604 | 0.026154 | 0.020894 | 0.013437 | 0.0003867 | 0.013704 | -0.0066799 | 0.00065592 | -0.064087 | 95.723 | -26.4496 | 95.636 | -1.0117 | -24.5316 | -25.2411 | 0.059509 | 0.16937 | -0.11902 | 0.054932 | 0.20142 | 0.068665 | 0.12817 | -0.12817 | -0.018311 | -0.10529 | -0.086975 | 25.3098 | 25.4379 | 1.0803 | 1.7899 | 1.2085 | 1.918 | 26.2024 | -95.5719 | -95.6863 | -0.11444 | 26.5228 | 0.32044 | 26.4496 | 26.3901 | 0.24719 | 0.18768
2620.1719 | 122.3602 | -26.3214 | 95.755 | -0.12291 | 0.018004 | 0.10166 | 0.090281 | 0.0014675 | 0.01277 | 0.0085876 | 0.17264 | 0.021749 | 0.01301 | -0.0085855 | 0.012422 | -0.0060391 | -0.0040439 | 0.022888 | 95.7367 | -26.4175 | 95.6497 | -0.97504 | -24.3256 | -25.2457 | 0.09613 | 0.13275 | -0.13275 | 0.041199 | 0.21973 | 0.059509 | 0.1236 | -0.096131 | -0.018311 | -0.10529 | -0.086975 | 25.3464 | 25.4425 | 1.0757 | 1.9958 | 1.1719 | 2.092 | 26.1887 | -95.6223 | -95.7138 | -0.091552 | 26.5411 | 0.35248 | 26.445 | 26.3809 | 0.25635 | 0.19226
2621.1719 | 122.2961 | -26.3123 | 95.7275 | -0.1992 | 0.016722 | 0.099734 | 0.093272 | -0.00045509 | 0.011916 | 0.0088012 | -0.03183 | 0.020254 | 0.012155 | -0.0004679 | 0.013277 | -0.0075345 | -0.0048983 | 0.068665 | 95.7321 | -26.4313 | 95.6406 | -1.0025 | -24.3256 | -25.3098 | 0.009155 | 0.15564 | -0.10986 | 0.050354 | 0.16479 | 0.032043 | 0.10986 | -0.11902 | 0.004578 | -0.086975 | -0.091553 | 25.3098 | 25.4288 | 1.0025 | 1.9867 | 1.1215 | 2.1057 | 26.2024 | -95.5719 | -95.6772 | -0.10529 | 26.4771 | 0.27466 | 26.4221 | 26.3443 | 0.21973 | 0.14191
2622.1719 | 122.3145 | -26.3718 | 95.6635 | -0.16869 | 0.017577 | 0.099948 | 0.095408 | -0.00045509 | 0.01277 | 0.0081604 | -0.028778 | 0.021535 | 0.01301 | -0.0011087 | 0.011995 | -0.0066799 | -0.00041228 | -0.11444 | 95.6589 | -26.4771 | 95.6039 | -1.0025 | -24.6918 | -25.2823 | 0.027466 | 0.16479 | -0.11444 | 0.064087 | 0.17853 | 0.036621 | 0.1236 | -0.10529 | -0.004577 | -0.059509 | -0.054932 | 25.3693 | 25.4745 | 1.0895 | 1.68 | 1.1948 | 1.7853 | 26.2573 | -95.4987 | -95.5994 | -0.10071 | 26.5503 | 0.29297 | 26.4954 | 26.4084 | 0.23804 | 0.15106
2623.1719 | 122.3419 | -26.3443 | 95.7138 | -0.20683 | 0.018218 | 0.1008 | 0.095195 | -0.0043003 | 0.014479 | 0.006665 | 0.023102 | 0.02004 | 0.01066 | 0.0003867 | 0.013704 | -0.0047573 | -0.0078891 | -0.036621 | 95.6909 | -26.4816 | 95.5856 | -1.0483 | -24.5911 | -25.3052 | 0.009155 | 0.14648 | -0.12817 | 0.041199 | 0.17395 | 0.059509 | 0.11444 | -0.13733 | -0.022888 | -0.12817 | -0.10529 | 25.296 | 25.4333 | 1.0391 | 1.7532 | 1.1765 | 1.8906 | 26.2161 | -95.5673 | -95.6726 | -0.10529 | 26.5182 | 0.30212 | 26.4587 | 26.4038 | 0.24262 | 0.18768
2624.1719 | 122.2961 | -26.4038 | 95.6635 | -0.21446 | 0.016509 | 0.10038 | 0.093486 | 0.00082671 | 0.011702 | 0.0090148 | 0.0078429 | 0.020894 | 0.01301 | 0.0020957 | 0.014345 | -0.0060391 | -0.0046847 | 0.036621 | 95.6177 | -26.5366 | 95.5399 | -1.0071 | -24.3851 | -25.3693 | 0.009155 | 0.15564 | -0.1236 | 0.045776 | 0.16022 | 0.027466 | 0.1236 | -0.13275 | -0.045776 | -0.1236 | -0.07782 | 25.3967 | 25.5295 | 1.0345 | 2.0187 | 1.1673 | 2.1515 | 26.2802 | -95.5078 | -95.6177 | -0.10986 | 26.564 | 0.28381 | 26.5274 | 26.4313 | 0.24719 | 0.15106
2625.1719 | 122.319 | -26.2802 | 95.7275 | -0.1458 | 0.016295 | 0.10144 | 0.09199 | -0.0043003 | 0.01213 | 0.008374 | -0.019623 | 0.020254 | 0.014719 | -0.0030313 | 0.016695 | -0.0060391 | -0.0027621 | 0.036621 | 95.7047 | -26.4725 | 95.5856 | -1.0483 | -24.5865 | -26.3672 | -0.059509 | 0.16479 | -0.12817 | 0.027466 | 0.13733 | -0.004578 | 0.086975 | -0.19226 | -0.022888 | -0.14191 | -0.11902 | 25.2319 | 25.4242 | -0.086975 | 1.6937 | 0.10529 | 1.886 | 26.152 | -95.5627 | -95.7001 | -0.13733 | 26.4175 | 0.2655 | 26.3672 | 26.2756 | 0.21515 | 0.1236
2626.1719 | 122.3282 | -28.8254 | 93.1686 | -0.58067 | 0.41556 | 0.10208 | 0.092204 | -0.00024129 | 0.40968 | 0.006665 | -0.0013123 | 0.021322 | 0.012369 | -0.0008951 | 0.013704 | -0.0049709 | -0.00062588 | -0.068665 | 93.0725 | -28.981 | 92.9764 | -1.0712 | 31.7093 | -27.7954 | -0.027466 | 0.15106 | -0.11902 | 0.036621 | 0.13275 | 0.009155 | 0.07782 | -0.15564 | -0.09613 | -0.19226 | -0.096131 | 27.7542 | 27.9099 | 1.03 | 60.5347 | 1.1856 | 60.6903 | 28.7064 | -93.0176 | -93.132 | -0.11444 | 28.9581 | 0.25177 | 28.9032 | 28.8345 | 0.19684 | 0.12817
2627.1719 | 122.2641 | -28.7155 | 93.2465 | -0.52727 | 0.37967 | 0.1023 | 0.088359 | -0.0043003 | 0.37208 | 0.0088012 | -0.025726 | 0.021535 | 0.01365 | 0.0014547 | 0.015199 | -0.0068935 | -0.0019077 | 1.6983 | 93.0725 | -28.9719 | 93.0496 | -0.98877 | 30.6427 | -27.6718 | 0 | 0.13733 | -0.10986 | 0.045776 | 0.13733 | 0.032043 | 0.09613 | -0.25635 | -0.17395 | -0.19684 | -0.022888 | 27.7267 | 27.9831 | 1.0437 | 59.3582 | 1.3 | 59.6146 | 28.6057 | -93.1091 | -93.2007 | -0.091553 | 28.8528 | 0.24719 | 28.8116 | 28.7476 | 0.20599 | 0.14191
2628.1719 | 122.3099 | -30.4001 | 91.5344 | -0.55016 | 0.44205 | 0.10166 | 0.087504 | -0.0023777 | 0.43425 | 0.0064514 | 0.023102 | 0.020894 | 0.011728 | 0.0016685 | 0.014986 | -0.0051845 | -0.0055393 | 0.027466 | 91.5344 | -30.5145 | 91.4658 | -2.7557 | 38.237 | -28.5507 | 0.15106 | 0.1236 | -0.14191 | 0.059509 | 0.20599 | 0.082397 | 0.11902 | -0.11444 | 0 | -0.068665 | -0.068665 | 27.6443 | 27.7588 | 1.8494 | 68.6371 | 1.9638 | 68.7515 | 30.2582 | -91.4108 | -91.4749 | -0.064087 | 30.6061 | 0.3479 | 30.5191 | 30.4825 | 0.26093 | 0.2243
2629.1719 | 122.2366 | -31.3934 | 90.5411 | -0.69511 | 0.54886 | 0.10443 | 0.089 | 0.0018949 | 0.54085 | 0.013287 | -0.019623 | 0.021108 | 0.011942 | -0.0030313 | 0.012849 | -0.0062527 | -0.0031895 | -0.032043 | 90.4861 | -31.4667 | 90.4358 | -3.1677 | 50.679 | -30.1895 | 0.013733 | 0.15106 | -0.10986 | 0.082397 | 0.15564 | 0.027466 | 0.12817 | -0.073242 | -0.054932 | -0.10529 | -0.050354 | 28.2257 | 28.2989 | 1.2039 | 82.0724 | 1.2772 | 82.1457 | 31.2836 | -90.39 | -90.4587 | -0.068665 | 31.5491 | 0.2655 | 31.5216 | 31.4209 | 0.23804 | 0.13733
2630.1719 | 122.2504 | -31.0226 | 90.8798 | -0.5349 | 0.51019 | 0.10337 | 0.088572 | -0.0025913 | 0.5041 | 0.008374 | -0.044037 | 0.021322 | 0.011087 | 0.0016685 | 0.014772 | -0.0068935 | -0.0046847 | -0.054932 | 90.802 | -31.1737 | 90.7562 | -2.2797 | 53.2883 | -29.8645 | 0.009155 | 0.13733 | -0.13275 | 0.091553 | 0.16937 | 0.045776 | 0.1236 | -0.15106 | -0.07782 | -0.1236 | -0.045776 | 28.743 | 28.894 | 1.1581 | 84.3109 | 1.3092 | 84.462 | 30.8899 | -90.7425 | -90.7883 | -0.045776 | 31.192 | 0.30212 | 31.1462 | 31.0684 | 0.25635 | 0.17853
2631.1719 | 122.2366 | -33.4305 | 88.4216 | -0.87822 | 0.74069 | 0.10443 | 0.087077 | -0.0015231 | 0.72905 | 0.011151 | -0.037933 | 0.020894 | 0.01066 | 0.0038047 | 0.015199 | -0.0036893 | -0.0070347 | -0.054932 | 88.353 | -33.5999 | 88.2614 | -2.7695 | 82.2601 | -32.1945 | 0.045776 | 0.10529 | -0.13275 | 0.07782 | 0.16479 | 0.041199 | 0.13733 | -0.16937 | -0.068665 | -0.16022 | -0.091552 | 30.661 | 30.8304 | 1.236 | 115.6906 | 1.4053 | 115.86 | 33.2977 | -88.3163 | -88.3438 | -0.027466 | 33.5953 | 0.29755 | 33.5678 | 33.4717 | 0.27008 | 0.17395
2632.1719 | 122.1909 | -33.0643 | 88.8245 | -0.86296 | 0.74219 | 0.10379 | 0.086436 | -0.0030185 | 0.73118 | 0.0096558 | -0.013519 | 0.021322 | 0.011942 | 0.0018821 | 0.013917 | -0.0062527 | -0.0023349 | -0.054932 | 88.7192 | -33.1833 | 88.6414 | -1.7487 | 83.2672 | -31.7871 | 0.009155 | 0.14191 | -0.12817 | 0.050354 | 0.15564 | 0.022888 | 0.10071 | -0.11902 | -0.10529 | -0.18311 | -0.07782 | 31.3156 | 31.4346 | 1.2772 | 116.3315 | 1.3962 | 116.4505 | 32.9361 | -88.6826 | -88.7741 | -0.091553 | 33.2199 | 0.28381 | 33.165 | 33.0872 | 0.22888 | 0.15106
2633.1719 | 122.2 | -33.2108 | 88.6322 | -0.85533 | 0.74133 | 0.10144 | 0.086222 | 0.0033901 | 0.72862 | 0.0075196 | 0.023102 | 0.021749 | 0.012155 | -0.0064493 | 0.012849 | -0.0060391 | -0.0019077 | 0.009155 | 88.5956 | -33.284 | 88.5773 | -1.7075 | 83.3359 | -31.81 | 0.14191 | 0.16022 | -0.11902 | 0.064087 | 0.19684 | 0.10071 | 0.13275 | -0.073243 | -0.036621 | -0.054931 | -0.01831 | 31.5033 | 31.5765 | 1.4008 | 116.5466 | 1.474 | 116.6199 | 33.0917 | -88.472 | -88.5681 | -0.09613 | 33.4076 | 0.31586 | 33.3435 | 33.3115 | 0.25177 | 0.21973
2634.1719 | 122.1405 | -33.3572 | 88.4491 | -0.85533 | 0.73962 | 0.10422 | 0.085795 | -0.0017367 | 0.72713 | 0.0088012 | -0.013519 | 0.023458 | 0.014291 | -0.0036721 | 0.013704 | -0.0047573 | -0.0053257 | -0.054932 | 88.4262 | -33.4671 | 88.353 | -1.7349 | 83.0795 | -31.8787 | 0.036621 | 0.15564 | -0.13275 | 0.059509 | 0.1831 | 0.036621 | 0.14191 | -0.10986 | -0.022889 | -0.096131 | -0.073242 | 31.6223 | 31.7322 | 1.4786 | 116.4368 | 1.5884 | 116.5466 | 33.2245 | -88.2935 | -88.3896 | -0.096131 | 33.5403 | 0.31586 | 33.4991 | 33.3939 | 0.27466 | 0.16937
2635.1719 | 122.0947 | -33.5403 | 88.2431 | -0.8477 | 0.73898 | 0.10465 | 0.088572 | -0.00024129 | 0.72648 | 0.0092286 | -0.080658 | 0.023458 | 0.0089508 | 0.0029501 | 0.014131 | -0.0032619 | -0.0025485 | -0.091553 | 88.2431 | -33.6456 | 88.2019 | -1.7212 | 82.9102 | -32.309 | 0 | 0.13275 | -0.13733 | 0.082397 | 0.16937 | 0.064087 | -0.013733 | -0.10529 | 0 | -0.041199 | -0.041199 | 31.8192 | 31.9244 | 1.2314 | 116.4505 | 1.3367 | 116.5558 | 33.403 | -88.1104 | -88.1607 | -0.050354 | 33.7097 | 0.3067 | 33.5266 | 33.6044 | 0.1236 | 0.20142
2636.1719 | 122.1359 | -34.053 | 87.767 | -0.89348 | 0.73578 | 0.10742 | 0.086009 | 0.0029629 | 0.72328 | 0.010297 | -0.053192 | 0.024526 | 0.010873 | -0.0021769 | 0.015199 | -0.0060391 | 1.4916e-05 | -0.059509 | 87.7579 | -34.1309 | 87.6251 | -1.7349 | 82.4432 | -32.7713 | -0.022888 | 0.13733 | -0.11902 | 0.086975 | 0.1831 | 0.050354 | 0.14648 | -0.077819 | -0.009155 | -0.14191 | -0.13275 | 32.3181 | 32.3959 | 1.2817 | 116.4963 | 1.3596 | 116.5741 | 33.934 | -87.6297 | -87.6801 | -0.050354 | 34.2361 | 0.30212 | 34.1995 | 34.1034 | 0.2655 | 0.16937
2637.1719 | 122.113 | -33.8699 | 87.9318 | -0.81718 | 0.73407 | 0.10572 | 0.088359 | 0.00061311 | 0.72157 | 0.0085876 | -0.016571 | 0.023031 | 0.0085236 | -0.0019631 | 0.014558 | -0.0053983 | -0.0051121 | -0.013733 | 87.9135 | -33.9157 | 87.8952 | -1.7349 | 82.6126 | -32.7988 | 0.027466 | 0.11902 | -0.13275 | 0.07782 | 0.17853 | 0.050354 | 0.14191 | -0.045776 | -0.018311 | -0.036621 | -0.01831 | 32.135 | 32.1808 | 1.0712 | 116.4825 | 1.1169 | 116.5283 | 33.7372 | -87.8128 | -87.854 | -0.041199 | 34.0485 | 0.31128 | 34.0118 | 33.9203 | 0.27466 | 0.18311
2638.1719 | 122.1359 | -33.8333 | 87.941 | -0.80193 | 0.73215 | 0.10422 | 0.084727 | -0.0036593 | 0.71815 | 0.0098694 | 0.013947 | 0.021108 | 0.011728 | -0.0062357 | 0.014558 | -0.0049709 | -0.0025485 | 0.027466 | 87.8952 | -33.9294 | 87.8769 | -1.5427 | 82.8049 | -32.4051 | 0.11444 | 0.14191 | -0.10986 | 0.050354 | 0.21973 | 0.082397 | 0.13733 | -0.09613 | -0.045776 | -0.064087 | -0.018311 | 32.2906 | 32.3868 | 1.4282 | 116.6382 | 1.5244 | 116.7343 | 33.7234 | -87.7991 | -87.8906 | -0.091553 | 34.053 | 0.32959 | 33.9706 | 33.9157 | 0.24719 | 0.19226
2639.1719 | 122.0993 | -34.0576 | 87.735 | -0.84007 | 0.72851 | 0.10465 | 0.085154 | 0.0010403 | 0.71687 | 0.0092286 | -0.0013123 | 0.023031 | 0.010446 | -0.0013223 | 0.013277 | -0.0056119 | -0.0046847 | -0.050354 | 87.7304 | -34.1125 | 87.6343 | -1.6937 | 82.5073 | -32.6569 | 0.013733 | 0.14191 | -0.10986 | 0.059509 | 0.14648 | 0.041199 | 0.13275 | -0.054932 | -0.004577 | -0.10071 | -0.096131 | 32.3639 | 32.4188 | 1.4008 | 116.5649 | 1.4557 | 116.6199 | 33.9478 | -87.5931 | -87.6755 | -0.082398 | 34.2041 | 0.25635 | 34.1904 | 34.0988 | 0.24261 | 0.15106
2640.1719 | 122.0856 | -34.1217 | 87.735 | -0.86296 | 0.72787 | 0.10636 | 0.086009 | -0.00045509 | 0.71367 | 0.006665 | 0.1116 | 0.02474 | 0.013223 | -0.0017495 | 0.014345 | -0.0064663 | -0.00083948 | -0.054932 | 87.6709 | -34.1492 | 87.5381 | -1.7441 | 82.5119 | -32.7759 | -0.064087 | 0.09613 | -0.1236 | 0.009155 | 0.14191 | 0.027466 | 0.10529 | -0.027466 | -0.064087 | -0.19684 | -0.13275 | 32.3776 | 32.4051 | 1.3458 | 116.6336 | 1.3733 | 116.6611 | 33.9981 | -87.6389 | -87.7258 | -0.086975 | 34.2636 | 0.2655 | 34.227 | 34.1492 | 0.22888 | 0.15106
2641.1719 | 122.113 | -34.053 | 87.7579 | -0.82481 | 0.72467 | 0.10507 | 0.088359 | 0.0031765 | 0.71174 | 0.0077332 | -0.019623 | 0.022603 | 0.012155 | 0.0010275 | 0.013917 | -0.0041165 | -0.0029757 | -0.032043 | 87.6022 | -34.1629 | 87.6068 | -1.6663 | 82.5577 | -32.8125 | 0.018311 | 0.15564 | -0.1236 | 0.07782 | 0.15564 | 0.054932 | 0.13275 | -0.10986 | -0.15564 | -0.15106 | 0.004578 | 32.3868 | 32.4966 | 1.2405 | 116.6107 | 1.3504 | 116.7206 | 33.9294 | -87.6022 | -87.6801 | -0.07782 | 34.2087 | 0.27924 | 34.1858 | 34.108 | 0.25635 | 0.17853
2642.1719 | 122.1222 | -34.053 | 87.6892 | -0.82481 | 0.72296 | 0.10742 | 0.089 | 0.00039951 | 0.71153 | 0.0088012 | -0.037933 | 0.023244 | 0.013223 | 0.0038047 | 0.014345 | -0.0041165 | -0.0042575 | -0.059509 | 87.6434 | -34.1675 | 87.6114 | -1.6891 | 82.5897 | -32.7896 | 0.013733 | 0.14191 | -0.11444 | 0.086975 | 0.16937 | 0.041199 | 0.15564 | -0.11444 | -0.045776 | -0.07782 | -0.032044 | 32.3639 | 32.4783 | 1.2634 | 116.6428 | 1.3779 | 116.7572 | 33.9386 | -87.5473 | -87.6022 | -0.054932 | 34.2224 | 0.28381 | 34.2087 | 34.0942 | 0.27008 | 0.15564
2643.1719 | 122.049 | -34.1537 | 87.6022 | -0.87822 | 0.72018 | 0.10657 | 0.087931 | 0.00039951 | 0.70961 | 0.005597 | -0.0013123 | 0.021963 | 0.011514 | 0.0023093 | 0.014345 | -0.0047573 | -0.0029757 | -0.050354 | 87.5473 | -34.2636 | 87.4603 | -1.6342 | 82.4661 | -32.9361 | -0.018311 | 0.16479 | -0.10071 | 0.032043 | 0.14191 | 0.032043 | 0.10986 | -0.10986 | -0.054932 | -0.14191 | -0.086975 | 32.5195 | 32.6294 | 1.2177 | 116.6199 | 1.3275 | 116.7297 | 34.053 | -87.4374 | -87.5702 | -0.13275 | 34.2957 | 0.24262 | 34.2636 | 34.1858 | 0.21057 | 0.13275
2644.1719 | 122.0764 | -34.0759 | 87.6297 | -0.90874 | 0.72061 | 0.107 | 0.088145 | 0.00082671 | 0.70576 | 0.0079468 | -0.10202 | 0.022817 | 0.01066 | 0.0035911 | 0.014558 | -0.0030483 | -0.0031895 | 0 | 87.6114 | -34.2178 | 87.529 | -1.6251 | 82.5531 | -33.2428 | 0.11444 | 0.10071 | -0.13275 | 0.059509 | 0.19684 | 0.036621 | 0.13733 | -0.14191 | -0.018311 | -0.10071 | -0.082397 | 32.4509 | 32.5928 | 0.83313 | 116.629 | 0.97504 | 116.7709 | 33.9432 | -87.529 | -87.5702 | -0.041199 | 34.2728 | 0.32959 | 34.2133 | 34.1125 | 0.27008 | 0.16937
2645.1719 | 122.1039 | -34.0439 | 87.7213 | -0.78667 | 0.71805 | 0.1055 | 0.086009 | -0.00045509 | 0.71004 | 0.010297 | 0.0078429 | 0.021535 | 0.013437 | 0.0027365 | 0.014772 | -0.0053983 | -0.00062588 | 0.013733 | 87.6297 | -34.1629 | 87.5885 | -1.5289 | 82.576 | -32.6477 | 0.10986 | 0.1236 | -0.11444 | 0.059509 | 0.19684 | 0.064087 | 0.11902 | -0.11902 | -0.091552 | -0.13275 | -0.041199 | 32.515 | 32.634 | 1.3962 | 116.6199 | 1.5152 | 116.7389 | 33.9294 | -87.5977 | -87.6617 | -0.064087 | 34.2407 | 0.31128 | 34.1629 | 34.108 | 0.23346 | 0.17853
2646.1719 | 122.0718 | -34.1263 | 87.6343 | -0.82481 | 0.71762 | 0.10572 | 0.088145 | -0.0019503 | 0.70534 | 0.010937 | -0.022675 | 0.021108 | 0.011728 | -0.0051675 | 0.013277 | -0.0062527 | 1.4916e-05 | -0.050354 | 87.5244 | -34.2636 | 87.5061 | -1.6113 | 82.5394 | -32.7942 | 0.022888 | 0.15564 | -0.1236 | 0.07782 | 0.19226 | 0.054932 | 0.14191 | -0.13733 | -0.10986 | -0.12817 | -0.01831 | 32.515 | 32.6523 | 1.3321 | 116.6656 | 1.4694 | 116.803 | 34.0027 | -87.4786 | -87.5565 | -0.07782 | 34.3185 | 0.31586 | 34.2682 | 34.1812 | 0.2655 | 0.17853
2647.1719 | 122.0764 | -34.0668 | 87.6938 | -0.84007 | 0.71484 | 0.10721 | 0.08964 | 0.00018591 | 0.70448 | 0.0090148 | -0.019623 | 0.021535 | 0.010873 | -0.0002543 | 0.014345 | -0.0049709 | -0.0025485 | -0.16937 | 87.5381 | -34.2178 | 87.5565 | -1.6067 | 82.6126 | -32.7621 | 0.022888 | 0.1236 | -0.13275 | 0.07782 | 0.17395 | 0.054932 | 0.13733 | -0.15106 | -0.15564 | -0.13733 | 0.018311 | 32.46 | 32.6111 | 1.3046 | 116.6794 | 1.4557 | 116.8304 | 33.934 | -87.5702 | -87.616 | -0.045776 | 34.2407 | 0.3067 | 34.2041 | 34.1217 | 0.27008 | 0.18768
2648.1719 | 122.049 | -34.1995 | 87.6068 | -0.85533 | 0.71356 | 0.10657 | 0.086222 | -0.0055819 | 0.70128 | 0.012006 | 0.02005 | 0.021963 | 0.015146 | 0.0001731 | 0.014772 | -0.0056119 | -0.0046847 | -0.059509 | 87.4649 | -34.3323 | 87.41 | -1.6708 | 82.4982 | -32.9178 | 0.004578 | 0.14648 | -0.12817 | 0.027466 | 0.16937 | 0.036621 | 0.10071 | -0.13275 | -0.14191 | -0.19684 | -0.054932 | 32.5287 | 32.6614 | 1.2817 | 116.6977 | 1.4145 | 116.8304 | 34.0713 | -87.4603 | -87.5793 | -0.11902 | 34.3689 | 0.29755 | 34.3002 | 34.2361 | 0.22888 | 0.1648
2649.1719 | 122.0581 | -34.1675 | 87.5748 | -0.84007 | 0.71185 | 0.10379 | 0.087291 | 0.0016813 | 0.6985 | 0.012219 | -0.080658 | 0.023244 | 0.010446 | 0.0020957 | 0.015199 | -0.0041165 | -0.0059665 | -0.064087 | 87.5336 | -34.2773 | 87.442 | -1.6205 | 82.5531 | -32.8033 | -0.050354 | 0.13275 | -0.11902 | 0.068665 | 0.17395 | 0.022888 | 0.13275 | -0.10986 | -0.041199 | -0.13275 | -0.091552 | 32.547 | 32.6569 | 1.3641 | 116.7206 | 1.474 | 116.8304 | 34.0485 | -87.442 | -87.5061 | -0.064086 | 34.3414 | 0.29297 | 34.3002 | 34.1904 | 0.25177 | 0.14191
2650.1875 | 121.9986 | -34.1721 | 87.6022 | -0.74852 | 0.70886 | 0.10422 | 0.083873 | 0.0025357 | 0.69658 | 0.0098694 | 0.0078429 | 0.020681 | 0.011087 | -0.0028177 | 0.015413 | -0.0066799 | -0.0044711 | 0.041199 | 87.5748 | -34.227 | 87.4557 | -1.6251 | 82.6172 | -32.9361 | 0.082397 | 0.16937 | -0.13275 | 0 | 0.11902 | -0.009155 | 0.059509 | -0.054932 | -0.027466 | -0.14648 | -0.11902 | 32.547 | 32.6019 | 1.236 | 116.7892 | 1.2909 | 116.8442 | 34.0393 | -87.4329 | -87.6022 | -0.16937 | 34.2911 | 0.25177 | 34.2316 | 34.1629 | 0.19226 | 0.1236
2651.1719 | 122.049 | -34.1675 | 87.5793 | -0.86296 | 0.70608 | 0.10572 | 0.085368 | 0.0027493 | 0.6953 | 0.007306 | -0.010467 | 0.021749 | 0.012155 | -0.0055947 | 0.015413 | -0.0049709 | -0.0029757 | -0.050354 | 87.5336 | -34.2499 | 87.4512 | -1.5884 | 82.6035 | -32.8308 | 0.027466 | 0.14191 | -0.12817 | 0.086975 | 0.16022 | 0.032043 | 0.11444 | -0.082398 | -0.045777 | -0.12817 | -0.082397 | 32.579 | 32.6614 | 1.3367 | 116.7709 | 1.4191 | 116.8533 | 34.0393 | -87.4374 | -87.4924 | -0.054932 | 34.3277 | 0.28839 | 34.2819 | 34.1995 | 0.24262 | 0.16022
2652.1719 | 122.0581 | -34.1904 | 87.5244 | -0.93163 | 0.70416 | 0.10337 | 0.087077 | -2.7692e-05 | 0.69444 | 0.0068786 | -0.016571 | 0.022176 | 0.013437 | 0.0008139 | 0.015626 | -0.0047573 | -0.0044711 | 0.027466 | 87.5336 | -34.2316 | 87.4969 | -1.5701 | 82.6401 | -32.753 | 0.013733 | 0.13733 | -0.13275 | 0.068665 | 0.17395 | 0.018311 | 0.12817 | -0.041198 | 0.009155 | -0.027466 | -0.036621 | 32.6202 | 32.6614 | 1.4374 | 116.8304 | 1.4786 | 116.8716 | 34.0576 | -87.3871 | -87.4557 | -0.068664 | 34.3643 | 0.3067 | 34.3185 | 34.2087 | 0.26093 | 0.15106
2653.1875 | 122.0444 | -34.2773 | 87.4924 | -0.80193 | 0.70587 | 0.1055 | 0.085154 | -0.0015231 | 0.69231 | 0.011578 | 0.0017395 | 0.019826 | 0.01365 | -0.0051675 | 0.01584 | -0.0066799 | -0.0031895 | 0.064087 | 87.4054 | -34.3781 | 87.323 | -1.593 | 82.5073 | -33.1604 | 0.11444 | 0.14648 | -0.11444 | 0.032043 | 0.16937 | 0.050354 | 0.10071 | -0.10071 | -0.086975 | -0.16937 | -0.082398 | 32.6843 | 32.785 | 1.1169 | 116.7847 | 1.2177 | 116.8854 | 34.1629 | -87.3459 | -87.4603 | -0.11444 | 34.4467 | 0.28381 | 34.3781 | 34.3277 | 0.21515 | 0.16479
2654.1719 | 122.0627 | -34.0942 | 87.6755 | -0.89348 | 0.70544 | 0.10572 | 0.087718 | 0.0031765 | 0.69252 | 0.010724 | -0.016571 | 0.021963 | 0.011514 | -0.0034585 | 0.012849 | -0.0028347 | -0.00041228 | -0.022888 | 87.5381 | -34.2133 | 87.5015 | -1.4236 | 82.782 | -32.8308 | 0.027466 | 0.10071 | -0.14191 | 0.064087 | 0.17853 | 0.050354 | 0.13275 | -0.11902 | -0.13733 | -0.17395 | -0.036621 | 32.6706 | 32.7896 | 1.2634 | 116.8762 | 1.3824 | 116.9952 | 33.9523 | -87.5748 | -87.6114 | -0.036621 | 34.2728 | 0.32044 | 34.227 | 34.1446 | 0.27466 | 0.19226
2655.1719 | 122.0673 | -34.2087 | 87.5198 | -0.87822 | 0.70181 | 0.10657 | 0.086436 | -0.00066869 | 0.69038 | 0.0092286 | 0.047516 | 0.021963 | 0.011942 | 0.0044455 | 0.016481 | -0.0047573 | -0.0027621 | -0.032043 | 87.4512 | -34.3094 | 87.3871 | -1.5884 | 82.3059 | -32.9544 | 0 | 0.086975 | -0.14191 | 0.073242 | 0.16479 | 0.032043 | 0.1236 | -0.10071 | -0.068664 | -0.13275 | -0.064087 | 32.6202 | 32.7209 | 1.2543 | 116.5146 | 1.355 | 116.6153 | 34.0668 | -87.4329 | -87.4466 | -0.013733 | 34.3735 | 0.3067 | 34.3323 | 34.2407 | 0.2655 | 0.17395
2656.1719 | 122.0627 | -34.2407 | 87.4924 | -0.80956 | 0.70032 | 0.10379 | 0.089 | -2.7692e-05 | 0.6906 | 0.01051 | -0.040985 | 0.022176 | 0.013437 | 0.0012411 | 0.014986 | -0.0039029 | -0.00019868 | -0.082397 | 87.4466 | -34.3369 | 87.3276 | -1.6067 | 82.5623 | -32.9361 | -0.036621 | 0.14191 | -0.13733 | 0.068665 | 0.17395 | 0.022888 | 0.13275 | -0.09613 | -0.045777 | -0.1648 | -0.11902 | 32.634 | 32.7301 | 1.3046 | 116.803 | 1.4008 | 116.8991 | 34.1034 | -87.3505 | -87.4237 | -0.073242 | 34.4147 | 0.31128 | 34.3735 | 34.2636 | 0.27008 | 0.16022
2657.1719 | 122.0764 | -34.2865 | 87.4466 | -0.82481 | 0.69968 | 0.10465 | 0.087504 | 0.00061311 | 0.6859 | 0.0096558 | -0.071503 | 0.024312 | 0.013223 | 0.0003867 | 0.01584 | -0.0041165 | -0.0016941 | -0.10986 | 87.3917 | -34.3964 | 87.2681 | -1.593 | 82.5119 | -33.0048 | -0.036621 | 0.14191 | -0.13275 | 0.091553 | 0.15564 | 0.036621 | 0.13733 | -0.10986 | -0.054931 | -0.17853 | -0.1236 | 32.6935 | 32.8033 | 1.2817 | 116.7984 | 1.3916 | 116.9083 | 34.1537 | -87.3047 | -87.355 | -0.050354 | 34.4421 | 0.28839 | 34.4238 | 34.3231 | 0.27008 | 0.16937
2658.1719 | 121.9986 | -34.1583 | 87.5793 | -0.8477 | 0.69519 | 0.10486 | 0.08665 | -0.00024129 | 0.68611 | 0.008374 | -0.0074159 | 0.02239 | 0.011728 | -0.0004679 | 0.015199 | -0.0066799 | -0.0023349 | -0.07782 | 87.5061 | -34.2407 | 87.4008 | -1.5244 | 82.6538 | -32.9086 | -0.059509 | 0.15564 | -0.1236 | 0.022888 | 0.14648 | 0.009155 | 0.09613 | -0.082398 | -0.073242 | -0.17853 | -0.10529 | 32.634 | 32.7164 | 1.2497 | 116.8121 | 1.3321 | 116.8945 | 34.0347 | -87.4237 | -87.5565 | -0.13275 | 34.3048 | 0.27008 | 34.2545 | 34.1675 | 0.21973 | 0.13275
2659.1719 | 122.0261 | -34.1583 | 87.5427 | -0.87059 | 0.69498 | 0.10593 | 0.085368 | 0.00018591 | 0.68419 | 0.0096558 | 0.0017395 | 0.020467 | 0.013437 | 0.0003867 | 0.01584 | -0.0053983 | -0.0031895 | -0.068665 | 87.4557 | -34.3002 | 87.4008 | -1.4786 | 82.6675 | -32.9086 | 0.013733 | 0.13275 | -0.10529 | 0.068665 | 0.16479 | 0.036621 | 0.14191 | -0.14191 | -0.086975 | -0.14191 | -0.054932 | 32.6797 | 32.8217 | 1.2497 | 116.8259 | 1.3916 | 116.9678 | 34.053 | -87.41 | -87.4741 | -0.064086 | 34.3231 | 0.27008 | 34.3002 | 34.1949 | 0.24719 | 0.14191
2660.1719 | 122.0123 | -34.2865 | 87.4374 | -0.81718 | 0.69519 | 0.10764 | 0.086222 | 0.00061311 | 0.68269 | 0.013501 | -0.013519 | 0.023244 | 0.01365 | -0.0034585 | 0.015413 | -0.0047573 | -0.0021213 | -0.027466 | 87.3688 | -34.3826 | 87.3184 | -1.5381 | 82.5714 | -32.9361 | 0.009155 | 0.15564 | -0.13275 | 0.064087 | 0.16937 | 0.050354 | 0.13275 | -0.09613 | -0.068665 | -0.11902 | -0.050354 | 32.7484 | 32.8445 | 1.3504 | 116.8579 | 1.4465 | 116.954 | 34.1537 | -87.2818 | -87.3734 | -0.091553 | 34.4559 | 0.30212 | 34.4192 | 34.3369 | 0.2655 | 0.18311
2661.1719 | 122.0352 | -34.2911 | 87.4146 | -0.80956 | 0.69177 | 0.10529 | 0.085154 | 0.00039951 | 0.6812 | 0.0088012 | 0.0017395 | 0.020894 | 0.01301 | -0.0047403 | 0.016054 | -0.0049709 | -0.0036167 | -0.064087 | 87.3413 | -34.3735 | 87.3047 | -1.5244 | 82.5623 | -32.9361 | 0.013733 | 0.13275 | -0.1236 | 0.073242 | 0.16937 | 0.045776 | 0.14191 | -0.082397 | -0.073242 | -0.10986 | -0.036622 | 32.7667 | 32.8491 | 1.355 | 116.8533 | 1.4374 | 116.9357 | 34.1675 | -87.2818 | -87.3413 | -0.059509 | 34.4605 | 0.29297 | 34.433 | 34.3369 | 0.2655 | 0.16937
2662.1719 | 122.0261 | -34.3643 | 87.355 | -0.80956 | 0.69198 | 0.10465 | 0.089 | 0.0036039 | 0.67885 | 0.012433 | -0.016571 | 0.021963 | 0.01301 | 0.0014547 | 0.012849 | -0.0051845 | -0.0012667 | -0.027466 | 87.3047 | -34.4284 | 87.2543 | -1.5106 | 82.5531 | -33.0734 | 0.004578 | 0.15106 | -0.1236 | 0.064087 | 0.16479 | 0.054932 | 0.13275 | -0.064087 | -0.050355 | -0.10071 | -0.050354 | 32.8537 | 32.9178 | 1.2909 | 116.9174 | 1.355 | 116.9815 | 34.2407 | -87.204 | -87.291 | -0.086975 | 34.5291 | 0.28839 | 34.4971 | 34.4193 | 0.25635 | 0.17853
2663.1719 | 122.0764 | -34.3002 | 87.4191 | -0.76378 | 0.68963 | 0.10465 | 0.085795 | -0.0015231 | 0.67778 | 0.01051 | 0.038361 | 0.020254 | 0.01365 | -0.0017495 | 0.015626 | -0.0073209 | -0.0038303 | 0.004578 | 87.41 | -34.3964 | 87.3276 | -1.4694 | 82.6401 | -32.9773 | 0.11444 | 0.13733 | -0.11444 | 0.091553 | 0.20599 | 0.09613 | 0.14191 | -0.09613 | -0.009155 | -0.091552 | -0.082397 | 32.8308 | 32.9269 | 1.3229 | 116.9403 | 1.4191 | 117.0364 | 34.1858 | -87.2818 | -87.3276 | -0.045776 | 34.5062 | 0.32044 | 34.4421 | 34.3964 | 0.25635 | 0.21057
2664.1719 | 122.0444 | -34.2819 | 87.442 | -0.80193 | 0.68857 | 0.10443 | 0.088145 | 0.0029629 | 0.67479 | 0.0094422 | -0.065399 | 0.02004 | 0.011728 | -0.0015359 | 0.013277 | -0.0036893 | -0.0083163 | -0.036621 | 87.3688 | -34.3689 | 87.3322 | -1.442 | 82.6675 | -32.9956 | -0.022888 | 0.13733 | -0.11902 | 0.091553 | 0.17853 | 0.041199 | 0.13275 | -0.086975 | -0.073243 | -0.10986 | -0.036621 | 32.84 | 32.9269 | 1.2863 | 116.9495 | 1.3733 | 117.0364 | 34.1629 | -87.3047 | -87.3505 | -0.045776 | 34.4604 | 0.29755 | 34.4147 | 34.3231 | 0.25177 | 0.16022
2665.1719 | 122.0215 | -34.3643 | 87.3367 | -0.74852 | 0.68985 | 0.10507 | 0.086222 | 0.0036039 | 0.67628 | 0.0079468 | -0.028778 | 0.022176 | 0.014505 | -0.0030313 | 0.014772 | -0.0062527 | -0.0034031 | -0.009155 | 87.2269 | -34.5062 | 87.1857 | -1.4374 | 82.5211 | -32.9956 | 0.018311 | 0.15106 | -0.1236 | 0.07782 | 0.16937 | 0.045776 | 0.13275 | -0.14191 | -0.10986 | -0.15106 | -0.041199 | 32.9269 | 33.0688 | 1.3687 | 116.8854 | 1.5106 | 117.0273 | 34.2407 | -87.1857 | -87.2589 | -0.073242 | 34.5337 | 0.29297 | 34.4971 | 34.4101 | 0.25635 | 0.16937
2666.1719 | 122.0078 | -34.4284 | 87.291 | -0.84007 | 0.68622 | 0.10614 | 0.087718 | -0.0043003 | 0.67265 | 0.0081604 | 0.0047913 | 0.022603 | 0.014719 | 0.0023093 | 0.014558 | -0.0047573 | -0.0036167 | -0.045776 | 87.2635 | -34.4696 | 87.1902 | -1.5152 | 82.5165 | -33.1009 | -0.027466 | 0.14648 | -0.091553 | 0.032043 | 0.14191 | 0.004578 | 0.091553 | -0.041198 | -0.027466 | -0.10071 | -0.073242 | 32.9132 | 32.9544 | 1.3275 | 116.9449 | 1.3687 | 116.9861 | 34.3369 | -87.1445 | -87.2589 | -0.11444 | 34.5703 | 0.23346 | 34.52 | 34.433 | 0.18311 | 0.096131
2667.1719 | 121.9986 | -34.3643 | 87.3505 | -0.80956 | 0.68686 | 0.10636 | 0.087931 | -0.0019503 | 0.6718 | 0.011792 | -0.0074159 | 0.021749 | 0.011301 | -0.0015359 | 0.014772 | -0.0060391 | 1.4916e-05 | -0.054932 | 87.3001 | -34.4467 | 87.1948 | -1.4465 | 82.5668 | -33.046 | -0.032043 | 0.15564 | -0.1236 | 0.032043 | 0.1236 | 0.018311 | 0.11902 | -0.082397 | -0.050354 | -0.15564 | -0.10529 | 32.9178 | 33.0002 | 1.3184 | 116.9312 | 1.4008 | 117.0135 | 34.2407 | -87.1948 | -87.3184 | -0.1236 | 34.4879 | 0.24719 | 34.4833 | 34.3826 | 0.24261 | 0.14191
2668.1719 | 122.0078 | -34.3506 | 87.4191 | -0.73326 | 0.68173 | 0.10614 | 0.086436 | -0.0051547 | 0.66902 | 0.0098694 | 0.013947 | 0.020254 | 0.012369 | -0.0002543 | 0.013704 | -0.0073209 | -0.0029757 | 0.036621 | 87.3276 | -34.4147 | 87.2635 | -1.4694 | 82.6538 | -32.9956 | 0.11444 | 0.14191 | -0.14648 | 0.032043 | 0.15564 | 0.045776 | 0.091553 | -0.064087 | -0.091552 | -0.15564 | -0.064087 | 32.8812 | 32.9453 | 1.355 | 117.0044 | 1.4191 | 117.0685 | 34.2041 | -87.2772 | -87.3871 | -0.10986 | 34.5062 | 0.30212 | 34.4421 | 34.3964 | 0.23804 | 0.19226
2669.1719 | 121.9574 | -34.3689 | 87.3917 | -0.82481 | 0.68237 | 0.10614 | 0.086222 | -0.0023777 | 0.66987 | 0.011578 | -0.010467 | 0.022176 | 0.011728 | -0.0026041 | 0.015199 | -0.0043301 | -0.0057529 | -0.054932 | 87.2681 | -34.4742 | 87.1994 | -1.4236 | 82.5806 | -33.0872 | 0 | 0.16022 | -0.11902 | 0.045776 | 0.15564 | 0.022888 | 0.11444 | -0.10529 | -0.1236 | -0.19226 | -0.068664 | 32.9453 | 33.0505 | 1.2817 | 116.9495 | 1.387 | 117.0547 | 34.2499 | -87.2314 | -87.3459 | -0.11444 | 34.5245 | 0.27466 | 34.4833 | 34.3918 | 0.23346 | 0.14191
2670.1719 | 122.049 | -34.3597 | 87.3734 | -0.81718 | 0.68152 | 0.107 | 0.088145 | 0.0012539 | 0.67116 | 0.011578 | 0.02005 | 0.023458 | 0.011942 | -0.0017495 | 0.014131 | -0.0051845 | -0.0048983 | -0.018311 | 87.2772 | -34.465 | 87.2681 | -1.6846 | 82.3837 | -32.9773 | -0.036621 | 0.10529 | -0.14191 | 0.09613 | 0.16937 | 0.045776 | 0.14191 | -0.10529 | -0.09613 | -0.10529 | -0.009156 | 32.6752 | 32.7805 | 1.3824 | 116.7435 | 1.4877 | 116.8488 | 34.2178 | -87.2681 | -87.2772 | -0.009156 | 34.5291 | 0.31128 | 34.5016 | 34.4055 | 0.28381 | 0.18768
2671.1719 | 122.0078 | -34.4101 | 87.2955 | -0.82481 | 0.68024 | 0.10614 | 0.0843 | 0.0018949 | 0.66603 | 0.0090148 | -0.0013123 | 0.021749 | 0.012369 | -0.0015359 | 0.015199 | -0.0066799 | -0.0023349 | -0.045776 | 87.2406 | -34.52 | 87.1857 | -1.3916 | 82.5394 | -33.1284 | 0.013733 | 0.14648 | -0.13275 | 0.054932 | 0.15106 | 0.027466 | 0.1236 | -0.10986 | -0.054931 | -0.10986 | -0.054932 | 33.0185 | 33.1284 | 1.2817 | 116.9495 | 1.3916 | 117.0593 | 34.2773 | -87.149 | -87.2406 | -0.091552 | 34.5612 | 0.28381 | 34.5337 | 34.4376 | 0.25635 | 0.16022
2672.1719 | 122.0123 | -34.4559 | 87.2452 | -0.83244 | 0.67682 | 0.107 | 0.087504 | 0.0033901 | 0.66475 | 0.0081604 | -0.022675 | 0.021108 | 0.010873 | -0.0017495 | 0.013917 | -0.0041165 | -0.0019077 | -0.036621 | 87.1719 | -34.5703 | 87.1078 | -1.3824 | 82.5165 | -33.1833 | 0.018311 | 0.13733 | -0.11902 | 0.064087 | 0.15564 | 0.027466 | 0.13275 | -0.11444 | -0.073242 | -0.13733 | -0.064087 | 33.0734 | 33.1879 | 1.2726 | 116.9724 | 1.387 | 117.0868 | 34.3369 | -87.1078 | -87.1811 | -0.073242 | 34.6115 | 0.27466 | 34.5886 | 34.4833 | 0.25177 | 0.14649
2673.1719 | 122.0444 | -34.3231 | 87.4008 | -0.77904 | 0.67767 | 0.10593 | 0.086436 | -0.00045509 | 0.66325 | 0.011365 | -0.010467 | 0.023031 | 0.012155 | -0.0032449 | 0.012636 | -0.0049709 | 0.00065592 | -0.050354 | 87.323 | -34.4376 | 87.2681 | -1.4328 | 82.6721 | -33.0048 | 0.12817 | 0.15106 | -0.11902 | 0.064087 | 0.17853 | 0.07782 | 0.12817 | -0.11444 | -0.07782 | -0.13275 | -0.054932 | 32.8903 | 33.0048 | 1.3184 | 116.9952 | 1.4328 | 117.1097 | 34.2041 | -87.2498 | -87.3367 | -0.086975 | 34.5016 | 0.29755 | 34.4513 | 34.4009 | 0.24719 | 0.19684
2674.1719 | 121.994 | -34.4284 | 87.3138 | -0.84007 | 0.67532 | 0.1055 | 0.085795 | -0.0015231 | 0.66432 | 0.0088012 | 0.0078429 | 0.023031 | 0.01365 | -0.0030313 | 0.014986 | -0.0051845 | 0.0010831 | -0.068665 | 87.2589 | -34.4833 | 87.149 | -1.4511 | 82.5714 | -33.1192 | -0.054932 | 0.16022 | -0.11444 | 0.022888 | 0.14648 | 0.018311 | 0.091553 | -0.054931 | -0.054932 | -0.1648 | -0.10986 | 32.9773 | 33.0322 | 1.3092 | 116.9998 | 1.3641 | 117.0547 | 34.314 | -87.1536 | -87.291 | -0.13733 | 34.5749 | 0.26093 | 34.52 | 34.4467 | 0.20599 | 0.13275
2675.1719 | 122.0169 | -34.52 | 87.1902 | -0.82481 | 0.67468 | 0.10443 | 0.087718 | -0.00045509 | 0.66347 | 0.0098694 | -0.013519 | 0.023244 | 0.010446 | 0.0035911 | 0.013917 | -0.0053983 | 0.00044212 | -0.07782 | 87.0941 | -34.6527 | 87.0438 | -1.4374 | 82.4387 | -33.4167 | -0.050354 | 0.16479 | -0.10986 | 0.022888 | 0.13275 | 0.018311 | 0.10986 | -0.13275 | -0.096131 | -0.14648 | -0.050354 | 33.0826 | 33.2153 | 1.1032 | 116.9586 | 1.236 | 117.0914 | 34.4101 | -87.0255 | -87.1674 | -0.14191 | 34.6527 | 0.24261 | 34.6298 | 34.5383 | 0.21973 | 0.12817
2676.1719 | 122.0032 | -34.5154 | 87.2223 | -0.718 | 0.67276 | 0.10678 | 0.085795 | 0.0027493 | 0.66026 | 0.0085876 | -0.0013123 | 0.022603 | 0.010019 | 0.0003867 | 0.015199 | -0.0039029 | -0.0019077 | -0.073242 | 87.1811 | -34.5749 | 87.0438 | -1.4511 | 82.4707 | -33.2245 | -0.022888 | 0.15106 | -0.11444 | 0.027466 | 0.13733 | 0.018311 | 0.09613 | -0.059509 | -0.041199 | -0.17853 | -0.13733 | 33.0643 | 33.1238 | 1.2909 | 116.9861 | 1.3504 | 117.0456 | 34.4009 | -87.0712 | -87.1948 | -0.1236 | 34.6527 | 0.25177 | 34.6115 | 34.5337 | 0.21057 | 0.13275
2677.1719 | 122.0123 | -34.4421 | 87.2681 | -0.88585 | 0.67276 | 0.10636 | 0.087718 | 0.0027493 | 0.6609 | 0.01051 | -0.022675 | 0.023671 | 0.012369 | 0.0085043 | 0.015199 | -0.0049709 | -0.0083163 | -0.013733 | 87.1994 | -34.5474 | 87.1307 | -1.387 | 82.5851 | -33.1284 | 0.022888 | 0.13275 | -0.09613 | 0.064087 | 0.1831 | 0.032043 | 0.13733 | -0.10529 | -0.068664 | -0.13733 | -0.068665 | 33.0551 | 33.1604 | 1.3138 | 117.0273 | 1.4191 | 117.1326 | 34.346 | -87.1353 | -87.204 | -0.068664 | 34.6252 | 0.27924 | 34.5795 | 34.4742 | 0.23346 | 0.12817
2678.1719 | 121.9711 | -34.4055 | 87.3047 | -0.81718 | 0.67105 | 0.10593 | 0.088359 | 0.00082671 | 0.65941 | 0.0090148 | -0.013519 | 0.021108 | 0.011087 | -0.0049539 | 0.012849 | -0.0041165 | -0.0040439 | -0.041199 | 87.1811 | -34.5474 | 87.1399 | -1.4786 | 82.5943 | -33.165 | -0.027466 | 0.15564 | -0.10986 | 0.022888 | 0.13733 | -0.004578 | 0.09613 | -0.14191 | -0.1236 | -0.16479 | -0.041198 | 32.9269 | 33.0688 | 1.2405 | 116.9998 | 1.3824 | 117.1417 | 34.2957 | -87.149 | -87.2818 | -0.13275 | 34.5428 | 0.24719 | 34.5016 | 34.4009 | 0.20599 | 0.10529
2679.1719 | 121.9711 | -34.4925 | 87.2314 | -0.80193 | 0.67148 | 0.1055 | 0.090281 | 0.0014675 | 0.65748 | 0.01051 | -0.010467 | 0.022817 | 0.010873 | 0.0038047 | 0.014345 | -0.0053983 | -0.0021213 | -0.036621 | 87.1719 | -34.5657 | 87.0621 | -1.3962 | 82.5027 | -33.165 | -0.022888 | 0.13733 | -0.10529 | 0.036621 | 0.13733 | 0.032043 | 0.091553 | -0.073242 | -0.059509 | -0.16937 | -0.10986 | 33.0963 | 33.1696 | 1.3275 | 116.9952 | 1.4008 | 117.0685 | 34.3872 | -87.0941 | -87.1948 | -0.10071 | 34.6298 | 0.24262 | 34.584 | 34.5245 | 0.19684 | 0.13733
2680.1719 | 121.9757 | -34.5337 | 87.1582 | -0.76378 | 0.67041 | 0.10572 | 0.087504 | -0.0032321 | 0.65556 | 0.0081604 | -0.016571 | 0.020894 | 0.012155 | -0.0043131 | 0.016481 | -0.0060391 | -0.0044711 | -0.086975 | 87.1216 | -34.6481 | 86.9888 | -1.5106 | 82.5119 | -33.2886 | -0.032043 | 0.14191 | -0.12817 | 0.022888 | 0.13733 | 0 | 0.068665 | -0.11444 | -0.036621 | -0.16937 | -0.13275 | 33.0231 | 33.1375 | 1.2451 | 117.0456 | 1.3596 | 117.16 | 34.4055 | -87.0163 | -87.1353 | -0.11902 | 34.671 | 0.2655 | 34.6024 | 34.5337 | 0.19684 | 0.12817
2681.1719 | 121.9894 | -34.4421 | 87.2543 | -0.7943 | 0.66656 | 0.10678 | 0.087931 | 0.00018591 | 0.65343 | 0.007306 | -0.0074159 | 0.023244 | 0.014078 | 0.0033773 | 0.017335 | -0.0049709 | -0.0070347 | -0.10529 | 87.2955 | -34.4879 | 87.0804 | -1.4374 | 82.5439 | -33.1604 | -0.041199 | 0.14191 | -0.12817 | 0.018311 | 0.13733 | 0.004578 | 0.10986 | -0.045776 | 0.041199 | -0.17395 | -0.21515 | 33.0048 | 33.0505 | 1.2817 | 116.9861 | 1.3275 | 117.0319 | 34.314 | -87.1124 | -87.236 | -0.1236 | 34.5795 | 0.2655 | 34.552 | 34.4467 | 0.23804 | 0.13275
2682.1719 | 121.9711 | -34.4009 | 87.3093 | -0.80956 | 0.66656 | 0.10572 | 0.088359 | -0.0021641 | 0.65428 | 0.010297 | 0.010895 | 0.022603 | 0.011942 | -0.0049539 | 0.013917 | -0.0062527 | -0.0055393 | -0.059509 | 87.2498 | -34.5016 | 87.1307 | -1.4465 | 82.5714 | -33.1604 | -0.032043 | 0.14648 | -0.12817 | 0.022888 | 0.14648 | 0 | 0.091553 | -0.10071 | -0.059509 | -0.17853 | -0.11902 | 32.9544 | 33.0551 | 1.2405 | 116.9724 | 1.3412 | 117.0731 | 34.2728 | -87.1628 | -87.2864 | -0.1236 | 34.5474 | 0.27466 | 34.4925 | 34.4009 | 0.21973 | 0.12817
2683.1719 | 122.0444 | -34.4513 | 87.236 | -0.7943 | 0.6672 | 0.10465 | 0.086009 | 0.0044583 | 0.65214 | 0.0079468 | 0.035309 | 0.021108 | 0.013437 | 0.0033773 | 0.014986 | -0.0053983 | -0.0057529 | 0.004578 | 87.2269 | -34.5062 | 87.1307 | -1.4099 | 82.6401 | -33.1421 | 0.16022 | 0.13275 | -0.10529 | 0.059509 | 0.20142 | 0.082397 | 0.13733 | -0.054932 | -0.009155 | -0.10529 | -0.096131 | 33.0414 | 33.0963 | 1.3092 | 117.0914 | 1.3641 | 117.1463 | 34.346 | -87.1033 | -87.1765 | -0.073242 | 34.6527 | 0.3067 | 34.5886 | 34.5337 | 0.24262 | 0.18768
2684.1719 | 121.9894 | -34.3552 | 87.3138 | -0.8477 | 0.66507 | 0.10636 | 0.085582 | -0.0030185 | 0.65428 | 0.0085876 | -0.03183 | 0.023885 | 0.011728 | 0.0010275 | 0.013063 | -0.0045437 | -0.00019868 | -0.036621 | 87.2681 | -34.4604 | 87.204 | -1.3824 | 82.6813 | -33.1558 | 0.018311 | 0.15106 | -0.10529 | 0.064087 | 0.16937 | 0.036621 | 0.14191 | -0.10528 | -0.045777 | -0.10986 | -0.064087 | 32.9727 | 33.078 | 1.1993 | 117.0364 | 1.3046 | 117.1417 | 34.2499 | -87.1628 | -87.2498 | -0.086975 | 34.5245 | 0.27466 | 34.4971 | 34.3918 | 0.24719 | 0.14191
2685.1719 | 121.962 | -34.4101 | 87.2818 | -0.77904 | 0.66443 | 0.10443 | 0.085154 | -0.0049411 | 0.65257 | 0.0092286 | -0.03183 | 0.023031 | 0.011728 | 0.0042319 | 0.016267 | -0.0049709 | -0.0036167 | -0.032043 | 87.2498 | -34.4833 | 87.149 | -1.4099 | 82.6492 | -33.1329 | -0.045776 | 0.15106 | -0.11902 | 0.018311 | 0.13733 | 0.018311 | 0.10071 | -0.073242 | -0.032043 | -0.13275 | -0.10071 | 33.0002 | 33.0734 | 1.2772 | 117.0593 | 1.3504 | 117.1326 | 34.2911 | -87.1307 | -87.2635 | -0.13275 | 34.5474 | 0.25635 | 34.5108 | 34.4284 | 0.21973 | 0.13733
2686.1719 | 121.9803 | -34.4696 | 87.2818 | -0.80193 | 0.66357 | 0.10572 | 0.088145 | 0.0025357 | 0.64979 | 0.0088012 | -0.03183 | 0.02239 | 0.011942 | -0.0019631 | 0.01349 | -0.0045437 | -0.0023349 | -0.082397 | 87.1948 | -34.5154 | 87.1033 | -1.4191 | 82.5989 | -33.1924 | -0.041199 | 0.15106 | -0.10986 | 0.018311 | 0.13733 | 0.009155 | 0.068665 | -0.045777 | -0.086975 | -0.17853 | -0.091553 | 33.0505 | 33.0963 | 1.2772 | 117.0685 | 1.3229 | 117.1143 | 34.3597 | -87.1307 | -87.2635 | -0.13275 | 34.6069 | 0.24719 | 34.5383 | 34.4788 | 0.17853 | 0.11902
2687.1719 | 121.962 | -34.5291 | 87.2131 | -0.80193 | 0.66315 | 0.10422 | 0.088572 | 0.00039951 | 0.64787 | 0.010724 | -0.010467 | 0.021322 | 0.013864 | -0.0015359 | 0.014772 | -0.0036893 | -0.0038303 | -0.082397 | 87.1857 | -34.552 | 87.0666 | -1.4328 | 82.5531 | -33.1558 | -0.027466 | 0.16022 | -0.13275 | -0.28839 | 0 | -0.10986 | 0.045776 | -0.022888 | -0.027466 | -0.14648 | -0.11902 | 33.0963 | 33.1192 | 1.3733 | 117.0822 | 1.3962 | 117.1051 | 34.3964 | -87.0529 | -87.5015 | -0.44861 | 34.5291 | 0.13275 | 34.5749 | 34.4193 | 0.17853 | 0.022888
2688.1719 | 121.9025 | -34.433 | 87.3459 | -0.77141 | 0.6608 | 0.10401 | 0.088572 | -0.00045509 | 0.64616 | 0.0081604 | -0.013519 | 0.021108 | 0.014078 | -0.0081583 | 0.013277 | -0.0043301 | -0.0046847 | 0.086975 | 87.323 | -34.465 | 87.1124 | -1.4511 | 82.6492 | -33.4396 | 0.022888 | 0.17853 | -0.15106 | -0.045776 | 0.073242 | -0.059509 | 0 | -0.032044 | -0.022888 | -0.23346 | -0.21057 | 32.9819 | 33.0139 | 0.99335 | 117.0822 | 1.0254 | 117.1143 | 34.2819 | -87.1674 | -87.3917 | -0.2243 | 34.5062 | 0.2243 | 34.433 | 34.3735 | 0.15106 | 0.091553
2689.1719 | 121.9482 | -34.4696 | 87.236 | -0.77141 | 0.6593 | 0.10358 | 0.087077 | -0.0030185 | 0.64638 | 0.0081604 | 0.0017395 | 0.021322 | 0.011514 | -0.0008951 | 0.013917 | -0.0066799 | -0.0042575 | -0.082397 | 87.1399 | -34.6069 | 87.0438 | -1.4236 | 82.5623 | -33.0963 | -0.050354 | 0.15106 | -0.12817 | 0.013733 | 0.12817 | 0.009155 | 0.09613 | -0.13733 | -0.09613 | -0.19226 | -0.096131 | 33.046 | 33.1833 | 1.3733 | 117.0319 | 1.5106 | 117.1692 | 34.3414 | -87.085 | -87.2223 | -0.13733 | 34.5978 | 0.25635 | 34.5657 | 34.4788 | 0.2243 | 0.13733
2690.1719 | 121.9711 | -34.5657 | 87.1216 | -0.84007 | 0.66293 | 0.10721 | 0.087718 | 0.0029629 | 0.64702 | 0.010083 | -0.040985 | 0.02239 | 0.0098052 | -0.0043131 | 0.014558 | -0.0062527 | 0.00065592 | -0.045776 | 87.0758 | -34.6664 | 86.9476 | -1.4557 | 82.4615 | -33.2657 | 0.009155 | 0.14191 | -0.10071 | 0.050354 | 0.14648 | 0.027466 | 0.10986 | -0.10071 | -0.045776 | -0.17395 | -0.12817 | 33.11 | 33.2108 | 1.3 | 117.0273 | 1.4008 | 117.128 | 34.465 | -86.9797 | -87.0712 | -0.091553 | 34.7122 | 0.24719 | 34.6756 | 34.5932 | 0.21057 | 0.12817
2691.1719 | 121.9849 | -34.6527 | 87.0209 | -0.78667 | 0.65909 | 0.10529 | 0.088572 | 0.0038175 | 0.64403 | 0.0088012 | -0.0013123 | 0.023458 | 0.010873 | 0.0042319 | 0.013917 | -0.0062527 | -0.0040439 | -0.064087 | 86.9797 | -34.7626 | 86.8881 | -1.4099 | 82.37 | -33.2794 | 0.013733 | 0.14191 | -0.1236 | 0.036621 | 0.14191 | 0.050354 | 0.11444 | -0.10986 | -0.041199 | -0.13275 | -0.091552 | 33.2428 | 33.3527 | 1.3733 | 117.0227 | 1.4832 | 117.1326 | 34.5291 | -86.879 | -86.9843 | -0.10529 | 34.7946 | 0.2655 | 34.7672 | 34.7031 | 0.23804 | 0.17395
2692.1719 | 121.9711 | -34.8175 | 86.9888 | -0.7943 | 0.65909 | 0.10849 | 0.087291 | 0.00061311 | 0.64403 | 0.0098694 | -0.025726 | 0.021963 | 0.011728 | -0.0026041 | 0.013277 | -0.0056119 | -0.0023349 | -0.082397 | 87.0804 | -34.6573 | 86.6547 | -1.4557 | 82.196 | -33.4808 | -0.059509 | 0.14648 | -0.1236 | 0.013733 | 0.13275 | 0.022888 | 0.07782 | 0.16022 | 0.091552 | -0.33417 | -0.42572 | 33.3618 | 33.2016 | 1.3367 | 117.0135 | 1.1765 | 116.8533 | 34.6939 | -86.8423 | -86.9751 | -0.13275 | 34.9503 | 0.25635 | 34.8953 | 34.8404 | 0.20142 | 0.14648
2693.1719 | 122.0398 | -34.5978 | 87.0758 | -0.73326 | 0.65546 | 0.10507 | 0.085582 | -0.0038729 | 0.6436 | 0.0068786 | 0.02005 | 0.020894 | 0.011301 | 0.0038047 | 0.014558 | -0.0062527 | -0.0053257 | -0.018311 | 87.0392 | -34.6985 | 86.9934 | -1.3458 | 82.5348 | -33.3252 | 0.14648 | 0.13733 | -0.11444 | 0.059509 | 0.20142 | 0.07782 | 0.11902 | -0.10071 | -0.036621 | -0.082398 | -0.045777 | 33.252 | 33.3527 | 1.2726 | 117.1326 | 1.3733 | 117.2333 | 34.4833 | -86.9385 | -87.0163 | -0.07782 | 34.7992 | 0.31586 | 34.7168 | 34.6756 | 0.23346 | 0.19226
2694.1719 | 121.9711 | -34.6115 | 87.0941 | -0.78667 | 0.65481 | 0.10337 | 0.086009 | 0.0012539 | 0.64232 | 0.0077332 | 0.0017395 | 0.020254 | 0.010233 | 0.0001731 | 0.014558 | -0.0047573 | -0.00019868 | -0.082397 | 87.0117 | -34.7305 | 86.8973 | -1.4374 | 82.4661 | -33.4076 | -0.27924 | 0.13275 | -0.16022 | -0.013733 | 0.12817 | 0.009155 | 0.07782 | -0.11902 | -0.082397 | -0.19684 | -0.11444 | 33.1741 | 33.2932 | 1.2039 | 117.0776 | 1.3229 | 117.1967 | 34.4513 | -86.9614 | -87.1078 | -0.14648 | 34.7397 | 0.28839 | 34.6893 | 34.6207 | 0.23804 | 0.16937
2695.1719 | 121.9574 | -34.5749 | 87.1262 | -0.77904 | 0.65353 | 0.10657 | 0.085795 | -0.0015231 | 0.64082 | 0.0079468 | 0.029205 | 0.022176 | 0.012369 | -0.0013223 | 0.016695 | -0.0039029 | -0.0036167 | -0.073242 | 87.1262 | -34.6436 | 86.9934 | -1.4511 | 82.5302 | -33.284 | -0.018311 | 0.15564 | -0.19226 | -0.032043 | 0.13275 | -0.013733 | 0.064087 | -0.068665 | 0 | -0.13275 | -0.13275 | 33.1238 | 33.1924 | 1.2909 | 117.1051 | 1.3596 | 117.1738 | 34.3826 | -86.9705 | -87.1582 | -0.18768 | 34.7076 | 0.32501 | 34.639 | 34.5612 | 0.25635 | 0.17853
2696.1719 | 122.0215 | -34.5108 | 87.1674 | -0.70274 | 0.65417 | 0.10764 | 0.086863 | -0.0040867 | 0.63954 | 0.0098694 | 0.16348 | 0.020894 | 0.013437 | -0.0008951 | 0.014345 | -0.0066799 | -0.0034031 | 0.036621 | 87.0987 | -34.6298 | 87.0346 | -1.4236 | 82.6355 | -33.1512 | 0.027466 | 0.14191 | -0.15564 | 0.068665 | 0.1831 | 0.068665 | 0.1236 | -0.11902 | -0.068664 | -0.13275 | -0.064087 | 33.0872 | 33.2062 | 1.3596 | 117.1463 | 1.4786 | 117.2653 | 34.3552 | -87.0255 | -87.0987 | -0.073242 | 34.6939 | 0.33874 | 34.6344 | 34.5795 | 0.27924 | 0.22431
2697.1719 | 121.962 | -34.6207 | 87.0712 | -0.7943 | 0.6514 | 0.10593 | 0.084941 | 0.0010403 | 0.64104 | 0.01051 | -0.03183 | 0.022603 | 0.011514 | 0.0003867 | 0.015199 | -0.0066799 | -0.0012667 | -0.050354 | 87.0255 | -34.6939 | 86.911 | -1.4053 | 82.457 | -33.2474 | -0.032043 | 0.15106 | -0.12817 | 0.013733 | 0.11902 | 0.018311 | 0.10986 | -0.073242 | -0.045776 | -0.16022 | -0.11444 | 33.2153 | 33.2886 | 1.3733 | 117.0776 | 1.4465 | 117.1509 | 34.4925 | -86.9202 | -87.0575 | -0.13733 | 34.7397 | 0.24719 | 34.7305 | 34.639 | 0.23804 | 0.14649
2698.1719 | 121.9986 | -34.5657 | 87.0941 | -0.73326 | 0.65054 | 0.10678 | 0.086436 | -0.0023777 | 0.63698 | 0.0085876 | 0.0047913 | 0.019826 | 0.01301 | -0.0034585 | 0.015413 | -0.0043301 | 0.00065592 | 0 | 87.0483 | -34.6664 | 86.9659 | -1.355 | 82.5348 | -33.1696 | 0.11902 | 0.16022 | -0.11444 | 0.07782 | 0.2243 | 0.091553 | 0.14648 | -0.10071 | -0.045776 | -0.12817 | -0.082398 | 33.2108 | 33.3115 | 1.3962 | 117.1005 | 1.4969 | 117.2012 | 34.4513 | -86.9339 | -87.0163 | -0.082397 | 34.79 | 0.33875 | 34.7122 | 34.6573 | 0.26093 | 0.20599
2699.1719 | 121.9299 | -34.3781 | 87.3367 | -0.74089 | 0.64926 | 0.10486 | 0.088786 | -0.0028049 | 0.63804 | 0.010297 | 0.010895 | 0.021322 | 0.012582 | 0.0050863 | 0.012208 | -0.0056119 | -0.0053257 | -0.064087 | 87.2818 | -34.433 | 87.1628 | -1.4236 | 82.7316 | -33.0734 | -0.036621 | 0.14191 | -0.1236 | -0.022888 | 0.11444 | 0.004578 | 0.086975 | -0.054931 | -0.054932 | -0.17395 | -0.11902 | 32.9544 | 33.0093 | 1.3046 | 117.1097 | 1.3596 | 117.1646 | 34.2545 | -87.1948 | -87.3596 | -0.16479 | 34.4925 | 0.23804 | 34.465 | 34.3826 | 0.21057 | 0.12817
2700.1719 | 121.9345 | -34.5291 | 87.1994 | -0.7943 | 0.6499 | 0.10486 | 0.085582 | 0.0025357 | 0.63591 | 0.012433 | -0.013519 | 0.022603 | 0.013864 | -0.0021769 | 0.013917 | -0.0058255 | -0.0048983 | -0.027466 | 87.1811 | -34.5657 | 87.0392 | -1.4328 | 82.6263 | -33.2062 | -0.041199 | 0.14648 | -0.11902 | -0.018311 | 0.10071 | -0.009155 | 0.09613 | -0.036621 | -0.018311 | -0.16022 | -0.14191 | 33.0963 | 33.1329 | 1.3229 | 117.1555 | 1.3596 | 117.1921 | 34.4101 | -87.0529 | -87.2177 | -0.16479 | 34.6298 | 0.21973 | 34.6252 | 34.52 | 0.21515 | 0.10986
2701.1875 | 121.875 | -34.3323 | 87.442 | -0.72563 | 0.64841 | 0.1055 | 0.089 | 0.0036039 | 0.6342 | 0.012219 | -0.13254 | 0.023671 | 0.012155 | 0.0072227 | 0.014345 | -0.010953 | -0.00062588 | -0.022888 | 87.3459 | -34.4238 | 87.1399 | -1.5106 | 82.695 | -33.046 | -0.054932 | 0.17853 | -0.16479 | -0.050354 | 0.027466 | -0.10986 | 0.009155 | -0.091553 | -0.096131 | -0.30212 | -0.20599 | 32.8217 | 32.9132 | 1.2863 | 117.0273 | 1.3779 | 117.1188 | 34.1675 | -87.2635 | -87.4924 | -0.22888 | 34.3597 | 0.19226 | 34.3414 | 34.2224 | 0.17395 | 0.054932
2702.1719 | 121.9986 | -34.3094 | 87.3505 | -0.76378 | 0.64606 | 0.10657 | 0.084513 | 0.0048855 | 0.6342 | 0.0085876 | 0.010895 | 0.020681 | 0.011942 | -0.0045267 | 0.015413 | -0.0051845 | -0.00083948 | -0.07782 | 87.3276 | -34.4101 | 87.1948 | -1.4008 | 82.8415 | -33.0643 | 0.022888 | 0.15106 | -0.11902 | 0.050354 | 0.18768 | 0.07782 | 0.1236 | -0.10071 | -0.022888 | -0.15564 | -0.13275 | 32.9086 | 33.0093 | 1.2451 | 117.1509 | 1.3458 | 117.2516 | 34.1904 | -87.1994 | -87.3001 | -0.10071 | 34.4971 | 0.3067 | 34.433 | 34.3872 | 0.24261 | 0.19684
2703.1719 | 121.907 | -34.4833 | 87.2452 | -0.70274 | 0.64648 | 0.10507 | 0.090068 | 0.0080899 | 0.63313 | 0.011792 | -0.11728 | 0.023458 | 0.012582 | 0.010213 | 0.015199 | -0.0092435 | -0.0038303 | 0.018311 | 87.1948 | -34.5749 | 87.0987 | -1.4191 | 82.6401 | -33.197 | 0.032043 | 0.21973 | -0.064087 | 0.004578 | 0.082397 | -0.041199 | 0.054932 | -0.091553 | -0.050354 | -0.14648 | -0.09613 | 33.0643 | 33.1558 | 1.2863 | 117.1234 | 1.3779 | 117.215 | 34.4192 | -87.0255 | -87.2406 | -0.21515 | 34.5657 | 0.14648 | 34.5383 | 34.4421 | 0.11902 | 0.022888
2704.1719 | 121.9666 | -34.3781 | 87.3413 | -0.75615 | 0.6467 | 0.10529 | 0.084941 | 0.0014675 | 0.63142 | 0.007306 | 0.029205 | 0.019826 | 0.01301 | -0.0087991 | 0.01349 | -0.0068935 | -0.0010531 | -0.041199 | 87.3413 | -34.3826 | 87.1857 | -1.442 | 82.8094 | -32.9864 | -0.045776 | 0.13733 | -0.15106 | 0.009155 | 0.10529 | -0.013733 | 0.09613 | -0.004577 | 0 | -0.15564 | -0.15564 | 32.9361 | 32.9407 | 1.3916 | 117.1875 | 1.3962 | 117.1921 | 34.227 | -87.204 | -87.3322 | -0.12817 | 34.4833 | 0.25635 | 34.4742 | 34.3643 | 0.24719 | 0.13733
2705.1719 | 121.9437 | -34.3323 | 87.3596 | -0.76378 | 0.6452 | 0.10636 | 0.086436 | 0.00061311 | 0.63335 | 0.013501 | 0.013947 | 0.02239 | 0.01301 | -0.0011087 | 0.014772 | -0.0058255 | -0.0025485 | -0.059509 | 87.2681 | -34.433 | 87.1674 | -1.4145 | 82.7499 | -32.9773 | -0.032043 | 0.13275 | -0.15106 | -0.013733 | 0.10986 | -0.004578 | 0.091553 | -0.10071 | -0.091553 | -0.19226 | -0.10071 | 32.9178 | 33.0185 | 1.355 | 117.0822 | 1.4557 | 117.1829 | 34.1812 | -87.2269 | -87.3734 | -0.14648 | 34.4421 | 0.26093 | 34.4238 | 34.3277 | 0.24261 | 0.14648
2706.1875 | 121.9391 | -34.4376 | 87.355 | -0.75615 | 0.642 | 0.10507 | 0.090281 | 0.0010403 | 0.63206 | 0.010083 | -0.086761 | 0.023031 | 0.014932 | 0.0076499 | 0.014772 | -0.0045437 | -0.0016941 | 0 | 87.2177 | -34.5795 | 87.1307 | -1.5381 | 82.6675 | -33.934 | 0.045776 | 0.16937 | -0.14191 | 0.009155 | 0.10986 | -0.013733 | 0.059509 | -0.14191 | -0.13733 | -0.22431 | -0.086975 | 32.8995 | 33.0414 | 0.50354 | 117.1051 | 0.64545 | 117.247 | 34.2957 | -87.1857 | -87.3459 | -0.16022 | 34.5474 | 0.25177 | 34.4971 | 34.4238 | 0.20142 | 0.12817
2707.1719 | 122.0306 | -34.2407 | 87.442 | -0.77904 | 0.64456 | 0.10401 | 0.085368 | -0.0036593 | 0.62971 | 0.013287 | 0.041412 | 0.022176 | 0.013864 | -0.0040995 | 0.013917 | -0.0075345 | -0.0027621 | 0.027466 | 87.4054 | -34.2957 | 87.3138 | -1.3367 | 82.9788 | -32.84 | 0.12817 | 0.12817 | -0.10071 | 0.073242 | 0.21973 | 0.091553 | 0.16022 | -0.054931 | -0.036621 | -0.12817 | -0.091553 | 32.9041 | 32.959 | 1.4008 | 117.2195 | 1.4557 | 117.2745 | 34.14 | -87.3138 | -87.3688 | -0.054932 | 34.4605 | 0.32044 | 34.4009 | 34.3323 | 0.26093 | 0.19226
2708.1719 | 121.9391 | -34.1171 | 87.5931 | -0.77141 | 0.64178 | 0.10614 | 0.086009 | 0.0025357 | 0.62758 | 0.011365 | 0.013947 | 0.022176 | 0.011301 | 0.0035911 | 0.015413 | -0.0068935 | -0.0038303 | 0.041199 | 87.5107 | -34.2041 | 87.4283 | -1.3367 | 83.02 | -32.7484 | 0.10986 | 0.15106 | -0.12817 | 0 | 0.13733 | 0 | 0.086975 | -0.086976 | -0.082398 | -0.16479 | -0.082397 | 32.7805 | 32.8674 | 1.3687 | 117.1371 | 1.4557 | 117.2241 | 33.989 | -87.442 | -87.5931 | -0.15106 | 34.2545 | 0.2655 | 34.2041 | 34.1171 | 0.21515 | 0.12817
2709.1719 | 121.9803 | -34.0942 | 87.6068 | -0.76378 | 0.64136 | 0.10678 | 0.087718 | 0.0012539 | 0.63121 | 0.0085876 | 0.047516 | 0.022176 | 0.012155 | -0.0066629 | 0.014131 | -0.0062527 | -0.0025485 | 0.054932 | 87.5427 | -34.1675 | 87.4649 | -1.2131 | 83.0154 | -32.7484 | 0.14191 | 0.13275 | -0.13733 | 0.054932 | 0.19684 | 0.073242 | 0.12817 | -0.073242 | -0.064087 | -0.14191 | -0.07782 | 32.8812 | 32.9544 | 1.3458 | 117.1097 | 1.4191 | 117.1829 | 33.9569 | -87.4741 | -87.5519 | -0.077819 | 34.2911 | 0.33417 | 34.2224 | 34.1675 | 0.2655 | 0.21057
2710.1719 | 122.0261 | -34.0805 | 87.6022 | -0.72563 | 0.63901 | 0.10379 | 0.088359 | -0.00024129 | 0.62822 | 0.0081604 | 0.013947 | 0.023458 | 0.012796 | 0.0003867 | 0.013063 | -0.0047573 | -0.0036167 | 0.022888 | 87.5198 | -34.1446 | 87.4786 | -1.2817 | 83.1299 | -32.6889 | 0.14648 | 0.13275 | -0.10529 | 0.059509 | 0.20599 | 0.086975 | 0.14648 | -0.064087 | -0.082398 | -0.1236 | -0.041198 | 32.7988 | 32.8629 | 1.3916 | 117.2104 | 1.4557 | 117.2745 | 33.9752 | -87.4695 | -87.5427 | -0.073242 | 34.2865 | 0.31128 | 34.227 | 34.1675 | 0.25177 | 0.19226
2711.1719 | 121.994 | -34.1858 | 87.4832 | -0.75615 | 0.63879 | 0.1055 | 0.084513 | 0.0040311 | 0.62779 | 0.0088012 | 0.026154 | 0.023244 | 0.01066 | 0.0003867 | 0.015626 | -0.0068935 | -0.0048983 | -0.009155 | 87.4466 | -34.2819 | 87.3642 | -1.3504 | 82.9788 | -32.7896 | 0.064087 | 0.15564 | -0.10071 | 0.054932 | 0.20142 | 0.073242 | 0.14191 | -0.09613 | -0.036621 | -0.11902 | -0.082397 | 32.8354 | 32.9315 | 1.3962 | 117.1646 | 1.4923 | 117.2607 | 34.0851 | -87.3276 | -87.4283 | -0.10071 | 34.3872 | 0.30212 | 34.3277 | 34.259 | 0.24262 | 0.17395
2712.1719 | 121.9757 | -34.1263 | 87.529 | -0.75615 | 0.6405 | 0.10657 | 0.089854 | 0.0018949 | 0.6263 | 0.010083 | 0.016998 | 0.020254 | 0.012369 | 0.0001731 | 0.014558 | -0.0047573 | -0.0042575 | -0.004578 | 87.4878 | -34.2316 | 87.3871 | -1.3916 | 83.0109 | -32.872 | 0.054932 | 0.15106 | -0.10071 | 0.064087 | 0.21057 | 0.086975 | 0.15106 | -0.10528 | -0.041199 | -0.14191 | -0.10071 | 32.7347 | 32.84 | 1.2543 | 117.1371 | 1.3596 | 117.2424 | 34.0256 | -87.3779 | -87.4649 | -0.086975 | 34.3369 | 0.31128 | 34.2773 | 34.2133 | 0.25177 | 0.18768
2713.1719 | 122.0673 | -34.2041 | 87.4374 | -0.74852 | 0.63708 | 0.10443 | 0.084941 | 0.0014675 | 0.6248 | 0.010724 | 0.029205 | 0.023458 | 0.010873 | 0.0018821 | 0.013704 | -0.0056119 | -0.0057529 | 0.036621 | 87.3688 | -34.3094 | 87.3642 | -1.2772 | 82.9422 | -32.3776 | 0.15564 | 0.13733 | -0.10986 | 0.073242 | 0.18768 | -0.022888 | 0.027466 | -0.10529 | -0.068665 | -0.073242 | -0.004577 | 32.9269 | 33.0322 | 1.8265 | 117.1463 | 1.9318 | 117.2516 | 34.0942 | -87.3001 | -87.3642 | -0.064087 | 34.3918 | 0.29755 | 34.2316 | 34.1812 | 0.13733 | 0.086975
2714.1719 | 121.9574 | -34.1537 | 87.5244 | -0.71037 | 0.63879 | 0.10572 | 0.086436 | 0.00018591 | 0.62416 | 0.010083 | -0.0043641 | 0.02239 | 0.013864 | -0.0017495 | 0.013917 | -0.0036893 | -0.0040439 | 0.045776 | 87.4786 | -34.2407 | 87.3779 | -1.3092 | 83.0383 | -32.753 | 0.13733 | 0.15106 | -0.10071 | 0.041199 | 0.17853 | 0.050354 | 0.10071 | -0.086975 | -0.045776 | -0.14648 | -0.10071 | 32.8445 | 32.9315 | 1.4008 | 117.1921 | 1.4877 | 117.2791 | 34.053 | -87.3734 | -87.4832 | -0.10986 | 34.3323 | 0.27924 | 34.2545 | 34.2041 | 0.20142 | 0.15106
2715.1719 | 121.9528 | -34.227 | 87.529 | -0.76378 | 0.63666 | 0.10486 | 0.084941 | 0.0018949 | 0.62288 | 0.0090148 | 0.010895 | 0.021322 | 0.012582 | 0.0016685 | 0.015413 | -0.0066799 | 0.00044212 | 0.045776 | 87.4191 | -34.2819 | 87.3367 | -1.3092 | 82.9697 | -32.8995 | 0.10986 | 0.15106 | -0.13733 | 0.022888 | 0.14191 | 0.036621 | 0.086975 | -0.054931 | -0.10986 | -0.19226 | -0.082397 | 32.9178 | 32.9727 | 1.3275 | 117.1967 | 1.3824 | 117.2516 | 34.0897 | -87.3779 | -87.5061 | -0.12817 | 34.3689 | 0.27924 | 34.314 | 34.2636 | 0.2243 | 0.17395
2716.1719 | 121.9299 | -34.2499 | 87.4695 | -0.718 | 0.6358 | 0.10593 | 0.086436 | 0.0031765 | 0.62245 | 0.0077332 | -0.013519 | 0.023031 | 0.012582 | -0.0064493 | 0.01349 | -0.0047573 | -0.00041228 | 0.041199 | 87.3871 | -34.3277 | 87.2818 | -1.3779 | 82.933 | -32.9864 | 0.10986 | 0.17853 | -0.12817 | -0.013733 | 0.14191 | 0.022888 | 0.082397 | -0.07782 | -0.082397 | -0.18768 | -0.10529 | 32.872 | 32.9498 | 1.2634 | 117.1829 | 1.3412 | 117.2607 | 34.1217 | -87.291 | -87.4832 | -0.19226 | 34.3918 | 0.27008 | 34.3323 | 34.2728 | 0.21057 | 0.15106
2717.1719 | 121.8658 | -34.1217 | 87.6434 | -0.718 | 0.63644 | 0.10572 | 0.087077 | 0.00082671 | 0.61861 | 0.012646 | -0.025726 | 0.022176 | 0.01365 | -0.0015359 | 0.013917 | -0.0053983 | -0.0036167 | 0.004578 | 87.5748 | -34.1949 | 87.3596 | -1.4786 | 83.0109 | -32.8171 | 0.013733 | 0.19684 | -0.15564 | -0.068665 | 0.064087 | -0.041199 | 0.004578 | -0.073242 | -0.068665 | -0.28381 | -0.21515 | 32.6431 | 32.7164 | 1.3046 | 117.1326 | 1.3779 | 117.2058 | 33.9661 | -87.4466 | -87.7121 | -0.2655 | 34.1858 | 0.21973 | 34.1263 | 34.0805 | 0.16022 | 0.11444
2718.1719 | 121.9208 | -33.9615 | 87.7533 | -0.70274 | 0.63409 | 0.10678 | 0.088786 | -0.00088229 | 0.62053 | 0.01051 | -0.077606 | 0.022603 | 0.01301 | 0.0074363 | 0.013917 | -0.0028347 | -0.0053257 | -0.009155 | 87.6755 | -34.108 | 87.5244 | -1.4236 | 83.1162 | -32.7164 | 0.054932 | 0.17395 | -0.12817 | 0.022888 | 0.20599 | -0.009155 | 0.059509 | -0.14648 | -0.07782 | -0.22888 | -0.15106 | 32.5378 | 32.6843 | 1.2451 | 117.0776 | 1.3916 | 117.2241 | 33.8333 | -87.5793 | -87.7304 | -0.15106 | 34.1675 | 0.33417 | 34.021 | 33.9523 | 0.18768 | 0.11902
2719.1719 | 121.9666 | -34.0439 | 87.6251 | -0.67223 | 0.6358 | 0.10678 | 0.087504 | 0.0012539 | 0.61967 | 0.0096558 | 0.041412 | 0.021963 | 0.015359 | -0.0013223 | 0.014558 | -0.0058255 | -0.0072483 | -0.10071 | 87.529 | -34.1812 | 87.4557 | -1.3733 | 83.107 | -32.6797 | 0.13733 | 0.14191 | -0.10986 | 0.032043 | 0.17853 | 0.064087 | 0.12817 | -0.13733 | -0.09613 | -0.16937 | -0.073242 | 32.6706 | 32.8079 | 1.3641 | 117.1509 | 1.5015 | 117.2882 | 33.934 | -87.4832 | -87.5931 | -0.10986 | 34.2224 | 0.28839 | 34.1721 | 34.108 | 0.23804 | 0.17395
2720.1719 | 121.875 | -34.2041 | 87.529 | -0.76378 | 0.63367 | 0.1055 | 0.086009 | 0.0023221 | 0.61903 | 0.010724 | -0.0074159 | 0.021322 | 0.013437 | 0.0003867 | 0.015626 | -0.0068935 | -0.00062588 | 0.027466 | 87.4649 | -34.2545 | 87.3184 | -1.4236 | 82.9376 | -32.9041 | 0.050354 | 0.20142 | -0.13733 | -0.045776 | 0.082397 | -0.054932 | 0.013733 | -0.050354 | -0.064087 | -0.21057 | -0.14648 | 32.7805 | 32.8308 | 1.3 | 117.1417 | 1.3504 | 117.1921 | 34.0668 | -87.3276 | -87.5748 | -0.24719 | 34.2865 | 0.21973 | 34.2178 | 34.1492 | 0.15106 | 0.082397
2721.1719 | 121.8613 | -34.0393 | 87.7167 | -0.77904 | 0.63174 | 0.10529 | 0.087718 | 0.00061311 | 0.61796 | 0.0075196 | -0.022675 | 0.022176 | 0.011514 | 0.0027365 | 0.015413 | -0.0041165 | 0.00086952 | -0.1236 | 87.4786 | -34.3185 | 87.4329 | -1.4786 | 83.1894 | -34.5932 | 0.21057 | 0.19684 | -0.11902 | -0.027466 | 0.027466 | -0.041199 | 0.018311 | -0.27923 | -0.23804 | -0.28381 | -0.045777 | 32.5607 | 32.84 | -0.55389 | 117.2287 | -0.27466 | 117.5079 | 33.9203 | -87.5198 | -87.7441 | -0.2243 | 34.0668 | 0.14649 | 34.0576 | 33.9981 | 0.13733 | 0.07782
2722.1719 | 121.8658 | -33.9249 | 87.8128 | -0.67223 | 0.63217 | 0.10572 | 0.088572 | 0.0014675 | 0.61626 | 0.010297 | -0.12033 | 0.023671 | 0.012582 | -0.0013223 | 0.015413 | -0.0032619 | -0.0048983 | -0.027466 | 87.7808 | -34.0302 | 87.4191 | -1.4236 | 83.3771 | -32.5287 | -0.013733 | 0.2243 | -0.11444 | -0.068665 | 0.059509 | -0.036621 | 0.032043 | -0.10528 | -0.032043 | -0.39368 | -0.36163 | 32.5012 | 32.6065 | 1.3962 | 117.3019 | 1.5015 | 117.4072 | 33.8104 | -87.5885 | -87.8815 | -0.29297 | 33.9844 | 0.17395 | 33.9569 | 33.8882 | 0.14648 | 0.07782
2723.1719 | 121.8704 | -34.3506 | 87.3779 | -0.67986 | 0.63238 | 0.10764 | 0.089 | -0.0025913 | 0.61626 | 0.01051 | -0.071503 | 0.02239 | 0.012155 | -0.0004679 | 0.016054 | -0.0041165 | -0.0036167 | -0.041199 | 87.323 | -34.4147 | 87.1216 | -1.4282 | 82.7271 | -33.1146 | -0.068665 | 0.17853 | -0.14191 | -0.064087 | 0.018311 | -0.09613 | 0.018311 | -0.064087 | -0.054932 | -0.25635 | -0.20142 | 32.9224 | 32.9864 | 1.236 | 117.0776 | 1.3 | 117.1417 | 34.2087 | -87.1994 | -87.442 | -0.24261 | 34.3689 | 0.16022 | 34.3689 | 34.2545 | 0.16022 | 0.045777
2724.1719 | 121.8842 | -33.9935 | 87.7213 | -0.6646 | 0.62918 | 0.10379 | 0.084941 | -0.0019503 | 0.61925 | 0.01051 | -0.03183 | 0.02239 | 0.014505 | 0.0008139 | 0.013917 | -0.0062527 | -0.0034031 | 0 | 87.7258 | -34.0668 | 87.41 | -1.4191 | 83.2352 | -32.7484 | -0.022888 | 0.21057 | -0.11444 | -0.068665 | 0.073242 | -0.054932 | 0.018311 | -0.073242 | 0.004578 | -0.31128 | -0.31586 | 32.5745 | 32.6477 | 1.2451 | 117.2287 | 1.3184 | 117.3019 | 33.8791 | -87.5107 | -87.7899 | -0.27924 | 34.0668 | 0.18768 | 34.0118 | 33.9386 | 0.13275 | 0.059509
2725.1719 | 121.907 | -34.2361 | 87.4969 | -0.64934 | 0.62875 | 0.10507 | 0.087931 | 0.0044583 | 0.61412 | 0.011365 | 0.0047913 | 0.023671 | 0.012582 | -0.0006815 | 0.015626 | -0.0049709 | -0.0021213 | 0.068665 | 87.5198 | -34.2545 | 87.2314 | -1.442 | 82.9056 | -33.0368 | -0.027466 | 0.19684 | -0.13733 | -0.064087 | 0.064087 | -0.054932 | 0.022888 | -0.018311 | 0.022888 | -0.2655 | -0.28839 | 32.7942 | 32.8125 | 1.1993 | 117.1417 | 1.2177 | 117.16 | 34.0988 | -87.3001 | -87.561 | -0.26093 | 34.3002 | 0.20142 | 34.259 | 34.1812 | 0.16022 | 0.082397
2726.1719 | 121.8567 | -34.0118 | 87.7625 | -0.71037 | 0.62768 | 0.10572 | 0.087504 | 0.00082671 | 0.61326 | 0.011578 | 0.038361 | 0.023885 | 0.016 | -0.0004679 | 0.013917 | -0.0053983 | 1.4916e-05 | -0.032043 | 87.6801 | -34.1034 | 87.4512 | -1.387 | 83.0978 | -32.7667 | 0.027466 | 0.18768 | -0.17395 | -0.073242 | 0.045776 | -0.054932 | 0.013733 | -0.091553 | -0.082397 | -0.31128 | -0.22888 | 32.6248 | 32.7164 | 1.2451 | 117.1097 | 1.3367 | 117.2012 | 33.8379 | -87.5748 | -87.8357 | -0.26092 | 34.0576 | 0.21973 | 34.0256 | 33.9569 | 0.18768 | 0.11902
2727.1719 | 121.8796 | -34.4284 | 87.323 | -0.72563 | 0.62662 | 0.107 | 0.086222 | 0.0010403 | 0.61348 | 0.0094422 | -0.0074159 | 0.022603 | 0.014291 | -0.0004679 | 0.015626 | -0.0075345 | -0.0051121 | 0.027466 | 87.0987 | -34.7076 | 86.9705 | -1.4145 | 82.576 | -33.2611 | -0.022888 | 0.16937 | -0.16479 | -0.068665 | 0.054932 | -0.050354 | 0.018311 | -0.27924 | -0.2243 | -0.35248 | -0.12817 | 33.0139 | 33.2932 | 1.1673 | 117.0044 | 1.4465 | 117.2836 | 34.2636 | -87.1536 | -87.3917 | -0.23804 | 34.4833 | 0.21973 | 34.4467 | 34.3781 | 0.18311 | 0.11444
2728.1719 | 121.8704 | -34.3918 | 87.3367 | -0.64934 | 0.62875 | 0.10764 | 0.088786 | 0.0040311 | 0.61134 | 0.013074 | -0.053192 | 0.024312 | 0.014078 | -0.0011087 | 0.015626 | -0.0053983 | -0.0023349 | -0.032043 | 87.2635 | -34.4833 | 87.0895 | -1.3733 | 82.6859 | -33.1467 | -0.022888 | 0.16022 | -0.13275 | -0.045776 | 0.041199 | -0.064087 | 0.022888 | -0.091552 | -0.073242 | -0.24719 | -0.17395 | 33.0185 | 33.11 | 1.2451 | 117.0776 | 1.3367 | 117.1692 | 34.259 | -87.1765 | -87.3825 | -0.20599 | 34.433 | 0.17395 | 34.4147 | 34.3277 | 0.15564 | 0.068664
2729.1719 | 121.8842 | -34.2773 | 87.4329 | -0.67223 | 0.62704 | 0.10614 | 0.086009 | -0.0038729 | 0.61391 | 0.013074 | -0.022675 | 0.022176 | 0.011728 | -0.0017495 | 0.014558 | -0.0034757 | -0.0010531 | -0.054932 | 87.3688 | -34.4147 | 87.149 | -1.3916 | 82.8049 | -33.6594 | 0.050354 | 0.14648 | -0.16022 | -0.018311 | 0.11444 | -0.018311 | 0.054932 | -0.13733 | -0.064087 | -0.28381 | -0.21973 | 32.8857 | 33.0231 | 0.61798 | 117.0822 | 0.75531 | 117.2195 | 34.1171 | -87.2864 | -87.4512 | -0.16479 | 34.3918 | 0.27466 | 34.3323 | 34.259 | 0.21515 | 0.14191
2730.1719 | 121.9254 | -34.3689 | 87.3459 | -0.718 | 0.62726 | 0.10678 | 0.086436 | -0.00088229 | 0.61284 | 0.0096558 | 0.013947 | 0.023458 | 0.011728 | 0.0010275 | 0.014986 | -0.0036893 | -0.00019868 | -0.004578 | 87.2772 | -34.52 | 87.1216 | -1.3962 | 82.6904 | -32.9864 | 0.064087 | 0.17395 | -0.14648 | -0.013733 | 0.091553 | -0.027466 | 0.068665 | -0.15106 | -0.068664 | -0.2243 | -0.15564 | 32.9727 | 33.1238 | 1.3824 | 117.0593 | 1.5335 | 117.2104 | 34.2224 | -87.1719 | -87.3596 | -0.18768 | 34.4604 | 0.23804 | 34.4376 | 34.3414 | 0.21515 | 0.11902
2731.1719 | 121.8933 | -34.0805 | 87.5977 | -0.71037 | 0.62448 | 0.107 | 0.088359 | 0.0012539 | 0.61049 | 0.010083 | -0.016571 | 0.021535 | 0.011301 | -0.0004679 | 0.015413 | -0.0043301 | -0.0061801 | 0.009155 | 87.5748 | -34.1995 | 87.3734 | -1.3733 | 82.9926 | -32.8537 | 0.059509 | 0.16022 | -0.15106 | -0.004578 | 0.10986 | -0.018311 | 0.068665 | -0.11902 | -0.022888 | -0.2243 | -0.20142 | 32.7072 | 32.8262 | 1.2268 | 117.0731 | 1.3458 | 117.1921 | 33.9294 | -87.4374 | -87.6022 | -0.16479 | 34.1904 | 0.26093 | 34.1492 | 34.0622 | 0.21973 | 0.13275
2732.1719 | 121.9162 | -34.0439 | 87.6984 | -0.67986 | 0.62555 | 0.10379 | 0.084941 | 0.0014675 | 0.6122 | 0.011578 | -0.0074159 | 0.022603 | 0.014719 | 0.0003867 | 0.016267 | -0.0034757 | -0.0023349 | -0.059509 | 87.6663 | -34.1309 | 87.4191 | -1.4694 | 83.0566 | -32.3181 | -0.018311 | 0.16479 | -0.16022 | -0.050354 | 0.082397 | -0.032043 | 0.045776 | -0.086975 | -0.032043 | -0.27924 | -0.24719 | 32.5745 | 32.6614 | 1.7258 | 117.1005 | 1.8127 | 117.1875 | 33.8837 | -87.5336 | -87.7487 | -0.21515 | 34.1263 | 0.24261 | 34.0897 | 34.0118 | 0.20599 | 0.12817
2733.1719 | 121.8613 | -34.0393 | 87.7213 | -0.71037 | 0.62448 | 0.10678 | 0.08964 | 0.0016813 | 0.61027 | 0.0094422 | -0.013519 | 0.022817 | 0.014291 | -0.0015359 | 0.013277 | -0.0047573 | 0.00022852 | -0.027466 | 87.6572 | -34.0988 | 87.4283 | -1.3412 | 83.0475 | -33.078 | -0.073242 | 0.14191 | -0.15106 | -0.086975 | 0.032043 | -0.10529 | 0.018311 | -0.059509 | -0.064086 | -0.29297 | -0.22888 | 32.6981 | 32.7576 | 0.9613 | 117.0868 | 1.0208 | 117.1463 | 33.8882 | -87.5793 | -87.8082 | -0.22888 | 34.0714 | 0.18311 | 34.0576 | 33.934 | 0.16937 | 0.045776
2734.1719 | 121.8887 | -33.989 | 87.7625 | -0.67986 | 0.62298 | 0.10507 | 0.087291 | 0.00061311 | 0.60878 | 0.011365 | -0.010467 | 0.022603 | 0.011942 | -4.0504e-05 | 0.015413 | -0.0049709 | -0.0061801 | 0.027466 | 87.6892 | -34.0759 | 87.4786 | -1.2863 | 83.1116 | -32.7164 | 0.022888 | 0.14648 | -0.12817 | -0.032043 | 0.082397 | -0.018311 | 0.064087 | -0.086975 | -0.073242 | -0.28381 | -0.21057 | 32.7026 | 32.7896 | 1.2726 | 117.1005 | 1.3596 | 117.1875 | 33.8608 | -87.616 | -87.7945 | -0.17853 | 34.0714 | 0.21057 | 34.053 | 33.9706 | 0.19226 | 0.10986
2735.1719 | 121.8704 | -34.0897 | 87.6572 | -0.76378 | 0.62149 | 0.10721 | 0.087931 | 0.0025357 | 0.60942 | 0.0090148 | -0.022675 | 0.021749 | 0.012155 | 0.0020957 | 0.014772 | -0.0032619 | -0.0019077 | -0.013733 | 87.6205 | -34.1125 | 87.442 | -1.3962 | 83.0292 | -32.7713 | -0.013733 | 0.15564 | -0.12817 | -0.050354 | 0.059509 | -0.073242 | 0.032043 | -0.022888 | -0.036622 | -0.21515 | -0.17853 | 32.6935 | 32.7164 | 1.3184 | 117.1188 | 1.3412 | 117.1417 | 33.9615 | -87.5015 | -87.7075 | -0.20599 | 34.1492 | 0.18768 | 34.1217 | 34.0164 | 0.16022 | 0.054932
2736.1719 | 121.8658 | -33.8791 | 87.854 | -0.74089 | 0.62063 | 0.10657 | 0.087077 | 0.0018949 | 0.60686 | 0.010724 | -0.0043641 | 0.023244 | 0.014291 | 0.0067953 | 0.014986 | -0.0019803 | -0.0046847 | -0.027466 | 87.822 | -33.9203 | 87.6617 | -1.3733 | 83.2535 | -32.6019 | -0.009155 | 0.16937 | -0.14648 | -0.059509 | 0.050354 | -0.059509 | 0.018311 | -0.041199 | -0.032044 | -0.19226 | -0.16022 | 32.5058 | 32.547 | 1.2772 | 117.1326 | 1.3184 | 117.1738 | 33.7326 | -87.6846 | -87.9135 | -0.22888 | 33.9294 | 0.19684 | 33.8974 | 33.8196 | 0.16479 | 0.086975
2737.1719 | 121.8521 | -33.8654 | 87.8586 | -0.73326 | 0.61999 | 0.10636 | 0.086863 | -0.00045509 | 0.60686 | 0.0094422 | -0.016571 | 0.022603 | 0.014291 | -0.0006815 | 0.014986 | -0.0036893 | -0.00041228 | -0.009155 | 87.8265 | -33.9386 | 87.6068 | -1.387 | 83.2169 | -32.6477 | -0.036621 | 0.14648 | -0.13275 | -0.050354 | 0.032043 | -0.07782 | 0.018311 | -0.073243 | -0.032044 | -0.25177 | -0.21973 | 32.4783 | 32.5516 | 1.2177 | 117.0822 | 1.2909 | 117.1555 | 33.7326 | -87.7121 | -87.9089 | -0.19684 | 33.8974 | 0.16479 | 33.8837 | 33.7875 | 0.15106 | 0.054931
2738.1719 | 121.9254 | -33.8516 | 87.8082 | -0.72563 | 0.61957 | 0.10614 | 0.089213 | 0.0023221 | 0.60622 | 0.011151 | 0.016998 | 0.022603 | 0.01365 | -0.0030313 | 0.015413 | -0.0060391 | -0.0010531 | -0.11444 | 87.7579 | -33.9752 | 87.5656 | -1.2955 | 83.1528 | -32.7484 | -0.10071 | 0.16022 | -0.14191 | 0.004578 | 0.10986 | -0.013733 | 0.09613 | -0.1236 | -0.050354 | -0.24262 | -0.19226 | 32.5562 | 32.6797 | 1.1032 | 117.0044 | 1.2268 | 117.128 | 33.7097 | -87.648 | -87.8037 | -0.15564 | 33.9615 | 0.25177 | 33.9478 | 33.8379 | 0.23804 | 0.12817
2739.1719 | 121.8796 | -33.7143 | 87.9868 | -0.67986 | 0.61828 | 0.10678 | 0.087931 | 0.0021085 | 0.60643 | 0.010083 | 0.026154 | 0.021963 | 0.012582 | -0.0002543 | 0.014986 | -0.0049709 | -0.0010531 | -0.018311 | 87.9456 | -33.8013 | 87.7258 | -1.3641 | 83.3221 | -32.3547 | 0.009155 | 0.23346 | 0.004578 | -0.013733 | 0.082397 | -0.045776 | 0.050354 | -0.086976 | -0.041198 | -0.26093 | -0.21973 | 32.3502 | 32.4371 | 1.3596 | 117.0364 | 1.4465 | 117.1234 | 33.7189 | -87.7533 | -88.0005 | -0.24719 | 33.7967 | 0.077819 | 33.7646 | 33.6685 | 0.045776 | -0.050354
2740.1719 | 121.8842 | -33.6914 | 87.9913 | -0.68748 | 0.61935 | 0.10785 | 0.089213 | -0.0017367 | 0.60707 | 0.0079468 | 0.0047913 | 0.022176 | 0.014719 | 0.0020957 | 0.015199 | -0.0047573 | -0.0029757 | -0.009155 | 87.9868 | -33.7646 | 87.9044 | -1.3138 | 83.4961 | -32.634 | 0 | 0.22888 | -0.022888 | -0.027466 | 0.082397 | -0.041199 | 0.054932 | -0.073242 | -0.004578 | -0.086975 | -0.082397 | 32.3776 | 32.4509 | 1.0574 | 117.1875 | 1.1307 | 117.2607 | 33.6685 | -87.7625 | -88.0188 | -0.25635 | 33.7738 | 0.10528 | 33.7463 | 33.6502 | 0.07782 | -0.018311
2741.1719 | 121.8521 | -33.8287 | 87.9135 | -0.75615 | 0.61807 | 0.10764 | 0.089 | 0.0025357 | 0.60365 | 0.012006 | -0.019623 | 0.020467 | 0.014291 | 0.0033773 | 0.016481 | -0.0049709 | -0.0016941 | -0.082397 | 87.8448 | -33.9432 | 87.5931 | -1.4008 | 83.1162 | -32.6477 | -0.13275 | 0.16937 | -0.16022 | -0.068665 | 0.004578 | -0.09613 | 0.009155 | -0.11444 | -0.068664 | -0.32043 | -0.25177 | 32.428 | 32.5424 | 1.181 | 116.9449 | 1.2955 | 117.0593 | 33.6685 | -87.7441 | -87.9822 | -0.23804 | 33.8333 | 0.16479 | 33.8379 | 33.7326 | 0.16937 | 0.064087
2742.1719 | 121.8567 | -33.7784 | 87.8906 | -0.80193 | 0.61764 | 0.10614 | 0.089 | 0.0014675 | 0.6028 | 0.0079468 | -0.016571 | 0.022603 | 0.014078 | -0.0036721 | 0.013917 | -0.0062527 | -0.0048983 | -0.11444 | 87.8815 | -33.8882 | 87.6205 | -1.5152 | 83.226 | -32.634 | -0.14191 | 0.16022 | -0.16479 | -0.068665 | 0.009155 | -0.091553 | 0.027466 | -0.10986 | -0.009155 | -0.27008 | -0.26093 | 32.2632 | 32.373 | 1.1444 | 117.0044 | 1.2543 | 117.1143 | 33.6136 | -87.7304 | -87.9593 | -0.22888 | 33.7875 | 0.17395 | 33.8058 | 33.6868 | 0.19226 | 0.073242
2743.1719 | 121.9254 | -33.4305 | 88.2019 | -0.7943 | 0.61679 | 0.10572 | 0.086436 | -0.00066869 | 0.60429 | 0.0098694 | 0.059723 | 0.020681 | 0.015359 | -0.0040995 | 0.015199 | -0.0068935 | -0.0031895 | -0.059509 | 88.1378 | -33.5312 | 88.0325 | -1.1719 | 83.6746 | -32.0755 | -0.009155 | 0.16022 | -0.10986 | 0.050354 | 0.16022 | 0.050354 | 0.13275 | -0.10071 | -0.064087 | -0.16937 | -0.10528 | 32.2586 | 32.3593 | 1.355 | 117.1051 | 1.4557 | 117.2058 | 33.3206 | -88.0417 | -88.1516 | -0.10986 | 33.5907 | 0.27008 | 33.5632 | 33.4808 | 0.24261 | 0.16022
2744.1719 | 121.7606 | -33.609 | 88.06 | -0.73326 | 0.61914 | 0.10593 | 0.088145 | -0.0017367 | 0.60301 | 0.0090148 | -0.022675 | 0.024099 | 0.011728 | -0.0043131 | 0.015626 | -0.0043301 | -0.0029757 | -0.13733 | 88.028 | -33.7418 | 87.7533 | -1.3687 | 83.3633 | -32.4738 | -0.13275 | 0.1831 | -0.15564 | -0.073242 | 0.027466 | -0.09613 | 0.1831 | -0.13275 | -0.032044 | -0.3067 | -0.27466 | 32.2403 | 32.373 | 1.1353 | 116.9724 | 1.268 | 117.1051 | 33.4534 | -87.8769 | -88.1332 | -0.25635 | 33.6365 | 0.18311 | 33.7921 | 33.5129 | 0.33874 | 0.05951
2745.1719 | 121.9345 | -33.6548 | 88.0371 | -0.67986 | 0.61401 | 0.10828 | 0.087718 | -2.7692e-05 | 0.6013 | 0.0070922 | -0.010467 | 0.022176 | 0.014719 | -0.0011087 | 0.013063 | -0.0060391 | -0.0046847 | -0.10986 | 87.9959 | -33.7784 | 87.7167 | -1.3596 | 83.313 | -32.5424 | -0.14191 | 0.1831 | -0.14648 | 0.1831 | -0.054932 | -0.09613 | 0.022888 | -0.1236 | -0.041198 | -0.32043 | -0.27924 | 32.2952 | 32.4188 | 1.1124 | 116.9678 | 1.236 | 117.0914 | 33.5083 | -87.854 | -87.854 | 0 | 33.5999 | 0.091552 | 33.6777 | 33.5587 | 0.16937 | 0.050354
2746.1719 | 121.8246 | -33.5678 | 88.1653 | -0.718 | 0.61359 | 0.10593 | 0.087931 | 0.0016813 | 0.60066 | 0.0070922 | -0.022675 | 0.022817 | 0.01066 | -0.0040995 | 0.014345 | -0.0039029 | -0.0053257 | -0.059509 | 88.1195 | -33.6365 | 87.8769 | -1.3138 | 83.4778 | -32.3364 | -0.10529 | 0.15564 | -0.14648 | -0.064087 | 0.022888 | -0.10529 | -0.004578 | -0.068665 | -0.045776 | -0.28839 | -0.24262 | 32.254 | 32.3227 | 1.2314 | 117.0456 | 1.3 | 117.1143 | 33.4213 | -88.0096 | -88.2294 | -0.21973 | 33.5907 | 0.16937 | 33.5632 | 33.4625 | 0.14191 | 0.041198
2747.1719 | 121.8201 | -33.5953 | 88.1241 | -0.718 | 0.61572 | 0.10593 | 0.086863 | 0.0042447 | 0.59874 | 0.0098694 | -0.0074159 | 0.022603 | 0.012155 | -0.0028177 | 0.015626 | -0.0064663 | -0.0059665 | -0.07782 | 88.0463 | -33.7234 | 87.8128 | -1.3229 | 83.5098 | -32.4371 | -0.10986 | 0.14648 | -0.16022 | -0.082397 | 0.022888 | -0.10986 | 0.018311 | -0.12817 | -0.077819 | -0.31128 | -0.23346 | 32.2723 | 32.4005 | 1.1581 | 117.1051 | 1.2863 | 117.2333 | 33.4351 | -87.9776 | -88.2065 | -0.22888 | 33.6182 | 0.1831 | 33.6136 | 33.4854 | 0.17853 | 0.050354
2748.1719 | 121.9666 | -33.4717 | 88.1607 | -0.74089 | 0.61401 | 0.10593 | 0.086222 | 0.0014675 | 0.60237 | 0.0077332 | 0.0078429 | 0.022603 | 0.014291 | -0.0004679 | 0.014558 | -0.0071071 | -0.0051121 | -0.086975 | 88.1287 | -33.5999 | 87.9456 | -1.2222 | 83.6838 | -32.2311 | -0.027466 | 0.15564 | -0.1236 | 0.054932 | 0.15564 | 0.041199 | 0.14191 | -0.12817 | -0.032044 | -0.21515 | -0.18311 | 32.2495 | 32.3776 | 1.2405 | 117.1555 | 1.3687 | 117.2836 | 33.3481 | -88.0051 | -88.1058 | -0.10071 | 33.6273 | 0.27924 | 33.6136 | 33.5129 | 0.2655 | 0.16479
2749.1719 | 121.8887 | -33.6136 | 88.0783 | -0.718 | 0.61166 | 0.10507 | 0.08665 | -0.0010959 | 0.59981 | 0.0088012 | -0.016571 | 0.021963 | 0.015146 | -4.0504e-05 | 0.017122 | -0.0045437 | -0.0053257 | -0.10986 | 88.0417 | -33.7189 | 87.8265 | -1.3275 | 83.4366 | -32.3868 | -0.10986 | 0.17395 | -0.16022 | -0.027466 | 0.073242 | -0.064087 | 0.041199 | -0.10529 | -0.036621 | -0.25177 | -0.21515 | 32.2861 | 32.3914 | 1.2268 | 117.0502 | 1.3321 | 117.1555 | 33.4534 | -87.9044 | -88.1058 | -0.20142 | 33.6868 | 0.23346 | 33.6548 | 33.5495 | 0.20142 | 0.09613
2750.1719 | 121.9025 | -33.6914 | 87.9776 | -0.71037 | 0.61252 | 0.10529 | 0.087291 | 0.0012539 | 0.60066 | 0.012433 | 0.0047913 | 0.02474 | 0.01301 | -0.0006815 | 0.014986 | -0.0073209 | -0.0031895 | -0.09613 | 87.9547 | -33.815 | 87.7121 | -1.2955 | 83.4091 | -32.5287 | -0.11902 | 0.17395 | -0.15106 | -0.036621 | 0.064087 | -0.07782 | 0.041199 | -0.1236 | -0.022888 | -0.2655 | -0.24262 | 32.3959 | 32.5195 | 1.1627 | 117.1005 | 1.2863 | 117.2241 | 33.5403 | -87.8036 | -88.0142 | -0.21057 | 33.7555 | 0.21515 | 33.7326 | 33.6136 | 0.19226 | 0.073242
2751.1719 | 121.8521 | -33.3893 | 88.3026 | -0.77141 | 0.60995 | 0.10572 | 0.087291 | 0.0016813 | 0.60066 | 0.0096558 | -0.0043641 | 0.022817 | 0.012155 | 0.0010275 | 0.013917 | -0.0049709 | -0.0083163 | -0.11444 | 88.2156 | -33.5312 | 87.973 | -1.3046 | 83.606 | -32.2449 | -0.11902 | 0.16937 | -0.15564 | -0.027466 | 0.059509 | -0.068665 | 0.032043 | -0.14191 | -0.086975 | -0.32959 | -0.24262 | 32.0847 | 32.2266 | 1.1444 | 116.9952 | 1.2863 | 117.1371 | 33.2336 | -88.1332 | -88.3301 | -0.19684 | 33.4488 | 0.21515 | 33.4213 | 33.3206 | 0.18768 | 0.086975
2752.1719 | 121.8658 | -33.4213 | 88.2248 | -0.72563 | 0.61081 | 0.10443 | 0.087291 | -0.0015231 | 0.59639 | 0.0096558 | 0.010895 | 0.023885 | 0.012369 | 0.0035911 | 0.01349 | -0.0064663 | -0.0031895 | -0.1236 | 88.1973 | -33.577 | 87.941 | -1.3184 | 83.7067 | -32.1808 | -0.10529 | 0.17395 | -0.16479 | -0.041199 | 0.082397 | -0.059509 | 0.036621 | -0.15564 | -0.027465 | -0.28381 | -0.25635 | 32.103 | 32.2586 | 1.2405 | 117.128 | 1.3962 | 117.2836 | 33.2565 | -88.0508 | -88.266 | -0.21515 | 33.5037 | 0.24719 | 33.4579 | 33.3618 | 0.20142 | 0.10529
2753.1719 | 121.9345 | -33.5175 | 88.0829 | -0.73326 | 0.61059 | 0.10636 | 0.085795 | -0.0010959 | 0.5981 | 0.0047424 | 0.062775 | 0.022176 | 0.010446 | -0.0058085 | 0.015413 | -0.0068935 | 1.4916e-05 | -0.045776 | 88.0646 | -33.5999 | 87.9593 | -1.1536 | 83.6151 | -32.2769 | 0 | 0.13733 | -0.10071 | 0.054932 | 0.16937 | 0.041199 | 0.1236 | -0.082398 | -0.018311 | -0.1236 | -0.10529 | 32.3639 | 32.4463 | 1.2405 | 117.1326 | 1.3229 | 117.215 | 33.4167 | -87.9456 | -88.028 | -0.082397 | 33.6868 | 0.27008 | 33.6411 | 33.5587 | 0.2243 | 0.14191
2754.1719 | 121.2799 | -33.4534 | 88.3026 | -0.69511 | 0.60931 | 0.107 | 0.091777 | -0.0013095 | 0.59276 | 0.011365 | -0.037933 | 0.023031 | 0.011942 | -4.0504e-05 | 0.015199 | -0.0056119 | -0.0025485 | -0.091553 | 88.2294 | -33.5541 | 87.9593 | -1.3458 | 83.5464 | -32.2952 | -0.13275 | 0.16022 | -0.14648 | -0.064087 | 0.027466 | -0.091553 | 0.23804 | -0.10071 | -0.073242 | -0.34332 | -0.27008 | 32.1075 | 32.2083 | 1.1581 | 116.9998 | 1.2589 | 117.1005 | 33.3069 | -88.1424 | -88.3667 | -0.2243 | 33.4808 | 0.17395 | 33.6914 | 33.3618 | 0.38452 | 0.054931
2755.1719 | 121.9162 | -33.4488 | 88.2385 | -0.718 | 0.60803 | 0.10465 | 0.087931 | -0.00066869 | 0.59617 | 0.0085876 | -0.0074159 | 0.021535 | 0.012369 | 0.0027365 | 0.013277 | -0.0090297 | -0.0025485 | -0.086975 | 88.1561 | -33.5632 | 88.0005 | -1.236 | 83.6426 | -32.1396 | -0.050354 | 0.1236 | -0.12817 | 0.032043 | 0.15564 | 0.013733 | 0.13733 | -0.11444 | -0.082397 | -0.23804 | -0.15564 | 32.2128 | 32.3273 | 1.3092 | 117.0914 | 1.4236 | 117.2058 | 33.3206 | -88.1149 | -88.2065 | -0.091553 | 33.6044 | 0.28381 | 33.5861 | 33.4625 | 0.2655 | 0.14191
2756.1719 | 121.9391 | -33.5266 | 88.1424 | -0.80193 | 0.60974 | 0.10636 | 0.088359 | 0.00039951 | 0.59553 | 0.0075196 | -0.0074159 | 0.021322 | 0.012796 | 0.0025229 | 0.01349 | -0.0053983 | -0.00062588 | -0.10071 | 88.0371 | -33.6868 | 87.9135 | -1.2589 | 83.519 | -32.309 | -0.059509 | 0.14648 | -0.12817 | 0.032043 | 0.14648 | 0.036621 | 0.14648 | -0.16022 | -0.10529 | -0.22888 | -0.1236 | 32.2678 | 32.428 | 1.2177 | 117.0456 | 1.3779 | 117.2058 | 33.3984 | -87.9959 | -88.1104 | -0.11444 | 33.6731 | 0.27466 | 33.6731 | 33.5632 | 0.27466 | 0.1648
2757.1875 | 121.8887 | -33.4625 | 88.2202 | -0.74089 | 0.61081 | 0.10593 | 0.084086 | -0.0017367 | 0.59383 | 0.007306 | 0.078034 | 0.021108 | 0.011942 | -0.0040995 | 0.014345 | -0.0086025 | -0.0031895 | 0.018311 | 88.1241 | -33.5678 | 88.0142 | -1.181 | 83.6746 | -32.1533 | 0.082397 | 0.15564 | -0.11902 | -0.027466 | 0.12817 | 0.018311 | 0.068665 | -0.10529 | -0.096131 | -0.20599 | -0.10986 | 32.2815 | 32.3868 | 1.3092 | 117.1371 | 1.4145 | 117.2424 | 33.3435 | -88.0646 | -88.2477 | -0.18311 | 33.5907 | 0.24719 | 33.5312 | 33.4808 | 0.18768 | 0.13733
2758.1719 | 121.9528 | -33.3344 | 88.3438 | -0.75615 | 0.60782 | 0.10358 | 0.085582 | 0.00039951 | 0.59425 | 0.010083 | 0.081085 | 0.020681 | 0.012582 | -0.0047403 | 0.014986 | -0.0049709 | -0.0029757 | 0 | 88.2431 | -33.4305 | 88.1378 | -1.1902 | 83.9081 | -32.0984 | 0.050354 | 0.14191 | -0.091553 | 0.032043 | 0.18768 | 0.059509 | 0.12817 | -0.09613 | -0.10071 | -0.20599 | -0.10529 | 32.1442 | 32.2403 | 1.236 | 117.2424 | 1.3321 | 117.3386 | 33.2428 | -88.2019 | -88.3118 | -0.10986 | 33.522 | 0.27924 | 33.4625 | 33.3939 | 0.21973 | 0.15106
2759.1719 | 121.9116 | -33.3069 | 88.3118 | -0.70274 | 0.60589 | 0.10358 | 0.090495 | 0.0023221 | 0.59404 | 0.0098694 | -0.037933 | 0.020894 | 0.012369 | 0.0040183 | 0.013063 | -0.0053983 | -0.0048983 | -0.1236 | 88.2523 | -33.4763 | 88.1012 | -1.2726 | 83.7112 | -32.1579 | -0.054932 | 0.14648 | -0.12817 | 0.018311 | 0.11902 | 0 | 0.13733 | -0.16937 | -0.05951 | -0.21057 | -0.15106 | 32.0343 | 32.2037 | 1.149 | 117.0181 | 1.3184 | 117.1875 | 33.1787 | -88.1653 | -88.2935 | -0.12817 | 33.4259 | 0.24719 | 33.4442 | 33.3069 | 0.2655 | 0.12817
2760.1719 | 121.9482 | -108.49 | 13.0875 | -0.8477 | 0.68002 | 0.1557 | 0.087718 | -2.7692e-05 | 0.58827 | 0.013501 | -0.071503 | 0.024312 | 0.010873 | -0.0015359 | 0.013063 | -0.0041165 | -0.0029757 | -0.24261 | 13.0692 | -108.6777 | 12.8082 | -3.8544 | 8.0933 | -106.4026 | -0.12817 | 0.17853 | -0.16937 | -0.054932 | 0.027466 | -0.082397 | 0.032043 | -0.18768 | -0.01831 | -0.27923 | -0.26093 | 104.6356 | 104.8233 | 2.0874 | 116.5833 | 2.2751 | 116.7709 | 108.3206 | -12.9089 | -13.1424 | -0.23346 | 108.5175 | 0.19684 | 108.522 | 108.4076 | 0.20142 | 0.086976
2761.1719 | 121.5408 | -6.4682 | 115.3656 | -0.74089 | 0.62042 | 0.1243 | 0.086863 | -0.0040867 | 0.58955 | 0.0077332 | -0.077606 | 0.023458 | 0.011514 | 0.0031637 | 0.013063 | -0.0049709 | -0.00062588 | -0.068665 | 115.2878 | -6.5414 | 114.9673 | -0.26093 | 110.8337 | -6.8298 | -0.12817 | 0.16479 | -0.12817 | -0.045776 | 0.041199 | -0.09613 | 0.027466 | -0.073242 | -0.07782 | -0.39825 | -0.32043 | 6.2073 | 6.2805 | -0.36163 | 117.3019 | -0.28839 | 117.3752 | 6.34 | -115.2008 | -115.4114 | -0.21057 | 6.5094 | 0.16937 | 6.4957 | 6.3721 | 0.15564 | 0.032044
2762.1875 | 121.9208 | -6.1844 | 115.5212 | -0.77904 | 0.61743 | 0.12451 | 0.085582 | 0.00018591 | 0.59532 | 0.0079468 | 0.07193 | 0.023031 | 0.014932 | 0.0018821 | 0.013277 | -0.0086025 | -0.00062588 | 0 | 115.5075 | -6.221 | 115.3244 | -0.022888 | 111.1588 | -4.953 | 0.091553 | 0.13733 | -0.10071 | 0.059509 | 0.21515 | 0.082397 | 0.15564 | -0.036621 | -0.013733 | -0.19684 | -0.1831 | 6.1615 | 6.1981 | 1.2314 | 117.3431 | 1.268 | 117.3798 | 6.0837 | -115.3839 | -115.4617 | -0.07782 | 6.3995 | 0.31586 | 6.34 | 6.2668 | 0.25635 | 0.18311
2763.1719 | 121.8704 | -6.1157 | 115.6036 | -0.75615 | 0.61828 | 0.12686 | 0.086436 | 0.0014675 | 0.59083 | 0.0062378 | 0.035309 | 0.021108 | 0.012796 | 0.0031637 | 0.015199 | -0.0053983 | -0.0012667 | -0.07782 | 115.5487 | -6.1707 | 115.3061 | -0.12817 | 111.1084 | -4.9164 | -0.041199 | 0.16937 | -0.11902 | -0.009155 | 0.12817 | 0 | 0.10071 | -0.054931 | -0.054932 | -0.29755 | -0.24262 | 5.9875 | 6.0425 | 1.1993 | 117.2241 | 1.2543 | 117.2791 | 5.9967 | -115.4343 | -115.6128 | -0.17853 | 6.2439 | 0.24719 | 6.2164 | 6.1157 | 0.21973 | 0.11902
2764.1719 | 121.9208 | -5.6351 | 116.0843 | -0.74089 | 0.61401 | 0.12729 | 0.088572 | 0.0036039 | 0.59489 | 0.010083 | -0.028778 | 0.022176 | 0.014078 | 0.0006003 | 0.013917 | -0.0043301 | -0.0014805 | -0.064087 | 116.0614 | -5.7175 | 115.8508 | -0.07782 | 111.7355 | -4.4678 | -0.009155 | 0.12817 | -0.11902 | 0.059509 | 0.16022 | 0.027466 | 0.13733 | -0.082397 | -0.022889 | -0.23346 | -0.21057 | 5.5573 | 5.6396 | 1.1673 | 117.3706 | 1.2497 | 117.453 | 5.5161 | -115.9561 | -116.0248 | -0.068665 | 5.7953 | 0.27924 | 5.7724 | 5.6625 | 0.25635 | 0.14649
2765.1719 | 121.9299 | -5.6625 | 116.1118 | -0.81718 | 0.61807 | 0.12451 | 0.088359 | -0.00088229 | 0.59383 | 0.0068786 | -0.03183 | 0.022176 | 0.012155 | 0.0038047 | 0.014131 | -0.0056119 | -0.0051121 | -0.082397 | 116.0614 | -5.6763 | 115.9012 | -0.050354 | 111.7722 | -4.2847 | -0.022888 | 0.11902 | -0.11902 | 0.041199 | 0.16479 | 0.036621 | 0.14191 | -0.013733 | -0.050354 | -0.21057 | -0.16022 | 5.6122 | 5.6259 | 1.3779 | 117.4347 | 1.3916 | 117.4484 | 5.5435 | -115.9927 | -116.0706 | -0.07782 | 5.8273 | 0.28381 | 5.8044 | 5.6992 | 0.26093 | 0.15564
2766.1719 | 121.9345 | -5.4016 | 116.3223 | -0.76378 | 0.61508 | 0.12686 | 0.087931 | -0.0010959 | 0.59105 | 0.006665 | -0.034882 | 0.019399 | 0.010019 | 0.0070089 | 0.014558 | -0.0051845 | -0.0012667 | -0.091553 | 116.2994 | -5.4886 | 116.0477 | -0.086975 | 111.9278 | -4.1794 | -0.045776 | 0.11444 | -0.11902 | 0.068665 | 0.16479 | 0.022888 | 0.12817 | -0.086975 | -0.022889 | -0.27466 | -0.25177 | 5.3146 | 5.4016 | 1.2222 | 117.3294 | 1.3092 | 117.4164 | 5.2826 | -116.2079 | -116.2537 | -0.045776 | 5.5664 | 0.28381 | 5.5298 | 5.4245 | 0.24719 | 0.14191
2767.1719 | 121.8201 | -5.4565 | 116.3132 | -0.718 | 0.61551 | 0.12751 | 0.089 | 0.0023221 | 0.58998 | 0.0088012 | -0.065399 | 0.020254 | 0.010873 | -0.0036721 | 0.01349 | -0.0047573 | -0.0025485 | -0.09613 | 116.2949 | -5.5115 | 115.9836 | -0.15106 | 111.9003 | -4.3762 | -0.11444 | 0.15564 | -0.16022 | -0.045776 | 0.054932 | -0.068665 | 0.041199 | -0.054932 | -0.01831 | -0.32959 | -0.31128 | 5.3055 | 5.3604 | 1.0803 | 117.3569 | 1.1353 | 117.4118 | 5.2963 | -116.1575 | -116.3589 | -0.20142 | 5.5115 | 0.21515 | 5.4977 | 5.3879 | 0.20142 | 0.091552
2768.1719 | 121.9025 | -5.3375 | 116.4185 | -0.72563 | 0.61743 | 0.12729 | 0.085795 | 0.00039951 | 0.58913 | 0.003247 | 0.029205 | 0.021963 | 0.010873 | -0.0034585 | 0.015626 | -0.0066799 | -0.0010531 | -0.064087 | 116.4001 | -5.3558 | 116.1346 | -0.059509 | 112.1201 | -4.0558 | -0.050354 | 0.14191 | -0.1236 | -0.004578 | 0.13275 | 0.009155 | 0.10071 | -0.018311 | -0.018311 | -0.28381 | -0.2655 | 5.278 | 5.2963 | 1.2817 | 117.4576 | 1.3 | 117.4759 | 5.2139 | -116.2766 | -116.423 | -0.14649 | 5.4703 | 0.25635 | 5.4382 | 5.3467 | 0.2243 | 0.13275
2769.1719 | 121.9208 | -5.3513 | 116.3589 | -0.73326 | 0.6138 | 0.12644 | 0.08964 | 0.0012539 | 0.58891 | 0.006665 | -0.040985 | 0.022176 | 0.011728 | -0.0038857 | 0.01349 | -0.0056119 | -0.00041228 | -0.091553 | 116.3177 | -5.452 | 116.0751 | -0.082397 | 112.0605 | -3.4332 | -0.064087 | 0.14191 | -0.1236 | 0.036621 | 0.14648 | 0.013733 | 0.12817 | -0.10071 | -0.041199 | -0.28381 | -0.24262 | 5.2689 | 5.3696 | 1.918 | 117.4118 | 2.0187 | 117.5125 | 5.2277 | -116.217 | -116.3223 | -0.10529 | 5.4977 | 0.27008 | 5.4794 | 5.365 | 0.25177 | 0.13733
2770.1719 | 121.9254 | -5.4703 | 116.2537 | -0.76378 | 0.61572 | 0.12729 | 0.087718 | -0.0019503 | 0.58977 | 0.0088012 | -0.016571 | 0.022176 | 0.011942 | -0.0030313 | 0.01584 | -0.0028347 | -0.0012667 | -0.050354 | 116.2537 | -5.4794 | 116.0843 | -0.013733 | 112.0239 | -4.1565 | -0.013733 | 0.1236 | -0.11902 | -0.086975 | 0.25635 | 0.10529 | 0.16937 | -0.009155 | 0 | -0.16937 | -0.16937 | 5.4565 | 5.4657 | 1.3138 | 117.4942 | 1.3229 | 117.5034 | 5.3513 | -116.1301 | -116.3406 | -0.21057 | 5.7266 | 0.37537 | 5.6396 | 5.5756 | 0.28839 | 0.22431
2771.1719 | 121.9894 | -5.3421 | 116.3773 | -0.71037 | 0.61273 | 0.12622 | 0.08665 | 0.0018949 | 0.58784 | 0.0058106 | -0.03183 | 0.020467 | 0.011301 | -0.0043131 | 0.015413 | -0.0047573 | -0.0016941 | -0.036621 | 116.3544 | -5.3787 | 116.1758 | -0.009155 | 112.0651 | -4.2297 | -0.009155 | 0.13275 | -0.11444 | 0.018311 | -0.086975 | -0.11444 | 0.073242 | -0.036621 | -0.022888 | -0.20142 | -0.17853 | 5.3329 | 5.3696 | 1.1124 | 117.4072 | 1.149 | 117.4438 | 5.2277 | -116.2445 | -116.3589 | -0.11444 | 5.2551 | 0.027466 | 5.4153 | 5.2277 | 0.18768 | 0
2772.1719 | 121.9894 | -5.3558 | 116.3681 | -0.76378 | 0.6138 | 0.12665 | 0.087504 | -0.0032321 | 0.58998 | 0.0047424 | -0.016571 | 0.020254 | 0.012582 | 0.0016685 | 0.013704 | -0.0045437 | 0.0015103 | -0.050354 | 116.3681 | -5.3696 | 116.1758 | -0.022888 | 112.0102 | -4.1245 | 0.004578 | 0.14648 | -0.10986 | 0.050354 | 0.10071 | -0.24261 | -0.027466 | -0.013733 | 0 | -0.19226 | -0.19226 | 5.3329 | 5.3467 | 1.2314 | 117.366 | 1.2451 | 117.3798 | 5.246 | -116.2216 | -116.3177 | -0.09613 | 5.4565 | 0.21057 | 5.3284 | 5.1132 | 0.082397 | -0.13275
2773.1719 | 121.9299 | -5.2963 | 116.4597 | -0.73326 | 0.61209 | 0.12537 | 0.084513 | -0.0015231 | 0.58827 | 0.0051696 | 0.053619 | 0.020467 | 0.010873 | 0.0001731 | 0.014772 | -0.0056119 | -0.0040439 | -0.032043 | 116.4001 | -5.3421 | 116.1621 | -0.054932 | 112.0972 | -4.0329 | 0.032043 | 0.15106 | -0.11444 | 0.022888 | 0.17853 | 0.064087 | 0.12817 | -0.045776 | -0.05951 | -0.29755 | -0.23804 | 5.2414 | 5.2872 | 1.2634 | 117.3935 | 1.3092 | 117.4393 | 5.1819 | -116.3086 | -116.4368 | -0.12817 | 5.4749 | 0.29297 | 5.4245 | 5.3604 | 0.24262 | 0.17853
2774.1719 | 121.9116 | -5.3146 | 116.4322 | -0.74089 | 0.61145 | 0.1243 | 0.089213 | -0.00088229 | 0.58742 | 0.004956 | -0.071503 | 0.021963 | 0.012582 | 0.0027365 | 0.014772 | -0.0060391 | -0.0031895 | -0.082397 | 116.3956 | -5.4062 | 116.1301 | -0.11444 | 112.0789 | -4.1199 | -0.07782 | 0.14191 | -0.14191 | 0.018311 | 0.12817 | 0 | 0.091553 | -0.091553 | -0.036621 | -0.30212 | -0.2655 | 5.2002 | 5.2917 | 1.1948 | 117.3935 | 1.2863 | 117.485 | 5.1727 | -116.2903 | -116.4139 | -0.1236 | 5.4428 | 0.27008 | 5.4062 | 5.3146 | 0.23346 | 0.14191
2775.1719 | 121.8384 | -5.2963 | 116.3773 | -0.77141 | 0.61294 | 0.12451 | 0.086436 | 0.0023221 | 0.58528 | 0.0079468 | -0.034882 | 0.019613 | 0.011942 | -0.0034585 | 0.014131 | -0.0051845 | 0.00022852 | -0.050354 | 116.4047 | -5.3329 | 116.2079 | -0.018311 | 112.0972 | -3.9642 | -0.004578 | 0.12817 | -0.11444 | 0.073242 | 0.17395 | 0.045776 | 0.15564 | -0.036621 | 0.027466 | -0.16937 | -0.19684 | 5.278 | 5.3146 | 1.3321 | 117.3935 | 1.3687 | 117.4301 | 5.1819 | -116.2491 | -116.304 | -0.054932 | 5.4703 | 0.28839 | 5.452 | 5.3421 | 0.27008 | 0.16022
2776.1719 | 121.9254 | -5.2826 | 116.4413 | -0.67986 | 0.6123 | 0.12857 | 0.08665 | -2.7692e-05 | 0.58678 | 0.0092286 | -0.016571 | 0.020894 | 0.014932 | -0.0038857 | 0.015413 | -0.0043301 | -0.0025485 | -0.068665 | 116.3773 | -5.3329 | 116.2216 | -0.013733 | 112.1429 | -3.9276 | 0.018311 | 0.1236 | -0.10986 | 0.054932 | 0.16479 | 0.036621 | 0.14191 | -0.050354 | -0.064087 | -0.21973 | -0.15564 | 5.2689 | 5.3192 | 1.355 | 117.4255 | 1.4053 | 117.4759 | 5.1727 | -116.3177 | -116.3864 | -0.068664 | 5.4474 | 0.27466 | 5.4245 | 5.3192 | 0.25177 | 0.14648
2777.1719 | 121.9025 | -5.0995 | 116.6199 | -0.81718 | 0.61188 | 0.12708 | 0.087931 | -0.0023777 | 0.58635 | 0.007306 | -0.028778 | 0.019826 | 0.014078 | -0.0026041 | 0.014772 | -0.0047573 | -0.0010531 | -0.045776 | 116.5649 | -5.2094 | 116.3773 | 0.009155 | 112.3398 | -3.7582 | 0.013733 | 0.14191 | -0.13275 | 0.068665 | 0.16479 | 0.059509 | 0.13275 | -0.10986 | -0.054932 | -0.24262 | -0.18768 | 5.1086 | 5.2185 | 1.3412 | 117.4393 | 1.4511 | 117.5491 | 4.9667 | -116.478 | -116.5512 | -0.073242 | 5.2643 | 0.29755 | 5.2322 | 5.159 | 0.2655 | 0.19226
2778.1719 | 121.9528 | -5.127 | 116.6016 | -0.74852 | 0.60931 | 0.1258 | 0.085795 | -0.0021641 | 0.58528 | 0.0036744 | 0.074982 | 0.023031 | 0.014078 | -0.0034585 | 0.016054 | -0.0079617 | 0.0010831 | 0.018311 | 116.5329 | -5.1865 | 116.3773 | 0.022888 | 112.3169 | -3.8544 | 0 | 0.12817 | -0.12817 | 0.054932 | 0.1831 | 0.073242 | 0.14648 | -0.059509 | -0.068664 | -0.2243 | -0.15564 | 5.1498 | 5.2093 | 1.2726 | 117.4438 | 1.3321 | 117.5034 | 4.9988 | -116.4734 | -116.5466 | -0.073242 | 5.3101 | 0.31128 | 5.2734 | 5.2002 | 0.27466 | 0.20142
2779.1719 | 121.9208 | -5.3101 | 116.4505 | -0.75615 | 0.61059 | 0.12815 | 0.087077 | 0.0010403 | 0.58528 | 0.0060242 | -0.0074159 | 0.02239 | 0.012369 | -0.0013223 | 0.014345 | -0.0058255 | -0.0025485 | -0.013733 | 116.3406 | -5.3558 | 116.185 | -0.009155 | 112.0834 | -3.9825 | 0.022888 | 0.15106 | -0.10071 | 0.068665 | 0.16937 | 0.045776 | 0.15106 | -0.045776 | -0.10986 | -0.2655 | -0.15564 | 5.3009 | 5.3467 | 1.3275 | 117.3935 | 1.3733 | 117.4393 | 5.2094 | -116.2994 | -116.3818 | -0.082397 | 5.4794 | 0.27008 | 5.4611 | 5.3558 | 0.25177 | 0.14648
2780.1719 | 121.9391 | -5.2139 | 116.4917 | -0.7943 | 0.60632 | 0.12793 | 0.088145 | 0.00061311 | 0.58336 | 0.008374 | 0.010895 | 0.023244 | 0.012369 | 0.0025229 | 0.012636 | -0.0064663 | -0.0076755 | -0.022888 | 116.4642 | -5.278 | 116.2766 | -0.009155 | 112.1567 | -3.186 | 0.013733 | 0.15564 | -0.09613 | 0.054932 | 0.15564 | 0.050354 | 0.13275 | -0.064087 | -0.027466 | -0.21515 | -0.18768 | 5.2048 | 5.2689 | 2.0279 | 117.3706 | 2.092 | 117.4347 | 5.1178 | -116.3361 | -116.4368 | -0.10071 | 5.3696 | 0.25177 | 5.3467 | 5.2643 | 0.22888 | 0.14648
2781.1719 | 121.8979 | -5.2322 | 116.5009 | -0.74089 | 0.60974 | 0.12665 | 0.086436 | 0.0010403 | 0.58378 | 0.007306 | 0.0078429 | 0.023671 | 0.011514 | 0.0048727 | 0.015413 | -0.0077481 | -0.0072483 | -0.036621 | 116.4322 | -5.3192 | 116.2582 | -0.004578 | 112.1567 | -4.248 | 0.009155 | 0.14648 | -0.10986 | 0.054932 | 0.16022 | 0.036621 | 0.12817 | -0.086975 | -0.068664 | -0.24261 | -0.17395 | 5.2277 | 5.3146 | 0.98419 | 117.3889 | 1.0712 | 117.4759 | 5.1224 | -116.3544 | -116.4459 | -0.091552 | 5.3925 | 0.27008 | 5.3604 | 5.2689 | 0.23804 | 0.14648
2782.1719 | 121.8887 | -5.2551 | 116.4642 | -0.73326 | 0.60931 | 0.12387 | 0.087291 | -0.0034457 | 0.58378 | 0.0081604 | 0.0078429 | 0.023671 | 0.012369 | 0.0029501 | 0.014772 | -0.0045437 | 0.00086952 | -0.054932 | 116.423 | -5.3055 | 116.2354 | -0.013733 | 112.1063 | -3.9185 | 0 | 0.15106 | -0.10071 | 0.045776 | 0.15106 | 0.022888 | 0.12817 | -0.050354 | -0.041198 | -0.22888 | -0.18768 | 5.2414 | 5.2917 | 1.3367 | 117.3615 | 1.387 | 117.4118 | 5.1544 | -116.3132 | -116.4185 | -0.10529 | 5.4062 | 0.25177 | 5.3833 | 5.278 | 0.22888 | 0.1236
2783.1719 | 121.8155 | -5.3238 | 116.4917 | -0.718 | 0.60654 | 0.12366 | 0.088359 | -0.0013095 | 0.58314 | 0.0047424 | 0.0078429 | 0.020254 | 0.011942 | -0.0021769 | 0.012849 | -0.0071071 | -0.0031895 | 0.032043 | 116.423 | -5.3329 | 116.1667 | -0.07782 | 112.0468 | -4.0466 | 0.027466 | 0.18768 | -0.14648 | -0.059509 | 0.082397 | -0.050354 | 0.009155 | -0.009155 | -0.068664 | -0.32501 | -0.25635 | 5.246 | 5.2551 | 1.2772 | 117.3706 | 1.2863 | 117.3798 | 5.1773 | -116.304 | -116.5512 | -0.24719 | 5.4062 | 0.22888 | 5.3329 | 5.2734 | 0.15564 | 0.09613
2784.1719 | 121.8887 | -5.3558 | 116.4322 | -0.74089 | 0.60654 | 0.12366 | 0.08665 | 0.00061311 | 0.58208 | 0.0079468 | -0.0013123 | 0.020467 | 0.01301 | -0.0062357 | 0.015413 | -0.0053983 | -0.00062588 | -0.068665 | 116.4322 | -5.3146 | 116.1804 | -0.073242 | 112.0834 | -3.891 | -0.036621 | 0.15564 | -0.1236 | -0.25177 | 0.064087 | -0.091553 | 0.045776 | 0.041199 | 0 | -0.25177 | -0.25177 | 5.2826 | 5.2414 | 1.4648 | 117.4393 | 1.4236 | 117.3981 | 5.2322 | -116.2766 | -116.684 | -0.40741 | 5.4199 | 0.18768 | 5.4016 | 5.2643 | 0.16937 | 0.032043
2785.1719 | 121.9116 | -5.3421 | 116.4185 | -0.71037 | 0.60547 | 0.12516 | 0.087504 | -0.0019503 | 0.57994 | 0.0085876 | 0.026154 | 0.020894 | 0.0098052 | -0.0038857 | 0.01349 | -0.0058255 | -0.0025485 | -0.054932 | 116.391 | -5.3741 | 116.1392 | -0.059509 | 112.0422 | -3.3188 | -0.16937 | 0.10071 | -0.13733 | -0.013733 | 0.11902 | 0.009155 | 0.086975 | -0.032044 | -0.027466 | -0.27924 | -0.25177 | 5.2826 | 5.3146 | 2.0233 | 117.3843 | 2.0554 | 117.4164 | 5.2048 | -116.3177 | -116.4322 | -0.11444 | 5.4611 | 0.25635 | 5.4291 | 5.3513 | 0.2243 | 0.14648
2786.1719 | 121.9025 | -5.2734 | 116.4642 | -0.718 | 0.60568 | 0.12238 | 0.087931 | -0.00088229 | 0.57951 | 0.0051696 | 0.026154 | 0.022817 | 0.01365 | 0.0016685 | 0.015199 | -0.0053983 | -0.0044711 | -0.054932 | 116.4185 | -5.3101 | 116.0614 | -0.086975 | 111.9278 | -3.9597 | -0.009155 | 0.091553 | -0.10529 | 0.009155 | 0.14191 | 0.032043 | 0.10071 | -0.036622 | -0.045776 | -0.40283 | -0.35706 | 5.1865 | 5.2231 | 1.3138 | 117.2012 | 1.3504 | 117.2379 | 5.1682 | -116.3727 | -116.4551 | -0.082398 | 5.4153 | 0.24719 | 5.3741 | 5.3055 | 0.20599 | 0.13733
2787.1719 | 121.8933 | -5.3329 | 116.3818 | -0.718 | 0.60611 | 0.12601 | 0.088786 | -0.0051547 | 0.57994 | 0.0053832 | -0.0043641 | 0.021535 | 0.012796 | -0.0043131 | 0.013063 | -0.0056119 | -0.0023349 | -0.032043 | 116.3086 | -5.4108 | 116.1438 | -0.073242 | 112.1017 | -4.01 | 0.018311 | 0.14648 | -0.10986 | 0.041199 | 0.15106 | 0.054932 | 0.10529 | -0.07782 | -0.073242 | -0.23804 | -0.16479 | 5.2597 | 5.3375 | 1.3229 | 117.4347 | 1.4008 | 117.5125 | 5.2231 | -116.2354 | -116.3406 | -0.10529 | 5.484 | 0.26093 | 5.4382 | 5.3879 | 0.21515 | 0.16479
2788.1719 | 121.8384 | -5.3787 | 116.391 | -0.72563 | 0.60504 | 0.12366 | 0.087504 | 0.0016813 | 0.57994 | 0.0077332 | -0.0074159 | 0.020894 | 0.015359 | -0.0019631 | 0.011995 | -0.0041165 | -0.0076755 | -0.07782 | 116.4047 | -5.3604 | 116.1713 | -0.068665 | 112.0377 | -4.0466 | -0.045776 | 0.14648 | -0.10529 | -0.032043 | -0.13733 | -0.09613 | 0.064087 | 0.01831 | 0.013733 | -0.21973 | -0.23346 | 5.3101 | 5.2917 | 1.3321 | 117.4164 | 1.3138 | 117.3981 | 5.2734 | -116.2445 | -116.423 | -0.17853 | 5.2414 | -0.032043 | 5.4428 | 5.2826 | 0.16937 | 0.009156
2789.1719 | 121.9391 | -5.0125 | 116.684 | -0.71037 | 0.60376 | 0.1243 | 0.088572 | -0.0015231 | 0.57866 | 0.0077332 | -0.010467 | 0.023031 | 0.010873 | -0.0034585 | 0.015199 | -0.0086025 | -0.0089573 | 0.032043 | 116.6702 | -5.0903 | 116.4688 | 0.041199 | 112.4542 | -3.5065 | 0.15106 | 0.13733 | -0.11902 | 0.068665 | 0.19684 | 0.07782 | 0.12817 | -0.07782 | -0.013733 | -0.21515 | -0.20142 | 5.0537 | 5.1315 | 1.506 | 117.4667 | 1.5839 | 117.5446 | 4.8935 | -116.5466 | -116.6153 | -0.068664 | 5.2093 | 0.31586 | 5.1407 | 5.0903 | 0.24719 | 0.19684
2790.1719 | 121.9666 | -5.0446 | 116.6977 | -0.73326 | 0.60504 | 0.12516 | 0.090281 | -0.0019503 | 0.57973 | 0.005597 | 0.016998 | 0.024312 | 0.011728 | 0.0082907 | 0.01349 | -0.0049709 | -0.0027621 | 0.050354 | 116.6977 | -5.0766 | 116.5146 | 0.073242 | 112.4634 | -3.7125 | 0.12817 | 0.15106 | -0.11444 | 0.064087 | 0.20599 | 0.082397 | 0.13275 | -0.032043 | 0 | -0.18311 | -0.18311 | 5.1178 | 5.1498 | 1.3321 | 117.5079 | 1.3641 | 117.54 | 4.9301 | -116.5466 | -116.6336 | -0.086975 | 5.2506 | 0.32044 | 5.1773 | 5.127 | 0.24719 | 0.19684
2791.1719 | 121.8887 | -5.0262 | 116.7664 | -0.65697 | 0.6029 | 0.12451 | 0.087504 | 0.0029629 | 0.57844 | 0.0064514 | -0.025726 | 0.021963 | 0.011728 | -0.0043131 | 0.014772 | -0.0073209 | -0.0044711 | 0.068665 | 116.6977 | -5.0446 | 116.51 | 0.004578 | 112.5366 | -3.6758 | 0.11902 | 0.15564 | -0.10986 | 0.027466 | 0.15564 | 0.045776 | 0.09613 | -0.018311 | -0.068664 | -0.25635 | -0.18768 | 5.0308 | 5.0491 | 1.3504 | 117.5629 | 1.3687 | 117.5812 | 4.9164 | -116.6107 | -116.7389 | -0.12817 | 5.1819 | 0.2655 | 5.1224 | 5.072 | 0.20599 | 0.15564
2792.1719 | 121.907 | -4.9759 | 116.7709 | -0.67223 | 0.60312 | 0.12451 | 0.087718 | -0.00066869 | 0.57481 | 0.008374 | -0.016571 | 0.022817 | 0.012155 | 0.0008139 | 0.012849 | -0.0071071 | -0.0023349 | 0.045776 | 116.7206 | -4.9942 | 116.5375 | -0.16022 | 112.4908 | -3.5202 | 0.12817 | 0.13275 | -0.1236 | 0.050354 | 0.1831 | 0.09613 | 0.13275 | -0.018311 | -0.050354 | -0.23346 | -0.1831 | 4.8157 | 4.834 | 1.4557 | 117.4667 | 1.474 | 117.485 | 4.8523 | -116.6382 | -116.7206 | -0.082397 | 5.159 | 0.3067 | 5.1086 | 5.072 | 0.25635 | 0.21973
2793.1719 | 121.8613 | -4.985 | 116.8213 | -0.63408 | 0.60269 | 0.12601 | 0.088786 | -0.00045509 | 0.57545 | 0.0081604 | -0.10202 | 0.02239 | 0.009378 | 0.0023093 | 0.015413 | -0.0028347 | -0.0044711 | 0.022888 | 116.7206 | -5.0491 | 116.4825 | -0.10071 | 112.4405 | -3.7582 | 0.073242 | 0.16479 | -0.15564 | -0.004578 | 0.09613 | -0.032043 | 0.041199 | -0.064087 | -0.10071 | -0.33874 | -0.23804 | 4.8843 | 4.9484 | 1.2268 | 117.4255 | 1.2909 | 117.4896 | 4.8294 | -116.6565 | -116.8259 | -0.16937 | 5.0812 | 0.25177 | 5.0262 | 4.953 | 0.19684 | 0.1236
2794.1719 | 121.9208 | -4.9164 | 116.8121 | -0.718 | 0.60162 | 0.12473 | 0.088786 | 0.0023221 | 0.57716 | 0.006665 | 0.02005 | 0.020894 | 0.017496 | 0.0067953 | 0.014986 | -0.0051845 | -0.00062588 | 0.041199 | 116.7755 | -4.9484 | 116.597 | 0.036621 | 112.587 | -3.5843 | 0.15564 | 0.13733 | -0.11902 | 0.086975 | 0.13733 | 0.10529 | 0.14191 | -0.032043 | -0.036621 | -0.21515 | -0.17853 | 4.953 | 4.985 | 1.3321 | 117.5034 | 1.3641 | 117.5354 | 4.7974 | -116.6748 | -116.7252 | -0.050354 | 5.0537 | 0.25635 | 5.0583 | 5.0217 | 0.26093 | 0.22431
2795.1719 | 121.9116 | -4.921 | 116.8167 | -0.68748 | 0.59991 | 0.12665 | 0.08964 | 0.0048855 | 0.57353 | 0.0043152 | 0.023102 | 0.021749 | 0.014291 | -0.0021769 | 0.01349 | -0.0053983 | -0.0040439 | 0.004578 | 116.7618 | -4.9942 | 116.6245 | 0.004578 | 112.4496 | -3.5614 | 0.1236 | 0.13733 | -0.11902 | 0.045776 | 0.17395 | 0.059509 | 0.13275 | -0.073243 | -0.054931 | -0.19226 | -0.13733 | 4.9255 | 4.9988 | 1.3596 | 117.3706 | 1.4328 | 117.4438 | 4.8019 | -116.6794 | -116.7709 | -0.091553 | 5.0949 | 0.29297 | 5.0537 | 4.9805 | 0.25177 | 0.17853
2796.1719 | 121.8292 | -5.0171 | 116.803 | -0.73326 | 0.60184 | 0.12665 | 0.088572 | -0.00066869 | 0.57802 | 0.0081604 | -0.03183 | 0.019613 | 0.012155 | 0.0057273 | 0.016267 | -0.0047573 | -0.0021213 | 0.041199 | 116.7435 | -5.0217 | 116.478 | -0.059509 | 112.5 | -3.6713 | 0.045776 | 0.16022 | -0.13733 | -0.064087 | 0.07782 | -0.041199 | 0.032043 | -0.004577 | -0.05951 | -0.32501 | -0.2655 | 4.9576 | 4.9622 | 1.3458 | 117.5171 | 1.3504 | 117.5217 | 4.8798 | -116.6428 | -116.8671 | -0.2243 | 5.0949 | 0.21515 | 5.0491 | 4.9759 | 0.16937 | 0.09613
2797.1719 | 121.8018 | -4.953 | 116.8945 | -0.6646 | 0.60034 | 0.12537 | 0.090068 | 0.0021085 | 0.57631 | 0.0062378 | -0.037933 | 0.024099 | 0.014505 | 0.0008139 | 0.016481 | -0.0043301 | -0.0019077 | 0.009155 | 116.8213 | -4.9988 | 116.5878 | -0.19226 | 112.3627 | -3.8132 | -0.018311 | 0.23346 | -0.11902 | -0.050354 | 0.054932 | -0.054932 | 0.032043 | -0.045776 | -0.073242 | -0.3067 | -0.23346 | 4.7607 | 4.8065 | 1.1398 | 117.3157 | 1.1856 | 117.3615 | 4.834 | -116.6611 | -116.9449 | -0.28381 | 5.0079 | 0.17395 | 4.985 | 4.8981 | 0.15106 | 0.064087
2798.1719 | 121.8246 | -4.9622 | 116.8213 | -0.72563 | 0.60034 | 0.12686 | 0.090922 | 0.0038175 | 0.5746 | 0.0088012 | -0.092865 | 0.024312 | 0.011728 | -0.0011087 | 0.015626 | -0.0049709 | -0.0025485 | -0.022888 | 116.7801 | -5.0034 | 116.4917 | -0.07782 | 112.4084 | -4.0009 | 0 | 0.16022 | -0.14191 | -0.041199 | 0.064087 | -0.068665 | 0.036621 | -0.041199 | -0.041199 | -0.32959 | -0.28839 | 4.8843 | 4.9255 | 0.9613 | 117.3706 | 1.0025 | 117.4118 | 4.8203 | -116.6611 | -116.8625 | -0.20142 | 5.0262 | 0.20599 | 4.9988 | 4.8935 | 0.17853 | 0.073242
2799.1719 | 121.8567 | -4.9438 | 116.8076 | -0.73326 | 0.59991 | 0.12644 | 0.088572 | 0.00082671 | 0.57588 | 0.0085876 | -0.053192 | 0.022817 | 0.014719 | 0.0008139 | 0.016054 | -0.0068935 | -0.00083948 | 0 | 116.8076 | -5.0125 | 116.5054 | -0.07782 | 112.4817 | -3.7079 | 0.054932 | 0.16022 | -0.13275 | 0 | 0.091553 | 0 | 0.054932 | -0.068664 | 0 | -0.30212 | -0.30212 | 4.866 | 4.9347 | 1.236 | 117.4255 | 1.3046 | 117.4942 | 4.8111 | -116.6473 | -116.8076 | -0.16022 | 5.0354 | 0.2243 | 4.9988 | 4.9438 | 0.18768 | 0.13275
2800.1719 | 121.8521 | -4.9988 | 116.7801 | -0.67223 | 0.59927 | 0.12409 | 0.08964 | 0.0044583 | 0.57524 | 0.0090148 | -0.044037 | 0.022603 | 0.012796 | 0.0014547 | 0.013063 | -0.0034757 | -0.0046847 | -0.027466 | 116.7664 | -5.0308 | 116.4368 | -0.11902 | 112.4359 | -3.6026 | -0.013733 | 0.17853 | -0.15106 | -0.045776 | 0.07782 | -0.054932 | 0.032043 | -0.032044 | -0.013733 | -0.34332 | -0.32959 | 4.8798 | 4.9118 | 1.3962 | 117.4347 | 1.4282 | 117.4667 | 4.8477 | -116.6016 | -116.8259 | -0.2243 | 5.0766 | 0.22888 | 5.0308 | 4.9438 | 0.18311 | 0.09613
2801.1719 | 121.8201 | -5.0491 | 116.8259 | -0.67986 | 0.59671 | 0.12451 | 0.084727 | -0.00066869 | 0.57289 | 0.0077332 | -0.010467 | 0.024526 | 0.012582 | 0.0008139 | 0.015413 | -0.0066799 | -0.0029757 | -0.018311 | 116.8259 | -4.9988 | 116.4001 | -0.13275 | 112.4542 | -3.7628 | -0.018311 | 0.20599 | -0.1236 | -0.068665 | 0.07782 | -0.054932 | 0.018311 | 0.050354 | 0 | -0.42572 | -0.42572 | 4.9164 | 4.866 | 1.2863 | 117.5034 | 1.236 | 117.453 | 4.9255 | -116.6199 | -116.8945 | -0.27466 | 5.127 | 0.20142 | 5.0674 | 4.9942 | 0.14191 | 0.068664
2802.1719 | 121.7972 | -4.9713 | 116.8442 | -0.51201 | 0.59799 | 0.12366 | 0.087504 | 0.0033901 | 0.57545 | 0.0062378 | 0.0017395 | 0.022817 | 0.014505 | -0.0060221 | 0.015199 | -0.0062527 | 1.4916e-05 | -0.004578 | 116.8716 | -4.9484 | 116.4459 | -0.10986 | 112.4313 | -3.6896 | -0.032043 | 0.16937 | -0.13733 | -0.068665 | 0.073242 | -0.059509 | 0.013733 | 0.022888 | 0.027466 | -0.39825 | -0.42572 | 4.8615 | 4.8386 | 1.2817 | 117.4026 | 1.2589 | 117.3798 | 4.834 | -116.6748 | -116.9128 | -0.23804 | 5.0446 | 0.21057 | 4.985 | 4.9118 | 0.15106 | 0.07782
2803.1719 | 121.8201 | -4.953 | 116.8579 | -0.74852 | 0.59842 | 0.12622 | 0.089213 | 0.0018949 | 0.5731 | 0.004956 | -0.08371 | 0.019826 | 0.012582 | 0.0029501 | 0.014131 | -0.0017667 | 1.4916e-05 | -0.10071 | 116.748 | -5.0949 | 116.4093 | -0.15106 | 112.4359 | -3.6942 | 0.004578 | 0.17853 | -0.13275 | -0.068665 | 0.050354 | -0.086975 | 0.018311 | -0.14191 | -0.10986 | -0.44861 | -0.33875 | 4.8019 | 4.9438 | 1.2588 | 117.3889 | 1.4008 | 117.5308 | 4.8203 | -116.6794 | -116.9266 | -0.24719 | 5.0034 | 0.18311 | 4.9713 | 4.866 | 0.15106 | 0.045776
2804.1719 | 121.8475 | -4.8981 | 116.8579 | -0.74852 | 0.59628 | 0.12345 | 0.089213 | 0.0040311 | 0.57417 | 0.005597 | -0.044037 | 0.024526 | 0.012369 | -0.0008951 | 0.015199 | -0.0045437 | 0.00065592 | 0.022888 | 116.8167 | -4.953 | 116.5237 | -0.064087 | 112.4359 | -3.5248 | 0.050354 | 0.19684 | -0.045776 | 0.004578 | 0.10529 | -0.009155 | 0.059509 | -0.054932 | -0.041199 | -0.33417 | -0.29297 | 4.834 | 4.8889 | 1.3733 | 117.334 | 1.4282 | 117.3889 | 4.8523 | -116.6611 | -116.8533 | -0.19226 | 5.0034 | 0.15106 | 4.9576 | 4.8889 | 0.10528 | 0.036621
2805.1719 | 121.8292 | -4.9164 | 116.8716 | -0.718 | 0.59564 | 0.12708 | 0.088359 | 0.0021085 | 0.57417 | 0.0079468 | -0.040985 | 0.023671 | 0.013437 | 0.0033773 | 0.014558 | -0.0041165 | -0.0038303 | 0.004578 | 116.8259 | -4.9438 | 116.6336 | 0.091553 | 112.5549 | -3.5934 | 0.009155 | 0.22888 | 0 | -0.032043 | 0.10071 | -0.045776 | 0.059509 | -0.027466 | -0.045776 | -0.23804 | -0.19226 | 5.0079 | 5.0354 | 1.3229 | 117.4713 | 1.3504 | 117.4988 | 4.9164 | -116.6428 | -116.9037 | -0.26093 | 5.0171 | 0.10071 | 4.9759 | 4.8706 | 0.059509 | -0.045776
2806.1719 | 121.8475 | -4.9301 | 116.835 | -0.718 | 0.59628 | 0.12622 | 0.089427 | 0.00082671 | 0.57204 | 0.0064514 | -0.040985 | 0.022817 | 0.013223 | 0.0042319 | 0.016481 | -0.0051845 | -0.0048983 | 0.009155 | 116.7847 | -5.0034 | 116.5329 | -0.068665 | 112.532 | -5.3238 | 0.009155 | 0.20599 | -0.07782 | -0.018311 | 0.086975 | -0.018311 | 0.050354 | -0.073242 | -0.050354 | -0.30212 | -0.25177 | 4.8614 | 4.9347 | -0.39368 | 117.4622 | -0.32043 | 117.5354 | 4.8523 | -116.629 | -116.8533 | -0.22431 | 5.0171 | 0.16479 | 4.9805 | 4.9118 | 0.12817 | 0.059509
2807.1719 | 121.7926 | -4.9393 | 116.8625 | -0.69511 | 0.59479 | 0.12601 | 0.087931 | 0.0048855 | 0.57332 | 0.010083 | -0.03183 | 0.023458 | 0.012155 | -0.0004679 | 0.016054 | -0.0043301 | 0.0012967 | -0.054932 | 116.8213 | -4.985 | 116.4825 | -0.11902 | 112.4817 | -3.6987 | -0.059509 | 0.14648 | -0.14191 | -0.064087 | 0.018311 | -0.10529 | 0.009155 | -0.045776 | -0.041199 | -0.37994 | -0.33874 | 4.8203 | 4.866 | 1.2405 | 117.421 | 1.2863 | 117.4667 | 4.7974 | -116.716 | -116.9266 | -0.21057 | 4.9576 | 0.16022 | 4.9484 | 4.834 | 0.15106 | 0.036621
2808.1719 | 121.8887 | -4.9438 | 116.7389 | -0.72563 | 0.59607 | 0.12494 | 0.087504 | -0.0032321 | 0.5746 | 0.010297 | -0.0013123 | 0.023458 | 0.014932 | -0.0002543 | 0.016267 | -0.0051845 | -0.0012667 | -0.091553 | 116.7572 | -5.0262 | 116.4597 | -0.082397 | 112.3764 | -3.717 | -0.041199 | 0.13733 | -0.13275 | 0.050354 | 0.16479 | 0.009155 | 0.13275 | -0.082397 | 0.01831 | -0.27924 | -0.29755 | 4.8615 | 4.9438 | 1.2268 | 117.3203 | 1.3092 | 117.4026 | 4.8111 | -116.6016 | -116.6885 | -0.086975 | 5.1086 | 0.29755 | 5.0766 | 4.953 | 0.2655 | 0.14191
2809.1719 | 121.8475 | -4.9805 | 116.7984 | -0.69511 | 0.59457 | 0.12238 | 0.088359 | 0.0010403 | 0.57225 | 0.0062378 | -0.013519 | 0.022176 | 0.016 | -0.0026041 | 0.013063 | -0.0030483 | -0.0016941 | 0.009155 | 116.7206 | -5.0583 | 116.4505 | -0.086975 | 112.4542 | -3.653 | 0.041199 | 0.15106 | -0.11902 | -0.018311 | 0.10071 | -0.013733 | 0.059509 | -0.07782 | -0.07782 | -0.3479 | -0.27008 | 4.8935 | 4.9713 | 1.3275 | 117.4347 | 1.4053 | 117.5125 | 4.8615 | -116.6473 | -116.8167 | -0.16937 | 5.0812 | 0.21973 | 5.04 | 4.9667 | 0.17853 | 0.10529
2810.1719 | 121.8018 | -5.0079 | 116.7847 | -0.71037 | 0.59671 | 0.12558 | 0.088145 | 0.0061673 | 0.5684 | 0.007306 | -0.03183 | 0.02239 | 0.01301 | 0.0095725 | 0.014772 | -0.0051845 | -0.0014805 | -0.082397 | 116.7389 | -5.04 | 116.423 | -0.15106 | 112.4359 | -3.7857 | -0.10529 | 0.14191 | -0.15564 | -0.082397 | 0.022888 | -0.082397 | 0.009155 | -0.032043 | -0.045776 | -0.36163 | -0.31586 | 4.8569 | 4.8889 | 1.2222 | 117.4438 | 1.2543 | 117.4759 | 4.8523 | -116.6428 | -116.8671 | -0.2243 | 5.0308 | 0.17853 | 5.0171 | 4.9255 | 0.16479 | 0.073243
2811.1719 | 121.7834 | -4.985 | 116.803 | -0.68748 | 0.59543 | 0.12323 | 0.089 | 0.0038175 | 0.57097 | 0.0090148 | -0.019623 | 0.025808 | 0.01365 | 0.0018821 | 0.012849 | -0.0036893 | -0.0027621 | -0.10529 | 116.7297 | -5.0629 | 116.4093 | -0.15106 | 112.4863 | -3.6118 | -0.11902 | 0.16479 | -0.15564 | -0.068665 | 0.013733 | -0.11444 | 0.013733 | -0.07782 | -0.073243 | -0.39368 | -0.32043 | 4.834 | 4.9118 | 1.3733 | 117.4713 | 1.4511 | 117.5491 | 4.8294 | -116.6382 | -116.8716 | -0.23346 | 4.9988 | 0.16937 | 4.9988 | 4.8706 | 0.16937 | 0.041199
2812.1719 | 121.7926 | -4.9713 | 116.803 | -0.71037 | 0.59372 | 0.1243 | 0.089854 | 0.0014675 | 0.56862 | 0.010083 | -0.010467 | 0.02239 | 0.014291 | 0.0038047 | 0.016267 | -0.0062527 | -0.0021213 | -0.068665 | 116.7938 | -5.0125 | 116.4368 | -0.16022 | 112.5 | -3.7308 | -0.11444 | 0.16479 | -0.13733 | -0.086975 | 0.013733 | -0.09613 | 0.004578 | -0.041199 | -0.009156 | -0.36621 | -0.35706 | 4.8111 | 4.8523 | 1.2405 | 117.4713 | 1.2817 | 117.5125 | 4.834 | -116.6382 | -116.89 | -0.25177 | 4.985 | 0.15106 | 4.9759 | 4.8752 | 0.14191 | 0.041199
2813.1875 | 121.8155 | -4.9255 | 116.8625 | -0.61119 | 0.59393 | 0.12302 | 0.087291 | 0.0023221 | 0.56947 | 0.0060242 | 0.065826 | 0.02239 | 0.012796 | -0.0002543 | 0.013917 | -0.0060391 | 0.0015103 | -0.041199 | 116.8167 | -4.9805 | 116.4688 | -0.13275 | 112.5641 | -3.6255 | 0.1831 | 0.19684 | -0.12817 | -0.059509 | 0.027466 | -0.036621 | 0.045776 | -0.054932 | -0.045777 | -0.39368 | -0.3479 | 4.7928 | 4.8477 | 1.3 | 117.4896 | 1.355 | 117.5446 | 4.7974 | -116.6656 | -116.922 | -0.25635 | 4.953 | 0.15564 | 4.9713 | 4.8889 | 0.17395 | 0.091553
2814.1719 | 121.788 | -4.9713 | 116.835 | -0.71037 | 0.59265 | 0.12537 | 0.08964 | 0.0016813 | 0.57097 | 0.0068786 | -0.03183 | 0.023671 | 0.01365 | 0.0029501 | 0.015413 | -0.0045437 | -0.0019077 | -0.09613 | 116.7572 | -5.0217 | 116.3498 | -0.16022 | 112.4176 | -3.6346 | -0.14191 | 0.16479 | -0.12817 | -0.086975 | 0.022888 | -0.1236 | 0.009155 | -0.050354 | -0.07782 | -0.48523 | -0.40741 | 4.8111 | 4.8614 | 1.3367 | 117.3889 | 1.387 | 117.4393 | 4.8431 | -116.6702 | -116.922 | -0.25177 | 4.9942 | 0.15106 | 4.9805 | 4.8477 | 0.13733 | 0.004578
2815.1719 | 121.7972 | -4.8294 | 116.9724 | -0.718 | 0.59372 | 0.1243 | 0.088145 | 0.00018591 | 0.57139 | 0.0053832 | -0.040985 | 0.023458 | 0.014505 | -0.0006815 | 0.01584 | -0.0058255 | -0.0023349 | -0.054932 | 116.9586 | -4.8889 | 116.5649 | -0.16022 | 112.5092 | -3.6072 | -0.14191 | 0.16937 | -0.16479 | -0.091553 | 0.022888 | -0.11444 | -0.004578 | -0.059509 | -0.013733 | -0.40741 | -0.39368 | 4.6692 | 4.7287 | 1.2222 | 117.3386 | 1.2817 | 117.3981 | 4.6646 | -116.803 | -117.0639 | -0.26093 | 4.8523 | 0.18768 | 4.8248 | 4.715 | 0.16022 | 0.050354
2816.1719 | 121.8384 | -4.5822 | 117.2012 | -0.71037 | 0.59201 | 0.12494 | 0.088786 | -0.00088229 | 0.56648 | 0.0034606 | -0.0074159 | 0.022176 | 0.014078 | -0.0038857 | 0.014131 | -0.0053983 | 1.4916e-05 | -0.045776 | 117.1646 | -4.6188 | 116.8854 | -0.086975 | 112.912 | -4.9393 | -0.041199 | 0.15564 | -0.15564 | -0.050354 | 0.018311 | -0.07782 | 0.032043 | -0.036621 | -0.036621 | -0.31586 | -0.27924 | 4.4952 | 4.5319 | -0.35706 | 117.4942 | -0.32043 | 117.5308 | 4.4266 | -117.0456 | -117.2516 | -0.20599 | 4.6005 | 0.17395 | 4.6143 | 4.5044 | 0.18768 | 0.07782
2817.1719 | 121.7972 | -4.6005 | 117.1738 | -0.718 | 0.59094 | 0.12409 | 0.088145 | 0.0016813 | 0.56669 | 0.0041016 | 0.0078429 | 0.021535 | 0.012796 | 0.0035911 | 0.015626 | -0.0051845 | -0.0010531 | -0.07782 | 117.128 | -4.6555 | 116.7755 | -0.15106 | 112.793 | -3.3646 | -0.13733 | 0.16022 | -0.15106 | -0.07782 | 0.004578 | -0.11444 | 0.004578 | -0.054932 | -0.045776 | -0.39825 | -0.35248 | 4.4495 | 4.5044 | 1.236 | 117.3935 | 1.2909 | 117.4484 | 4.4495 | -117.0135 | -117.2516 | -0.23804 | 4.6051 | 0.15564 | 4.6051 | 4.4861 | 0.15564 | 0.036621
2818.1875 | 121.9345 | -4.5273 | 117.1692 | -0.71037 | 0.59265 | 0.12409 | 0.083445 | -0.0030185 | 0.56776 | 0.006665 | 0.11465 | 0.021322 | 0.013223 | -0.0053811 | 0.017549 | -0.0086025 | -0.0034031 | -0.009155 | 117.1738 | -4.5776 | 116.9449 | 0.036621 | 113.0081 | -2.5909 | 0.086975 | 0.14648 | -0.10529 | 0.027466 | 0.19226 | 0.086975 | 0.14191 | -0.050354 | 0.004578 | -0.2243 | -0.22888 | 4.5639 | 4.6143 | 1.9363 | 117.5354 | 1.9867 | 117.5858 | 4.422 | -117.0227 | -117.1417 | -0.11902 | 4.7195 | 0.29755 | 4.6692 | 4.6143 | 0.24719 | 0.19226
2819.1719 | 121.8109 | -4.5959 | 117.1783 | -0.68748 | 0.59201 | 0.12451 | 0.090709 | 0.0012539 | 0.56734 | 0.0081604 | -0.053192 | 0.020467 | 0.011728 | 0.0035911 | 0.013917 | -0.0045437 | -0.0021213 | -0.10071 | 117.1188 | -4.66 | 116.7389 | -0.25635 | 113.1683 | -3.2181 | -0.10529 | 0.18768 | -0.13733 | -0.064087 | 0.027466 | -0.10071 | 0.018311 | -0.064087 | -0.05951 | -0.43945 | -0.37994 | 4.3396 | 4.4037 | 1.3779 | 117.7643 | 1.442 | 117.8284 | 4.4586 | -116.9907 | -117.2424 | -0.25177 | 4.6234 | 0.16479 | 4.6143 | 4.4952 | 0.15564 | 0.036621
2820.1719 | 121.8109 | -4.6051 | 117.2012 | -0.78667 | 0.5918 | 0.12494 | 0.088786 | 0.0023221 | 0.57075 | 0.0088012 | -0.044037 | 0.018758 | 0.01365 | -0.0047403 | 0.011995 | -0.0045437 | -0.0038303 | -0.09613 | 117.1371 | -4.6692 | 116.7435 | -0.1831 | 113.1454 | -3.3417 | -0.091553 | 0.17853 | -0.12817 | -0.073242 | 0.027466 | -0.082397 | 0.027466 | -0.064086 | -0.064087 | -0.45776 | -0.39368 | 4.422 | 4.4861 | 1.2634 | 117.7506 | 1.3275 | 117.8146 | 4.4769 | -117.0227 | -117.2745 | -0.25177 | 4.6326 | 0.15564 | 4.6326 | 4.5227 | 0.15564 | 0.045777
2821.1719 | 121.8292 | -4.6646 | 117.1005 | -0.72563 | 0.58945 | 0.12409 | 0.089854 | 0.0040311 | 0.56627 | 0.004956 | 0.02005 | 0.021322 | 0.01301 | 0.0044455 | 0.015626 | -0.0051845 | -0.0027621 | -0.10529 | 117.0593 | -4.7333 | 116.8396 | 0.036621 | 112.6648 | -3.3508 | -0.16479 | 0.1831 | -0.13275 | -0.050354 | 0.045776 | -0.091553 | 0.022888 | -0.068664 | -0.041199 | -0.26093 | -0.21973 | 4.7012 | 4.7699 | 1.3138 | 117.3294 | 1.3824 | 117.3981 | 4.5319 | -116.9174 | -117.1509 | -0.23346 | 4.7104 | 0.17853 | 4.6875 | 4.5731 | 0.15564 | 0.041198
2822.1719 | 121.8292 | -4.6188 | 117.1371 | -0.71037 | 0.59009 | 0.12494 | 0.09199 | 0.0050991 | 0.56328 | 0.0096558 | 0.016998 | 0.021749 | 0.012369 | 0.0020957 | 0.014345 | -0.0047573 | -0.0059665 | -0.091553 | 117.1417 | -4.6875 | 116.8671 | -0.10071 | 112.7197 | -3.6392 | -0.15564 | 0.18768 | -0.12817 | -0.050354 | 0.032043 | -0.07782 | 0.004578 | -0.068665 | 0.004578 | -0.27008 | -0.27466 | 4.5181 | 4.5868 | 0.97961 | 117.3386 | 1.0483 | 117.4072 | 4.4907 | -116.9495 | -117.1875 | -0.23804 | 4.6509 | 0.16022 | 4.6234 | 4.541 | 0.13275 | 0.050354
2823.1875 | 121.8109 | -4.6097 | 117.1555 | -0.67986 | 0.58603 | 0.12451 | 0.088786 | 0.0018949 | 0.56499 | 0.0058106 | 0.07193 | 0.020467 | 0.014505 | -0.0032449 | 0.014986 | -0.0049709 | -0.0034031 | -0.022888 | 117.1463 | -4.6738 | 116.9357 | -0.013733 | 112.6236 | -3.3371 | -0.036621 | 0.21057 | -0.10529 | -0.073242 | 0.059509 | -0.059509 | 0.022888 | -0.064087 | -0.009156 | -0.21973 | -0.21057 | 4.5959 | 4.66 | 1.2726 | 117.2333 | 1.3367 | 117.2974 | 4.5044 | -116.9449 | -117.2287 | -0.28381 | 4.6692 | 0.16479 | 4.6326 | 4.5502 | 0.12817 | 0.045777
2824.1719 | 121.8109 | -4.6234 | 117.1646 | -0.71037 | 0.58923 | 0.1258 | 0.089 | -0.0047275 | 0.56648 | 0.0070922 | -0.028778 | 0.021535 | 0.01301 | -0.0038857 | 0.014986 | -0.0034757 | -0.0021213 | -0.10529 | 117.1097 | -4.7012 | 116.716 | -0.1236 | 112.9074 | -3.4195 | -0.11902 | 0.16479 | -0.13733 | -0.059509 | 0.050354 | -0.09613 | 0.032043 | -0.07782 | -0.054932 | -0.44861 | -0.39368 | 4.4998 | 4.5776 | 1.2039 | 117.5308 | 1.2817 | 117.6086 | 4.4861 | -116.9998 | -117.2241 | -0.2243 | 4.6738 | 0.18768 | 4.6555 | 4.5273 | 0.16937 | 0.041199
2825.1719 | 121.9162 | -4.6188 | 117.0364 | -0.72563 | 0.58859 | 0.12686 | 0.088786 | 0.0012539 | 0.56605 | 0.0075196 | 0.0078429 | 0.023458 | 0.011087 | -0.0026041 | 0.014131 | -0.0058255 | 0.00086952 | -0.064087 | 117.0456 | -4.6646 | 116.8442 | 0 | 112.825 | -3.3188 | -0.013733 | 0.14191 | -0.11444 | 0.064087 | 0.16022 | 0.036621 | 0.14648 | -0.045777 | 0.009155 | -0.19226 | -0.20142 | 4.6188 | 4.6646 | 1.3 | 117.4438 | 1.3458 | 117.4896 | 4.5044 | -116.8945 | -116.9724 | -0.07782 | 4.7791 | 0.27466 | 4.7653 | 4.6555 | 0.26093 | 0.15106
2826.1719 | 121.8933 | -4.5639 | 117.1692 | -0.68748 | 0.58752 | 0.12451 | 0.089854 | -0.00045509 | 0.56477 | 0.0053832 | 0.0017395 | 0.022176 | 0.011514 | 0.0031637 | 0.012636 | -0.0068935 | -0.0031895 | -0.10529 | 117.1646 | -4.628 | 116.8442 | -0.07782 | 112.8616 | -3.36 | -0.082397 | 0.15106 | -0.14191 | 0.009155 | 0.11444 | 0.004578 | 0.091553 | -0.064087 | -0.004577 | -0.32501 | -0.32044 | 4.4861 | 4.5502 | 1.2039 | 117.4255 | 1.268 | 117.4896 | 4.422 | -117.0181 | -117.16 | -0.14191 | 4.6783 | 0.25635 | 4.6555 | 4.5685 | 0.23346 | 0.14649
2827.1719 | 121.8979 | -4.4128 | 117.3157 | -0.69511 | 0.58667 | 0.12686 | 0.088572 | 0.0023221 | 0.56456 | 0.0096558 | 0.013947 | 0.023031 | 0.012155 | -0.0028177 | 0.013277 | -0.0051845 | -0.00083948 | -0.09613 | 117.3248 | -4.4815 | 116.9861 | -0.07782 | 113.0127 | -3.1723 | -0.086975 | 0.14648 | -0.14648 | 0.013733 | 0.10529 | -0.018311 | 0.10529 | -0.068664 | 0.009155 | -0.32959 | -0.33874 | 4.335 | 4.4037 | 1.2405 | 117.4255 | 1.3092 | 117.4942 | 4.2664 | -117.1692 | -117.3019 | -0.13275 | 4.5181 | 0.25177 | 4.5181 | 4.3945 | 0.25177 | 0.12817
2828.1719 | 121.8796 | -4.4266 | 117.2974 | -0.72563 | 0.58774 | 0.1243 | 0.085582 | 0.0010403 | 0.56306 | 0.0077332 | 0.059723 | 0.021535 | 0.011514 | -0.0011087 | 0.01584 | -0.0058255 | -0.0010531 | -0.064087 | 117.2745 | -4.4769 | 117.0044 | 0.009155 | 113.0356 | -3.0121 | -0.036621 | 0.15564 | -0.10986 | 0.022888 | 0.14191 | 0.027466 | 0.10529 | -0.050354 | -0.022888 | -0.29297 | -0.27008 | 4.4357 | 4.4861 | 1.4145 | 117.4622 | 1.4648 | 117.5125 | 4.3167 | -117.1417 | -117.2745 | -0.13275 | 4.5685 | 0.25177 | 4.5319 | 4.454 | 0.21515 | 0.13733
2829.1719 | 121.8109 | -4.4678 | 117.3477 | -0.74089 | 0.58517 | 0.1243 | 0.091777 | 0.0010403 | 0.56584 | 0.0051696 | 0.041412 | 0.021535 | 0.012582 | 0.0023093 | 0.013704 | -0.0086025 | -0.0012667 | -0.086975 | 117.334 | -4.4952 | 116.9678 | -0.10529 | 112.8067 | -3.0991 | -0.21515 | 0.17853 | -0.12817 | -0.050354 | 0.018311 | -0.10529 | 0.009155 | -0.027466 | -0.013733 | -0.37994 | -0.36621 | 4.3625 | 4.39 | 1.3687 | 117.2745 | 1.3962 | 117.3019 | 4.3396 | -117.1692 | -117.3981 | -0.22888 | 4.4861 | 0.14649 | 4.4769 | 4.3625 | 0.13733 | 0.022888
2830.1719 | 121.8109 | -4.454 | 117.334 | -0.70274 | 0.58752 | 0.12345 | 0.090495 | -0.00024129 | 0.56306 | 0.0085876 | -0.089813 | 0.02538 | 0.012582 | 0.0003867 | 0.014986 | -0.0056119 | -0.0010531 | -0.091553 | 117.3157 | -4.4952 | 116.954 | -0.1236 | 112.9623 | -3.2593 | -0.21973 | 0.19226 | -0.11444 | -0.050354 | 0.009155 | -0.086975 | 0.027466 | -0.041198 | -0.01831 | -0.37994 | -0.36163 | 4.3304 | 4.3716 | 1.1948 | 117.4164 | 1.236 | 117.4576 | 4.3396 | -117.1417 | -117.3843 | -0.24261 | 4.4632 | 0.1236 | 4.4815 | 4.3671 | 0.14191 | 0.027466
2831.1719 | 121.9254 | -4.3854 | 117.3203 | -0.71037 | 0.58645 | 0.12686 | 0.08964 | -2.7692e-05 | 0.56264 | 0.004956 | 0.0017395 | 0.020681 | 0.012369 | -0.0015359 | 0.013063 | -0.0058255 | -0.0034031 | -0.082397 | 117.3111 | -4.4586 | 117.0639 | -0.013733 | 113.0768 | -3.0899 | -0.036621 | 0.12817 | -0.10986 | 0.041199 | 0.16937 | 0.041199 | 0.15106 | -0.073242 | -0.009155 | -0.25635 | -0.24719 | 4.3716 | 4.4449 | 1.2955 | 117.4622 | 1.3687 | 117.5354 | 4.2755 | -117.1921 | -117.2791 | -0.086975 | 4.5547 | 0.27924 | 4.5364 | 4.4266 | 0.26093 | 0.15106
2832.1719 | 121.9162 | -4.4128 | 117.2882 | -0.75615 | 0.58603 | 0.12622 | 0.088145 | -0.0015231 | 0.56306 | 0.0077332 | 0.0078429 | 0.021108 | 0.0091644 | -0.0032449 | 0.013917 | -0.0058255 | -0.00019868 | -0.054932 | 117.2379 | -4.4495 | 117.0731 | 0.018311 | 113.0859 | -3.3188 | 0.1236 | 0.073242 | -0.11902 | 0.041199 | 0.16937 | 0.059509 | 0.13275 | -0.036621 | -0.050354 | -0.21515 | -0.16479 | 4.4312 | 4.4678 | 1.0941 | 117.4988 | 1.1307 | 117.5354 | 4.2938 | -117.215 | -117.247 | -0.032043 | 4.5822 | 0.28839 | 4.5456 | 4.4724 | 0.25177 | 0.17853
2833.1719 | 121.8613 | -4.3991 | 117.3065 | -0.76378 | 0.5824 | 0.12494 | 0.087504 | 0.0012539 | 0.56135 | 0.0058106 | 0.035309 | 0.02239 | 0.01301 | -0.0075173 | 0.014131 | -0.0051845 | 0.0019375 | -0.041199 | 117.2699 | -4.454 | 117.0593 | 0.022888 | 113.0768 | -3.0624 | 0 | 0.13733 | -0.11444 | 0.032043 | 0.14191 | 0.027466 | 0.1236 | -0.054932 | -0.036622 | -0.24719 | -0.21057 | 4.422 | 4.4769 | 1.3367 | 117.4759 | 1.3916 | 117.5308 | 4.2847 | -117.1692 | -117.2745 | -0.10529 | 4.541 | 0.25635 | 4.5227 | 4.4266 | 0.23804 | 0.14191
2834.1719 | 121.8979 | -4.4312 | 117.2699 | -0.68748 | 0.58475 | 0.12729 | 0.089854 | 0.00082671 | 0.56178 | 0.006665 | -0.019623 | 0.021108 | 0.01365 | 0.0031637 | 0.015199 | -0.0051845 | -0.0012667 | -0.082397 | 117.2379 | -4.5364 | 116.9998 | -0.036621 | 113.063 | -3.1998 | -0.022888 | 0.15106 | -0.14191 | 0.041199 | 0.14191 | 0.027466 | 0.12817 | -0.10529 | -0.032043 | -0.27008 | -0.23804 | 4.3945 | 4.4998 | 1.2314 | 117.4942 | 1.3367 | 117.5995 | 4.2892 | -117.1188 | -117.2287 | -0.10986 | 4.5731 | 0.28381 | 4.5593 | 4.4586 | 0.27008 | 0.16937
2835.1719 | 121.8887 | -4.4266 | 117.2745 | -0.71037 | 0.58432 | 0.12537 | 0.090495 | 0.0010403 | 0.56264 | 0.008374 | -0.040985 | 0.021963 | 0.011942 | -0.0036721 | 0.013063 | -0.0066799 | 0.0023647 | -0.091553 | 117.2516 | -4.5227 | 116.9769 | -0.036621 | 112.9761 | -3.2364 | -0.050354 | 0.14648 | -0.10986 | 0.050354 | 0.15564 | 0.036621 | 0.13275 | -0.09613 | -0.022888 | -0.29755 | -0.27466 | 4.39 | 4.4861 | 1.1902 | 117.4026 | 1.2863 | 117.4988 | 4.3167 | -117.128 | -117.2241 | -0.09613 | 4.5822 | 0.2655 | 4.5593 | 4.4632 | 0.24261 | 0.14648
2836.1719 | 121.8246 | -4.4403 | 117.2974 | -0.72563 | 0.58133 | 0.12558 | 0.088572 | 0.0021085 | 0.56007 | 0.0090148 | -0.056244 | 0.024099 | 0.012582 | -0.0021769 | 0.014772 | -0.0045437 | -0.0051121 | -0.07782 | 117.2974 | -4.5044 | 116.9586 | -0.09613 | 112.9898 | -3.0853 | -0.11902 | 0.16022 | -0.13275 | -0.041199 | 0.059509 | -0.068665 | 0.036621 | -0.064087 | 0 | -0.33875 | -0.33875 | 4.3442 | 4.4083 | 1.355 | 117.4301 | 1.4191 | 117.4942 | 4.3076 | -117.1371 | -117.3386 | -0.20142 | 4.4998 | 0.19226 | 4.4769 | 4.3716 | 0.16937 | 0.064086
2837.1719 | 121.8063 | -4.4495 | 117.3386 | -0.78667 | 0.58282 | 0.12387 | 0.091136 | -0.0019503 | 0.55986 | 0.0094422 | -0.040985 | 0.022603 | 0.012796 | -0.0002543 | 0.014558 | -0.0053983 | -0.00062588 | -0.082397 | 117.3248 | -4.4632 | 116.954 | -0.11444 | 113.0127 | -3.3875 | -0.11902 | 0.17853 | -0.14648 | -0.045776 | 0.050354 | -0.07782 | 0.027466 | -0.013733 | -0.013733 | -0.38452 | -0.37079 | 4.335 | 4.3488 | 1.062 | 117.4622 | 1.0757 | 117.4759 | 4.303 | -117.16 | -117.3843 | -0.2243 | 4.4998 | 0.19684 | 4.4769 | 4.3716 | 0.17395 | 0.068664
2838.1719 | 121.8887 | -4.422 | 117.3111 | -0.70274 | 0.58304 | 0.12644 | 0.085582 | 0.0016813 | 0.559 | 0.0094422 | -0.022675 | 0.020894 | 0.011728 | -4.0504e-05 | 0.011995 | -0.0026211 | -0.0029757 | -0.041199 | 117.3019 | -4.4449 | 117.1097 | 0.032043 | 113.0859 | -3.1265 | 0 | 0.13275 | -0.10529 | -0.059509 | 0.27924 | 0.059509 | 0.11444 | -0.022888 | -0.009155 | -0.20142 | -0.19226 | 4.454 | 4.4769 | 1.2955 | 117.5079 | 1.3184 | 117.5308 | 4.3167 | -117.1783 | -117.3706 | -0.19226 | 4.7012 | 0.38452 | 4.5364 | 4.4815 | 0.21973 | 0.16479
2839.1719 | 121.8933 | -4.4037 | 117.3203 | -0.71037 | 0.5824 | 0.12387 | 0.089 | -0.0025913 | 0.55858 | 0.0075196 | -0.0013123 | 0.020894 | 0.011514 | 0.0029501 | 0.016695 | -0.0032619 | -0.0025485 | -0.027466 | 117.3157 | -4.3854 | 117.1326 | 0.027466 | 113.1638 | -3.2913 | -0.018311 | 0.1236 | -0.10986 | -0.068665 | 0.16937 | 0.16479 | 0.20142 | 0.018311 | -0.004577 | -0.18768 | -0.18311 | 4.4312 | 4.4128 | 1.1124 | 117.5674 | 1.0941 | 117.5491 | 4.2938 | -117.1967 | -117.3889 | -0.19226 | 4.5731 | 0.27924 | 4.6051 | 4.5685 | 0.31128 | 0.27466
2840.1719 | 121.8979 | -4.3671 | 117.3477 | -0.71037 | 0.58325 | 0.12451 | 0.087291 | -0.0032321 | 0.55986 | 0.0085876 | -0.013519 | 0.021535 | 0.014078 | 0.0046591 | 0.012208 | -0.0060391 | -0.0019077 | -0.022888 | 117.3065 | -4.3808 | 117.1234 | 0.018311 | 113.1592 | -3.0487 | -0.009155 | 0.13733 | -0.10529 | -0.050354 | 0.14648 | 0.16022 | 0.20599 | -0.013733 | -0.041198 | -0.2243 | -0.18311 | 4.3854 | 4.3991 | 1.3184 | 117.5262 | 1.3321 | 117.54 | 4.2618 | -117.2104 | -117.3981 | -0.18768 | 4.5135 | 0.25177 | 4.5731 | 4.5273 | 0.31128 | 0.2655
2841.1719 | 121.8933 | -4.4266 | 117.2699 | -0.76378 | 0.58197 | 0.12494 | 0.089 | -0.0023777 | 0.55943 | 0.0085876 | -0.0074159 | 0.021749 | 0.011087 | -0.0004679 | 0.014558 | -0.0058255 | -0.0014805 | -0.041199 | 117.2974 | -4.4266 | 117.1005 | 0.027466 | 113.0539 | -3.0945 | -0.009155 | 0.15106 | -0.10986 | -0.068665 | 0.045776 | 0.15106 | 0.20142 | 0 | 0.027466 | -0.16937 | -0.19684 | 4.454 | 4.454 | 1.3321 | 117.4805 | 1.3321 | 117.4805 | 4.3167 | -117.1188 | -117.3386 | -0.21973 | 4.4724 | 0.15564 | 4.628 | 4.5776 | 0.31128 | 0.26093
2842.1719 | 121.907 | -4.3991 | 117.3111 | -0.77141 | 0.58304 | 0.12644 | 0.090281 | -0.0028049 | 0.559 | 0.0079468 | -0.010467 | 0.02004 | 0.011514 | 0.0065817 | 0.013063 | -0.0073209 | -0.0053257 | -0.032043 | 117.2745 | -4.4403 | 117.1097 | 0 | 113.0814 | -2.7374 | -0.009155 | 0.13733 | -0.11902 | -0.09613 | -0.041199 | 0.12817 | 0.21057 | -0.041199 | -0.036621 | -0.20142 | -0.16479 | 4.3991 | 4.4403 | 1.6617 | 117.4805 | 1.7029 | 117.5217 | 4.2801 | -117.1738 | -117.4072 | -0.23346 | 4.3579 | 0.07782 | 4.6097 | 4.5273 | 0.32959 | 0.24719
2843.1719 | 121.9254 | -4.3945 | 117.2974 | -0.67223 | 0.58176 | 0.12302 | 0.086863 | -0.0047275 | 0.55687 | 0.0075196 | 0.074982 | 0.021108 | 0.013437 | -0.0060221 | 0.013277 | -0.0081753 | -0.00062588 | 0 | 117.3248 | -4.4037 | 117.1188 | 0.045776 | 113.1683 | -3.0991 | 0.1236 | 0.13733 | -0.10071 | 0.068665 | 0.21973 | 0.10529 | 0.15564 | -0.009156 | 0.027466 | -0.17853 | -0.20599 | 4.4403 | 4.4495 | 1.2955 | 117.5629 | 1.3046 | 117.572 | 4.2938 | -117.16 | -117.2287 | -0.068664 | 4.6143 | 0.32044 | 4.5502 | 4.4998 | 0.25635 | 0.20599
2844.1719 | 121.9116 | -4.3488 | 117.3615 | -0.718 | 0.57898 | 0.12323 | 0.090068 | -2.7692e-05 | 0.55943 | 0.0096558 | -0.062347 | 0.023458 | 0.012582 | 0.0050863 | 0.014131 | -0.0047573 | 0.0025785 | -0.073242 | 117.3386 | -4.4312 | 117.0364 | -0.059509 | 112.9715 | -3.067 | -0.059509 | 0.16022 | -0.10986 | 0.027466 | 0.14191 | 0.009155 | 0.11444 | -0.082397 | -0.022888 | -0.32501 | -0.30212 | 4.2892 | 4.3716 | 1.2817 | 117.3203 | 1.3641 | 117.4026 | 4.2389 | -117.2012 | -117.334 | -0.13275 | 4.4907 | 0.25177 | 4.4632 | 4.3579 | 0.2243 | 0.11902
2845.1719 | 121.9574 | -4.3808 | 117.3157 | -0.75615 | 0.58154 | 0.12387 | 0.090068 | 0.0025357 | 0.55836 | 0.0081604 | -0.019623 | 0.020254 | 0.012796 | -0.0019631 | 0.012636 | -0.0068935 | -0.0038303 | -0.036621 | 117.3248 | -4.4083 | 117.1371 | 0.013733 | 113.1271 | -3.1677 | 0 | 0.13733 | -0.09613 | 0.009155 | -0.086975 | -0.10529 | 0.086975 | -0.027466 | 0.009155 | -0.17853 | -0.18768 | 4.3945 | 4.422 | 1.2131 | 117.5079 | 1.2405 | 117.5354 | 4.2847 | -117.1783 | -117.3065 | -0.12817 | 4.2938 | 0.009155 | 4.4678 | 4.2755 | 0.1831 | -0.009156
2846.1719 | 121.8796 | -3.9459 | 117.7643 | -0.73326 | 0.57919 | 0.12131 | 0.088145 | -0.00024129 | 0.559 | 0.0051696 | 0.010895 | 0.020254 | 0.014078 | 0.0014547 | 0.014772 | -0.0062527 | -0.0044711 | -0.041199 | 117.7322 | -4.0237 | 117.5171 | 0.045776 | 113.475 | -2.7374 | 0.009155 | 0.12817 | -0.10986 | 0.050354 | 0.17395 | 0.032043 | 0.14648 | -0.07782 | -0.032043 | -0.24719 | -0.21515 | 3.9917 | 4.0695 | 1.2085 | 117.421 | 1.2863 | 117.4988 | 3.8361 | -117.6361 | -117.7139 | -0.07782 | 4.1199 | 0.28381 | 4.0924 | 3.978 | 0.25635 | 0.14191
2847.1719 | 121.8933 | -3.9917 | 117.7185 | -0.71037 | 0.58026 | 0.1258 | 0.088359 | 0.00018591 | 0.5558 | 0.0045288 | 0.0017395 | 0.02239 | 0.011728 | -0.0049539 | 0.011781 | -0.0047573 | -0.00083948 | -0.045776 | 117.6773 | -4.0237 | 117.4896 | 0.041199 | 113.475 | -2.7786 | 0.004578 | 0.13733 | -0.10071 | 0.045776 | 0.16479 | 0.050354 | 0.12817 | -0.032044 | -0.041199 | -0.22888 | -0.18768 | 4.0329 | 4.0649 | 1.2131 | 117.4667 | 1.2451 | 117.4988 | 3.891 | -117.5812 | -117.6727 | -0.091553 | 4.1565 | 0.2655 | 4.1199 | 4.0421 | 0.22888 | 0.15106
2848.1719 | 121.9437 | -3.9597 | 117.7094 | -0.718 | 0.58111 | 0.12387 | 0.087077 | 0.0012539 | 0.5558 | 0.0081604 | 0.050568 | 0.020681 | 0.01365 | -0.0026041 | 0.01349 | -0.0088161 | -0.0063937 | -0.009155 | 117.7185 | -4.0237 | 117.4988 | 0.07782 | 113.5345 | -5.4749 | 0.11902 | 0.14191 | -0.09613 | 0.064087 | 0.22888 | 0.10986 | 0.15564 | -0.064087 | 0.009155 | -0.21057 | -0.21973 | 4.0375 | 4.1016 | -1.5152 | 117.4942 | -1.4511 | 117.5583 | 3.8635 | -117.5674 | -117.6453 | -0.07782 | 4.1885 | 0.32501 | 4.1153 | 4.0695 | 0.25177 | 0.20599
2849.1719 | 121.8842 | -3.9688 | 117.7505 | -0.73326 | 0.5777 | 0.12558 | 0.088786 | -0.0032321 | 0.55772 | 0.0062378 | -0.013519 | 0.021963 | 0.014078 | 0.0029501 | 0.013917 | -0.0062527 | -0.0027621 | -0.036621 | 117.6636 | -4.0512 | 117.5125 | 0.045776 | 113.4567 | -2.7191 | 0.022888 | 0.14191 | -0.10986 | 0.059509 | 0.16937 | 0.059509 | 0.14191 | -0.082397 | -0.086975 | -0.23804 | -0.15106 | 4.0146 | 4.097 | 1.2497 | 117.4255 | 1.3321 | 117.5079 | 3.8589 | -117.6086 | -117.691 | -0.082398 | 4.1382 | 0.27924 | 4.1107 | 4.0283 | 0.25177 | 0.16937
2850.1719 | 122.1405 | -4.01 | 117.6086 | -0.76378 | 0.57876 | 0.12387 | 0.088145 | 0.0012539 | 0.55665 | 0.0085876 | -0.044037 | 0.021963 | 0.0098052 | -0.0023905 | 0.013277 | -0.0047573 | -0.0016941 | -0.027466 | 117.659 | -4.0604 | 117.4393 | 0.022888 | 113.4613 | -2.7603 | 0.013733 | 0.13275 | -0.11902 | 0.073242 | 0.17395 | 0.054932 | 0.15106 | -0.050354 | 0.050354 | -0.16937 | -0.21973 | 4.0329 | 4.0833 | 1.2497 | 117.4713 | 1.3 | 117.5217 | 3.891 | -117.4759 | -117.5354 | -0.059509 | 4.184 | 0.29297 | 4.1611 | 4.0649 | 0.27008 | 0.17395
2851.1719 | 121.9162 | -3.9871 | 117.6727 | -0.71037 | 0.57876 | 0.12473 | 0.089 | 0.00039951 | 0.55708 | 0.0075196 | -0.047089 | 0.020894 | 0.013223 | 0.0025229 | 0.013277 | -0.0043301 | 0.00065592 | -0.059509 | 117.6773 | -4.0604 | 117.4805 | 0.041199 | 113.4521 | -2.7374 | 0.004578 | 0.13733 | -0.10071 | 0.068665 | 0.17395 | 0.041199 | 0.16479 | -0.073242 | 0.004578 | -0.19226 | -0.19684 | 4.0283 | 4.1016 | 1.2497 | 117.4393 | 1.3229 | 117.5125 | 3.8864 | -117.5354 | -117.6041 | -0.068664 | 4.1611 | 0.27466 | 4.1519 | 4.0283 | 0.2655 | 0.14191
2852.1719 | 121.875 | -3.9917 | 117.7048 | -0.61119 | 0.57919 | 0.12473 | 0.088145 | -2.7692e-05 | 0.55388 | 0.0053832 | -0.019623 | 0.021322 | 0.012369 | 0.0044455 | 0.01349 | -0.0045437 | 0.00086952 | -0.027466 | 117.6498 | -4.0558 | 117.4576 | 0.054932 | 113.4247 | -2.6001 | 0.004578 | 0.13275 | -0.09613 | 0.054932 | 0.16937 | 0.045776 | 0.14191 | -0.064087 | -0.054932 | -0.24719 | -0.19226 | 4.0466 | 4.1107 | 1.3916 | 117.4164 | 1.4557 | 117.4805 | 3.8956 | -117.572 | -117.6498 | -0.077819 | 4.1611 | 0.2655 | 4.1336 | 4.0375 | 0.23804 | 0.14191
2853.1719 | 121.8613 | -4.0375 | 117.7002 | -0.71037 | 0.5777 | 0.12451 | 0.086863 | 0.0023221 | 0.55623 | 0.0068786 | 0.026154 | 0.020254 | 0.01365 | 0.0001731 | 0.014345 | -0.0053983 | -0.0070347 | 0.027466 | 117.691 | -4.0649 | 117.4713 | 0.036621 | 113.4659 | -2.7603 | 0.10071 | 0.15106 | -0.1236 | 0.013733 | 0.15564 | 0.032043 | 0.10529 | -0.027465 | -0.009155 | -0.22888 | -0.21973 | 4.0741 | 4.1016 | 1.2772 | 117.5034 | 1.3046 | 117.5308 | 3.9139 | -117.5491 | -117.6865 | -0.13733 | 4.1931 | 0.27924 | 4.1428 | 4.0695 | 0.22888 | 0.15564
2854.1719 | 121.8796 | -3.9368 | 117.7643 | -0.74852 | 0.57855 | 0.12409 | 0.087504 | 0.0025357 | 0.55452 | 0.01051 | -0.013519 | 0.020681 | 0.013223 | -0.0006815 | 0.014772 | -0.0058255 | -0.0046847 | -0.041199 | 117.7139 | -4.01 | 117.4896 | 0.032043 | 113.4247 | -2.7008 | -0.027466 | 0.13733 | -0.10529 | 0.032043 | 0.12817 | 0.022888 | 0.10071 | -0.073242 | -0.050354 | -0.27466 | -0.2243 | 3.9688 | 4.0421 | 1.236 | 117.3615 | 1.3092 | 117.4347 | 3.8315 | -117.627 | -117.7322 | -0.10529 | 4.0649 | 0.23346 | 4.0375 | 3.9597 | 0.20599 | 0.12817
2855.1719 | 121.8613 | -3.9871 | 117.7551 | -0.71037 | 0.57855 | 0.12345 | 0.089213 | -0.00024129 | 0.55601 | 0.0064514 | 0.010895 | 0.021963 | 0.011942 | 0.0055137 | 0.014131 | -0.0051845 | -0.0023349 | -0.050354 | 117.6773 | -4.0192 | 117.4622 | 0.045776 | 113.4567 | -2.6459 | -0.027466 | 0.15106 | -0.11444 | 0.018311 | 0.13275 | 0.018311 | 0.11444 | -0.032043 | -0.07782 | -0.29297 | -0.21515 | 4.0329 | 4.0649 | 1.3412 | 117.4438 | 1.3733 | 117.4759 | 3.8727 | -117.6041 | -117.7368 | -0.13275 | 4.1199 | 0.24719 | 4.1016 | 4.0054 | 0.22888 | 0.13275
2856.1719 | 121.8658 | -3.9871 | 117.7185 | -0.718 | 0.57748 | 0.12494 | 0.086863 | -0.00066869 | 0.55324 | 0.0034606 | -0.0013123 | 0.021535 | 0.011728 | -0.0019631 | 0.014131 | -0.0062527 | -0.0053257 | -0.036621 | 117.6682 | -4.0329 | 117.4622 | 0.045776 | 113.4064 | -2.5864 | -0.013733 | 0.15564 | -0.10986 | 0.027466 | 0.13275 | 0.022888 | 0.10529 | -0.045776 | -0.050354 | -0.25635 | -0.20599 | 4.0329 | 4.0787 | 1.4008 | 117.3935 | 1.4465 | 117.4393 | 3.8773 | -117.5629 | -117.691 | -0.12817 | 4.1199 | 0.24261 | 4.0924 | 4.01 | 0.21515 | 0.13275
2857.1719 | 121.8887 | -4.0009 | 117.6819 | -0.71037 | 0.5762 | 0.12451 | 0.090495 | 0.0010403 | 0.55366 | 0.0094422 | -0.065399 | 0.018331 | 0.011942 | 0.0014547 | 0.013704 | -0.0034757 | -0.0061801 | -0.027466 | 117.6727 | -4.0375 | 117.3431 | -0.004578 | 113.6398 | -2.7191 | 0.054932 | 0.10071 | -0.12817 | 0.068665 | 0.16479 | 0.045776 | 0.13275 | -0.036622 | -0.009156 | -0.33874 | -0.32959 | 3.9963 | 4.0329 | 1.2817 | 117.6407 | 1.3184 | 117.6773 | 3.8727 | -117.5812 | -117.6132 | -0.032043 | 4.1656 | 0.29297 | 4.1336 | 4.0466 | 0.26093 | 0.17395
2858.1719 | 121.907 | -4.0009 | 117.6819 | -0.74852 | 0.57748 | 0.12409 | 0.088786 | 0.0010403 | 0.55601 | 0.0090148 | 0.047516 | 0.021108 | 0.01066 | -0.0040995 | 0.014772 | -0.0053983 | -0.0057529 | -0.032043 | 117.6498 | -4.0558 | 117.4576 | -0.13733 | 113.1409 | -2.655 | -0.036621 | 0.091553 | -0.12817 | 0.073242 | 0.15564 | 0.032043 | 0.15106 | -0.054932 | -0.032044 | -0.2243 | -0.19226 | 3.8635 | 3.9185 | 1.3458 | 117.1417 | 1.4008 | 117.1967 | 3.8727 | -117.5903 | -117.6086 | -0.018311 | 4.1565 | 0.28381 | 4.1519 | 4.0329 | 0.27924 | 0.16022
2859.1719 | 121.8613 | -4.0741 | 117.6636 | -0.74852 | 0.5762 | 0.12473 | 0.087504 | 0.00061311 | 0.55366 | 0.008374 | 0.023102 | 0.022176 | 0.01365 | -0.0002543 | 0.016267 | -0.0066799 | -0.0053257 | -0.073242 | 117.6407 | -4.1016 | 117.3569 | -0.013733 | 113.356 | -2.8198 | -0.068665 | 0.14191 | -0.10529 | 0.013733 | 0.12817 | 0.022888 | 0.086975 | -0.027465 | -0.022888 | -0.3067 | -0.28381 | 4.0604 | 4.0878 | 1.2543 | 117.4301 | 1.2817 | 117.4576 | 3.9688 | -117.5217 | -117.6498 | -0.12817 | 4.2023 | 0.23346 | 4.1611 | 4.097 | 0.19226 | 0.12817
2860.1719 | 121.875 | -4.0604 | 117.6865 | -0.73326 | 0.57449 | 0.12088 | 0.089427 | -2.7692e-05 | 0.55067 | 0.007306 | 0.02005 | 0.022817 | 0.012369 | -0.0011087 | 0.013917 | -0.0071071 | -0.0029757 | -0.073242 | 117.6453 | -4.0695 | 117.4026 | -0.027466 | 113.4247 | -2.7878 | -0.036621 | 0.15564 | -0.10986 | 0.009155 | 0.14191 | 0.013733 | 0.10529 | -0.009155 | -0.041198 | -0.28381 | -0.24262 | 4.0329 | 4.0421 | 1.2726 | 117.485 | 1.2817 | 117.4942 | 3.9505 | -117.5308 | -117.6773 | -0.14649 | 4.2023 | 0.25177 | 4.1657 | 4.0741 | 0.21515 | 0.1236
2861.1719 | 121.8704 | -4.0604 | 117.6636 | -0.70274 | 0.57406 | 0.12366 | 0.088786 | 0.0016813 | 0.5526 | 0.0096558 | -0.0074159 | 0.019826 | 0.011728 | -0.0043131 | 0.014986 | -0.0043301 | -0.0019077 | -0.082397 | 117.659 | -4.0878 | 117.3706 | -0.022888 | 113.3835 | -2.7145 | -0.036621 | 0.11902 | -0.11902 | 0.018311 | 0.13275 | 0.009155 | 0.073242 | -0.027466 | -0.004577 | -0.29297 | -0.28839 | 4.0375 | 4.0649 | 1.3458 | 117.4438 | 1.3733 | 117.4713 | 3.9413 | -117.5446 | -117.6453 | -0.10071 | 4.1931 | 0.25177 | 4.1336 | 4.0695 | 0.19226 | 0.12817
2862.1719 | 121.8658 | -4.0192 | 117.7002 | -0.74852 | 0.57428 | 0.12281 | 0.087504 | 0.0040311 | 0.55217 | 0.0068786 | 0.013947 | 0.020681 | 0.011942 | 0.0025229 | 0.015199 | -0.0064663 | -0.0044711 | -0.07782 | 117.6682 | -4.0604 | 117.4026 | -0.004578 | 113.443 | -5.2139 | -0.054932 | 0.16022 | -0.10529 | 0.013733 | 0.12817 | 0 | 0.09613 | -0.041199 | -0.032043 | -0.29755 | -0.2655 | 4.0146 | 4.0558 | -1.1948 | 117.4622 | -1.1536 | 117.5034 | 3.9139 | -117.54 | -117.6865 | -0.14648 | 4.1473 | 0.23346 | 4.1153 | 4.0192 | 0.20142 | 0.10529
2863.1719 | 121.8613 | -4.0237 | 117.7048 | -0.76378 | 0.57641 | 0.1258 | 0.088786 | -0.0021641 | 0.5543 | 0.0079468 | 0.0078429 | 0.021322 | 0.010233 | 0.0038047 | 0.012422 | -0.0064663 | -0.0012667 | -0.068665 | 117.6773 | -4.0512 | 117.4301 | -0.018311 | 113.4293 | -2.655 | -0.050354 | 0.13275 | -0.11902 | -0.004578 | 0.1236 | 0.022888 | 0.086975 | -0.027465 | -0.027466 | -0.27466 | -0.24719 | 4.0054 | 4.0329 | 1.3687 | 117.453 | 1.3962 | 117.4805 | 3.9047 | -117.572 | -117.7094 | -0.13733 | 4.1473 | 0.24261 | 4.1107 | 4.0466 | 0.20599 | 0.14191
2864.1719 | 121.8704 | -4.0924 | 117.5949 | -0.74089 | 0.57535 | 0.12387 | 0.088359 | -0.00066869 | 0.54961 | 0.007306 | -0.03183 | 0.022817 | 0.011301 | -0.0043131 | 0.013704 | -0.0058255 | -0.0051121 | -0.027466 | 117.5995 | -4.1016 | 117.4026 | 0.036621 | 113.3743 | -2.7603 | 0.027466 | 0.15106 | -0.10529 | 0.068665 | 0.16479 | 0.050354 | 0.13733 | -0.009155 | 0.004577 | -0.19226 | -0.19684 | 4.129 | 4.1382 | 1.3321 | 117.4667 | 1.3412 | 117.4759 | 3.9871 | -117.4438 | -117.5262 | -0.082397 | 4.2572 | 0.27008 | 4.2297 | 4.1428 | 0.24262 | 0.15564
2865.1719 | 121.8567 | -22.4075 | 98.8541 | -0.69511 | 0.58389 | 0.12473 | 0.088572 | 0.0010403 | 0.5511 | 0.010297 | 0.013947 | 0.023671 | 0.012369 | 0.0016685 | 0.015199 | -0.0060391 | -0.0010531 | 16.2415 | 89.6393 | -34.8404 | 82.3105 | -1.9775 | 66.7831 | -43.5287 | 0.059509 | 0.13733 | -0.10529 | 0.064087 | 0.17853 | 0.041199 | 0.1236 | -12.4329 | -9.2148 | -16.5436 | -7.3288 | 20.43 | 32.8629 | -21.1212 | 89.1907 | -8.6884 | 101.6235 | 22.3022 | -98.7167 | -98.79 | -0.073242 | 22.5861 | 0.28381 | 22.5311 | 22.4487 | 0.22888 | 0.14649
2866.1719 | 121.907 | -121.0144 | 0.57678 | -0.038986 | 0.018645 | 0.012576 | -0.00029472 | -0.00066869 | 0.00038017 | 0.001538 | 0.0017395 | 0.019826 | 0.013864 | -0.0002543 | 0.015199 | -0.0081753 | -0.00041228 | -0.09613 | 0.59052 | -0.009155 | 0.35706 | 0.1831 | 0.2655 | 0.11902 | -0.11444 | 0.10986 | -0.11902 | 0 | 0.11902 | 0.013733 | 0.082397 | 121.0052 | 0.013733 | -0.21973 | -0.23346 | 121.1975 | 0.19226 | 121.1334 | 121.2799 | 0.12817 | 0.27466 | 120.8954 | -0.46692 | -0.57678 | -0.10986 | 121.1334 | 0.23804 | 121.0968 | 121.0281 | 0.20142 | 0.13275
2867.1719 | 121.9254 | -121.1426 | 0.52643 | 0.01442 | 0.024626 | 0.011722 | 0.0014141 | 0.00082671 | 0.0012348 | 0.0030334 | -0.016571 | 0.020681 | 0.009378 | 0.0018821 | 0.016695 | -0.0045437 | -0.0014805 | -0.10071 | 0.57678 | -0.009155 | 0.32501 | 0.16022 | 0.24261 | 0.10986 | -0.022888 | 0.11902 | -0.14648 | 0.004578 | 0.10071 | -0.004578 | 0.073242 | 121.1334 | 0.050354 | -0.20142 | -0.25177 | 121.3028 | 0.16937 | 121.2524 | 121.3852 | 0.11902 | 0.25177 | 120.9961 | -0.40741 | -0.52185 | -0.11444 | 121.2433 | 0.24719 | 121.2158 | 121.138 | 0.21973 | 0.14191
2868.1719 | 121.9299 | -121.1517 | 0.60425 | -0.016098 | 0.023131 | 0.011935 | -0.0011493 | -0.0049411 | 0.00038017 | 0.0028198 | -0.010467 | 0.021322 | 0.011728 | -0.0030313 | 0.01584 | -0.0045437 | 0.0021511 | 0.027466 | 0.56763 | -0.022888 | 0.37537 | 0.20599 | 0.29755 | 0.14191 | 0.1236 | 0.13275 | -0.11444 | 0.027466 | 0.1236 | 0.018311 | 0.073242 | 121.1288 | -0.036621 | -0.22888 | -0.19226 | 121.3577 | 0.22888 | 121.2936 | 121.4493 | 0.16479 | 0.32043 | 121.0373 | -0.4715 | -0.57678 | -0.10529 | 121.2753 | 0.23804 | 121.225 | 121.17 | 0.18768 | 0.13275
2869.1719 | 121.875 | -121.2341 | 0.60883 | -0.00083923 | 0.023772 | 0.011294 | -0.00093572 | 0.00018591 | 0.0014484 | -0.00038459 | -0.034882 | 0.020467 | 0.012369 | 0.0042319 | 0.015626 | -0.0071071 | -0.0036167 | 0 | 0.59967 | -0.041199 | 0.29297 | 0.11902 | 0.19226 | 0.068665 | 0.018311 | 0.18768 | -0.13733 | -0.032043 | 0.054932 | -0.068665 | -0.018311 | 121.1929 | -0.009156 | -0.31586 | -0.3067 | 121.3531 | 0.16022 | 121.3028 | 121.4264 | 0.10986 | 0.23346 | 121.0968 | -0.42114 | -0.64087 | -0.21973 | 121.2891 | 0.19226 | 121.2158 | 121.1655 | 0.11902 | 0.068664
2870.1719 | 121.9116 | -121.1884 | 0.63171 | -0.046616 | 0.022277 | 0.011722 | 0.00077328 | -0.0023777 | 0.0023028 | 0.00068361 | -0.037933 | 0.021749 | 0.012582 | 0.0020957 | 0.015626 | -0.0062527 | 0.00044212 | 0.022888 | 0.64545 | -0.027466 | 0.32959 | 0.13733 | 0.2655 | 0.07782 | 0.013733 | 0.16022 | -0.13275 | -0.045776 | 0.045776 | -0.059509 | 0.018311 | 121.1609 | 0.013733 | -0.30212 | -0.31586 | 121.3257 | 0.16479 | 121.2662 | 121.4539 | 0.10529 | 0.29297 | 121.0556 | -0.4715 | -0.67749 | -0.20599 | 121.2341 | 0.17853 | 121.2067 | 121.1288 | 0.15106 | 0.073242
2871.1719 | 121.9711 | -121.2387 | 0.57678 | 0.022049 | 0.022704 | 0.01044 | 0.0014141 | 0.00018591 | 0.00016657 | -0.0016662 | -0.03183 | 0.021535 | 0.010873 | -0.0004679 | 0.014558 | -0.0060391 | -0.0044711 | -0.068665 | 0.62714 | 0.009155 | 0.40741 | 0.1831 | 0.29297 | 0.14191 | -0.041199 | 0.14648 | -0.10986 | 0.032043 | 0.036621 | -0.17395 | -0.018311 | 121.2479 | 0.050354 | -0.16937 | -0.21973 | 121.4218 | 0.17395 | 121.3806 | 121.5317 | 0.13275 | 0.28381 | 121.1288 | -0.4303 | -0.54474 | -0.11444 | 121.2753 | 0.14648 | 121.2204 | 121.0648 | 0.091552 | -0.064087
2872.1719 | 122.0169 | -121.1288 | 0.67749 | -0.023727 | 0.024199 | 0.011081 | 0.0024823 | -2.7692e-05 | 0.0012348 | -0.00017099 | -0.013519 | 0.021963 | 0.010019 | -0.0004679 | 0.016267 | -0.0062527 | 0.0015103 | -0.07782 | 0.66376 | 0.018311 | 0.43945 | 0.20142 | 0.3479 | 0.14648 | -0.032043 | 0.12817 | -0.1236 | 0.054932 | 0.1236 | 0.009155 | 0.09613 | 121.1472 | -0.013733 | -0.23804 | -0.2243 | 121.3303 | 0.18311 | 121.2753 | 121.4767 | 0.12817 | 0.32959 | 121.0052 | -0.54932 | -0.62256 | -0.073242 | 121.2524 | 0.24719 | 121.225 | 121.138 | 0.21973 | 0.13275
2873.1719 | 122.0398 | -121.0693 | 0.74158 | -0.023727 | 0.023558 | 0.01044 | -0.0015765 | -0.0013095 | -0.00047423 | 0.00025641 | 0.0078429 | 0.017904 | 0.00831 | 0.0042319 | 0.016054 | -0.0068935 | -0.0051121 | 0.027466 | 0.71869 | 0.022888 | 0.55847 | 0.28381 | 0.47607 | 0.22888 | 0.11902 | 0.14648 | -0.10529 | 0.050354 | 0.18768 | 0.045776 | 0.09613 | 121.0922 | -0.022888 | -0.18311 | -0.16022 | 121.3531 | 0.26092 | 121.2982 | 121.5454 | 0.20599 | 0.45319 | 120.9641 | -0.59509 | -0.69122 | -0.09613 | 121.257 | 0.29297 | 121.1655 | 121.1151 | 0.20142 | 0.15106
2874.1875 | 121.9894 | -121.0464 | 0.83771 | -0.016098 | 0.024199 | 0.011294 | 0.0059003 | 0.0038175 | 0.00059397 | 0.0017518 | -0.10202 | 0.021108 | 0.011087 | 0.0052999 | 0.013704 | -0.0049709 | -0.0016941 | 0 | 0.79193 | 0.009155 | 0.51727 | 0.22888 | 0.42114 | 0.12817 | 0.1236 | 0.1831 | -0.11444 | 0.036621 | 0.11902 | -0.036621 | 0.054932 | 121.0556 | -0.045777 | -0.32044 | -0.27466 | 121.2753 | 0.21973 | 121.1746 | 121.4676 | 0.11902 | 0.41199 | 120.932 | -0.6546 | -0.80109 | -0.14648 | 121.1655 | 0.23346 | 121.1014 | 121.0098 | 0.16937 | 0.07782
2875.1719 | 122.0261 | -121.0281 | 0.79651 | -0.00083923 | 0.021849 | 0.012149 | -8.1116e-05 | -0.0019503 | 0.0012348 | 0.00025641 | 0.038361 | 0.018758 | 0.012369 | -0.0006815 | 0.016908 | -0.0051845 | -0.0044711 | -0.054932 | 0.80109 | 0.059509 | 0.56305 | 0.23804 | 0.47607 | 0.10529 | -0.022888 | 0.09613 | -0.10986 | 0.004578 | 0.14191 | 0.022888 | 0.091553 | 121.0876 | 0.004577 | -0.23346 | -0.23804 | 121.2662 | 0.17853 | 121.1334 | 121.5042 | 0.045777 | 0.41657 | 120.9183 | -0.70038 | -0.79193 | -0.091552 | 121.17 | 0.25177 | 121.1197 | 121.051 | 0.20142 | 0.13275
2876.1719 | 122.0306 | -121.0464 | 0.78278 | -0.00083923 | 0.022063 | 0.01044 | -0.00029472 | -2.7692e-05 | 0.0027302 | -0.0018798 | 0.026154 | 0.020681 | 0.0098052 | -0.0011087 | 0.015626 | -0.0056119 | -0.0040439 | -0.064087 | 0.83313 | 0.050354 | 0.59509 | 0.24719 | 0.49896 | 0.19226 | -0.027466 | 0.14191 | -0.14648 | 0.018311 | 0.10986 | -0.004578 | 0.07782 | 121.0968 | 0.050354 | -0.18768 | -0.23804 | 121.2936 | 0.19684 | 121.2387 | 121.5454 | 0.14191 | 0.44861 | 120.9 | -0.64087 | -0.76447 | -0.1236 | 121.1563 | 0.25635 | 121.1243 | 121.0419 | 0.2243 | 0.14191
2877.1719 | 122.0261 | -121.0327 | 0.78735 | 0.022049 | 0.023345 | 0.01279 | -8.1116e-05 | 0.00082671 | 0.003371 | -0.0010254 | -0.0013123 | 0.020681 | 0.011514 | -0.0026041 | 0.01349 | -0.0060391 | -0.0048983 | -0.073242 | 0.82398 | 0.073242 | 0.60425 | 0.24261 | 0.48981 | 0.20142 | -0.036621 | 0.13275 | -0.11444 | -0.21057 | 0.082397 | -0.064087 | 0.064087 | 121.106 | 0.036621 | -0.18311 | -0.21973 | 121.2753 | 0.16937 | 121.2341 | 121.5225 | 0.12817 | 0.41656 | 120.9183 | -0.6546 | -0.99792 | -0.34332 | 121.1151 | 0.19684 | 121.0968 | 120.9686 | 0.17853 | 0.050354
2878.1719 | 122.0673 | -120.9915 | 0.82855 | -0.00083923 | 0.023986 | 0.010013 | -8.1116e-05 | -0.0034457 | 0.0010212 | 0.0011108 | 0.026154 | 0.019399 | 0.012582 | -0.0008951 | 0.013277 | -0.0053983 | -0.0057529 | 0.050354 | 0.82855 | 0.064087 | 0.64545 | 0.32959 | 0.56305 | 0.27008 | 0.13275 | 0.14648 | -0.10071 | 0.041199 | 0.1236 | 0.15564 | 0.1831 | 121.0556 | 0 | -0.1831 | -0.1831 | 121.3211 | 0.2655 | 121.2616 | 121.5546 | 0.20599 | 0.49896 | 120.8908 | -0.68207 | -0.78735 | -0.10529 | 121.1151 | 0.2243 | 121.1746 | 121.1472 | 0.28381 | 0.25635
2879.1719 | 122.049 | -121.0052 | 0.84228 | -0.069504 | 0.022277 | 0.0091581 | -0.00029472 | 0.0023221 | 0.00080757 | 0.00047001 | -0.034882 | 0.018331 | 0.012155 | -0.0066629 | 0.014772 | -0.0077481 | -0.0025485 | -0.050354 | 0.82398 | 0.073242 | 0.6546 | 0.3067 | 0.51727 | 0.24719 | -0.004578 | 0.13733 | -0.07782 | 0.045776 | 0.14191 | 0.018311 | 0.10071 | 121.0785 | -0.01831 | -0.18768 | -0.16937 | 121.312 | 0.23346 | 121.2524 | 121.5225 | 0.17395 | 0.44403 | 120.9274 | -0.70496 | -0.79651 | -0.091553 | 121.1472 | 0.21973 | 121.106 | 121.0236 | 0.17853 | 0.096131
2880.1719 | 122.049 | -120.9961 | 0.83771 | -0.046616 | 0.023345 | 0.012362 | 0.00013248 | 0.0012539 | 0.003371 | -0.00059819 | -0.013519 | 0.019185 | 0.01365 | -0.0021769 | 0.013917 | -0.0066799 | -0.0036167 | -0.068665 | 0.83771 | 0.064087 | 0.61798 | 0.2243 | 0.45319 | 0.21515 | 0.009155 | 0.10071 | -0.1236 | 0.027466 | 0.14648 | 0.018311 | 0.073242 | 121.0602 | 0 | -0.21973 | -0.21973 | 121.2204 | 0.16022 | 121.2112 | 121.4493 | 0.15106 | 0.3891 | 120.8725 | -0.737 | -0.81024 | -0.073242 | 121.1426 | 0.27008 | 121.0693 | 121.0144 | 0.19684 | 0.14191
2881.1719 | 122.0993 | -121.0007 | 0.86975 | -0.016098 | 0.020781 | 0.010013 | -0.0015765 | -0.0010959 | 0.0010212 | -0.0016662 | -0.013519 | 0.018758 | 0.011728 | -0.0055947 | 0.014131 | -0.0036893 | -0.0016941 | -0.009155 | 0.86517 | 0.07782 | 0.64545 | 0.28381 | 0.54474 | 0.26093 | 0.073242 | 0.13275 | -0.091553 | 0.068665 | 0.19226 | 0.073242 | 0.14648 | 121.0785 | -0.004578 | -0.2243 | -0.21973 | 121.2845 | 0.20599 | 121.2616 | 121.5454 | 0.18311 | 0.46692 | 120.9091 | -0.737 | -0.80109 | -0.064086 | 121.1929 | 0.28381 | 121.1472 | 121.0739 | 0.23804 | 0.16479
2882.1719 | 122.0718 | -121.0281 | 0.83313 | -0.10002 | 0.023131 | 0.011294 | 0.00055968 | 0.00061311 | 0.0010212 | 0.002179 | -0.0074159 | 0.020681 | 0.011728 | -0.0098673 | 0.014345 | -0.0053983 | -0.0070347 | -0.054932 | 0.85144 | 0.068665 | 0.61798 | 0.25177 | 0.49896 | 0.073242 | -0.022888 | 0.073242 | -0.091553 | 0.004578 | 0.14191 | 0.027466 | 0.086975 | 121.0968 | 0.01831 | -0.21515 | -0.23346 | 121.2799 | 0.1831 | 121.1014 | 121.5271 | 0.004577 | 0.4303 | 120.9366 | -0.75989 | -0.82855 | -0.068664 | 121.17 | 0.23346 | 121.1151 | 121.0556 | 0.17853 | 0.11902
2883.1719 | 121.994 | -121.0693 | 0.88348 | 0.022049 | 0.022917 | 0.0097989 | 0.0014141 | -0.0019503 | 0.00038017 | -0.0027344 | -0.056244 | 0.019613 | 0.010873 | 0.0033773 | 0.014131 | -0.0049709 | -0.0010531 | -0.013733 | 0.85602 | 0.041199 | 0.56763 | 0.1831 | 0.44861 | 0.14191 | 0.004578 | 0.1831 | -0.14191 | -0.064087 | 0.082397 | -0.059509 | 0.004578 | 121.1105 | -0.027466 | -0.31586 | -0.28839 | 121.2524 | 0.14191 | 121.2112 | 121.5179 | 0.10071 | 0.40741 | 120.9274 | -0.70038 | -0.94757 | -0.24719 | 121.1517 | 0.2243 | 121.0739 | 121.0098 | 0.14649 | 0.082398
2884.1719 | 122.0581 | -121.0464 | 0.79651 | -0.077133 | 0.023558 | 0.0091581 | -0.00029472 | -0.0017367 | 0.00038017 | 4.2808e-05 | 0.013947 | 0.021108 | 0.011728 | 0.0038047 | 0.0098587 | -0.0056119 | -0.0034031 | -0.07782 | 0.84228 | 0.07782 | 0.61798 | 0.23346 | 0.48981 | 0.21973 | -0.036621 | 0.14648 | -0.10986 | 0.018311 | 0.09613 | 0 | 0.10071 | 121.1243 | 0.045776 | -0.17853 | -0.2243 | 121.2799 | 0.15564 | 121.2662 | 121.5363 | 0.14191 | 0.41199 | 120.9366 | -0.65003 | -0.7782 | -0.12817 | 121.1426 | 0.20599 | 121.1472 | 121.0464 | 0.21057 | 0.10986
2885.1719 | 122.0718 | -120.9915 | 0.85144 | -0.016098 | 0.022277 | 0.011294 | -0.0011493 | 0.0014675 | 0.0037982 | -0.00081179 | -0.022675 | 0.019826 | 0.01066 | 0.0025229 | 0.013704 | -0.0047573 | -0.0051121 | -0.041199 | 0.84686 | 0.091553 | 0.66834 | 0.29297 | 0.5127 | 0.23804 | -0.036621 | 0.14648 | -0.086975 | 0.045776 | 0.13733 | 0.013733 | 0.07782 | 121.0831 | -0.004577 | -0.1831 | -0.17853 | 121.2845 | 0.20142 | 121.2296 | 121.5042 | 0.14648 | 0.42114 | 120.9045 | -0.70496 | -0.80566 | -0.10071 | 121.1288 | 0.2243 | 121.0693 | 121.0052 | 0.16479 | 0.10071
2886.1719 | 122.0947 | -120.9915 | 0.88348 | -0.054245 | 0.021209 | 0.010013 | -0.00093572 | -0.00066869 | 0.001662 | -0.0020936 | -0.037933 | 0.020467 | 0.014505 | -0.0015359 | 0.014345 | -0.0041165 | -0.0025485 | -0.045776 | 0.84686 | 0.086975 | 0.68665 | 0.28839 | 0.54016 | 0.2243 | -0.018311 | 0.14191 | -0.07782 | 0.050354 | 0.14648 | 0.027466 | 0.11444 | 121.0785 | -0.036621 | -0.19684 | -0.16022 | 121.2799 | 0.20142 | 121.2158 | 121.5317 | 0.13733 | 0.45319 | 120.9137 | -0.74158 | -0.83313 | -0.091553 | 121.138 | 0.2243 | 121.106 | 121.019 | 0.19226 | 0.10529
2887.1719 | 122.113 | -121.0098 | 0.88348 | -0.10765 | 0.023772 | 0.010226 | -0.0015765 | -0.0010959 | 0.00059397 | -0.00038459 | 0.010895 | 0.020254 | 0.0091644 | -0.0017495 | 0.014131 | -0.0068935 | -0.0021213 | -0.050354 | 0.80109 | 0.045776 | 0.6546 | 0.23346 | 0.52643 | 0.22888 | 0.009155 | 0.14191 | -0.11902 | 0.036621 | 0.16022 | 0.041199 | 0.09613 | 121.0556 | -0.082398 | -0.22888 | -0.14648 | 121.2433 | 0.18768 | 121.2387 | 121.5363 | 0.18311 | 0.48065 | 120.8908 | -0.74158 | -0.84686 | -0.10529 | 121.17 | 0.27924 | 121.106 | 121.051 | 0.21515 | 0.16022
2888.1719 | 122.0993 | -120.9961 | 0.88806 | 0.075455 | 0.020568 | 0.01044 | -0.0013629 | -0.00024129 | 0.0012348 | -0.0027344 | -0.0074159 | 0.020681 | 0.01066 | 0.0038047 | 0.012422 | -0.0066799 | -0.0046847 | 0.009155 | 0.82855 | 0.073242 | 0.71869 | 0.29297 | 0.59967 | 0.2655 | 0.1236 | 0.14191 | -0.11444 | 0.036621 | 0.16937 | 0.059509 | 0.11902 | 121.0693 | -0.05951 | -0.16937 | -0.10986 | 121.2891 | 0.21973 | 121.2616 | 121.5958 | 0.19226 | 0.52643 | 120.8817 | -0.74616 | -0.85144 | -0.10529 | 121.1655 | 0.28381 | 121.1151 | 121.0556 | 0.23346 | 0.17395
2889.1719 | 122.1313 | -120.9641 | 0.88806 | -0.046616 | 0.023345 | 0.0095853 | -0.00029472 | -0.00066869 | 0.0023028 | -0.001239 | 0.010895 | 0.020681 | 0.0098052 | 0.0003867 | 0.016481 | -0.0060391 | -0.0076755 | -0.009155 | 0.89264 | 0.091553 | 0.70496 | 0.25635 | 0.58594 | 0.26093 | 0.073242 | 0.16022 | -0.09613 | 0.064087 | 0.19684 | 0.09613 | 0.15564 | 121.0556 | 0.004577 | -0.18311 | -0.18768 | 121.2204 | 0.1648 | 121.225 | 121.55 | 0.16937 | 0.49438 | 120.8679 | -0.72785 | -0.82398 | -0.09613 | 121.1609 | 0.29297 | 121.1197 | 121.0602 | 0.25177 | 0.19226
2890.1719 | 122.1313 | -120.9732 | 0.91553 | -0.069504 | 0.020995 | 0.010013 | 0.0012005 | 0.00061311 | 0.0018756 | 0.0019654 | -0.010467 | 0.019826 | 0.011087 | -0.0008951 | 0.014131 | -0.0056119 | -0.0034031 | -0.022888 | 0.86975 | 0.10071 | 0.70496 | 0.2655 | 0.59509 | 0.27924 | 0.064087 | 0.14648 | -0.10529 | 0.073242 | 0.19684 | 0.073242 | 0.15564 | 121.0739 | -0.045776 | -0.21057 | -0.1648 | 121.2387 | 0.16479 | 121.2524 | 121.5683 | 0.17853 | 0.49438 | 120.8679 | -0.76904 | -0.84228 | -0.073242 | 121.17 | 0.30212 | 121.1288 | 121.0464 | 0.26093 | 0.17853
2891.1719 | 122.1176 | -121.0144 | 0.90179 | -0.046616 | 0.023131 | 0.0089443 | 0.00077328 | -0.00024129 | 0.0027302 | 0.0011108 | -0.034882 | 0.020254 | 0.013437 | -0.0019631 | 0.012849 | -0.0058255 | -0.0014805 | -0.054932 | 0.86517 | 0.082397 | 0.59052 | 0.07782 | 0.44861 | 0.12817 | 0 | 0.082397 | -0.11444 | 0.032043 | 0.15106 | 0.018311 | 0.09613 | 121.0968 | -0.036621 | -0.31128 | -0.27466 | 121.0922 | -0.004577 | 121.1426 | 121.463 | 0.045777 | 0.36621 | 120.9 | -0.8194 | -0.86975 | -0.050354 | 121.1655 | 0.2655 | 121.1105 | 121.0327 | 0.21057 | 0.13275
2892.1719 | 122.1039 | -121.019 | 0.89264 | -0.046616 | 0.023345 | 0.0091581 | 0.0014141 | 0.00018591 | 0.0020892 | -0.0016662 | -0.034882 | 0.021535 | 0.01365 | -0.0023905 | 0.014131 | -0.0045437 | -0.0048983 | -0.082397 | 0.87433 | 0.086975 | 0.69122 | 0.23346 | 0.54016 | 0.22888 | -0.045776 | 0.14648 | -0.10986 | 0 | 0.15106 | 0.027466 | 0.082397 | 121.106 | -0.01831 | -0.20142 | -0.18311 | 121.2524 | 0.14648 | 121.2479 | 121.5591 | 0.14191 | 0.45319 | 120.9091 | -0.74615 | -0.89264 | -0.14648 | 121.17 | 0.26093 | 121.1014 | 121.0464 | 0.19226 | 0.13733
2893.1719 | 122.1313 | -120.9961 | 0.90179 | 0.0067903 | 0.023986 | 0.010653 | -8.1116e-05 | -0.0025913 | -0.0019696 | -0.00081179 | -0.013519 | 0.019399 | 0.01066 | -0.0013223 | 0.014345 | -0.0060391 | -0.0048983 | 0.032043 | 0.89264 | 0.09613 | 0.81482 | 0.14648 | 0.4303 | 0.31586 | 0.10071 | 0.13275 | -0.11444 | 0.068665 | 0.19226 | 0.082397 | 0.12817 | 121.0922 | -0.009155 | -0.086975 | -0.07782 | 121.1426 | 0.050354 | 121.312 | 121.4264 | 0.21973 | 0.33417 | 120.8817 | -0.76904 | -0.83313 | -0.064086 | 121.1884 | 0.3067 | 121.1243 | 121.0785 | 0.24262 | 0.19684
2894.1719 | 122.1313 | -120.9457 | 0.99335 | -0.023727 | 0.022277 | 0.0095853 | -0.00029472 | 0.0014675 | 0.0035846 | -0.00059819 | -0.0013123 | 0.020681 | 0.012369 | -0.0032449 | 0.015199 | -0.0064663 | -0.0040439 | -0.018311 | 0.95673 | 0.11444 | 0.76904 | 0.26093 | 0.65918 | 0.29297 | 0.022888 | 0.15106 | -0.09613 | 0.050354 | 0.17395 | 0.059509 | 0.14191 | 121.0602 | -0.036621 | -0.2243 | -0.18768 | 121.2067 | 0.14648 | 121.2387 | 121.6049 | 0.17853 | 0.54474 | 120.8496 | -0.84228 | -0.94299 | -0.10071 | 121.1197 | 0.27008 | 121.0876 | 121.0052 | 0.23804 | 0.15564
2895.1719 | 122.1222 | -120.9274 | 0.97504 | -0.023727 | 0.021422 | 0.0083035 | -0.0017901 | -0.0023777 | 0.0027302 | 0.001538 | 0.02005 | 0.021535 | 0.011942 | 0.0031637 | 0.016054 | -0.0068935 | -0.0019077 | -0.027466 | 0.92926 | 0.11444 | 0.78278 | 0.24261 | 0.6546 | 0.28381 | 0.009155 | 0.13275 | -0.10986 | 0.059509 | 0.17395 | 0.059509 | 0.1236 | 121.0419 | -0.045777 | -0.19226 | -0.14648 | 121.17 | 0.12817 | 121.2112 | 121.582 | 0.16937 | 0.54016 | 120.8176 | -0.84229 | -0.91553 | -0.073242 | 121.1014 | 0.28381 | 121.051 | 120.9869 | 0.23346 | 0.16937
2896.1719 | 122.1359 | -120.8588 | 1.0437 | -0.054245 | 0.021422 | 0.0087307 | -0.0020037 | -0.0021641 | 0.003371 | -0.0014526 | -0.013519 | 0.021963 | 0.012369 | -0.0011087 | 0.011781 | -0.0036893 | -0.0044711 | -0.041199 | 0.98419 | 0.12817 | 0.83313 | 0.2655 | 0.71411 | 0.3067 | 0.036621 | 0.15106 | -0.09613 | 0.045776 | 0.16479 | 0.073242 | 0.12817 | 120.9869 | -0.059509 | -0.21057 | -0.15106 | 121.1243 | 0.13733 | 121.1655 | 121.5729 | 0.17853 | 0.58594 | 120.7626 | -0.89264 | -0.99792 | -0.10529 | 121.0236 | 0.26092 | 120.9869 | 120.932 | 0.2243 | 0.16937
2897.1719 | 122.1497 | -120.8679 | 1.0391 | 0.0067903 | 0.022063 | 0.011081 | -8.1116e-05 | -0.0017367 | 0.0018756 | -0.0016662 | -0.019623 | 0.021108 | 0.010873 | 0.0020957 | 0.012208 | -0.0058255 | 0.00044212 | 0.004578 | 1.0391 | 0.16937 | 0.88348 | 0.31128 | 0.75531 | 0.3479 | 0.10986 | 0.13733 | -0.082397 | 0.064087 | 0.20599 | 0.091553 | 0.14191 | 121.0373 | 0 | -0.15564 | -0.15564 | 121.1792 | 0.14191 | 121.2158 | 121.6232 | 0.17853 | 0.58594 | 120.7855 | -0.90179 | -0.97504 | -0.073242 | 121.0739 | 0.28839 | 121.0098 | 120.9595 | 0.2243 | 0.17395
2898.1719 | 122.0398 | -120.8862 | 1.0986 | 0.029678 | 0.020354 | 0.01044 | -0.00093572 | 0.0012539 | 0.0012348 | 0.0034606 | -0.025726 | 0.020254 | 0.011087 | 0.0016685 | 0.01349 | -0.0073209 | -0.0051121 | 0.032043 | 1.0574 | 0.14648 | 0.80109 | 0.20599 | 0.70038 | 0.25177 | 0.050354 | 0.17395 | -0.10986 | -0.041199 | 0.054932 | -0.045776 | 0.009155 | 121.0327 | -0.041199 | -0.29755 | -0.25635 | 121.0922 | 0.05951 | 121.138 | 121.5866 | 0.10529 | 0.55389 | 120.7764 | -0.92468 | -1.1398 | -0.21515 | 120.9412 | 0.16479 | 120.8954 | 120.8405 | 0.11902 | 0.064087
2899.1719 | 122.1588 | -120.8496 | 1.0757 | -0.031357 | 0.021636 | 0.010226 | 0.00013248 | 0.0018949 | 0.0031574 | 0.00047001 | 0.0078429 | 0.019613 | 0.010446 | -0.0013223 | 0.014772 | -0.0058255 | -0.0066073 | -0.004578 | 1.0391 | 0.16022 | 0.90637 | 0.32044 | 0.77362 | 0.3479 | 0.10986 | 0.15564 | -0.10529 | 0.064087 | 0.20599 | 0.086975 | 0.14191 | 121.0098 | -0.036621 | -0.16937 | -0.13275 | 121.17 | 0.16022 | 121.1975 | 121.6232 | 0.18768 | 0.6134 | 120.7443 | -0.9201 | -1.0117 | -0.091553 | 121.0556 | 0.31128 | 120.9915 | 120.9366 | 0.24719 | 0.19226
2900.1719 | 122.1405 | -120.845 | 1.0849 | 0.10597 | 0.020781 | 0.01044 | -0.00093572 | -0.0023777 | 0.00016657 | 0.0019654 | -0.019623 | 0.019613 | 0.011087 | 0.0003867 | 0.013917 | -0.0075345 | -0.0063937 | -0.073242 | 1.0208 | 0.10071 | 0.87433 | 0.23804 | 0.75073 | 0.32044 | 0.004578 | 0.13733 | -0.11444 | 0.027466 | 0.16479 | 0.041199 | 0.11902 | 120.9457 | -0.064087 | -0.21057 | -0.14648 | 121.0831 | 0.13733 | 121.1655 | 121.5958 | 0.21973 | 0.65002 | 120.7306 | -0.94757 | -1.0574 | -0.10986 | 121.0098 | 0.27924 | 120.9641 | 120.8862 | 0.23346 | 0.15564
2901.1719 | 122.1817 | -120.8267 | 1.0895 | -0.031357 | 0.023345 | 0.0095853 | -0.0017901 | -0.0010959 | 0.0025166 | -0.0020936 | 0.010895 | 0.020681 | 0.011301 | 0.0038047 | 0.014558 | -0.0062527 | -0.0023349 | 0.013733 | 1.0895 | 0.16479 | 0.92468 | 0.30212 | 0.81024 | 0.35706 | 0.082397 | 0.15106 | -0.091553 | 0.082397 | 0.21057 | 0.082397 | 0.15106 | 120.9915 | 0 | -0.16479 | -0.16479 | 121.1288 | 0.13733 | 121.1838 | 121.637 | 0.19226 | 0.64545 | 120.7352 | -0.93842 | -1.0071 | -0.068665 | 121.0373 | 0.30212 | 120.9778 | 120.9091 | 0.24261 | 0.17395
2902.1719 | 122.1771 | -120.8038 | 1.1078 | -0.084763 | 0.023131 | 0.0097989 | -0.00050852 | 0.00082671 | 0.0029438 | 0.001538 | 0.02005 | 0.019399 | 0.011514 | 0.0003867 | 0.01349 | -0.0071071 | -0.0046847 | 0.004578 | 1.0849 | 0.16937 | 0.92468 | 0.28839 | 0.80109 | 0.35706 | 0.068665 | 0.15564 | -0.10529 | 0.07782 | 0.21057 | 0.09613 | 0.14648 | 120.9732 | -0.022888 | -0.1831 | -0.16022 | 121.0922 | 0.11902 | 121.1609 | 121.6049 | 0.18768 | 0.63171 | 120.6985 | -0.95215 | -1.03 | -0.07782 | 121.0144 | 0.31586 | 120.9503 | 120.9 | 0.25177 | 0.20142
2903.1719 | 121.9711 | -120.7993 | 1.1673 | 0.022049 | 0.020995 | 0.0091581 | -0.00050852 | -0.00024129 | 0.0040118 | -0.001239 | -0.056244 | 0.021322 | 0.012796 | 0.0006003 | 0.01349 | -0.0081753 | -0.0042575 | -0.032043 | 1.1307 | 0.13733 | 0.83313 | 0.18768 | 0.71411 | 0.2243 | -0.004578 | 0.17853 | -0.13733 | 0.009155 | 0.31586 | 0.10071 | 0.064087 | 120.9366 | -0.036621 | -0.33417 | -0.29755 | 120.9869 | 0.050354 | 121.0236 | 121.5134 | 0.086975 | 0.57678 | 120.6619 | -0.98877 | -1.1581 | -0.16937 | 121.1151 | 0.45319 | 120.8633 | 120.9 | 0.20142 | 0.23804
2904.1719 | 122.1909 | -120.7809 | 1.1307 | -0.031357 | 0.023986 | 0.0097989 | 0.00098688 | -2.7692e-05 | 0.0042254 | 0.001538 | 0.0078429 | 0.020467 | 0.012369 | -0.0034585 | 0.014345 | -0.0053983 | -0.0031895 | -0.013733 | 1.1307 | 0.19226 | 0.95215 | 0.29297 | 0.84228 | 0.36163 | 0.064087 | 0.16022 | -0.10529 | 0.07782 | 0.18768 | 0.073242 | 0.14191 | 120.9732 | 0 | -0.17853 | -0.17853 | 121.0739 | 0.10071 | 121.1426 | 121.6232 | 0.16937 | 0.65002 | 120.6757 | -0.97046 | -1.0529 | -0.082397 | 120.9686 | 0.29297 | 120.9229 | 120.8542 | 0.24719 | 0.17853
2905.1719 | 122.2183 | -120.7901 | 1.1444 | 0.0067903 | 0.023986 | 0.0057401 | 0.00013248 | -0.00024129 | 0.0048664 | 0.0011108 | 0.016998 | 0.02004 | 0.01301 | 0.0038047 | 0.015413 | -0.0060391 | -0.0023349 | 0.009155 | 1.1078 | 0.19226 | 0.99792 | 0.36163 | 0.86975 | 0.3891 | 0.14648 | 0.12817 | -0.09613 | 0.068665 | 0.20599 | 0.10529 | 0.14191 | 120.9824 | -0.036621 | -0.14648 | -0.10986 | 121.1517 | 0.16937 | 121.1792 | 121.6599 | 0.19684 | 0.67749 | 120.694 | -1.0162 | -1.0757 | -0.059509 | 120.9961 | 0.30212 | 120.932 | 120.8954 | 0.23804 | 0.20142
2906.1719 | 122.1359 | -120.7809 | 1.1536 | -0.016098 | 0.022277 | 0.0085171 | 0.00055968 | 0.0025357 | 0.003371 | -0.0020936 | 0.010895 | 0.021322 | 0.010873 | 0.0035911 | 0.014131 | -0.0036893 | -0.0021213 | -0.068665 | 1.1353 | 0.17853 | 0.94757 | 0.25635 | 0.8194 | 0.33875 | -0.24719 | 0.13733 | -0.14191 | 0 | 0.10986 | 0.009155 | 0.07782 | 120.9595 | -0.01831 | -0.20599 | -0.18768 | 121.0373 | 0.07782 | 121.1197 | 121.6003 | 0.16022 | 0.64087 | 120.639 | -1.0162 | -1.1536 | -0.13733 | 120.8908 | 0.25177 | 120.8588 | 120.7901 | 0.21973 | 0.15106
2907.1719 | 122.1954 | -120.7581 | 1.181 | -0.031357 | 0.021209 | 0.0085171 | -0.00050852 | -0.00045509 | 0.0029438 | 4.2808e-05 | 0.0017395 | 0.019399 | 0.014719 | 0.0027365 | 0.016695 | -0.0073209 | 0.00044212 | 0 | 1.1765 | 0.21057 | 1.03 | 0.35706 | 0.91553 | 0.4303 | 0.12817 | 0.14191 | -0.082397 | 0.07782 | 0.20599 | 0.09613 | 0.13275 | 120.9686 | -0.004577 | -0.15106 | -0.14648 | 121.1151 | 0.14648 | 121.1884 | 121.6736 | 0.21973 | 0.70496 | 120.6757 | -1.0391 | -1.1032 | -0.064087 | 120.9641 | 0.28839 | 120.8908 | 120.8542 | 0.21515 | 0.17853
2908.1719 | 122.0627 | -120.7626 | 1.2268 | -0.031357 | 0.02249 | 0.0093717 | 0.00013248 | 0.00018591 | 0.001662 | 0.0013244 | -0.056244 | 0.020254 | 0.013437 | -0.0006815 | 0.013704 | -0.0079617 | -0.0038303 | 0.027466 | 1.2039 | 0.18768 | 0.94757 | 0.24719 | 0.85144 | 0.29297 | 0.018311 | 0.19684 | -0.14191 | -0.045776 | 0.07782 | -0.041199 | 0 | 120.9503 | -0.022889 | -0.27924 | -0.25635 | 121.0098 | 0.059509 | 121.0556 | 121.6141 | 0.10529 | 0.66376 | 120.6207 | -1.03 | -1.2726 | -0.24261 | 120.8405 | 0.21973 | 120.7626 | 120.7214 | 0.14191 | 0.10071
2909.1719 | 122.1954 | -120.7169 | 1.1993 | -0.023727 | 0.023772 | 0.0091581 | 0.00055968 | 0.0027493 | 0.0025166 | -0.0016662 | -0.0013123 | 0.021535 | 0.01301 | -0.0045267 | 0.012636 | -0.0086025 | -0.0027621 | 0 | 1.2039 | 0.2243 | 1.03 | 0.3479 | 0.92468 | 0.40741 | 0.10529 | 0.15564 | -0.091553 | 0.073242 | 0.20599 | 0.091553 | 0.15106 | 120.9412 | 0.004577 | -0.16937 | -0.17395 | 121.0648 | 0.1236 | 121.1243 | 121.6415 | 0.18311 | 0.70038 | 120.6253 | -1.0437 | -1.1261 | -0.082398 | 120.9229 | 0.29755 | 120.8679 | 120.8084 | 0.24261 | 0.18311
2910.1719 | 122.2046 | -120.7214 | 1.2222 | -0.077133 | 0.021422 | 0.0078763 | -8.1116e-05 | 0.0031765 | 0.0027302 | 0.00089721 | 0.0047913 | 0.022176 | 0.011301 | -0.0011087 | 0.014772 | -0.0051845 | -0.0021213 | 0 | 1.1948 | 0.23346 | 1.0483 | 0.34332 | 0.9201 | 0.40741 | 0.10071 | 0.15106 | -0.09613 | 0.07782 | 0.21515 | 0.09613 | 0.14191 | 120.9549 | -0.027466 | -0.17395 | -0.14648 | 121.0648 | 0.10986 | 121.1288 | 121.6415 | 0.17395 | 0.68665 | 120.6253 | -1.0712 | -1.1444 | -0.073242 | 120.9366 | 0.31128 | 120.8633 | 120.8176 | 0.23804 | 0.19226
2911.1719 | 122.2 | -120.7077 | 1.2085 | -0.0084687 | 0.023558 | 0.0089443 | 0.00034608 | 0.0012539 | 0.0020892 | -0.0020936 | 0.0078429 | 0.019399 | 0.011514 | -0.0032449 | 0.013917 | -0.0068935 | -0.0040439 | -0.004578 | 1.1856 | 0.21973 | 1.0574 | 0.32959 | 0.91553 | 0.40741 | 0.082397 | 0.16479 | -0.086975 | 0.07782 | 0.19684 | 0.10529 | 0.13733 | 120.9274 | -0.022888 | -0.15106 | -0.12817 | 121.0373 | 0.10986 | 121.1151 | 121.6232 | 0.18768 | 0.6958 | 120.6207 | -1.0437 | -1.1307 | -0.086975 | 120.9045 | 0.28381 | 120.845 | 120.813 | 0.2243 | 0.19226
2912.1719 | 122.1725 | -120.7123 | 1.236 | 0.0067903 | 0.022063 | 0.010013 | 0.0024823 | 0.00082671 | 0.0035846 | -0.0052978 | -0.0013123 | 0.021322 | 0.013223 | 0.0014547 | 0.015413 | -0.0068935 | -0.0036167 | -0.045776 | 1.2085 | 0.21515 | 1.0071 | 0.2655 | 0.87433 | 0.35706 | 0.013733 | 0.10986 | -0.10071 | 0.027466 | 0.15106 | 0.032043 | 0.10529 | 120.9274 | -0.027466 | -0.22888 | -0.20142 | 120.9778 | 0.050354 | 121.0693 | 121.5866 | 0.14191 | 0.65918 | 120.6116 | -1.1261 | -1.2085 | -0.082397 | 120.8633 | 0.25177 | 120.8176 | 120.7443 | 0.20599 | 0.13275
2913.1719 | 122.1771 | -120.6985 | 1.2543 | 0.083084 | 0.023772 | 0.011722 | -0.00093572 | 0.0012539 | 0.001662 | -0.0014526 | -0.0043641 | 0.021322 | 0.011942 | -0.0034585 | 0.013917 | -0.0068935 | -0.0042575 | -0.082397 | 1.2085 | 0.19684 | 1.0849 | 0.35248 | 0.95215 | 0.42572 | 0.13275 | 0.1236 | -0.1236 | 0.045776 | 0.16937 | 0.059509 | 0.1236 | 120.8954 | -0.045776 | -0.16937 | -0.1236 | 121.051 | 0.15564 | 121.1243 | 121.6507 | 0.22888 | 0.75531 | 120.575 | -1.1307 | -1.2085 | -0.07782 | 120.8679 | 0.29297 | 120.8221 | 120.7581 | 0.24719 | 0.1831
2914.1719 | 122.1954 | -120.6894 | 1.2634 | -0.069504 | 0.022704 | 0.0089443 | -0.0015765 | 0.0033901 | 0.0025166 | -0.00081179 | -0.013519 | 0.019826 | 0.01301 | 0.0003867 | 0.015199 | -0.0083889 | 0.00065592 | -0.032043 | 1.2589 | 0.22888 | 1.0574 | 0.34332 | 0.96588 | 0.41199 | 0.064087 | 0.15106 | -0.09613 | 0.073242 | 0.20142 | 0.082397 | 0.14191 | 120.9183 | -0.004578 | -0.20599 | -0.20142 | 121.0327 | 0.11444 | 121.1014 | 121.6553 | 0.1831 | 0.737 | 120.5933 | -1.1124 | -1.1902 | -0.07782 | 120.8908 | 0.29755 | 120.8313 | 120.7718 | 0.23804 | 0.17853
2915.1719 | 122.1222 | -120.7489 | 1.2543 | 0.029678 | 0.024413 | 0.011508 | 0.00098688 | 0.00039951 | 0.0040118 | 0.00047001 | -0.098968 | 0.021535 | 0.011942 | 0.0029501 | 0.014772 | -0.0045437 | -0.0036167 | 0.004578 | 1.2451 | 0.1831 | 0.99335 | 0.23804 | 0.87433 | 0.29755 | 0.041199 | 0.16022 | -0.15106 | -0.009155 | 0.10986 | -0.018311 | 0.041199 | 120.932 | -0.009155 | -0.26093 | -0.25177 | 120.9869 | 0.054932 | 121.0464 | 121.6232 | 0.11444 | 0.69122 | 120.5978 | -1.0941 | -1.2634 | -0.16937 | 120.8588 | 0.26093 | 120.7901 | 120.7306 | 0.19226 | 0.13275
2916.1719 | 122.2183 | -120.6665 | 1.2497 | -0.023727 | 0.023345 | 0.0053129 | -0.00029472 | -0.0025913 | 0.0048664 | -0.00059819 | 0.013947 | 0.02004 | 0.0095916 | 0.0048727 | 0.01584 | -0.0071071 | -0.0036167 | 0.004578 | 1.2497 | 0.25635 | 1.1169 | 0.40741 | 0.98419 | 0.43945 | 0.13275 | 0.13275 | -0.10529 | 0.086975 | 0.2243 | 0.10986 | 0.16937 | 120.9229 | 0 | -0.13275 | -0.13275 | 121.0739 | 0.15106 | 121.106 | 121.6507 | 0.1831 | 0.72784 | 120.5612 | -1.1169 | -1.1627 | -0.045776 | 120.8908 | 0.32959 | 120.8359 | 120.7764 | 0.27466 | 0.21515
2917.1719 | 122.2366 | -120.6482 | 1.2543 | -0.0084687 | 0.021849 | 0.0089443 | -0.00029472 | 0.0014675 | 0.0018756 | -0.0010254 | 0.013947 | 0.021535 | 0.011514 | -0.0083719 | 0.014772 | -0.0092435 | -0.0025485 | 0 | 1.2405 | 0.25635 | 1.1398 | 0.42572 | 1.0071 | 0.43945 | 0.13733 | 0.14191 | -0.10071 | 0.07782 | 0.2243 | 0.086975 | 0.13733 | 120.9045 | -0.013732 | -0.11444 | -0.10071 | 121.0739 | 0.16937 | 121.0876 | 121.6553 | 0.1831 | 0.75073 | 120.5475 | -1.1124 | -1.1765 | -0.064087 | 120.8725 | 0.32501 | 120.7855 | 120.7352 | 0.23804 | 0.18768
2918.1719 | 122.0947 | -120.694 | 1.2909 | 0.029678 | 0.020995 | 0.0070217 | 0.00055968 | -0.0021641 | 0.0025166 | 0.0036744 | -0.089813 | 0.021322 | 0.012155 | 0.0023093 | 0.014772 | -0.0058255 | -0.00083948 | -0.022888 | 1.2909 | 0.18768 | 0.94299 | 0.17395 | 0.97046 | 0.32501 | 0 | 0.21057 | -0.1236 | -0.050354 | 0.064087 | -0.064087 | 0 | 120.8817 | 0 | -0.3479 | -0.3479 | 120.8679 | -0.013733 | 121.019 | 121.6644 | 0.13733 | 0.78278 | 120.5704 | -1.0803 | -1.3412 | -0.26093 | 120.7581 | 0.18768 | 120.694 | 120.6299 | 0.1236 | 0.059509
2919.1719 | 122.2229 | -120.6848 | 1.2817 | 0.0067903 | 0.020995 | 0.0087307 | -0.00093572 | -0.0023777 | 0.0042254 | -0.0018798 | 0.026154 | 0.021322 | 0.012582 | 0.0012411 | 0.014558 | -0.0053983 | -0.0051121 | -0.013733 | 1.2268 | 0.23804 | 1.0757 | 0.33417 | 0.94299 | 0.37994 | 0.022888 | 0.14648 | -0.091553 | 0.041199 | 0.18768 | 0.050354 | 0.12817 | 120.9229 | -0.054931 | -0.20599 | -0.15106 | 121.019 | 0.09613 | 121.0648 | 121.6278 | 0.14191 | 0.70496 | 120.5933 | -1.1353 | -1.2405 | -0.10529 | 120.8725 | 0.27924 | 120.813 | 120.7352 | 0.21973 | 0.14191
2920.1719 | 122.1542 | -120.7214 | 1.3 | -0.061874 | 0.019286 | 0.0076627 | 0.0022687 | 0.0046719 | 0.0040118 | -0.0014526 | -0.065399 | 0.021963 | 0.013437 | 0.0076499 | 0.013917 | -0.0051845 | -0.0019077 | -0.013733 | 1.2955 | 0.19226 | 1.0208 | 0.24719 | 0.88348 | 0.31128 | 0.004578 | 0.17853 | -0.14191 | -0.032043 | 0.082397 | -0.032043 | 0.032043 | 120.9137 | -0.004578 | -0.27924 | -0.27466 | 120.9686 | 0.054931 | 121.0327 | 121.6049 | 0.11902 | 0.69122 | 120.5795 | -1.1215 | -1.3321 | -0.21057 | 120.8038 | 0.2243 | 120.7535 | 120.6894 | 0.17395 | 0.10986
2921.1719 | 122.1725 | -120.6711 | 1.2497 | -0.054245 | 0.021849 | 0.0078763 | -0.00093572 | 0.0023221 | 0.0037982 | -0.0027344 | 0.038361 | 0.020894 | 0.012369 | -0.0062357 | 0.016267 | -0.0045437 | -0.00041228 | 0.045776 | 1.2726 | 0.22888 | 1.1398 | 0.4303 | 1.0254 | 0.44403 | 0.14191 | 0.14648 | -0.09613 | -0.004578 | 0.17853 | 0.13733 | 0.16022 | 120.9 | 0.022888 | -0.10986 | -0.13275 | 121.1014 | 0.20142 | 121.1151 | 121.6965 | 0.21515 | 0.79651 | 120.575 | -1.1032 | -1.2543 | -0.15106 | 120.8496 | 0.27466 | 120.8313 | 120.8084 | 0.25635 | 0.23346
2922.1719 | 122.2687 | -120.6528 | 1.2955 | -0.016098 | 0.021636 | 0.0085171 | -0.0020037 | 0.00018591 | 0.0031574 | -0.00081179 | 0.041412 | 0.020467 | 0.01301 | -0.0013223 | 0.012636 | -0.0083889 | -0.0031895 | 0.032043 | 1.2863 | 0.2655 | 1.1673 | 0.4303 | 1.0529 | 0.4715 | 0.15564 | 0.14648 | -0.09613 | 0.086975 | 0.10071 | -0.09613 | 0.045776 | 120.9183 | -0.009155 | -0.12817 | -0.11902 | 121.0831 | 0.1648 | 121.1243 | 121.7056 | 0.20599 | 0.78735 | 120.5566 | -1.149 | -1.2085 | -0.059509 | 120.7535 | 0.19684 | 120.6985 | 120.5566 | 0.14191 | 0
2923.1719 | 122.1451 | -120.6482 | 1.3412 | -0.0084687 | 0.02249 | 0.0085171 | 0.0016279 | 0.0014675 | 0.003371 | -0.0023072 | -0.068451 | 0.022176 | 0.012369 | 0.0008139 | 0.014131 | -0.0041165 | -0.0029757 | 0.004578 | 1.3229 | 0.22888 | 1.1032 | 0.33417 | 0.95215 | 0.35706 | 0.059509 | 0.15106 | -0.1236 | 0.004578 | 0.10529 | -0.027466 | 0.059509 | 120.8771 | -0.018311 | -0.23804 | -0.21973 | 120.9824 | 0.10528 | 121.0052 | 121.6003 | 0.12817 | 0.72327 | 120.5246 | -1.1902 | -1.3367 | -0.14648 | 120.7535 | 0.22888 | 120.7077 | 120.6207 | 0.1831 | 0.09613
2924.1719 | 122.1954 | -120.6528 | 1.3138 | -0.038986 | 0.023558 | 0.0083035 | -0.0026447 | 0.0014675 | 0.0035846 | 0.0017518 | 0.026154 | 0.020681 | 0.012796 | -0.0028177 | 0.014772 | -0.0073209 | -0.00019868 | 0.045776 | 1.2772 | 0.25635 | 1.1536 | 0.43488 | 1.0574 | 0.45776 | 0.15106 | 0.14648 | -0.10071 | -0.054932 | 0.33417 | -0.11902 | 0.013733 | 120.9091 | -0.036621 | -0.16022 | -0.1236 | 121.0876 | 0.17853 | 121.1105 | 121.7102 | 0.20142 | 0.80109 | 120.5521 | -1.1673 | -1.3687 | -0.20142 | 120.9869 | 0.43488 | 120.6665 | 120.5338 | 0.11444 | -0.018311
2925.1719 | 122.0764 | -120.6894 | 1.3596 | -0.023727 | 0.020995 | 0.010867 | 0.0031231 | 0.0010403 | 0.0048664 | 0.0013244 | -0.092865 | 0.021963 | 0.011942 | -0.0017495 | 0.01349 | -0.0047573 | -0.0057529 | -0.050354 | 1.3138 | 0.21057 | 1.0345 | 0.24261 | 0.88348 | 0.28839 | -0.11902 | 0.16022 | -0.16022 | -0.068665 | 0 | -0.10986 | -0.004578 | 120.9 | -0.045776 | -0.32501 | -0.27924 | 120.932 | 0.032044 | 120.9778 | 121.5729 | 0.07782 | 0.67291 | 120.5292 | -1.1993 | -1.4282 | -0.22888 | 120.6894 | 0.16022 | 120.6848 | 120.5795 | 0.15564 | 0.050354
2926.1719 | 122.2458 | -120.5933 | 1.3824 | -0.092392 | 0.02249 | 0.0093717 | -0.00029472 | 0.0018949 | 0.0023028 | -0.0010254 | 0.053619 | 0.019185 | 0.014291 | -0.0036721 | 0.014986 | -0.0064663 | -0.0048983 | -0.018311 | 1.3824 | 0.28839 | 1.2222 | 0.41657 | 1.1078 | 0.48523 | 0.07782 | 0.13275 | -0.082397 | 0.064087 | 0.21057 | 0.086975 | 0.13733 | 120.8817 | 0 | -0.16022 | -0.16022 | 121.0098 | 0.12817 | 121.0785 | 121.701 | 0.19684 | 0.8194 | 120.5109 | -1.2497 | -1.3184 | -0.068664 | 120.8038 | 0.29297 | 120.7306 | 120.6802 | 0.21973 | 0.16937
2927.1719 | 122.2595 | -120.5475 | 1.4145 | -0.061874 | 0.024413 | 0.0076627 | -0.0020037 | -0.0032321 | 0.0042254 | -0.0016662 | 0.038361 | 0.019826 | 0.014291 | -0.0043131 | 0.014986 | -0.0071071 | -0.0042575 | -0.022888 | 1.4145 | 0.30212 | 1.2863 | 0.44403 | 1.149 | 0.50354 | 0.10529 | 0.14191 | -0.086975 | 0.059509 | 0.19226 | 0.09613 | 0.13275 | 120.8496 | 0 | -0.12817 | -0.12817 | 120.9915 | 0.14191 | 121.051 | 121.6965 | 0.20142 | 0.84686 | 120.4605 | -1.2726 | -1.355 | -0.082398 | 120.7397 | 0.27924 | 120.6802 | 120.6436 | 0.21973 | 0.1831
2928.1719 | 122.1405 | -120.5521 | 1.4511 | 0.0067903 | 0.019713 | 0.0085171 | 0.0012005 | -0.0021641 | 0.001662 | -0.00038459 | -0.0043641 | 0.021535 | 0.011514 | 0.0016685 | 0.012208 | -0.0081753 | -0.0031895 | 0.027466 | 1.4328 | 0.29297 | 1.2085 | 0.37537 | 1.0895 | 0.40283 | 0.036621 | 0.16937 | -0.12817 | -0.054932 | 0.045776 | -0.041199 | 0.013733 | 120.845 | -0.018311 | -0.24262 | -0.2243 | 120.9274 | 0.082397 | 120.9549 | 121.6415 | 0.10986 | 0.79651 | 120.4239 | -1.2817 | -1.506 | -0.22431 | 120.5978 | 0.17395 | 120.5658 | 120.5109 | 0.14191 | 0.086975
2929.1719 | 122.1909 | -120.5704 | 1.4374 | -0.046616 | 0.023131 | 0.01044 | -0.0015765 | -0.0030185 | 0.0052936 | -0.00038459 | -0.0013123 | 0.01769 | 0.012582 | 0.0008139 | 0.013063 | -0.0092435 | -0.0078891 | 0.041199 | 1.4236 | 0.31128 | 1.2634 | 0.42572 | 1.1398 | 0.49439 | 0.1236 | 0.14648 | -0.086975 | 0.013733 | 0.15564 | 0.054932 | 0.086975 | 120.8817 | -0.013733 | -0.17395 | -0.16022 | 120.9961 | 0.11444 | 121.0648 | 121.7102 | 0.18311 | 0.82855 | 120.4834 | -1.2909 | -1.4236 | -0.13275 | 120.726 | 0.24261 | 120.6573 | 120.6253 | 0.17395 | 0.14191
2930.1719 | 122.1451 | -120.5566 | 1.4694 | -0.061874 | 0.022063 | 0.0083035 | -8.1116e-05 | 0.0038175 | 0.0046528 | -0.0010254 | -0.077606 | 0.02239 | 0.011514 | 0.0008139 | 0.014772 | -0.0017667 | -0.0044711 | 0.009155 | 1.4282 | 0.28839 | 1.2085 | 0.33417 | 1.0803 | 0.39825 | 0.018311 | 0.17853 | -0.11902 | 0 | 0.07782 | -0.036621 | 0.050354 | 120.845 | -0.041198 | -0.26093 | -0.21973 | 120.8908 | 0.045776 | 120.9549 | 121.637 | 0.10986 | 0.79193 | 120.4376 | -1.2909 | -1.4694 | -0.17853 | 120.6345 | 0.19684 | 120.607 | 120.52 | 0.16937 | 0.082398

## Thermal Data
Fire Time | North 25 C | Cable 1 | Cable 2
-0.5115 | 22.0642 | 20.8138 | 20.9555
0.4885 | 22.0556 | 20.8252 | 20.9615
1.4885 | 22.0616 | 20.8291 | 20.9635
2.4885 | 22.0596 | 20.8212 | 20.9591
3.4885 | 22.0859 | 20.828 | 20.959
4.4885 | 22.419 | 20.822 | 20.9623
5.4885 | 23.5014 | 20.8126 | 20.9622
6.4885 | 25.3813 | 20.8286 | 20.9644
7.4885 | 27.9037 | 20.8258 | 20.9583
8.4885 | 30.9894 | 20.8236 | 20.961
9.4885 | 34.525 | 20.8252 | 20.961
10.4885 | 38.4213 | 20.8156 | 20.9661
11.4885 | 42.5965 | 20.8166 | 20.9602
12.4885 | 47.0937 | 20.8155 | 20.9608
13.4885 | 51.8028 | 20.8121 | 20.962
14.4885 | 56.7304 | 20.8252 | 20.9664
15.4885 | 61.7795 | 20.8176 | 20.9634
16.4885 | 67.0223 | 20.8224 | 20.9651
17.4885 | 72.2915 | 20.8321 | 20.9703
18.4885 | 77.6682 | 20.8286 | 20.9713
19.4885 | 83.2282 | 20.8253 | 20.9796
20.4885 | 88.8108 | 20.8312 | 20.976
21.4885 | 94.5377 | 20.8462 | 20.986
22.4885 | 100.3464 | 20.8454 | 20.9822
23.4885 | 106.1414 | 20.8634 | 20.9864
24.4885 | 112.0394 | 20.8654 | 20.9894
25.4885 | 118.046 | 20.8652 | 20.9903
26.4885 | 124.0217 | 20.8859 | 20.9973
27.4885 | 130.0228 | 20.8951 | 21.007
28.4885 | 135.9653 | 20.9067 | 21.0143
29.4885 | 141.8941 | 20.9208 | 21.0222
30.4885 | 147.7688 | 20.9405 | 21.0379
31.4885 | 153.6592 | 20.966 | 21.053
32.4885 | 159.5549 | 20.9861 | 21.0631
33.4885 | 165.3751 | 21.0018 | 21.076
34.4885 | 171.3124 | 21.0528 | 21.0905
35.4885 | 177.3542 | 21.0693 | 21.1207
36.4885 | 183.2893 | 21.1223 | 21.1339
37.4885 | 189.209 | 21.1576 | 21.1483
38.4885 | 194.9984 | 21.2075 | 21.1754
39.4885 | 200.9115 | 21.2557 | 21.2097
40.4885 | 206.6786 | 21.3034 | 21.2389
41.4885 | 212.5087 | 21.3541 | 21.273
42.4885 | 218.3586 | 21.4287 | 21.3033
43.4885 | 224.1876 | 21.4968 | 21.3437
44.4885 | 229.9072 | 21.5666 | 21.3848
45.4885 | 235.6014 | 21.6338 | 21.4353
46.4885 | 241.2839 | 21.7298 | 21.4914
47.4885 | 246.9655 | 21.8127 | 21.5457
48.4885 | 252.5975 | 21.9143 | 21.6053
49.4885 | 258.3215 | 22.0135 | 21.6752
50.4885 | 264.0142 | 22.1225 | 21.7409
51.4885 | 269.4835 | 22.2416 | 21.814
52.4885 | 274.8425 | 22.3616 | 21.8932
53.4885 | 280.3276 | 22.48 | 21.9813
54.4885 | 285.8739 | 22.6176 | 22.0791
55.4885 | 291.4419 | 22.7534 | 22.1743
56.4885 | 296.9838 | 22.8957 | 22.2937
57.4885 | 302.5984 | 23.0674 | 22.4098
58.4885 | 307.9905 | 23.2237 | 22.5133
59.4885 | 312.8809 | 23.3818 | 22.6506
60.4885 | 317.2375 | 23.5786 | 22.7789
61.4885 | 321.3951 | 23.7711 | 22.9117
62.4885 | 325.1724 | 23.9716 | 23.0631
63.4885 | 328.7314 | 24.1826 | 23.214
64.4885 | 332.1231 | 24.3915 | 23.3689
65.4885 | 335.2458 | 24.611 | 23.543
66.4885 | 338.3061 | 24.8626 | 23.7103
67.4885 | 341.2721 | 25.1088 | 23.8925
68.4885 | 343.8299 | 25.3563 | 24.0824
69.4885 | 346.2514 | 25.6187 | 24.2818
70.4885 | 348.261 | 25.9166 | 24.5
71.4885 | 350.1062 | 26.1867 | 24.7184
72.4885 | 351.553 | 26.477 | 24.9631
73.4885 | 352.9021 | 26.7831 | 25.1953
74.4885 | 353.9317 | 27.0827 | 25.4372
75.4885 | 354.6307 | 27.4172 | 25.7135
76.4885 | 355.2203 | 27.7313 | 25.992
77.4885 | 355.5101 | 28.0849 | 26.2702
78.4885 | 355.7695 | 28.4356 | 26.5555
79.4885 | 356.1938 | 28.7723 | 26.8624
80.4885 | 356.2417 | 29.1343 | 27.1803
81.4885 | 356.16 | 29.5194 | 27.4907
82.4885 | 356.0225 | 29.9016 | 27.812
83.4885 | 355.8958 | 30.2853 | 28.164
84.4885 | 355.7503 | 30.6563 | 28.5084
85.4885 | 355.6245 | 31.079 | 28.8441
86.4885 | 355.4015 | 31.4872 | 29.2217
87.4885 | 355.0044 | 31.9072 | 29.5898
88.4885 | 354.6688 | 32.3339 | 29.9659
89.4885 | 354.3924 | 32.766 | 30.3826
90.4885 | 354.0579 | 33.2099 | 30.7699
91.4885 | 353.7571 | 33.637 | 31.1711
92.4885 | 353.4388 | 34.0884 | 31.5832
93.4885 | 352.9323 | 34.5397 | 32.0079
94.4885 | 352.3166 | 34.994 | 32.434
95.4885 | 351.7388 | 35.4781 | 32.8748
96.4885 | 351.3812 | 35.944 | 33.3092
97.4885 | 351.025 | 36.4101 | 33.7688
98.4885 | 350.5647 | 36.8947 | 34.2059
99.4885 | 350.1498 | 37.3678 | 34.6763
100.4885 | 349.9132 | 37.8515 | 35.1284
101.4885 | 349.8021 | 38.3359 | 35.6105
102.4885 | 349.5076 | 38.825 | 36.0745
103.4885 | 349.1556 | 39.3123 | 36.5514
104.4885 | 348.9 | 39.8123 | 37.0387
105.4885 | 348.7643 | 40.2999 | 37.5296
106.4885 | 348.5912 | 40.8127 | 38.0136
107.4885 | 348.3867 | 41.3015 | 38.5165
108.4885 | 348.2511 | 41.7877 | 38.9994
109.4885 | 348.1433 | 42.3174 | 39.4957
110.4885 | 348.239 | 42.8297 | 39.9949
111.4885 | 348.2324 | 43.3339 | 40.5065
112.4885 | 348.2933 | 43.8489 | 41.015
113.4885 | 348.3567 | 44.3562 | 41.5284
114.4885 | 348.422 | 44.8665 | 42.0355
115.4885 | 348.5297 | 45.3941 | 42.5402
116.4885 | 348.6769 | 45.909 | 43.0657
117.4885 | 348.7397 | 46.4275 | 43.5921
118.4885 | 348.8577 | 46.9424 | 44.0996
119.4885 | 348.8976 | 47.4645 | 44.622
120.4885 | 349.0543 | 47.9803 | 45.1503
121.4885 | 349.2501 | 48.496 | 45.6609
122.4885 | 349.5216 | 49.022 | 46.1782
123.4885 | 349.4985 | 49.544 | 46.7072
124.4885 | 349.6418 | 50.0671 | 47.2273
125.4885 | 349.9115 | 50.5856 | 47.7506
126.4885 | 350.1658 | 51.109 | 48.272
127.4885 | 350.2614 | 51.6267 | 48.7859
128.4885 | 350.3414 | 52.1494 | 49.3194
129.4885 | 350.4636 | 52.6765 | 49.865
130.4885 | 350.512 | 53.1939 | 50.3779
131.4885 | 350.609 | 53.7215 | 50.9092
132.4885 | 350.7357 | 54.2369 | 51.4261
133.4885 | 350.7883 | 54.7739 | 51.9543
134.4885 | 350.8805 | 55.2953 | 52.48
135.4885 | 351.0573 | 55.8162 | 53.0296
136.4885 | 351.2617 | 56.3303 | 53.5424
137.4885 | 351.453 | 56.8599 | 54.0728
138.4885 | 351.5198 | 57.3724 | 54.6163
139.4885 | 351.4482 | 57.8913 | 55.1527
140.4885 | 351.4606 | 58.4174 | 55.6701
141.4885 | 351.5223 | 58.9351 | 56.204
142.4885 | 351.6104 | 59.441 | 56.7402
143.4885 | 351.7477 | 59.9585 | 57.2712
144.4885 | 351.8802 | 60.4603 | 57.8001
145.4885 | 351.9204 | 60.9781 | 58.3259
146.4885 | 352.0533 | 61.4969 | 58.868
147.4885 | 352.131 | 61.991 | 59.3947
148.4885 | 352.0319 | 62.5166 | 59.9291
149.4885 | 352.1646 | 63.0331 | 60.4593
150.4885 | 352.3726 | 63.5461 | 60.9869
151.4885 | 352.5392 | 64.0735 | 61.5205
152.4885 | 352.5557 | 64.582 | 62.0532
153.4885 | 352.5424 | 65.0947 | 62.5907
154.4885 | 352.5808 | 65.5974 | 63.1216
155.4885 | 352.6412 | 66.0958 | 63.6556
156.4885 | 352.7379 | 66.6112 | 64.1998
157.4885 | 352.8147 | 67.1156 | 64.7192
158.4885 | 352.823 | 67.6291 | 65.2512
159.4885 | 352.8097 | 68.1362 | 65.7772
160.4885 | 352.8515 | 68.6292 | 66.3203
161.4885 | 352.6971 | 69.1343 | 66.8544
162.4885 | 352.6293 | 69.6376 | 67.3815
163.4885 | 352.6299 | 70.1456 | 67.9015
164.4885 | 352.5883 | 70.6472 | 68.4359
165.4885 | 352.6576 | 71.1431 | 68.9554
166.4885 | 352.532 | 71.6478 | 69.4812
167.4885 | 352.6117 | 72.1328 | 70.0154
168.4885 | 352.7176 | 72.6156 | 70.5428
169.4885 | 352.7771 | 73.122 | 71.0576
170.4885 | 352.8507 | 73.6187 | 71.5803
171.4885 | 352.6254 | 74.0993 | 72.1044
172.4885 | 352.5852 | 74.6017 | 72.6222
173.4885 | 352.5783 | 75.0886 | 73.1507
174.4885 | 352.5666 | 75.5705 | 73.6796
175.4885 | 352.5158 | 76.0737 | 74.1932
176.4885 | 352.4832 | 76.557 | 74.7155
177.4885 | 352.3242 | 77.0602 | 75.2224
178.4885 | 352.2761 | 77.5488 | 75.7385
179.4885 | 352.2565 | 78.0446 | 76.2582
180.4885 | 352.3066 | 78.5479 | 76.7748
181.4885 | 352.1991 | 79.042 | 77.2874
182.4885 | 352.1685 | 79.5522 | 77.8022
183.4885 | 352.1551 | 80.0556 | 78.3166
184.4885 | 352.0472 | 80.5595 | 78.8442
185.4885 | 352.1725 | 81.0623 | 79.3576
186.4885 | 352.3535 | 81.5696 | 79.8811
187.4885 | 352.4057 | 82.0777 | 80.4127
188.4885 | 352.2742 | 82.5859 | 80.9486
189.4885 | 352.1229 | 83.1054 | 81.482
190.4885 | 351.9664 | 83.6074 | 82.0341
191.4885 | 351.9097 | 84.13 | 82.6188
192.4885 | 351.898 | 84.6453 | 83.1734
193.4885 | 351.9043 | 85.1608 | 83.7622
194.4885 | 351.8497 | 85.6774 | 84.3588
195.4885 | 351.8283 | 86.2215 | 84.9849
196.4885 | 351.7978 | 86.7512 | 85.6436
197.4885 | 351.7516 | 87.2811 | 86.3566
198.4885 | 351.8411 | 87.841 | 87.1391
199.4885 | 351.8642 | 88.411 | 88.0445
200.4885 | 352.0718 | 88.984 | 89.1502
201.4885 | 352.2029 | 89.6119 | 90.5146
202.4885 | 352.3707 | 90.3158 | 92.3133
203.4885 | 352.497 | 91.1138 | 94.0902
204.4885 | 352.5714 | 91.9737 | 95.3172
205.4885 | 352.6318 | 92.8419 | 96.0733
206.4885 | 352.7382 | 93.6676 | 96.4958
207.4885 | 352.7827 | 94.397 | 96.7324
208.4885 | 352.8817 | 94.9907 | 96.873
209.4885 | 353.0087 | 95.3589 | 96.9791
210.4885 | 352.9917 | 95.6422 | 97.0732
211.4885 | 352.9901 | 95.9173 | 97.1402
212.4885 | 353.1405 | 96.198 | 97.2132
213.4885 | 353.2263 | 96.4884 | 97.2747
214.4885 | 353.3728 | 96.7696 | 97.3512
215.4885 | 353.5598 | 97.0523 | 97.4121
216.4885 | 353.5636 | 97.305 | 97.4882
217.4885 | 353.6239 | 97.5599 | 97.5721
218.4885 | 353.6403 | 97.8191 | 97.6582
219.4885 | 353.6507 | 98.098 | 97.7608
220.4885 | 353.754 | 98.376 | 97.8694
221.4885 | 353.7234 | 98.6416 | 97.999
222.4885 | 353.6192 | 98.8971 | 98.121
223.4885 | 353.5438 | 99.1507 | 98.271
224.4885 | 353.5668 | 99.3966 | 98.4205
225.4885 | 353.5931 | 99.6445 | 98.582
226.4885 | 353.61 | 99.8949 | 98.7547
227.4885 | 353.5626 | 100.1455 | 98.9383
228.4885 | 353.5382 | 100.4045 | 99.1289
229.4885 | 353.5379 | 100.6299 | 99.3161
230.4885 | 353.5911 | 100.8959 | 99.5195
231.4885 | 353.6493 | 101.1666 | 99.719
232.4885 | 353.7119 | 101.4223 | 99.9319
233.4885 | 353.6839 | 101.703 | 100.1466
234.4885 | 353.6951 | 101.9717 | 100.3715
235.4885 | 353.5734 | 102.242 | 100.607
236.4885 | 353.5285 | 102.5089 | 100.8255
237.4885 | 353.5176 | 102.783 | 101.0692
238.4885 | 353.5233 | 103.0507 | 101.2996
239.4885 | 353.5714 | 103.3088 | 101.5472
240.4885 | 353.3621 | 103.602 | 101.7952
241.4885 | 353.242 | 103.867 | 102.047
242.4885 | 353.1403 | 104.1444 | 102.2856
243.4885 | 353.0179 | 104.4268 | 102.5332
244.4885 | 353.1123 | 104.6931 | 102.7866
245.4885 | 353.2268 | 104.9791 | 103.0469
246.4885 | 353.2808 | 105.2567 | 103.2933
247.4885 | 353.1355 | 105.525 | 103.5403
248.4885 | 353.1936 | 105.8159 | 103.7995
249.4885 | 353.2801 | 106.0872 | 104.0645
250.4885 | 353.1692 | 106.368 | 104.3097
251.4885 | 353.0295 | 106.6427 | 104.5712
252.4885 | 353.0092 | 106.9356 | 104.8228
253.4885 | 353.1886 | 107.2137 | 105.084
254.4885 | 353.195 | 107.4905 | 105.3492
255.4885 | 353.048 | 107.7732 | 105.6168
256.4885 | 353.0225 | 108.0529 | 105.8709
257.4885 | 353.0028 | 108.329 | 106.1229
258.4885 | 353.044 | 108.589 | 106.388
259.4885 | 353.0528 | 108.8918 | 106.6515
260.4885 | 353.0337 | 109.1647 | 106.919
261.4885 | 353.014 | 109.4521 | 107.1762
262.4885 | 352.8628 | 109.7354 | 107.446
263.5198 | 352.7333 | 110.0222 | 107.7204
264.4885 | 352.803 | 110.2855 | 107.9832
265.4885 | 352.786 | 110.5692 | 108.2402
266.4885 | 352.916 | 110.8551 | 108.5051
267.4885 | 352.9634 | 111.1459 | 108.7798
268.4885 | 353.008 | 111.4163 | 109.0316
269.4885 | 352.9992 | 111.7009 | 109.3141
270.4885 | 353.0728 | 111.9688 | 109.5799
271.4885 | 353.0345 | 112.2576 | 109.8449
272.4885 | 352.9181 | 112.5348 | 110.1122
273.4885 | 352.9809 | 112.8105 | 110.398
274.4885 | 352.922 | 113.0948 | 110.6618
275.4885 | 352.8578 | 113.3725 | 110.9317
276.4885 | 352.8528 | 113.643 | 111.2096
277.4885 | 352.8934 | 113.9289 | 111.4846
278.4885 | 353.0367 | 114.2111 | 111.757
279.4885 | 352.7918 | 114.4907 | 112.0333
280.4885 | 352.7993 | 114.7579 | 112.3031
281.4885 | 352.8769 | 115.039 | 112.5798
282.4885 | 352.8567 | 115.3144 | 112.8541
283.4885 | 352.8393 | 115.5974 | 113.1308
284.4885 | 352.7417 | 115.8687 | 113.4068
285.4885 | 352.7671 | 116.1545 | 113.6818
286.4885 | 352.6932 | 116.4176 | 113.9535
287.4885 | 352.7532 | 116.6929 | 114.2189
288.4885 | 352.7437 | 116.9801 | 114.5092
289.4885 | 352.6959 | 117.24 | 114.7513
290.551 | 352.5443 | 117.5259 | 115.0288
291.4885 | 352.5449 | 117.7885 | 115.2803
292.4885 | 352.5599 | 118.0551 | 115.5489
293.4885 | 352.5166 | 118.3307 | 115.8194
294.4885 | 352.6483 | 118.5977 | 116.0928
295.4885 | 352.7332 | 118.8487 | 116.3816
296.4885 | 352.6603 | 119.0982 | 116.6454
297.4885 | 352.4764 | 119.3504 | 116.9388
298.4885 | 352.4939 | 119.5969 | 117.2138
299.4885 | 352.5667 | 119.8541 | 117.5072
300.4885 | 352.6187 | 120.1329 | 117.7897
301.4885 | 352.5272 | 120.4031 | 118.0803
302.4885 | 352.375 | 120.6586 | 118.3619
303.4885 | 352.3964 | 120.948 | 118.666
304.4885 | 352.5299 | 121.2321 | 118.9499
305.4885 | 352.6073 | 121.5062 | 119.2419
306.4885 | 352.6313 | 121.7787 | 119.5441
307.4885 | 352.5808 | 122.0636 | 119.8431
308.4885 | 352.6153 | 122.357 | 120.1471
309.4885 | 352.5482 | 122.6497 | 120.4621
310.4885 | 352.6359 | 122.9376 | 120.7644
311.4885 | 352.745 | 123.2349 | 121.0674
312.4885 | 352.7464 | 123.5151 | 121.3791
313.4885 | 352.6137 | 123.8169 | 121.6877
314.4885 | 352.4693 | 124.1142 | 121.9959
315.5198 | 352.3906 | 124.4256 | 122.3261
316.4885 | 352.3277 | 124.7195 | 122.6215
317.4885 | 352.3841 | 125.0064 | 122.9473
318.4885 | 352.4503 | 125.3035 | 123.2842
319.4885 | 352.3955 | 125.6113 | 123.6009
320.4885 | 352.343 | 125.9125 | 123.9235
321.4885 | 352.137 | 126.2192 | 124.2402
322.4885 | 352.077 | 126.5124 | 124.5651
323.4885 | 352.1419 | 126.8161 | 124.886
324.4885 | 352.097 | 127.1283 | 125.2253
325.4885 | 352.0206 | 127.4431 | 125.547
326.4885 | 351.9527 | 127.7412 | 125.8884
327.4885 | 351.9869 | 128.0517 | 126.2224
328.4885 | 351.978 | 128.357 | 126.5543
329.4885 | 351.9105 | 128.6662 | 126.9
330.4885 | 351.9746 | 128.9877 | 127.2367
331.4885 | 351.9776 | 129.2961 | 127.571
332.4885 | 351.9047 | 129.5953 | 127.9082
333.4885 | 352.007 | 129.9147 | 128.2488
334.4885 | 351.9846 | 130.2371 | 128.5755
335.4885 | 352.1059 | 130.5503 | 128.9359
336.4885 | 352.008 | 130.8631 | 129.2736
337.4885 | 351.8978 | 131.2037 | 129.6187
338.4885 | 351.7597 | 131.5082 | 129.9794
339.4885 | 351.8116 | 131.8272 | 130.3301
340.4885 | 351.9453 | 132.1447 | 130.6713
341.4885 | 351.7913 | 132.475 | 131.0229
342.4885 | 351.7015 | 132.7724 | 131.3591
343.4885 | 351.7763 | 133.1098 | 131.7228
344.4885 | 351.7648 | 133.4261 | 132.0586
345.4885 | 351.5771 | 133.7548 | 132.4048
346.4885 | 351.5495 | 134.0739 | 132.7549
347.4885 | 351.466 | 134.4039 | 133.1082
348.4885 | 351.5345 | 134.7227 | 133.4557
349.4885 | 351.6552 | 135.044 | 133.8165
350.4885 | 351.4033 | 135.3671 | 134.1697
351.4885 | 351.3796 | 135.6777 | 134.5229
352.4885 | 351.3556 | 136.0097 | 134.8642
353.4885 | 351.3207 | 136.334 | 135.2201
354.4885 | 351.3744 | 136.651 | 135.5574
355.4885 | 351.5616 | 136.9772 | 135.9157
356.4885 | 351.6427 | 137.2938 | 136.266
357.4885 | 351.5562 | 137.6094 | 136.6269
358.4885 | 351.6234 | 137.9495 | 136.9737
359.4885 | 351.739 | 138.2683 | 137.3185
360.4885 | 351.7638 | 138.5803 | 137.6822
361.4885 | 351.7748 | 138.9246 | 138.022
362.4885 | 351.8056 | 139.2415 | 138.3785
363.4885 | 351.8202 | 139.5679 | 138.7251
364.4885 | 351.8887 | 139.9039 | 139.0749
365.4885 | 351.8505 | 140.2046 | 139.4335
366.4885 | 351.8081 | 140.5447 | 139.7874
367.4885 | 351.8484 | 140.8613 | 140.131
368.4885 | 351.9963 | 141.2053 | 140.4853
369.4885 | 352.1057 | 141.5293 | 140.8342
370.4885 | 352.1806 | 141.8749 | 141.1821
371.4885 | 352.2415 | 142.1942 | 141.5226
372.4885 | 352.2078 | 142.5235 | 141.8723
373.4885 | 352.0563 | 142.8429 | 142.2238
374.4885 | 352.0095 | 143.1705 | 142.5634
375.4885 | 352.0889 | 143.5006 | 142.9078
376.4885 | 352.0348 | 143.8404 | 143.2502
377.4885 | 352.0472 | 144.1568 | 143.5877
378.4885 | 351.9435 | 144.4866 | 143.9458
379.4885 | 351.9795 | 144.7954 | 144.2808
380.4885 | 352.0325 | 145.1239 | 144.6242
381.4885 | 351.8899 | 145.4449 | 144.9607
382.4885 | 351.7521 | 145.7575 | 145.3007
383.4885 | 351.755 | 146.0762 | 145.6227
384.4885 | 351.7107 | 146.3972 | 145.9706
385.4885 | 351.7584 | 146.721 | 146.2984
386.4885 | 351.7597 | 147.0369 | 146.6401
387.4885 | 351.8501 | 147.353 | 146.9763
388.4885 | 351.9043 | 147.6834 | 147.3071
389.4885 | 351.9064 | 148.0036 | 147.6372
390.4885 | 351.8866 | 148.3183 | 147.9741
391.4885 | 351.9762 | 148.6325 | 148.3056
392.4885 | 351.9909 | 148.9398 | 148.6294
393.4885 | 351.9529 | 149.2767 | 148.9704
394.4885 | 351.9488 | 149.5897 | 149.2874
395.4885 | 351.893 | 149.8947 | 149.6284
396.5042 | 351.8818 | 150.2146 | 149.9625
397.4885 | 351.8529 | 150.528 | 150.2849
398.4885 | 351.893 | 150.8359 | 150.6089
399.4885 | 351.7638 | 151.1357 | 150.919
400.4885 | 351.6065 | 151.4562 | 151.2427
401.4885 | 351.5223 | 151.7589 | 151.5639
402.4885 | 351.3901 | 152.0614 | 151.8885
403.4885 | 351.3913 | 152.356 | 152.2049
404.4885 | 351.3973 | 152.6848 | 152.5173
405.4885 | 351.3276 | 152.9832 | 152.8414
406.4885 | 351.3872 | 153.288 | 153.1533
407.4885 | 351.354 | 153.5898 | 153.4753
408.4885 | 351.2442 | 153.8837 | 153.7859
409.4885 | 351.2965 | 154.1754 | 154.1046
410.4885 | 351.1124 | 154.4849 | 154.4143
411.4885 | 350.9349 | 154.765 | 154.7241
412.4885 | 350.9529 | 155.0515 | 155.033
413.4885 | 350.8084 | 155.3586 | 155.3415
414.4885 | 350.7954 | 155.6551 | 155.6552
415.4885 | 350.7738 | 155.8984 | 155.9574
416.4885 | 350.8881 | 156.1907 | 156.2576
417.4885 | 351.0164 | 156.4722 | 156.5578
418.4885 | 350.9347 | 156.7289 | 156.8696
419.4885 | 350.9856 | 156.9804 | 157.1734
420.4885 | 350.9079 | 157.2746 | 157.4753
421.4885 | 350.8607 | 157.5569 | 157.7752
422.4885 | 350.9543 | 157.8459 | 158.0869
423.4885 | 350.9604 | 158.1527 | 158.3853
424.4885 | 351.0583 | 158.4593 | 158.6894
425.4885 | 351.1418 | 158.7492 | 158.9861
426.4885 | 351.1725 | 159.06 | 159.2935
427.4885 | 351.2797 | 159.3633 | 159.5812
428.4885 | 351.336 | 159.6616 | 159.886
429.4885 | 351.1981 | 159.9594 | 160.1803
430.4885 | 351.1235 | 160.2639 | 160.4725
431.4885 | 351.119 | 160.5605 | 160.7669
432.4885 | 351.0456 | 160.8617 | 161.0675
433.4885 | 350.9383 | 161.1547 | 161.352
434.4885 | 350.9478 | 161.4671 | 161.6425
435.4885 | 350.9351 | 161.7608 | 161.9284
436.4885 | 350.9531 | 162.0595 | 162.2247
437.4885 | 351.0573 | 162.3625 | 162.5164
438.4885 | 351.142 | 162.6287 | 162.8033
439.4885 | 351.2341 | 162.9393 | 163.0967
440.4885 | 351.103 | 163.2431 | 163.3795
441.4885 | 350.9769 | 163.5373 | 163.6556
442.4885 | 350.8573 | 163.8372 | 163.9442
443.4885 | 350.8361 | 164.1268 | 164.2359
444.4885 | 350.8498 | 164.4216 | 164.5242
445.4885 | 350.9453 | 164.7261 | 164.7977
446.4885 | 351.0098 | 165.0119 | 165.0819
447.4885 | 351.0658 | 165.2951 | 165.3599
448.4885 | 351.0144 | 165.5963 | 165.6394
449.4885 | 351.0499 | 165.8803 | 165.9182
450.4885 | 351.2028 | 166.1594 | 166.2048
451.4885 | 351.2475 | 166.4628 | 166.4858
452.4885 | 351.2585 | 166.7519 | 166.7583
453.4885 | 351.277 | 167.0359 | 167.0358
454.4885 | 351.3104 | 167.3076 | 167.3159
455.4885 | 351.4416 | 167.596 | 167.5948
456.4885 | 351.2786 | 167.8968 | 167.8699
457.4885 | 351.1515 | 168.1866 | 168.1433
458.4885 | 351.1582 | 168.4574 | 168.42
459.4885 | 351.0949 | 168.7667 | 168.7007
460.4885 | 351.193 | 169.0386 | 168.9779
461.4885 | 351.3272 | 169.3178 | 169.2622
462.4885 | 351.376 | 169.6098 | 169.5184
463.4885 | 351.3154 | 169.8762 | 169.7886
464.4885 | 351.2848 | 170.1513 | 170.0662
465.4885 | 351.3646 | 170.4127 | 170.339
466.4885 | 351.4174 | 170.669 | 170.6033
467.4885 | 351.2835 | 170.9425 | 170.8765
468.4885 | 351.3007 | 171.2124 | 171.1381
469.4885 | 351.3737 | 171.5048 | 171.4112
470.4885 | 351.4516 | 171.7613 | 171.6713
471.4885 | 351.4083 | 172.0292 | 171.9315
472.4885 | 351.4419 | 172.301 | 172.2048
473.4885 | 351.3673 | 172.5602 | 172.4727
474.4885 | 351.3338 | 172.8473 | 172.7375
475.4885 | 351.2196 | 173.1184 | 172.998
476.4885 | 351.1459 | 173.3742 | 173.2636
477.4885 | 351.0558 | 173.6387 | 173.5291
478.4885 | 351.0792 | 173.899 | 173.7732
479.4885 | 350.9797 | 174.1776 | 174.0367
480.4885 | 351.0618 | 174.4345 | 174.2992
481.4885 | 351.1514 | 174.6879 | 174.5636
482.4885 | 351.0487 | 174.9559 | 174.8179
483.4885 | 351.0738 | 175.2137 | 175.0799
484.4885 | 351.0881 | 175.4762 | 175.329
485.4885 | 351.0423 | 175.7356 | 175.5838
486.4885 | 351.0133 | 175.9902 | 175.8454
487.4885 | 350.8662 | 176.2581 | 176.1007
488.4885 | 350.6743 | 176.5198 | 176.3467
489.551 | 350.6202 | 176.7841 | 176.6277
490.4885 | 350.7323 | 177.0188 | 176.8603
491.4885 | 350.7337 | 177.2742 | 177.1185
492.4885 | 350.7596 | 177.5338 | 177.3651
493.4885 | 350.7872 | 177.7818 | 177.6226
494.4885 | 350.7984 | 178.0502 | 177.8636
495.4885 | 350.8114 | 178.2868 | 178.1127
496.4885 | 350.7966 | 178.5445 | 178.366
497.4885 | 350.786 | 178.7935 | 178.6227
498.4885 | 350.8257 | 179.0636 | 178.8715
499.4885 | 350.9149 | 179.3035 | 179.1196
500.4885 | 351.0269 | 179.5652 | 179.3557
501.4885 | 351.254 | 179.806 | 179.6117
502.4885 | 351.1651 | 180.0507 | 179.8539
503.4885 | 351.1396 | 180.2956 | 180.0908
504.4885 | 351.164 | 180.5489 | 180.3403
505.4885 | 351.0939 | 180.781 | 180.5756
506.4885 | 351.2306 | 181.0217 | 180.8255
507.4885 | 351.2387 | 181.2808 | 181.0656
508.4885 | 351.2882 | 181.5173 | 181.3114
509.4885 | 351.3451 | 181.7523 | 181.5502
510.4885 | 351.2629 | 181.9894 | 181.7873
511.4885 | 351.329 | 182.2302 | 182.0196
512.4885 | 351.4527 | 182.4739 | 182.2559
513.4885 | 351.4776 | 182.7134 | 182.5071
514.4885 | 351.5242 | 182.9492 | 182.7386
515.4885 | 351.4964 | 183.1748 | 182.9809
516.4885 | 351.4614 | 183.4136 | 183.2155
517.4885 | 351.4989 | 183.6557 | 183.4443
518.4885 | 351.6215 | 183.8806 | 183.6913
519.4885 | 351.6925 | 184.1303 | 183.9165
520.4885 | 351.6742 | 184.3505 | 184.1577
521.4885 | 351.5653 | 184.5892 | 184.3867
522.4885 | 351.5968 | 184.8278 | 184.626
523.4885 | 351.7227 | 185.05 | 184.8529
524.4885 | 351.7364 | 185.2961 | 185.0939
525.4885 | 351.7722 | 185.5244 | 185.3184
526.4885 | 351.8684 | 185.7467 | 185.5527
527.4885 | 351.852 | 185.9686 | 185.7786
528.4885 | 351.871 | 186.2186 | 186.016
529.4885 | 351.8892 | 186.4437 | 186.2411
530.4885 | 351.8953 | 186.6676 | 186.476
531.4885 | 351.8523 | 186.8808 | 186.6931
532.4885 | 351.7152 | 187.1247 | 186.9339
533.4885 | 351.6589 | 187.3467 | 187.1616
534.4885 | 351.6933 | 187.5811 | 187.3822
535.4885 | 351.6637 | 187.8052 | 187.6127
536.4885 | 351.5616 | 188.0377 | 187.8336
537.4885 | 351.4951 | 188.2752 | 188.0706
538.4885 | 351.4986 | 188.5068 | 188.3035
539.4885 | 351.4589 | 188.7196 | 188.534
540.4885 | 351.5595 | 188.9594 | 188.7487
541.4885 | 351.5767 | 189.1861 | 188.9778
542.4885 | 351.6411 | 189.409 | 189.2029
543.4885 | 351.5739 | 189.626 | 189.4209
544.4885 | 351.5483 | 189.8479 | 189.644
545.4885 | 351.5004 | 190.0746 | 189.869
546.4885 | 351.4401 | 190.3066 | 190.0862
547.4885 | 351.4808 | 190.5154 | 190.3206
548.4885 | 351.5703 | 190.7602 | 190.5249
549.4885 | 351.6041 | 190.9707 | 190.75
550.4885 | 351.5804 | 191.1931 | 190.9607
551.4885 | 351.4166 | 191.4102 | 191.1806
552.4885 | 351.249 | 191.6464 | 191.4033
553.4885 | 351.1245 | 191.8635 | 191.6089
554.4885 | 351.0268 | 192.0828 | 191.843
555.4885 | 351.0665 | 192.3172 | 192.0529
556.4885 | 351.1329 | 192.5312 | 192.2788
557.4885 | 351.1869 | 192.7401 | 192.489
558.4885 | 351.2276 | 192.9595 | 192.694
559.4885 | 351.1387 | 193.1853 | 192.9194
560.4885 | 351.1292 | 193.4032 | 193.1237
561.4885 | 351.0958 | 193.6196 | 193.3357
562.4885 | 351.1117 | 193.8361 | 193.5538
563.4885 | 351.187 | 194.059 | 193.762
564.4885 | 351.2617 | 194.2673 | 193.9729
565.4885 | 351.3041 | 194.4893 | 194.1778
566.4885 | 351.3823 | 194.7038 | 194.3904
567.4885 | 351.3449 | 194.9154 | 194.6108
568.4885 | 351.2057 | 195.1336 | 194.8133
569.4885 | 351.1172 | 195.3419 | 195.0164
570.4885 | 351.0348 | 195.5553 | 195.2105
571.4885 | 350.9277 | 195.7745 | 195.428
572.4885 | 350.8383 | 195.9931 | 195.6297
573.4885 | 350.8991 | 196.1943 | 195.8302
756.4885 | 350.701 | 226.953 | 224.8315
757.1177 | 350.7225 | 227.03 | 224.9019
757.852 | 350.7451 | 227.1052 | 224.9839
758.852 | 350.8402 | 227.2421 | 225.1019
759.852 | 350.9072 | 227.387 | 225.2226
760.852 | 350.7922 | 227.5161 | 225.3455
761.852 | 350.6882 | 227.6389 | 225.465
762.852 | 350.6639 | 227.7503 | 225.5738
763.852 | 350.6646 | 227.8787 | 225.6908
764.852 | 350.5909 | 228.0088 | 225.8051
765.852 | 350.544 | 228.1316 | 225.9253
766.852 | 350.661 | 228.2534 | 226.0385
767.852 | 350.7623 | 228.3842 | 226.1546
768.852 | 350.683 | 228.495 | 226.2735
769.852 | 350.5901 | 228.607 | 226.3879
770.852 | 350.6341 | 228.7376 | 226.5033
771.852 | 350.6863 | 228.8563 | 226.6173
772.852 | 350.7458 | 228.9692 | 226.7334
773.852 | 350.7671 | 229.0947 | 226.8392
774.852 | 350.8999 | 229.2167 | 226.9589
775.852 | 350.8435 | 229.3454 | 227.0804
776.852 | 350.7022 | 229.4547 | 227.1836
777.852 | 350.6494 | 229.5765 | 227.3112
778.8833 | 350.4844 | 229.6837 | 227.4207
779.852 | 350.4427 | 229.8001 | 227.5391
780.852 | 350.4598 | 229.9141 | 227.6425
781.852 | 350.4882 | 230.026 | 227.7542
782.852 | 350.4047 | 230.1426 | 227.8688
783.852 | 350.3523 | 230.2461 | 227.9814
784.852 | 350.1463 | 230.3683 | 228.0993
785.852 | 350.1474 | 230.475 | 228.2077
786.852 | 350.0565 | 230.5999 | 228.3131
787.852 | 350.046 | 230.6985 | 228.4208
788.852 | 349.9892 | 230.8089 | 228.5291
789.852 | 350.0236 | 230.9186 | 228.6422
790.852 | 350.0322 | 231.0316 | 228.7476
791.852 | 350.0382 | 231.1472 | 228.8536
792.852 | 350.08 | 231.2429 | 228.9558
793.852 | 350.1926 | 231.3748 | 229.0696
794.852 | 350.1795 | 231.4919 | 229.1726
795.852 | 350.122 | 231.604 | 229.2916
796.852 | 350.1294 | 231.7131 | 229.3976
797.852 | 350.1198 | 231.8357 | 229.5048
798.852 | 350.0437 | 231.9398 | 229.6091
799.852 | 350.1295 | 232.0496 | 229.7126
800.852 | 350.2058 | 232.1466 | 229.8243
801.852 | 350.2376 | 232.2699 | 229.9218
802.852 | 350.1752 | 232.3845 | 230.0296
803.852 | 350.0915 | 232.4901 | 230.1342
804.852 | 350.0594 | 232.5887 | 230.2425
805.852 | 350.1131 | 232.7067 | 230.3464
806.852 | 350.1382 | 232.8104 | 230.4503
807.852 | 350.1285 | 232.922 | 230.5589
808.852 | 350.0569 | 233.0194 | 230.6573
809.852 | 350.0054 | 233.1156 | 230.7549
810.852 | 350.0501 | 233.2241 | 230.8618
811.852 | 350.0647 | 233.3337 | 230.9639
812.852 | 350.1108 | 233.4489 | 231.0591
813.852 | 350.1449 | 233.55 | 231.17
814.852 | 350.0618 | 233.6579 | 231.2658
815.852 | 350.1257 | 233.7537 | 231.368
816.852 | 350.1052 | 233.8613 | 231.4783
817.852 | 350.1311 | 233.9573 | 231.5652
818.852 | 350.1877 | 234.0818 | 231.6693
819.852 | 350.1519 | 234.1874 | 231.7662
820.852 | 350.1804 | 234.2833 | 231.8657
821.852 | 350.2555 | 234.3808 | 231.9726
822.852 | 350.3102 | 234.4898 | 232.0701
823.852 | 350.3297 | 234.5919 | 232.1681
824.852 | 350.3389 | 234.694 | 232.2748
825.852 | 350.3115 | 234.7836 | 232.3725
826.852 | 350.42 | 234.8965 | 232.4626
827.852 | 350.522 | 235.0017 | 232.563
828.852 | 350.6026 | 235.1057 | 232.6581
829.852 | 350.6376 | 235.2072 | 232.7583
830.852 | 350.6237 | 235.3024 | 232.8472
831.852 | 350.5971 | 235.4107 | 232.9468
832.852 | 350.5184 | 235.5107 | 233.0484
833.852 | 350.5006 | 235.6154 | 233.1343
834.852 | 350.3923 | 235.7124 | 233.2321
835.852 | 350.4382 | 235.8073 | 233.3366
836.852 | 350.3413 | 235.9109 | 233.4365
837.852 | 350.4739 | 236.0033 | 233.5219
838.852 | 350.5771 | 236.0903 | 233.6201
839.852 | 350.5799 | 236.2086 | 233.7119
840.852 | 350.6281 | 236.3001 | 233.8006
841.852 | 350.6631 | 236.392 | 233.8964
842.852 | 350.6451 | 236.5057 | 233.9852
843.852 | 350.5667 | 236.5994 | 234.0849
844.852 | 350.608 | 236.7066 | 234.1776
845.852 | 350.6088 | 236.7971 | 234.2703
846.852 | 350.6939 | 236.8958 | 234.3674
847.852 | 350.788 | 236.9961 | 234.4608
848.852 | 350.7668 | 237.0932 | 234.5539
849.852 | 350.8281 | 237.1825 | 234.6422
850.852 | 350.6909 | 237.2984 | 234.7403
851.852 | 350.5271 | 237.3868 | 234.8299
852.852 | 350.5211 | 237.4931 | 234.9288
853.852 | 350.502 | 237.5698 | 235.0106
854.852 | 350.5626 | 237.6811 | 235.1074
855.852 | 350.4575 | 237.7608 | 235.2025
856.852 | 350.4924 | 237.8574 | 235.2879
857.852 | 350.4094 | 237.9599 | 235.3801
858.852 | 350.4315 | 238.0465 | 235.4642
859.852 | 350.4158 | 238.1437 | 235.5643
860.852 | 350.5338 | 238.2315 | 235.6513
861.852 | 350.4877 | 238.341 | 235.7422
862.852 | 350.4559 | 238.4327 | 235.8286
863.852 | 350.297 | 238.5247 | 235.9283
864.852 | 350.3655 | 238.6263 | 236.0154
865.852 | 350.4292 | 238.7035 | 236.1116
866.852 | 350.3327 | 238.7914 | 236.1899
867.852 | 350.2097 | 238.8978 | 236.2898
868.852 | 350.0577 | 238.9867 | 236.3689
869.852 | 350.0434 | 239.0848 | 236.455
870.852 | 349.9539 | 239.1766 | 236.5477
871.852 | 349.942 | 239.2734 | 236.6434
872.852 | 349.9799 | 239.3558 | 236.7297
873.852 | 349.9463 | 239.4397 | 236.8217
874.852 | 349.9208 | 239.5386 | 236.9134
875.852 | 350.0433 | 239.6299 | 236.9944
876.852 | 350.0805 | 239.7273 | 237.0882
877.852 | 350.0712 | 239.8045 | 237.1773
878.852 | 350.0595 | 239.8772 | 237.2694
879.852 | 350.016 | 239.9935 | 237.353
880.852 | 350.0334 | 240.084 | 237.4404
881.852 | 350.0525 | 240.1731 | 237.5263
882.852 | 350.1002 | 240.2549 | 237.6147
883.852 | 350.1222 | 240.3372 | 237.6995
884.852 | 350.0711 | 240.4142 | 237.7905
885.852 | 350.0131 | 240.5152 | 237.8805
886.852 | 349.9944 | 240.5908 | 237.9685
887.852 | 349.8128 | 240.6867 | 238.0458
888.852 | 349.7701 | 240.7623 | 238.1336
889.852 | 349.7575 | 240.8601 | 238.2127
890.852 | 349.8066 | 240.9554 | 238.2975
891.852 | 349.9179 | 241.0191 | 238.3784
892.852 | 349.977 | 241.1159 | 238.4733
893.852 | 350.0557 | 241.1903 | 238.5541
894.852 | 350.0089 | 241.2785 | 238.6424
895.852 | 350.203 | 241.3578 | 238.7168
896.852 | 350.2467 | 241.4483 | 238.8044
897.852 | 350.2697 | 241.5266 | 238.8841
898.852 | 350.2202 | 241.6151 | 238.9674
899.852 | 350.1854 | 241.6983 | 239.0577
900.852 | 350.2185 | 241.7806 | 239.1279
901.852 | 350.2152 | 241.8599 | 239.2166
902.852 | 350.2792 | 241.9337 | 239.3003
903.852 | 350.4202 | 242.0278 | 239.3806
904.852 | 350.4372 | 242.1163 | 239.4611
905.852 | 350.5259 | 242.1898 | 239.5512
906.852 | 350.4905 | 242.2667 | 239.6278
907.852 | 350.5102 | 242.3556 | 239.7061
908.852 | 350.4988 | 242.4424 | 239.7888
909.852 | 350.5876 | 242.5165 | 239.8701
910.852 | 350.6153 | 242.5912 | 239.9533
911.852 | 350.6646 | 242.6791 | 240.0336
912.852 | 350.7995 | 242.7527 | 240.1115
913.852 | 350.9419 | 242.841 | 240.1809
914.852 | 351.0767 | 242.8975 | 240.2687
915.852 | 351.0744 | 242.981 | 240.341
916.852 | 350.9518 | 243.0684 | 240.424
917.852 | 351.033 | 243.1602 | 240.5056
918.852 | 350.9989 | 243.2298 | 240.5828
919.852 | 350.9928 | 243.3227 | 240.6676
920.852 | 351.0604 | 243.4056 | 240.7461
921.852 | 351.0871 | 243.4791 | 240.8246
922.852 | 351.0282 | 243.5428 | 240.8931
923.852 | 350.972 | 243.6398 | 240.9707
924.852 | 350.8739 | 243.7113 | 241.052
925.852 | 350.868 | 243.7847 | 241.1275
926.852 | 350.7137 | 243.8481 | 241.203
927.852 | 350.5326 | 243.9479 | 241.2806
928.852 | 350.3606 | 244.0266 | 241.3472
929.852 | 350.3178 | 244.0987 | 241.4395
930.852 | 350.2705 | 244.1779 | 241.5108
931.852 | 350.1759 | 244.2406 | 241.5936
932.852 | 350.1509 | 244.316 | 241.6576
933.852 | 350.0466 | 244.4083 | 241.7302
934.852 | 350.0112 | 244.4807 | 241.8091
935.852 | 349.9439 | 244.5582 | 241.884
936.852 | 349.8242 | 244.6293 | 241.9563
937.852 | 349.7405 | 244.7106 | 242.0311
938.852 | 349.8229 | 244.7709 | 242.0984
939.852 | 349.8687 | 244.8414 | 242.1852
940.852 | 349.9474 | 244.9128 | 242.2568
941.852 | 349.974 | 244.9981 | 242.3306
942.852 | 349.9644 | 245.0745 | 242.4048
943.852 | 349.8127 | 245.1413 | 242.4743
944.852 | 349.8386 | 245.2086 | 242.5553
945.852 | 349.8641 | 245.2957 | 242.6279
946.852 | 349.7646 | 245.3597 | 242.7007
947.852 | 349.6596 | 245.432 | 242.7712
948.852 | 349.6713 | 245.4987 | 242.8479
949.852 | 349.7504 | 245.5697 | 242.917
950.852 | 349.8094 | 245.649 | 242.991
951.852 | 349.8385 | 245.7261 | 243.0632
952.852 | 349.8237 | 245.7944 | 243.1319
953.852 | 349.9387 | 245.8638 | 243.2064
954.852 | 349.9709 | 245.9327 | 243.28
955.852 | 349.9551 | 246.0011 | 243.3466
956.852 | 349.9172 | 246.0714 | 243.422
957.852 | 349.9579 | 246.1369 | 243.4835
958.852 | 349.9846 | 246.2098 | 243.5572
959.852 | 350.1472 | 246.2764 | 243.6321
960.852 | 350.1626 | 246.344 | 243.7022
961.852 | 350.1871 | 246.4017 | 243.7609
962.852 | 350.2415 | 246.4843 | 243.8335
963.852 | 350.2423 | 246.5478 | 243.9083
964.852 | 350.267 | 246.6222 | 243.9781
965.852 | 350.2762 | 246.6696 | 244.0487
966.852 | 350.2876 | 246.7608 | 244.1211
967.852 | 350.3459 | 246.8274 | 244.1861
968.852 | 350.4141 | 246.8767 | 244.2513
969.852 | 350.5621 | 246.9489 | 244.3259
970.852 | 350.5501 | 247.0154 | 244.388
971.852 | 350.581 | 247.0869 | 244.4592
972.852 | 350.4165 | 247.1407 | 244.5311
973.852 | 350.3526 | 247.204 | 244.5922
974.852 | 350.3757 | 247.2701 | 244.6559
975.852 | 350.4007 | 247.3485 | 244.7217
976.852 | 350.4263 | 247.4002 | 244.7836
977.852 | 350.4459 | 247.4595 | 244.8449
978.852 | 350.5126 | 247.5307 | 244.9162
979.852 | 350.5398 | 247.5893 | 244.9836
980.852 | 350.5367 | 247.6586 | 245.0461
981.852 | 350.5627 | 247.7079 | 245.1122
982.852 | 350.497 | 247.7818 | 245.181
983.852 | 350.5247 | 247.8362 | 245.2421
984.852 | 350.5068 | 247.9025 | 245.3166
985.852 | 350.5039 | 247.9715 | 245.3725
986.852 | 350.4661 | 248.026 | 245.4368
987.852 | 350.414 | 248.0852 | 245.5005
988.852 | 350.3458 | 248.1433 | 245.5692
989.852 | 350.3123 | 248.1866 | 245.6292
990.852 | 350.3195 | 248.2502 | 245.6913
991.852 | 350.3087 | 248.3197 | 245.7494
992.852 | 350.5375 | 248.3816 | 245.817
993.852 | 350.5673 | 248.4557 | 245.8717
994.852 | 350.5509 | 248.5097 | 245.9408
995.852 | 350.5695 | 248.5697 | 246.0081
996.852 | 350.558 | 248.6361 | 246.0681
997.852 | 350.5996 | 248.6898 | 246.1253
998.852 | 350.499 | 248.7394 | 246.1879
999.852 | 350.4946 | 248.7974 | 246.2445
1000.852 | 350.6346 | 248.8635 | 246.3058
1001.852 | 350.6234 | 248.9124 | 246.3625
1002.852 | 350.5855 | 248.9716 | 246.4305
1003.852 | 350.4043 | 249.0355 | 246.4914
1004.852 | 350.3818 | 249.1008 | 246.5525
1005.852 | 350.4901 | 249.1548 | 246.6143
1006.852 | 350.5202 | 249.2122 | 246.6725
1007.852 | 350.3723 | 249.2661 | 246.7204
1008.852 | 350.2775 | 249.325 | 246.7892
1009.852 | 350.3083 | 249.3725 | 246.8386
1010.852 | 350.2628 | 249.443 | 246.8965
1011.852 | 350.3206 | 249.4835 | 246.9523
1012.852 | 350.3199 | 249.5545 | 247.0103
1013.852 | 350.2431 | 249.6049 | 247.0669
1014.852 | 350.0712 | 249.663 | 247.1236
1015.852 | 349.9766 | 249.7203 | 247.1733
1016.852 | 349.9779 | 249.7635 | 247.229
1017.852 | 350.0665 | 249.8171 | 247.2854
1018.852 | 350.1628 | 249.8798 | 247.3402
1019.852 | 350.1323 | 249.9513 | 247.4019
1020.852 | 350.2399 | 249.9971 | 247.4572
1021.852 | 350.1841 | 250.0504 | 247.5108
1022.852 | 350.1022 | 250.0973 | 247.5685
1023.852 | 350.0945 | 250.1586 | 247.6181
1024.852 | 350.0451 | 250.2166 | 247.6735
1025.852 | 349.959 | 250.2647 | 247.7223
1026.852 | 349.9688 | 250.3095 | 247.7889
1027.852 | 350.0072 | 250.3641 | 247.843
1028.852 | 350.0857 | 250.4182 | 247.8978
1029.852 | 350.1973 | 250.4754 | 247.9508
1030.852 | 350.2007 | 250.5175 | 247.9972
1031.852 | 350.1231 | 250.5655 | 248.0601
1032.852 | 350.0932 | 250.6374 | 248.1114
1033.852 | 349.9808 | 250.6899 | 248.1682
1034.852 | 349.9545 | 250.7269 | 248.2168
1035.852 | 349.9502 | 250.7862 | 248.2634
1036.852 | 349.8809 | 250.847 | 248.3229
1037.852 | 349.8759 | 250.8938 | 248.3714
1038.852 | 349.9817 | 250.9332 | 248.4242
1039.852 | 349.9332 | 250.9973 | 248.4759
1040.852 | 349.8954 | 251.0382 | 248.5392
1041.852 | 349.9385 | 251.0741 | 248.5787
1042.852 | 350.0304 | 251.1271 | 248.6299
1043.852 | 350.1271 | 251.183 | 248.682
1044.852 | 350.146 | 251.2426 | 248.737
1045.852 | 350.1534 | 251.2952 | 248.7854
1046.852 | 350.2042 | 251.3452 | 248.8292
1047.852 | 350.2854 | 251.4017 | 248.8794
1048.852 | 350.2697 | 251.4401 | 248.9364
1049.852 | 350.1597 | 251.4812 | 248.986
1050.852 | 350.0786 | 251.5371 | 249.0288
1051.852 | 350.017 | 251.5788 | 249.0956
1052.852 | 350.0092 | 251.637 | 249.1343
1053.852 | 350.0171 | 251.69 | 249.1777
1054.852 | 350.0635 | 251.7355 | 249.2277
1055.852 | 350.1164 | 251.7665 | 249.2775
1056.852 | 350.1153 | 251.8298 | 249.332
1057.852 | 350.0995 | 251.8695 | 249.3774
1058.852 | 350.1074 | 251.9244 | 249.4328
1059.852 | 350.1353 | 251.9692 | 249.4771
1060.852 | 350.2304 | 252.0069 | 249.5257
1061.852 | 350.211 | 252.0603 | 249.5752
1062.852 | 350.1943 | 252.1052 | 249.6295
1063.852 | 350.238 | 252.1406 | 249.6781
1064.852 | 350.1175 | 252.1743 | 249.724
1065.852 | 350.0118 | 252.2378 | 249.7676
1066.852 | 349.9373 | 252.2937 | 249.8205
1067.852 | 349.9168 | 252.3311 | 249.8749
1068.852 | 349.9058 | 252.3627 | 249.9204
1069.852 | 349.8604 | 252.425 | 249.9736
1070.852 | 349.8941 | 252.4688 | 250.0229
1071.852 | 349.9394 | 252.5133 | 250.0636
1072.852 | 350.002 | 252.5549 | 250.1193
1073.852 | 350.1232 | 252.5949 | 250.1646
1074.852 | 350.1223 | 252.6551 | 250.2134
1075.852 | 350.2591 | 252.6955 | 250.2578
1076.852 | 350.3161 | 252.73 | 250.3116
1077.852 | 350.3114 | 252.7788 | 250.3566
1078.852 | 350.2562 | 252.8154 | 250.4032
1079.852 | 350.2467 | 252.8744 | 250.457
1080.852 | 350.401 | 252.9149 | 250.4945
1081.852 | 350.4916 | 252.9616 | 250.5537
1082.852 | 350.5148 | 252.9986 | 250.6015
1083.852 | 350.5742 | 253.0397 | 250.646
1084.852 | 350.4966 | 253.0821 | 250.6908
1085.852 | 350.4494 | 253.12 | 250.7356
1086.852 | 350.3979 | 253.1711 | 250.7754
1087.852 | 350.4082 | 253.2077 | 250.8243
1088.852 | 350.3671 | 253.2504 | 250.8693
1089.852 | 350.2995 | 253.286 | 250.9228
1090.852 | 350.3566 | 253.3329 | 250.9719
1091.852 | 350.3242 | 253.3748 | 251.0115
1092.852 | 350.2796 | 253.4267 | 251.0578
1093.852 | 350.1104 | 253.4472 | 251.1048
1094.852 | 350.0714 | 253.5093 | 251.1471
1095.852 | 349.9958 | 253.551 | 251.1834
1096.852 | 349.9552 | 253.5836 | 251.2158
1097.852 | 349.9629 | 253.6307 | 251.2773
1098.852 | 349.8935 | 253.6658 | 251.3155
1099.852 | 349.8844 | 253.7003 | 251.3645
1100.852 | 349.9047 | 253.7301 | 251.4054
1101.852 | 349.9735 | 253.7738 | 251.4405
1102.852 | 349.9653 | 253.8143 | 251.4817
1103.852 | 349.8353 | 253.8654 | 251.5216
1104.852 | 349.8796 | 253.9049 | 251.5661
1105.852 | 349.9114 | 253.9438 | 251.598
1106.852 | 349.9652 | 253.988 | 251.6507
1107.852 | 349.9377 | 254.0304 | 251.6959
1108.852 | 349.9356 | 254.0646 | 251.733
1109.852 | 349.9327 | 254.1138 | 251.7768
1110.852 | 349.9073 | 254.1281 | 251.8167
1111.852 | 349.8338 | 254.1891 | 251.8562
1112.852 | 349.8367 | 254.2273 | 251.8982
1113.852 | 349.8772 | 254.2642 | 251.9429
1114.852 | 349.9423 | 254.2903 | 251.9887
1115.852 | 349.939 | 254.3425 | 252.0257
1116.852 | 349.886 | 254.3857 | 252.0662
1117.852 | 349.948 | 254.4215 | 252.109
1118.852 | 350.0371 | 254.4627 | 252.1414
1119.852 | 350.0286 | 254.4808 | 252.1869
1120.852 | 350.1109 | 254.5414 | 252.2347
1121.852 | 350.217 | 254.5767 | 252.267
1122.852 | 350.2314 | 254.6209 | 252.3115
1123.852 | 350.1416 | 254.6559 | 252.3425
1124.852 | 350.0283 | 254.6804 | 252.3944
1125.852 | 349.9267 | 254.7129 | 252.4291
1126.852 | 349.9348 | 254.7619 | 252.4655
1127.852 | 349.8574 | 254.7957 | 252.5129
1128.852 | 349.8239 | 254.8248 | 252.5501
1129.852 | 349.7657 | 254.8678 | 252.5872
1130.852 | 349.7053 | 254.9052 | 252.6289
1131.852 | 349.7388 | 254.9418 | 252.6641
1132.852 | 349.6239 | 254.975 | 252.6929
1133.852 | 349.7449 | 255.0165 | 252.73
1134.852 | 349.715 | 255.0446 | 252.7774
1135.852 | 349.6803 | 255.0677 | 252.8218
1136.852 | 349.5578 | 255.1179 | 252.8566
1137.852 | 349.5337 | 255.1641 | 252.8983
1138.852 | 349.5966 | 255.1821 | 252.9242
1139.852 | 349.6332 | 255.2164 | 252.9636
1140.852 | 349.6866 | 255.2635 | 252.9991
1141.852 | 349.7965 | 255.2943 | 253.0432
1142.852 | 349.7087 | 255.3266 | 253.077
1143.852 | 349.7431 | 255.3423 | 253.1056
1144.852 | 349.7118 | 255.3663 | 253.1429
1145.852 | 349.6121 | 255.4195 | 253.1857
1146.852 | 349.5988 | 255.4436 | 253.2131
1147.852 | 349.6979 | 255.4797 | 253.2536
1148.852 | 349.8092 | 255.5284 | 253.2936
1149.852 | 349.7795 | 255.5545 | 253.3177
1150.852 | 349.7765 | 255.6012 | 253.3596
1151.852 | 349.5865 | 255.6442 | 253.4019
1152.852 | 349.4823 | 255.6723 | 253.4375
1153.852 | 349.5026 | 255.7119 | 253.4638
1154.852 | 349.6106 | 255.7584 | 253.5011
1155.852 | 349.5743 | 255.785 | 253.5418
1156.852 | 349.6232 | 255.8335 | 253.5783
1157.852 | 349.6088 | 255.8484 | 253.6165
1158.852 | 349.6663 | 255.8842 | 253.6471
1159.852 | 349.6837 | 255.9072 | 253.6764
1160.852 | 349.7513 | 255.9476 | 253.7174
1161.852 | 349.7324 | 255.9714 | 253.7526
1162.852 | 349.7448 | 256.0278 | 253.7787
1163.852 | 349.7598 | 256.0641 | 253.8218
1164.852 | 349.7681 | 256.0939 | 253.8563
1165.852 | 349.7575 | 256.1441 | 253.8836
1166.852 | 349.769 | 256.1685 | 253.9197
1167.852 | 349.8507 | 256.2007 | 253.9557
1168.852 | 349.989 | 256.2491 | 253.9981
1169.852 | 350.0096 | 256.2862 | 254.0242
1170.852 | 350.0511 | 256.3232 | 254.0599
1171.852 | 350.018 | 256.358 | 254.0944
1172.852 | 350.0313 | 256.3883 | 254.1202
1173.852 | 349.884 | 256.4305 | 254.1614
1174.852 | 349.7728 | 256.4623 | 254.1904
1175.852 | 349.7243 | 256.5078 | 254.229
1176.852 | 349.7644 | 256.5337 | 254.2474
1177.852 | 349.755 | 256.5639 | 254.291
1178.852 | 349.7951 | 256.5992 | 254.3233
1179.852 | 349.8198 | 256.6359 | 254.3584
1180.852 | 349.8183 | 256.6687 | 254.3893
1181.852 | 349.8107 | 256.6978 | 254.4206
1182.852 | 349.7912 | 256.7388 | 254.4527
1183.852 | 349.8314 | 256.7635 | 254.492
1184.852 | 349.8371 | 256.7974 | 254.5241
1185.852 | 349.8725 | 256.828 | 254.5559
1186.852 | 349.9218 | 256.8788 | 254.591
1187.852 | 349.829 | 256.9053 | 254.6239
1188.852 | 349.7402 | 256.9389 | 254.6596
1189.852 | 349.7394 | 256.9644 | 254.6931
1190.852 | 349.7486 | 256.9958 | 254.7271
1191.852 | 349.6924 | 257.0336 | 254.7634
1192.852 | 349.6912 | 257.0616 | 254.7885
1193.852 | 349.6903 | 257.0916 | 254.8242
1194.852 | 349.6462 | 257.1272 | 254.865
1195.852 | 349.5249 | 257.1523 | 254.8931
1196.852 | 349.423 | 257.1879 | 254.9339
1197.852 | 349.4149 | 257.2197 | 254.977
1198.852 | 349.5206 | 257.2578 | 254.9934
1199.852 | 349.6422 | 257.2869 | 255.0393
1200.852 | 349.6563 | 257.3121 | 255.0733
1201.852 | 349.6542 | 257.3595 | 255.1027
1202.852 | 349.6631 | 257.3932 | 255.1271
1203.852 | 349.7527 | 257.4244 | 255.1641
1204.852 | 349.8008 | 257.4553 | 255.1971
1205.852 | 349.8618 | 257.4862 | 255.2283
1206.852 | 349.951 | 257.5146 | 255.2718
1207.852 | 349.8167 | 257.5529 | 255.2958
1208.852 | 349.8702 | 257.5818 | 255.3324
1209.852 | 349.9039 | 257.6017 | 255.3705
1210.852 | 349.9181 | 257.6359 | 255.3915
1211.852 | 349.9329 | 257.6848 | 255.43
1212.852 | 349.8333 | 257.7038 | 255.4536
1213.852 | 349.8863 | 257.743 | 255.4976
1214.852 | 349.8829 | 257.7767 | 255.5227
1215.852 | 349.9298 | 257.8131 | 255.5615
1216.852 | 349.9546 | 257.8408 | 255.5992
1217.852 | 349.8067 | 257.8733 | 255.6257
1218.852 | 349.7661 | 257.8949 | 255.6578
1219.852 | 349.8639 | 257.93 | 255.6993
1220.852 | 349.8569 | 257.9693 | 255.7262
1221.852 | 349.8222 | 257.9929 | 255.7686
1222.852 | 349.8082 | 258.0195 | 255.8015
1223.852 | 349.7602 | 258.0658 | 255.8318
1224.852 | 349.6206 | 258.0901 | 255.8638
1225.852 | 349.5653 | 258.1297 | 255.8873
1226.852 | 349.6453 | 258.1464 | 255.9246
1227.852 | 349.5998 | 258.1852 | 255.9625
1228.852 | 349.6147 | 258.2226 | 255.9931
1229.852 | 349.6486 | 258.2544 | 256.0299
1230.852 | 349.6099 | 258.2871 | 256.0598
1231.852 | 349.6263 | 258.3104 | 256.1073
1232.852 | 349.6507 | 258.3444 | 256.1405
1233.852 | 349.7995 | 258.3762 | 256.1687
1234.852 | 350.0535 | 258.4035 | 256.1976
1235.852 | 350.0291 | 258.4375 | 256.234
1236.852 | 350.0304 | 258.4693 | 256.2663
1237.852 | 350.0876 | 258.4954 | 256.298
1238.852 | 349.9043 | 258.5259 | 256.3295
1239.852 | 349.8377 | 258.5561 | 256.3501
1240.852 | 349.7812 | 258.5819 | 256.3849
1241.852 | 349.7929 | 258.6173 | 256.4224
1242.852 | 349.8571 | 258.6566 | 256.4454
1243.852 | 349.8915 | 258.6756 | 256.4742
1244.852 | 349.842 | 258.7069 | 256.5123
1245.852 | 349.6691 | 258.7469 | 256.5538
1246.852 | 349.5226 | 258.7666 | 256.5819
1247.852 | 349.6182 | 258.8069 | 256.615
1248.852 | 349.6931 | 258.8439 | 256.653
1249.852 | 349.6533 | 258.8694 | 256.6773
1250.852 | 349.6793 | 258.8914 | 256.7127
1251.852 | 349.7538 | 258.9419 | 256.7436
1252.852 | 349.7383 | 258.9511 | 256.7829
1253.852 | 349.7569 | 258.9916 | 256.8208
1254.852 | 349.7768 | 259.0292 | 256.8506
1255.852 | 349.6697 | 259.0432 | 256.8854
1256.852 | 349.6834 | 259.0822 | 256.9158
1257.852 | 349.8127 | 259.0987 | 256.9492
1258.852 | 349.8159 | 259.1314 | 256.9859
1259.852 | 349.7577 | 259.161 | 257.0188
1260.852 | 349.7045 | 259.1875 | 257.0596
1261.852 | 349.7066 | 259.2018 | 257.0889
1262.852 | 349.8215 | 259.2415 | 257.127
1263.852 | 349.8875 | 259.2844 | 257.1576
1264.852 | 349.9141 | 259.3033 | 257.1919
1265.852 | 349.8725 | 259.3402 | 257.2249
1266.852 | 349.7934 | 259.3595 | 257.2573
1267.852 | 349.8929 | 259.4011 | 257.2877
1268.852 | 349.9673 | 259.4314 | 257.3227
1269.852 | 350.0794 | 259.4589 | 257.3619
1270.852 | 350.147 | 259.4728 | 257.4013
1271.852 | 350.1564 | 259.5067 | 257.4334
1272.852 | 350.1282 | 259.5401 | 257.4676
1273.852 | 350.1477 | 259.5653 | 257.5066
1274.852 | 350.218 | 259.585 | 257.5326
1275.852 | 350.427 | 259.6244 | 257.5729
1276.852 | 350.4902 | 259.6486 | 257.6087
1277.852 | 350.5917 | 259.6831 | 257.6439
1278.852 | 350.5573 | 259.6912 | 257.681
1279.852 | 350.4565 | 259.7362 | 257.7078
1280.852 | 350.3797 | 259.7586 | 257.7468
1281.852 | 350.2477 | 259.7909 | 257.785
1282.852 | 350.186 | 259.8079 | 257.8127
1283.852 | 350.0583 | 259.8388 | 257.8439
1284.852 | 350.0332 | 259.8697 | 257.8896
1285.852 | 350.1149 | 259.8743 | 257.9187
1286.852 | 349.9415 | 259.9168 | 257.9516
1287.852 | 349.9533 | 259.9554 | 257.9919
1288.852 | 349.9505 | 259.9829 | 258.0281
1289.852 | 349.9115 | 259.9955 | 258.0569
1290.852 | 349.8078 | 260.0321 | 258.0888
1291.852 | 349.7624 | 260.0553 | 258.1343
1292.852 | 349.6356 | 260.0846 | 258.1715
1293.852 | 349.5925 | 260.1037 | 258.2049
1294.852 | 349.5889 | 260.1409 | 258.234
1295.852 | 349.5181 | 260.1602 | 258.2619
1296.852 | 349.4251 | 260.1815 | 258.3052
1297.852 | 349.4147 | 260.2023 | 258.3452
1298.852 | 349.4088 | 260.2387 | 258.3764
1299.852 | 349.3841 | 260.2688 | 258.4206
1300.852 | 349.3917 | 260.296 | 258.4486
1301.852 | 349.332 | 260.3202 | 258.4864
1302.852 | 349.3126 | 260.3574 | 258.5205
1303.852 | 349.3558 | 260.3674 | 258.5574
1304.852 | 349.3812 | 260.3969 | 258.5876
1305.852 | 349.4043 | 260.4251 | 258.6217
1306.852 | 349.4511 | 260.4517 | 258.6523
1307.852 | 349.506 | 260.4752 | 258.6914
1308.852 | 349.4072 | 260.4879 | 258.7309
1309.852 | 349.418 | 260.5186 | 258.762
1310.852 | 349.438 | 260.5504 | 258.7887
1311.852 | 349.3967 | 260.5772 | 258.8223
1312.852 | 349.431 | 260.5913 | 258.8544
1313.852 | 349.4544 | 260.6304 | 258.8902
1314.852 | 349.3914 | 260.6581 | 258.9273
1315.852 | 349.4695 | 260.6791 | 258.9561
1316.852 | 349.5078 | 260.7048 | 258.9893
1317.852 | 349.4705 | 260.7291 | 259.0193
1318.852 | 349.5512 | 260.7358 | 259.0586
1319.852 | 349.6878 | 260.7755 | 259.0937
1320.852 | 349.6157 | 260.8048 | 259.1321
1321.852 | 349.6955 | 260.818 | 259.165
1322.852 | 349.7138 | 260.8522 | 259.2111
1323.852 | 349.7326 | 260.8723 | 259.2422
1324.852 | 349.7113 | 260.9011 | 259.268
1325.852 | 349.7415 | 260.9252 | 259.3174
1326.852 | 349.7996 | 260.9477 | 259.3448
1327.852 | 349.8097 | 260.9673 | 259.3903
1328.852 | 349.8202 | 260.9869 | 259.4226
1329.852 | 349.8168 | 261.008 | 259.4488
1330.852 | 349.7839 | 261.0442 | 259.4834
1331.852 | 349.7944 | 261.0649 | 259.5308
1332.852 | 349.9242 | 261.0784 | 259.5585
1333.852 | 349.9622 | 261.1081 | 259.5901
1334.852 | 349.9155 | 261.1234 | 259.6231
1335.852 | 349.965 | 261.1496 | 259.6559
1336.852 | 350.0196 | 261.1868 | 259.6943
1337.852 | 350.0917 | 261.1894 | 259.7235
1338.852 | 350.244 | 261.2258 | 259.7646
1339.852 | 350.3502 | 261.2418 | 259.7989
1340.852 | 350.329 | 261.2619 | 259.8297
1341.852 | 350.4247 | 261.2807 | 259.8614
1342.852 | 350.4046 | 261.3198 | 259.8964
1343.852 | 350.3395 | 261.3413 | 259.9263
1344.852 | 350.2033 | 261.3636 | 259.964
1345.852 | 350.1782 | 261.3756 | 260.0022
1346.852 | 350.0998 | 261.408 | 260.0329
1347.852 | 350.0648 | 261.4409 | 260.0698
1348.852 | 349.9627 | 261.4619 | 260.0892
1349.852 | 349.8301 | 261.4608 | 260.1175
1350.852 | 349.6839 | 261.5062 | 260.1328
1351.852 | 349.6509 | 261.524 | 260.1572
1352.852 | 349.7514 | 261.5395 | 260.1781
1353.852 | 349.6608 | 261.5689 | 260.2023
1354.852 | 349.6028 | 261.5886 | 260.236
1355.852 | 349.5593 | 261.6085 | 260.2717
1356.852 | 349.5357 | 261.6415 | 260.2925
1357.852 | 349.6691 | 261.6603 | 260.3165
1358.852 | 349.699 | 261.6873 | 260.3577
1359.852 | 349.674 | 261.7006 | 260.3874
1360.852 | 349.7569 | 261.7366 | 260.4183
1361.852 | 349.7521 | 261.7583 | 260.4535
1362.852 | 349.6757 | 261.7777 | 260.4767
1363.852 | 349.6602 | 261.8063 | 260.516
1364.852 | 349.7099 | 261.8126 | 260.5417
1365.852 | 349.8451 | 261.8514 | 260.5817
1366.852 | 349.8003 | 261.8812 | 260.6251
1367.852 | 349.8577 | 261.8958 | 260.6537
1368.852 | 349.9172 | 261.9176 | 260.6895
1369.852 | 349.914 | 261.9438 | 260.7232
1370.852 | 349.9922 | 261.9752 | 260.7652
1371.852 | 349.9304 | 261.9978 | 260.795
1372.852 | 349.9676 | 262.0211 | 260.8328
1373.852 | 349.8365 | 262.0458 | 260.8701
1374.852 | 349.7488 | 262.0645 | 260.907
1375.852 | 349.863 | 262.097 | 260.9482
1376.852 | 350.0014 | 262.1122 | 260.9872
1377.852 | 350.0362 | 262.1286 | 261.0216
1378.852 | 350.0072 | 262.1576 | 261.059
1379.852 | 349.8018 | 262.1893 | 261.0989
1380.852 | 349.7551 | 262.1978 | 261.1394
1381.852 | 349.7043 | 262.2462 | 261.1729
1382.852 | 349.7279 | 262.2579 | 261.2122
1383.852 | 349.7592 | 262.2711 | 261.2491
1384.852 | 349.7609 | 262.3045 | 261.2947
1385.852 | 349.772 | 262.3344 | 261.326
1386.852 | 349.6939 | 262.3463 | 261.3685
1387.852 | 349.6649 | 262.3858 | 261.4055
1388.852 | 349.7199 | 262.4064 | 261.452
1389.852 | 349.6349 | 262.4306 | 261.4947
1390.852 | 349.6828 | 262.4563 | 261.532
1391.852 | 349.8186 | 262.4978 | 261.5582
1392.852 | 349.8713 | 262.506 | 261.6018
1393.852 | 349.872 | 262.5476 | 261.6385
1394.852 | 349.7902 | 262.5403 | 261.687
1395.852 | 349.7941 | 262.6019 | 261.7321
1396.852 | 349.7696 | 262.6172 | 261.7663
1397.852 | 349.6419 | 262.6376 | 261.7972
1398.852 | 349.6263 | 262.6723 | 261.8326
1399.852 | 349.7461 | 262.6962 | 261.8694
1400.852 | 349.7248 | 262.7189 | 261.9051
1401.852 | 349.8319 | 262.7462 | 261.9544
1402.852 | 349.9084 | 262.7674 | 261.9882
1403.852 | 349.9919 | 262.7974 | 262.0329
1404.852 | 349.9764 | 262.8366 | 262.071
1405.852 | 349.8904 | 262.8654 | 262.1018
1406.852 | 349.8442 | 262.8802 | 262.147
1407.852 | 349.7777 | 262.9067 | 262.1871
1408.852 | 349.6711 | 262.9287 | 262.2252
1409.852 | 349.76 | 262.9475 | 262.2743
1410.852 | 349.8562 | 262.9755 | 262.3121
1411.852 | 349.7639 | 262.9937 | 262.35
1412.852 | 349.6792 | 263.0056 | 262.3929
1413.852 | 349.6814 | 263.0354 | 262.4304
1414.852 | 349.7496 | 263.0507 | 262.4677
1415.852 | 349.7763 | 263.0847 | 262.5056
1416.852 | 349.7732 | 263.0994 | 262.5439
1417.852 | 349.7644 | 263.1418 | 262.5828
1418.852 | 349.7145 | 263.1582 | 262.6239
1419.852 | 349.7284 | 263.1827 | 262.6645
1420.852 | 349.6356 | 263.1987 | 262.7088
1421.852 | 349.5997 | 263.2272 | 262.7456
1422.852 | 349.6107 | 263.2458 | 262.7822
1423.852 | 349.6508 | 263.267 | 262.8286
1424.852 | 349.6514 | 263.3086 | 262.8521
1425.852 | 349.5929 | 263.3142 | 262.8877
1426.852 | 349.5565 | 263.3384 | 262.9206
1427.852 | 349.4788 | 263.3792 | 262.9391
1428.852 | 349.3312 | 263.3817 | 262.9609
1429.852 | 349.3382 | 263.4143 | 262.9869
1430.852 | 349.395 | 263.4237 | 263.0084
1431.852 | 349.3184 | 263.4548 | 263.0219
1432.852 | 349.3387 | 263.4761 | 263.0476
1433.852 | 349.3206 | 263.4731 | 263.0626
1434.852 | 349.2178 | 263.5181 | 263.0854
1435.852 | 349.1975 | 263.5314 | 263.0959
1436.852 | 349.1846 | 263.554 | 263.1227
1437.852 | 349.0141 | 263.5861 | 263.1458
1438.852 | 348.9142 | 263.6077 | 263.1562
1439.852 | 349.0004 | 263.6313 | 263.1751
1440.852 | 349.0481 | 263.6311 | 263.208
1441.852 | 349.1109 | 263.6766 | 263.2341
1442.852 | 349.1051 | 263.7007 | 263.256
1443.852 | 349.0486 | 263.7297 | 263.2873
1444.852 | 349.1074 | 263.7422 | 263.3156
1445.852 | 348.982 | 263.7833 | 263.3509
1446.852 | 348.9845 | 263.8078 | 263.3892
1447.852 | 349.0539 | 263.8276 | 263.4152
1448.852 | 349.1087 | 263.841 | 263.4534
1449.852 | 349.188 | 263.8693 | 263.4986
1450.852 | 349.1228 | 263.8816 | 263.5376
1451.852 | 349.1116 | 263.9129 | 263.5722
1452.852 | 349.1859 | 263.9426 | 263.6236
1453.852 | 349.2973 | 263.9513 | 263.6649
1454.852 | 349.4116 | 263.9796 | 263.7079
1455.852 | 349.5451 | 264.0091 | 263.7556
1456.852 | 349.5204 | 264.0245 | 263.803
1457.852 | 349.572 | 264.036 | 263.846
1458.852 | 349.5402 | 264.0619 | 263.8787
1459.852 | 349.5928 | 264.0838 | 263.9194
1460.852 | 349.6037 | 264.1092 | 263.9523
1461.852 | 349.507 | 264.1409 | 263.9857
1462.852 | 349.5782 | 264.1558 | 264.0302
1463.852 | 349.7214 | 264.1663 | 264.0586
1464.852 | 349.8221 | 264.1942 | 264.0886
1465.852 | 349.9075 | 264.2141 | 264.098
1466.852 | 349.939 | 264.249 | 264.1299
1467.852 | 349.8438 | 264.2724 | 264.1549
1468.852 | 349.7557 | 264.3039 | 264.1755
1469.852 | 349.6735 | 264.327 | 264.2
1470.852 | 349.7252 | 264.3605 | 264.2181
1471.852 | 349.7968 | 264.3781 | 264.2359
1472.852 | 349.9219 | 264.3982 | 264.2617
1473.852 | 350.0022 | 264.4265 | 264.276
1474.852 | 350.0202 | 264.4533 | 264.2967
1475.852 | 349.938 | 264.4728 | 264.3113
1476.852 | 349.8656 | 264.5079 | 264.3285
1477.852 | 349.8655 | 264.5191 | 264.342
1478.852 | 349.9069 | 264.548 | 264.3566
1479.852 | 349.981 | 264.5726 | 264.372
1480.852 | 349.9209 | 264.5879 | 264.3915
1481.852 | 349.9115 | 264.6065 | 264.4152
1482.852 | 349.9244 | 264.6174 | 264.4293
1483.852 | 349.9225 | 264.633 | 264.4466
1484.852 | 349.8409 | 264.6457 | 264.466
1485.852 | 349.7286 | 264.6541 | 264.4873
1486.852 | 349.7255 | 264.6536 | 264.5108
1487.852 | 349.624 | 264.6704 | 264.5418
1488.852 | 349.6593 | 264.6709 | 264.5558
1489.852 | 349.6502 | 264.6783 | 264.5747
1490.852 | 349.7735 | 264.6709 | 264.6122
1491.852 | 349.7344 | 264.6765 | 264.6423
1492.852 | 349.8642 | 264.674 | 264.6722
1493.852 | 349.8325 | 264.6841 | 264.6949
1494.852 | 349.8004 | 264.6885 | 264.7199
1495.852 | 349.7263 | 264.6837 | 264.7381
1496.852 | 349.6999 | 264.6859 | 264.7678
1497.852 | 349.6839 | 264.6833 | 264.7944
1498.852 | 349.5832 | 264.6889 | 264.8256
1499.852 | 349.4939 | 264.6883 | 264.8434
1500.852 | 349.5018 | 264.6966 | 264.8657
1501.852 | 349.6156 | 264.6992 | 264.8799
1502.852 | 349.7622 | 264.7033 | 264.8978
1503.852 | 349.8219 | 264.6996 | 264.912
1504.852 | 349.9539 | 264.7069 | 264.914
1505.852 | 350.0547 | 264.7146 | 264.9259
1506.852 | 350.0279 | 264.6869 | 264.945
1507.852 | 350.0719 | 264.6941 | 264.9516
1508.852 | 350.0565 | 264.7054 | 264.9651
1509.852 | 350.0442 | 264.7057 | 264.9703
1510.852 | 350.0895 | 264.7053 | 264.9816
1511.852 | 350.0926 | 264.7049 | 264.9914
1512.852 | 350.105 | 264.71 | 265.005
1513.852 | 350.0779 | 264.715 | 265.0085
1514.852 | 350.0984 | 264.7231 | 265.0124
1515.852 | 350.1051 | 264.7287 | 265.0057
1516.852 | 350.0408 | 264.7166 | 265.0106
1517.852 | 349.9546 | 264.7037 | 265.0182
1518.852 | 349.9116 | 264.7866 | 265.0193
1519.852 | 349.9186 | 264.7109 | 265.0081
1520.852 | 350.0482 | 264.6842 | 264.9796
1521.852 | 350.1028 | 264.6826 | 264.9048
1522.852 | 350.1572 | 264.6723 | 264.8395
1523.852 | 350.2034 | 264.6665 | 264.7948
1524.852 | 350.1539 | 264.6604 | 264.739
1525.852 | 350.2024 | 264.6702 | 264.6736
1526.852 | 350.2436 | 264.6748 | 264.6485
1527.852 | 350.2472 | 264.6237 | 264.624
1528.852 | 350.2284 | 264.6796 | 264.5853
1529.852 | 350.1896 | 264.6194 | 264.5738
1530.852 | 350.1911 | 264.6853 | 264.578
1531.852 | 350.245 | 264.6634 | 264.5563
1532.852 | 350.1207 | 264.6582 | 264.5618
1533.852 | 350.1414 | 264.6361 | 264.5557
1534.852 | 350.2165 | 264.6429 | 264.5521
1535.852 | 350.1666 | 264.6355 | 264.5519
1536.852 | 350.1524 | 264.6993 | 264.5588
1537.852 | 350.1707 | 264.7085 | 264.5594
1538.852 | 350.1027 | 264.6898 | 264.5559
1539.852 | 350.2038 | 264.7083 | 264.5581
1540.852 | 350.2222 | 264.6894 | 264.5589
1541.852 | 350.1954 | 264.6934 | 264.5541
1542.852 | 350.193 | 264.6349 | 264.5576
1543.852 | 350.1634 | 264.6355 | 264.5528
1544.852 | 350.0566 | 264.6379 | 264.5585
1545.852 | 349.9943 | 264.6434 | 264.5635
1546.852 | 349.9508 | 264.6202 | 264.5473
1547.852 | 349.8837 | 264.6276 | 264.5512
1548.852 | 349.8464 | 264.6333 | 264.5539
1549.852 | 349.7983 | 264.6363 | 264.5511
1550.852 | 349.8063 | 264.6317 | 264.5539
1551.852 | 349.7363 | 264.6023 | 264.5534
1552.852 | 349.7563 | 264.6067 | 264.5595
1553.852 | 349.7657 | 264.6072 | 264.5588
1554.852 | 349.7405 | 264.5977 | 264.5491
1555.852 | 349.6158 | 264.5765 | 264.546
1556.852 | 349.5534 | 264.6118 | 264.5323
1557.852 | 349.5863 | 264.6042 | 264.5311
1558.852 | 349.544 | 264.5855 | 264.5134
1559.852 | 349.4948 | 264.581 | 264.5077
1560.852 | 349.4176 | 264.5604 | 264.4996
1561.852 | 349.4316 | 264.5493 | 264.4927
1562.852 | 349.4135 | 264.5459 | 264.4856
1563.852 | 349.355 | 264.5295 | 264.4713
1564.852 | 349.2922 | 264.4997 | 264.4625
1565.852 | 349.2412 | 264.4763 | 264.4547
1566.852 | 349.3041 | 264.4722 | 264.4415
1567.852 | 349.2611 | 264.4413 | 264.4248
1568.852 | 349.1748 | 264.4409 | 264.4196
1569.852 | 349.1125 | 264.4212 | 264.4061
1570.852 | 348.9139 | 264.4261 | 264.384
1571.852 | 348.8489 | 264.4262 | 264.3717
1572.852 | 348.8962 | 264.4254 | 264.3684
1573.852 | 349.0274 | 264.4307 | 264.3494
1574.852 | 349.0146 | 264.4306 | 264.3328
1575.852 | 349.0754 | 264.4266 | 264.3206
1576.852 | 349.0299 | 264.4064 | 264.3044
1577.852 | 349.1128 | 264.416 | 264.2983
1578.852 | 349.0985 | 264.4143 | 264.2803
1579.852 | 349.0111 | 264.3998 | 264.2658
1580.852 | 349.0921 | 264.3933 | 264.2532
1581.852 | 349.0416 | 264.3998 | 264.2346
1582.852 | 349.1828 | 264.3811 | 264.2266
1583.852 | 349.2563 | 264.3723 | 264.2133
1584.852 | 349.3202 | 264.364 | 264.1964
1585.852 | 349.2805 | 264.3687 | 264.1766
1586.852 | 349.2667 | 264.3514 | 264.1643
1587.852 | 349.1751 | 264.3604 | 264.1537
1588.852 | 349.0299 | 264.3482 | 264.1417
1589.852 | 348.9741 | 264.3365 | 264.1221
1590.852 | 349.0419 | 264.3426 | 264.1048
1591.852 | 349.1367 | 264.3279 | 264.0951
1592.852 | 349.2256 | 264.3305 | 264.0764
1593.852 | 349.2321 | 264.3291 | 264.0704
1594.852 | 349.3044 | 264.3229 | 264.0404
1595.852 | 349.3113 | 264.3157 | 264.0332
1596.852 | 349.2644 | 264.3238 | 264.0187
1597.852 | 349.3244 | 264.3239 | 264.0083
1598.852 | 349.4013 | 264.3212 | 263.9944
1599.852 | 349.4943 | 264.3274 | 263.9721
1600.852 | 349.5611 | 264.3138 | 263.9576
1601.852 | 349.5133 | 264.3182 | 263.9591
1602.852 | 349.5212 | 264.3056 | 263.9354
1603.852 | 349.4918 | 264.3136 | 263.9199
1604.852 | 349.5271 | 264.3158 | 263.9139
1605.852 | 349.5958 | 264.31 | 263.9046
1606.852 | 349.638 | 264.321 | 263.8895
1607.852 | 349.5858 | 264.3256 | 263.8836
1608.852 | 349.6811 | 264.3294 | 263.8671
1609.852 | 349.7651 | 264.3334 | 263.8532
1610.852 | 349.76 | 264.3245 | 263.8447
1611.852 | 349.7046 | 264.3263 | 263.8309
1612.852 | 349.6155 | 264.327 | 263.812
1613.852 | 349.6124 | 264.331 | 263.7981
1614.852 | 349.5337 | 264.3369 | 263.7814
1615.852 | 349.5996 | 264.3278 | 263.7722
1616.852 | 349.5168 | 264.3509 | 263.7631
1617.852 | 349.4669 | 264.3578 | 263.7645
1618.852 | 349.4417 | 264.3551 | 263.7578
1619.852 | 349.6107 | 264.3528 | 263.742
1620.852 | 349.7376 | 264.3643 | 263.7404
1621.852 | 349.7868 | 264.3616 | 263.7279
1622.852 | 349.8072 | 264.3652 | 263.7201
1623.852 | 349.9311 | 264.3765 | 263.7191
1624.852 | 350.01 | 264.3895 | 263.7155
1625.852 | 349.9673 | 264.387 | 263.7087
1626.852 | 349.9773 | 264.3906 | 263.7135
1627.852 | 350.0764 | 264.3958 | 263.7106
1628.852 | 350.0771 | 264.3985 | 263.7028
1629.852 | 350.0475 | 264.3959 | 263.7068
1630.852 | 350.0065 | 264.405 | 263.7
1631.852 | 349.867 | 264.4255 | 263.6999
1632.852 | 349.8355 | 264.4165 | 263.7028
1633.852 | 349.827 | 264.4354 | 263.7085
1634.852 | 349.7914 | 264.4431 | 263.7043
1635.852 | 349.7962 | 264.4321 | 263.7106
1636.852 | 349.8359 | 264.4462 | 263.7092
1637.852 | 349.8857 | 264.4647 | 263.7124
1638.852 | 349.9546 | 264.4722 | 263.7091
1639.852 | 350.0101 | 264.4611 | 263.7178
1640.852 | 350.0195 | 264.475 | 263.7217
1641.852 | 349.9981 | 264.4974 | 263.725
1642.852 | 350.0122 | 264.5032 | 263.7243
1643.852 | 349.9737 | 264.5093 | 263.7295
1644.852 | 349.7296 | 264.5146 | 263.7352
1645.852 | 349.75 | 264.5127 | 263.734
1646.852 | 349.6542 | 264.5156 | 263.7423
1647.852 | 349.6327 | 264.5305 | 263.7435
1648.852 | 349.6502 | 264.5458 | 263.7529
1649.852 | 349.7475 | 264.5566 | 263.7637
1650.852 | 349.8381 | 264.5642 | 263.7629
1651.852 | 349.8546 | 264.5766 | 263.77
1652.852 | 349.8182 | 264.5809 | 263.7783
1653.852 | 349.7285 | 264.5888 | 263.7742
1654.852 | 349.7323 | 264.5863 | 263.7795
1655.852 | 349.7899 | 264.6105 | 263.7752
1656.852 | 349.7662 | 264.6177 | 263.7861
1657.852 | 349.6295 | 264.6338 | 263.7882
1658.852 | 349.5768 | 264.6363 | 263.7964
1659.852 | 349.5865 | 264.6376 | 263.7995
1660.852 | 349.535 | 264.652 | 263.8107
1661.852 | 349.3658 | 264.6551 | 263.8105
1662.852 | 349.3633 | 264.6703 | 263.8135
1663.852 | 349.3634 | 264.6836 | 263.8243
1664.852 | 349.3369 | 264.6945 | 263.8235
1665.852 | 349.2733 | 264.7047 | 263.834
1666.852 | 349.1761 | 264.7079 | 263.8405
1667.852 | 349.2411 | 264.7233 | 263.8505
1668.852 | 349.3985 | 264.7343 | 263.8464
1669.852 | 349.466 | 264.7361 | 263.8589
1670.852 | 349.481 | 264.7493 | 263.8702
1671.852 | 349.3425 | 264.7496 | 263.8719
1672.852 | 349.1738 | 264.767 | 263.873
1673.852 | 349.2171 | 264.7873 | 263.8852
1674.852 | 349.3781 | 264.794 | 263.8818
1675.852 | 349.455 | 264.8074 | 263.8854
1676.852 | 349.4594 | 264.8176 | 263.897
1677.852 | 349.4606 | 264.8266 | 263.9021
1678.852 | 349.3416 | 264.8349 | 263.9106
1679.852 | 349.2223 | 264.8505 | 263.9144
1680.852 | 349.2062 | 264.8604 | 263.9226
1681.852 | 349.2389 | 264.8366 | 263.9268
1682.852 | 349.3198 | 264.8719 | 263.9318
1683.852 | 349.2508 | 264.8984 | 263.9312
1684.852 | 349.3566 | 264.8939 | 263.945
1685.852 | 349.3773 | 264.9042 | 263.9492
1686.852 | 349.383 | 264.9169 | 263.95
1687.852 | 349.5297 | 264.9271 | 263.9532
1688.852 | 349.6372 | 264.9274 | 263.9617
1689.852 | 349.531 | 264.9272 | 263.9708
1690.852 | 349.5673 | 264.9528 | 263.98
1691.852 | 349.5731 | 264.9654 | 263.9771
1692.852 | 349.6234 | 264.9597 | 263.9898
1693.852 | 349.6347 | 264.9675 | 263.9996
1694.852 | 349.7131 | 264.9825 | 264.0028
1695.852 | 349.8727 | 264.9913 | 264.006
1696.852 | 349.936 | 265.0019 | 264.0127
1697.852 | 349.9526 | 265.0103 | 264.0248
1698.852 | 349.9469 | 265.0249 | 264.0187
1699.852 | 349.8387 | 265.0258 | 264.0265
1700.852 | 349.7672 | 265.0276 | 264.0339
1701.852 | 349.7905 | 265.0479 | 264.0409
1702.852 | 349.785 | 265.0597 | 264.0477
1703.852 | 349.7543 | 265.0629 | 264.0536
1704.852 | 349.6975 | 265.0745 | 264.0685
1705.852 | 349.723 | 265.0786 | 264.0651
1706.852 | 349.7749 | 265.0785 | 264.0737
1707.852 | 349.7257 | 265.0902 | 264.0768
1708.852 | 349.7174 | 265.1115 | 264.0839
1709.852 | 349.6562 | 265.0998 | 264.0973
1710.852 | 349.5599 | 265.1221 | 264.0958
1711.852 | 349.3948 | 265.1334 | 264.1027
1712.852 | 349.323 | 265.1322 | 264.1171
1713.852 | 349.387 | 265.1501 | 264.125
1714.852 | 349.43 | 265.155 | 264.1337
1715.852 | 349.4338 | 265.1605 | 264.1321
1716.852 | 349.5014 | 265.1603 | 264.1422
1717.852 | 349.5811 | 265.1783 | 264.1474
1718.852 | 349.6992 | 265.1777 | 264.1532
1719.852 | 349.7279 | 265.1768 | 264.1617
1720.852 | 349.6849 | 265.1922 | 264.1678
1721.852 | 349.601 | 265.2084 | 264.1748
1722.852 | 349.5813 | 265.205 | 264.176
1723.852 | 349.6567 | 265.2134 | 264.19
1724.852 | 349.8104 | 265.2229 | 264.1907
1725.852 | 349.8317 | 265.2259 | 264.2069
1726.852 | 349.7338 | 265.226 | 264.2099
1727.852 | 349.7191 | 265.2511 | 264.2154
1728.852 | 349.5584 | 265.2357 | 264.2235
1729.852 | 349.4249 | 265.2528 | 264.2322
1730.852 | 349.4575 | 265.2564 | 264.2426
1731.852 | 349.3667 | 265.2691 | 264.2497
1732.852 | 349.1657 | 265.2729 | 264.2561
1733.852 | 349.0391 | 265.2757 | 264.2631
1734.852 | 348.9881 | 265.2864 | 264.2686
1735.852 | 348.9934 | 265.2909 | 264.2785
1736.852 | 349.033 | 265.3004 | 264.2787
1737.852 | 349.2078 | 265.2883 | 264.2796
1738.852 | 349.3476 | 265.3166 | 264.2933
1739.852 | 349.3322 | 265.3054 | 264.3011
1740.852 | 349.4097 | 265.3121 | 264.3056
1741.852 | 349.4779 | 265.3236 | 264.3133
1742.852 | 349.4955 | 265.3285 | 264.318
1743.852 | 349.4158 | 265.3265 | 264.3188
1744.852 | 349.4436 | 265.3334 | 264.3313
1745.852 | 349.3867 | 265.339 | 264.3362
1746.852 | 349.3238 | 265.3504 | 264.3483
1747.852 | 349.3612 | 265.328 | 264.3501
1748.852 | 349.3761 | 265.3425 | 264.3602
1749.852 | 349.4003 | 265.3528 | 264.3602
1750.852 | 349.5245 | 265.3595 | 264.3667
1751.852 | 349.6956 | 265.3641 | 264.3769
1752.852 | 349.8064 | 265.3688 | 264.379
1753.852 | 349.8032 | 265.3653 | 264.3853
1754.852 | 349.8013 | 265.3816 | 264.3919
1755.852 | 349.8679 | 265.3821 | 264.404
1756.852 | 349.9029 | 265.3879 | 264.4089
1757.852 | 349.992 | 265.399 | 264.4165
1758.852 | 350.0418 | 265.396 | 264.4226
1759.852 | 350.1364 | 265.4011 | 264.4291
1760.852 | 350.2278 | 265.3987 | 264.4325
1761.852 | 350.3124 | 265.4027 | 264.4428
1762.852 | 350.3871 | 265.407 | 264.4443
1763.852 | 350.378 | 265.4113 | 264.4549
1764.852 | 350.2758 | 265.4149 | 264.4545
1765.852 | 350.2854 | 265.4182 | 264.4634
1766.852 | 350.3176 | 265.4158 | 264.4706
1767.852 | 350.4905 | 265.4344 | 264.4812
1768.852 | 350.4833 | 265.43 | 264.4846
1769.852 | 350.4016 | 265.4386 | 264.492
1770.852 | 350.2621 | 265.4491 | 264.4971
1771.852 | 350.1427 | 265.4426 | 264.5053
1772.852 | 350.0933 | 265.4385 | 264.5085
1773.852 | 350.1035 | 265.4677 | 264.5176
1774.852 | 349.9601 | 265.4593 | 264.5265
1775.852 | 349.9456 | 265.4721 | 264.5225
1776.852 | 349.9588 | 265.4769 | 264.532
1777.852 | 349.8892 | 265.4729 | 264.5368
1778.852 | 349.8155 | 265.4865 | 264.5471
1779.852 | 349.74 | 265.4619 | 264.5563
1780.852 | 349.7796 | 265.5054 | 264.5635
1781.852 | 349.7826 | 265.4961 | 264.5702
1782.852 | 349.6968 | 265.5158 | 264.5798
1783.852 | 349.6423 | 265.5152 | 264.5835
1784.852 | 349.6132 | 265.5206 | 264.5943
1785.852 | 349.6674 | 265.523 | 264.5925
1786.852 | 349.6594 | 265.5238 | 264.605
1787.852 | 349.6949 | 265.5422 | 264.5999
1788.852 | 349.6471 | 265.5096 | 264.6044
1789.8833 | 349.6322 | 265.54 | 264.6133
1790.8677 | 349.5662 | 265.5418 | 264.627
1791.852 | 349.5025 | 265.5558 | 264.6375
1792.852 | 349.5472 | 265.5595 | 264.6414
1793.852 | 349.6092 | 265.565 | 264.6532
1794.852 | 349.6404 | 265.575 | 264.6518
1795.852 | 349.6741 | 265.5861 | 264.6664
1796.852 | 349.6911 | 265.5796 | 264.6738
1797.852 | 349.7011 | 265.5874 | 264.6861
1798.852 | 349.7569 | 265.6026 | 264.6918
1799.852 | 349.7397 | 265.6024 | 264.6897
1800.852 | 349.7216 | 265.6001 | 264.7055
1801.852 | 349.6474 | 265.6113 | 264.713
1802.852 | 349.5917 | 265.6187 | 264.7217
1803.852 | 349.5602 | 265.6153 | 264.7252
1804.852 | 349.5829 | 265.6205 | 264.7304
1805.852 | 349.5628 | 265.6331 | 264.7434
1806.852 | 349.4322 | 265.6356 | 264.7532
1807.852 | 349.4378 | 265.6312 | 264.7444
1808.852 | 349.43 | 265.6413 | 264.7722
1809.852 | 349.444 | 265.644 | 264.7714
1810.852 | 349.4979 | 265.6534 | 264.7845
1811.852 | 349.6193 | 265.6563 | 264.7867
1812.852 | 349.641 | 265.6558 | 264.7988
1813.852 | 349.5469 | 265.6535 | 264.8033
1814.852 | 349.5452 | 265.6654 | 264.8147
1815.852 | 349.4958 | 265.6743 | 264.829
1816.852 | 349.3659 | 265.6713 | 264.8301
1817.852 | 349.3116 | 265.6804 | 264.8465
1818.852 | 349.164 | 265.6666 | 264.8481
1819.852 | 349.0373 | 265.6911 | 264.8601
1820.852 | 348.9797 | 265.6916 | 264.8679
1821.852 | 349.0709 | 265.7015 | 264.8809
1822.8677 | 349.1033 | 265.6933 | 264.8828
1823.8989 | 349.1949 | 265.7046 | 264.8912
1824.8833 | 349.2086 | 265.6874 | 264.8952
1825.852 | 349.2449 | 265.7013 | 264.9036
1826.852 | 349.2905 | 265.7083 | 264.9026
1827.852 | 349.3677 | 265.7187 | 264.913
1828.852 | 349.3238 | 265.7073 | 264.9196
1829.852 | 349.4084 | 265.7076 | 264.9339
1830.852 | 349.4067 | 265.7136 | 264.9331
1831.852 | 349.2844 | 265.728 | 264.9473
1832.852 | 349.3211 | 265.7316 | 264.9502
1833.852 | 349.4424 | 265.7315 | 264.958
1834.852 | 349.5752 | 265.7408 | 264.9647
1835.852 | 349.6156 | 265.7352 | 264.9743
1836.852 | 349.6127 | 265.7421 | 264.9787
1837.852 | 349.6167 | 265.7401 | 264.9777
1838.852 | 349.5812 | 265.7465 | 264.99
1839.852 | 349.6023 | 265.7412 | 264.9943
1840.852 | 349.6045 | 265.7608 | 264.9836
1841.852 | 349.6301 | 265.745 | 264.9866
1842.852 | 349.5282 | 265.7621 | 264.9888
1843.852 | 349.4515 | 265.7539 | 265.0037
1844.852 | 349.3367 | 265.7668 | 265.001
1845.852 | 349.2089 | 265.7639 | 265.0044
1846.852 | 349.1889 | 265.7745 | 265.0232
1847.852 | 349.3498 | 265.7766 | 265.0208
1848.852 | 349.3981 | 265.7752 | 265.029
1849.852 | 349.4257 | 265.7707 | 265.0257
1850.852 | 349.3479 | 265.7932 | 265.0393
1851.852 | 349.2907 | 265.7933 | 265.0453
1852.852 | 349.303 | 265.796 | 265.0523
1853.852 | 349.3873 | 265.791 | 265.0595
1854.852 | 349.4239 | 265.7992 | 265.0607
1855.852 | 349.4887 | 265.791 | 265.0681
1856.852 | 349.5848 | 265.8055 | 265.0784
1857.852 | 349.672 | 265.7908 | 265.0766
1858.852 | 349.7527 | 265.8108 | 265.0861
1859.852 | 349.8896 | 265.8083 | 265.0877
1860.852 | 349.8808 | 265.8081 | 265.0932
1861.852 | 349.8879 | 265.8202 | 265.0936
1862.852 | 349.8796 | 265.8114 | 265.112
1863.852 | 349.7962 | 265.8196 | 265.1111
1864.852 | 349.7044 | 265.8154 | 265.1226
1865.852 | 349.6793 | 265.8246 | 265.1276
1866.852 | 349.7753 | 265.8251 | 265.1335
1867.852 | 349.7441 | 265.832 | 265.1308
1868.852 | 349.7304 | 265.8312 | 265.1463
1869.852 | 349.6437 | 265.8374 | 265.1479
1870.852 | 349.678 | 265.8311 | 265.1544
1871.852 | 349.7342 | 265.8388 | 265.1546
1872.852 | 349.7753 | 265.8523 | 265.1681
1873.852 | 349.7853 | 265.835 | 265.1678
1874.852 | 349.8858 | 265.8516 | 265.1763
1875.852 | 349.9281 | 265.8448 | 265.1776
1876.852 | 349.9361 | 265.8395 | 265.1903
1877.852 | 349.9621 | 265.8582 | 265.1955
1878.852 | 350.0385 | 265.8654 | 265.2089
1879.852 | 350.0219 | 265.8565 | 265.2075
1880.852 | 349.9711 | 265.8661 | 265.2176
1881.852 | 349.9139 | 265.8648 | 265.2217
1882.852 | 349.8826 | 265.8693 | 265.2233
1883.852 | 349.8566 | 265.8738 | 265.2315
1884.852 | 349.821 | 265.8644 | 265.2385
1885.852 | 349.7472 | 265.878 | 265.2507
1886.852 | 349.7759 | 265.8722 | 265.2549
1887.852 | 349.7825 | 265.867 | 265.2635
1888.852 | 349.822 | 265.8884 | 265.2691
1889.852 | 349.8429 | 265.8738 | 265.277
1890.852 | 349.8755 | 265.8877 | 265.2788
1891.852 | 349.8411 | 265.8874 | 265.289
1892.852 | 349.8976 | 265.8871 | 265.2934
1893.852 | 349.8813 | 265.8969 | 265.2957
1894.852 | 349.824 | 265.8882 | 265.3047
1895.852 | 349.8457 | 265.8934 | 265.3148
1896.852 | 349.7793 | 265.9056 | 265.3224
1897.852 | 349.636 | 265.8926 | 265.3285
1898.852 | 349.5528 | 265.9031 | 265.3296
1899.852 | 349.4658 | 265.9042 | 265.3382
1900.852 | 349.4052 | 265.9062 | 265.3469
1901.852 | 349.3148 | 265.9029 | 265.3458
1902.852 | 349.4118 | 265.9117 | 265.3568
1903.852 | 349.4711 | 265.9081 | 265.358
1904.852 | 349.496 | 265.9342 | 265.3757
1905.852 | 349.5314 | 265.9212 | 265.379
1906.852 | 349.5448 | 265.937 | 265.3847
1907.852 | 349.5076 | 265.9387 | 265.3937
1908.852 | 349.5527 | 265.9386 | 265.4081
1909.852 | 349.4459 | 265.9324 | 265.4074
1910.852 | 349.3199 | 265.9518 | 265.4159
1911.852 | 349.2441 | 265.9372 | 265.4234
1912.852 | 349.2679 | 265.9491 | 265.4283
1913.852 | 349.3646 | 265.949 | 265.4396
1914.852 | 349.3982 | 265.9549 | 265.4428
1915.852 | 349.3959 | 265.9575 | 265.4579
1916.852 | 349.4214 | 265.9701 | 265.4619
1917.852 | 349.4623 | 265.9638 | 265.4647
1918.852 | 349.3781 | 265.9785 | 265.4731
1919.852 | 349.3885 | 265.9724 | 265.4862
1920.852 | 349.2989 | 265.9833 | 265.4893
1921.852 | 349.1659 | 265.9931 | 265.5015
1922.852 | 349.1863 | 265.9923 | 265.4955
1923.852 | 349.21 | 265.9874 | 265.5084
1924.852 | 349.1975 | 265.9948 | 265.5116
1925.852 | 349.2493 | 266.0033 | 265.5273
1926.852 | 349.1716 | 266.0098 | 265.5394
1927.852 | 349.2341 | 266.0134 | 265.5381
1928.852 | 349.195 | 266.02 | 265.5445
1929.852 | 349.2549 | 266.0238 | 265.548
1930.852 | 349.1816 | 266.0177 | 265.5638
1931.852 | 349.0314 | 266.0283 | 265.5662
1932.852 | 349.0278 | 266.0232 | 265.5745
1933.852 | 349.1877 | 266.0434 | 265.579
1934.852 | 349.3099 | 266.0332 | 265.5879
1935.852 | 349.2883 | 266.0465 | 265.5886
1936.852 | 349.3716 | 266.041 | 265.5941
1937.852 | 349.413 | 266.0568 | 265.5997
1938.852 | 349.4766 | 266.0474 | 265.611
1939.852 | 349.4739 | 266.0571 | 265.6228
1940.852 | 349.5172 | 266.0649 | 265.6164
1941.852 | 349.6266 | 266.0672 | 265.6308
1942.852 | 349.6389 | 266.063 | 265.6248
1943.852 | 349.6583 | 266.0611 | 265.6336
1944.852 | 349.6721 | 266.053 | 265.6325
1945.852 | 349.5903 | 266.0558 | 265.6455
1946.852 | 349.432 | 266.0647 | 265.6481
1947.852 | 349.4342 | 266.0774 | 265.6521
1948.852 | 349.5173 | 266.0771 | 265.6542
1949.852 | 349.6766 | 266.0666 | 265.6607
1950.852 | 349.8486 | 266.077 | 265.6657
1951.852 | 349.8298 | 266.0675 | 265.6676
1952.8677 | 349.7275 | 266.0813 | 265.6747
1953.852 | 349.7059 | 266.0742 | 265.6713
1954.852 | 349.6602 | 266.0774 | 265.6788
1955.852 | 349.5053 | 266.0729 | 265.6966
1956.852 | 349.6083 | 266.0831 | 265.6961
1957.852 | 349.7188 | 266.0758 | 265.6986
1958.852 | 349.7042 | 266.0835 | 265.7009
1959.852 | 349.741 | 266.083 | 265.709
1960.852 | 349.8089 | 266.0817 | 265.7139
1961.852 | 349.8452 | 266.0925 | 265.712
1962.852 | 349.8745 | 266.0824 | 265.7274
1963.852 | 349.8892 | 266.102 | 265.7255
1964.852 | 349.975 | 266.0946 | 265.7269
1965.852 | 350.064 | 266.0811 | 265.7423
1966.852 | 350.0869 | 266.0874 | 265.7369
1967.852 | 350.0704 | 266.0886 | 265.7489
1968.852 | 349.9646 | 266.094 | 265.7511
1969.852 | 349.9707 | 266.0926 | 265.7552
1970.852 | 349.9594 | 266.0914 | 265.7587
1971.852 | 349.9686 | 266.095 | 265.7621
1972.852 | 350.0189 | 266.0956 | 265.7699
1973.852 | 349.8995 | 266.0901 | 265.7717
1974.852 | 349.958 | 266.0995 | 265.777
1975.852 | 350.0713 | 266.087 | 265.781
1976.852 | 349.9813 | 266.099 | 265.783
1977.852 | 349.8815 | 266.095 | 265.7968
1978.852 | 349.8797 | 266.0958 | 265.7996
1979.852 | 349.8178 | 266.103 | 265.8034
1980.852 | 349.8694 | 266.0959 | 265.8021
1981.852 | 349.824 | 266.0856 | 265.8102
1982.852 | 349.6275 | 266.0983 | 265.8089
1983.852 | 349.4723 | 266.0968 | 265.817
1984.852 | 349.3735 | 266.0877 | 265.8235
1985.852 | 349.2234 | 266.094 | 265.8291
1986.852 | 349.1403 | 266.0949 | 265.8321
1987.852 | 349.1507 | 266.1044 | 265.8404
1988.852 | 349.0545 | 266.1063 | 265.8393
1989.852 | 349.0669 | 266.1072 | 265.8449
1990.852 | 349.1468 | 266.1116 | 265.8504
1991.852 | 349.1436 | 266.1126 | 265.8498
1992.852 | 349.2364 | 266.1128 | 265.8665
1993.852 | 349.214 | 266.1157 | 265.8659
1994.852 | 349.3415 | 266.1173 | 265.8694
1995.852 | 349.3654 | 266.1095 | 265.8769
1996.852 | 349.3618 | 266.1198 | 265.8828
1997.852 | 349.3406 | 266.1264 | 265.8845
1998.852 | 349.3385 | 266.1369 | 265.8876
1999.852 | 349.3286 | 266.1374 | 265.8877
2000.852 | 349.2966 | 266.1261 | 265.9004
2001.852 | 349.4052 | 266.1277 | 265.9022
2002.852 | 349.5151 | 266.1221 | 265.9078
2003.852 | 349.5812 | 266.1233 | 265.9139
2004.852 | 349.5989 | 266.1333 | 265.9211
2005.852 | 349.6162 | 266.1464 | 265.9272
2006.852 | 349.6469 | 266.1314 | 265.9252
2007.852 | 349.6191 | 266.1368 | 265.9318
2008.852 | 349.6071 | 266.1446 | 265.9349
2009.852 | 349.6811 | 266.1421 | 265.9359
2010.852 | 349.7167 | 266.1523 | 265.94
2011.852 | 349.7116 | 266.1491 | 265.9476
2012.852 | 349.6074 | 266.1565 | 265.958
2013.852 | 349.5204 | 266.1542 | 265.9669
2014.852 | 349.5458 | 266.1365 | 265.9665
2015.852 | 349.6259 | 266.1599 | 265.9682
2016.852 | 349.6989 | 266.1618 | 265.9749
2017.852 | 349.6814 | 266.1657 | 265.9886
2018.852 | 349.6495 | 266.1566 | 265.9781
2019.852 | 349.6448 | 266.1558 | 265.9895
2020.852 | 349.7226 | 266.1641 | 265.9962
2021.852 | 349.6426 | 266.1697 | 266.0008
2022.852 | 349.598 | 266.1709 | 266.003
2023.852 | 349.5727 | 266.1725 | 266.0041
2024.852 | 349.5386 | 266.1772 | 266.0104
2025.852 | 349.5183 | 266.1718 | 266.0164
2026.852 | 349.5134 | 266.1879 | 266.0194
2027.852 | 349.5715 | 266.1853 | 266.0276
2028.852 | 349.6374 | 266.1851 | 266.0255
2029.852 | 349.7485 | 266.1905 | 266.0314
2030.852 | 349.8106 | 266.1939 | 266.0317
2031.852 | 349.8559 | 266.1923 | 266.0434
2032.852 | 349.8201 | 266.2117 | 266.0504
2033.852 | 349.8322 | 266.1961 | 266.0449
2034.852 | 349.857 | 266.2013 | 266.0562
2035.852 | 349.9069 | 266.2063 | 266.0625
2036.852 | 349.8233 | 266.2003 | 266.0617
2037.852 | 349.7668 | 266.207 | 266.0656
2038.852 | 349.7887 | 266.2107 | 266.0677
2039.852 | 349.7493 | 266.2134 | 266.0724
2040.852 | 349.7124 | 266.2082 | 266.0806
2041.852 | 349.8049 | 266.2095 | 266.0916
2042.852 | 349.8366 | 266.2223 | 266.0958
2043.852 | 349.9291 | 266.2192 | 266.0892
2044.852 | 350.0031 | 266.2278 | 266.1034
2045.852 | 349.9747 | 266.2179 | 266.1152
2046.852 | 349.9725 | 266.2293 | 266.117
2047.852 | 350.0735 | 266.2298 | 266.124
2048.852 | 350.117 | 266.2373 | 266.1301
2049.852 | 350.1546 | 266.2442 | 266.1331
2050.852 | 350.2006 | 266.2388 | 266.1368
2051.852 | 350.0965 | 266.2445 | 266.1518
2052.852 | 349.9397 | 266.2449 | 266.1646
2053.852 | 349.8191 | 266.2569 | 266.156
2054.852 | 349.7351 | 266.2523 | 266.1661
2055.852 | 349.692 | 266.2534 | 266.1686
2056.852 | 349.6922 | 266.2624 | 266.1804
2057.852 | 349.6131 | 266.2573 | 266.1861
2058.852 | 349.5857 | 266.2526 | 266.1911
2059.852 | 349.4037 | 266.2641 | 266.2073
2060.852 | 349.1416 | 266.2783 | 266.2075
2061.852 | 349.1273 | 266.282 | 266.2117
2062.852 | 349.1565 | 266.2908 | 266.2239
2063.852 | 349.2525 | 266.2923 | 266.2308
2064.852 | 349.3038 | 266.299 | 266.2338
2065.852 | 349.3203 | 266.2798 | 266.2409
2066.852 | 349.2459 | 266.2945 | 266.2435
2067.852 | 349.3222 | 266.3012 | 266.2562
2068.852 | 349.4716 | 266.2997 | 266.2643
2069.852 | 349.3169 | 266.3035 | 266.2716
2070.852 | 349.3517 | 266.319 | 266.2857
2071.852 | 349.3721 | 266.3148 | 266.2866
2072.852 | 349.2165 | 266.3238 | 266.2982
2073.852 | 349.0426 | 266.3268 | 266.3063
2074.852 | 348.9517 | 266.3293 | 266.3088
2075.852 | 349.0239 | 266.3378 | 266.3176
2076.852 | 349.1255 | 266.3312 | 266.3277
2077.852 | 349.1931 | 266.3359 | 266.3345
2078.852 | 349.2089 | 266.3469 | 266.3369
2079.852 | 349.2265 | 266.3467 | 266.3461
2080.852 | 349.3276 | 266.3416 | 266.3526
2081.852 | 349.3221 | 266.3538 | 266.3554
2082.852 | 349.3981 | 266.3544 | 266.353
2083.852 | 349.4402 | 266.3521 | 266.3638
2084.852 | 349.5436 | 266.3624 | 266.3727
2085.852 | 349.5126 | 266.3539 | 266.3744
2086.852 | 349.5837 | 266.3613 | 266.3878
2087.852 | 349.564 | 266.3715 | 266.3899
2088.852 | 349.5248 | 266.3676 | 266.3997
2089.852 | 349.4956 | 266.3683 | 266.4023
2090.852 | 349.5338 | 266.3767 | 266.4052
2091.852 | 349.6481 | 266.372 | 266.4123
2092.852 | 349.6015 | 266.3791 | 266.4147
2093.852 | 349.5822 | 266.3698 | 266.4272
2094.852 | 349.7138 | 266.3857 | 266.4243
2095.852 | 349.6367 | 266.3927 | 266.4363
2096.852 | 349.6598 | 266.3872 | 266.4398
2097.852 | 349.5581 | 266.383 | 266.4417
2098.852 | 349.5361 | 266.389 | 266.4484
2099.852 | 349.4206 | 266.3635 | 266.4536
2100.852 | 349.5256 | 266.385 | 266.457
2101.852 | 349.6877 | 266.3995 | 266.4722
2102.852 | 349.6765 | 266.3996 | 266.4782
2103.852 | 349.6725 | 266.4022 | 266.4786
2104.852 | 349.645 | 266.401 | 266.4816
2105.852 | 349.626 | 266.3989 | 266.4884
2106.852 | 349.6173 | 266.4057 | 266.4965
2107.852 | 349.5536 | 266.3977 | 266.4997
2108.852 | 349.6143 | 266.398 | 266.5135
2109.852 | 349.5677 | 266.4034 | 266.5106
2110.852 | 349.6515 | 266.4034 | 266.5183
2111.852 | 349.5062 | 266.4178 | 266.5275
2112.852 | 349.4753 | 266.402 | 266.5281
2113.852 | 349.453 | 266.432 | 266.5245
2114.852 | 349.3179 | 266.4145 | 266.5333
2115.852 | 349.2323 | 266.4184 | 266.5445
2116.852 | 349.1577 | 266.4326 | 266.544
2117.852 | 349.2184 | 266.4272 | 266.5539
2118.852 | 349.2827 | 266.4267 | 266.5604
2119.852 | 349.3077 | 266.4148 | 266.5701
2120.852 | 349.1364 | 266.4439 | 266.5721
2121.852 | 349.1944 | 266.4401 | 266.5843
2122.852 | 349.209 | 266.4429 | 266.588
2123.852 | 349.1055 | 266.4447 | 266.5948
2124.852 | 349.0807 | 266.4374 | 266.5966
2125.852 | 349.0803 | 266.4456 | 266.605
2126.852 | 349.0275 | 266.4596 | 266.6117
2127.852 | 348.9222 | 266.4374 | 266.6237
2128.852 | 348.701 | 266.4528 | 266.6226
2129.852 | 348.6751 | 266.4662 | 266.6344
2130.852 | 348.6662 | 266.4701 | 266.6338
2131.852 | 348.6797 | 266.4631 | 266.6378
2132.852 | 348.6929 | 266.4525 | 266.6528
2133.852 | 348.7575 | 266.4734 | 266.6549
2134.852 | 348.9344 | 266.4719 | 266.6625
2135.852 | 349.0473 | 266.4821 | 266.6641
2136.852 | 349.067 | 266.4848 | 266.6639
2137.852 | 349.1487 | 266.488 | 266.6739
2138.852 | 349.2727 | 266.4864 | 266.6779
2139.852 | 349.2747 | 266.491 | 266.6843
2140.852 | 349.2884 | 266.5055 | 266.6847
2141.852 | 349.3206 | 266.4988 | 266.6898
2142.852 | 349.3487 | 266.4845 | 266.6998
2143.852 | 349.3175 | 266.4986 | 266.7038
2144.852 | 349.483 | 266.5114 | 266.7076
2145.852 | 349.5947 | 266.5101 | 266.7088
2146.852 | 349.66 | 266.4988 | 266.7224
2147.852 | 349.7145 | 266.5102 | 266.7217
2148.852 | 349.7578 | 266.5089 | 266.7225
2149.852 | 349.7082 | 266.5232 | 266.7314
2150.852 | 349.8156 | 266.5203 | 266.7393
2151.852 | 349.7609 | 266.5317 | 266.7411
2152.852 | 349.82 | 266.5274 | 266.7541
2153.852 | 349.8891 | 266.5394 | 266.7563
2154.852 | 349.9271 | 266.5381 | 266.7571
2155.852 | 349.9696 | 266.5289 | 266.7632
2156.852 | 350.0007 | 266.5508 | 266.7665
2157.852 | 350.1093 | 266.5401 | 266.7698
2158.852 | 350.2387 | 266.5537 | 266.7792
2159.852 | 350.1477 | 266.5506 | 266.7817
2160.852 | 350.0796 | 266.5624 | 266.7825
2161.852 | 350.1051 | 266.5646 | 266.7943
2162.852 | 349.9602 | 266.5636 | 266.8001
2163.852 | 349.8799 | 266.5728 | 266.802
2164.852 | 349.9366 | 266.5773 | 266.8111
2165.852 | 349.8363 | 266.5679 | 266.812
2166.852 | 349.8541 | 266.5781 | 266.818
2167.852 | 349.8276 | 266.5792 | 266.8212
2168.852 | 349.7824 | 266.5917 | 266.8272
2169.852 | 349.7144 | 266.5848 | 266.8345
2170.852 | 349.6809 | 266.5978 | 266.8403
2171.852 | 349.6129 | 266.5893 | 266.8483
2172.852 | 349.652 | 266.5885 | 266.8489
2173.852 | 349.6152 | 266.6102 | 266.8608
2174.852 | 349.6161 | 266.6118 | 266.8645
2175.852 | 349.5363 | 266.6157 | 266.8668
2176.852 | 349.5722 | 266.6279 | 266.8744
2177.852 | 349.8086 | 266.6107 | 266.8851
2178.852 | 349.7796 | 266.64 | 266.8853
2179.852 | 349.7692 | 266.6245 | 266.8954
2180.852 | 349.7678 | 266.6197 | 266.9018
2181.852 | 349.725 | 266.6404 | 266.9127
2182.852 | 349.6551 | 266.6351 | 266.9058
2183.852 | 349.6094 | 266.6461 | 266.9133
2184.852 | 349.4864 | 266.6422 | 266.9185
2185.852 | 349.4832 | 266.6535 | 266.9256
2186.852 | 349.6034 | 266.6557 | 266.9394
2187.852 | 349.6452 | 266.6482 | 266.9442
2188.852 | 349.5462 | 266.6479 | 266.9516
2189.852 | 349.436 | 266.6724 | 266.9475
2190.852 | 349.4852 | 266.6594 | 266.9613
2191.852 | 349.477 | 266.6708 | 266.9645
2192.852 | 349.5139 | 266.6671 | 266.9746
2193.852 | 349.4256 | 266.6767 | 266.9751
2194.852 | 349.4514 | 266.6721 | 266.9817
2195.852 | 349.4683 | 266.6682 | 266.9862
2196.852 | 349.4312 | 266.688 | 266.9934
2197.852 | 349.4719 | 266.6898 | 266.9965
2198.852 | 349.3922 | 266.6885 | 267.007
2199.852 | 349.4868 | 266.6938 | 267.0162
2200.852 | 349.4974 | 266.7013 | 267.0212
2201.852 | 349.5843 | 266.7013 | 267.0286
2202.852 | 349.538 | 266.7088 | 267.0391
2203.852 | 349.5347 | 266.7125 | 267.0472
2204.852 | 349.5245 | 266.7102 | 267.0475
2205.852 | 349.4238 | 266.6969 | 267.0533
2206.852 | 349.5365 | 266.7106 | 267.0618
2207.852 | 349.5987 | 266.7056 | 267.068
2208.852 | 349.6619 | 266.7331 | 267.0811
2209.852 | 349.6214 | 266.7096 | 267.0801
2210.852 | 349.6465 | 266.7285 | 267.0849
2211.852 | 349.641 | 266.7259 | 267.0942
2212.852 | 349.6638 | 266.7289 | 267.0922
2213.852 | 349.7397 | 266.7427 | 267.1023
2214.852 | 349.7647 | 266.7321 | 267.1112
2215.852 | 349.639 | 266.7387 | 267.1169
2216.852 | 349.6278 | 266.7385 | 267.1147
2217.852 | 349.6584 | 266.7447 | 267.1325
2218.852 | 349.6296 | 266.7434 | 267.1349
2219.852 | 349.539 | 266.7318 | 267.1405
2220.852 | 349.5618 | 266.7256 | 267.1432
2221.852 | 349.6912 | 266.7427 | 267.1485
2222.852 | 349.778 | 266.7381 | 267.1515
2223.852 | 349.8497 | 266.7372 | 267.1622
2224.852 | 349.908 | 266.7398 | 267.1684
2225.852 | 349.8896 | 266.7377 | 267.173
2226.852 | 349.8735 | 266.7405 | 267.1749
2227.852 | 349.8169 | 266.7462 | 267.1857
2228.852 | 349.8824 | 266.7316 | 267.1839
2229.852 | 349.9176 | 266.7433 | 267.1935
2230.852 | 349.9019 | 266.7397 | 267.1995
2231.852 | 349.868 | 266.7421 | 267.2053
2232.852 | 349.882 | 266.7398 | 267.2103
2233.852 | 349.9382 | 266.7525 | 267.2218
2234.852 | 349.9578 | 266.7394 | 267.2224
2235.852 | 349.9386 | 266.7476 | 267.2257
2236.852 | 350.0928 | 266.75 | 267.2263
2237.852 | 350.0335 | 266.741 | 267.2399
2238.852 | 349.9847 | 266.752 | 267.2411
2239.852 | 349.9768 | 266.7503 | 267.2492
2240.852 | 349.9995 | 266.7414 | 267.2489
2241.852 | 349.9963 | 266.7513 | 267.256
2242.852 | 349.947 | 266.7374 | 267.257
2243.852 | 349.8372 | 266.7394 | 267.2643
2244.852 | 349.8216 | 266.7519 | 267.2675
2245.852 | 349.7542 | 266.7487 | 267.269
2246.852 | 349.7202 | 266.748 | 267.2664
2247.852 | 349.7783 | 266.7427 | 267.273
2248.852 | 349.768 | 266.7391 | 267.2757
2249.852 | 349.7513 | 266.7483 | 267.2833
2250.852 | 349.6938 | 266.7518 | 267.2828
2251.852 | 349.713 | 266.7438 | 267.2869
2252.852 | 349.6824 | 266.7551 | 267.2917
2253.852 | 349.6004 | 266.7343 | 267.2984
2254.852 | 349.6245 | 266.7476 | 267.3
2255.852 | 349.6609 | 266.7432 | 267.308
2256.852 | 349.6706 | 266.7258 | 267.312
2257.852 | 349.6042 | 266.7444 | 267.3134
2258.852 | 349.5916 | 266.7485 | 267.3161
2259.852 | 349.6753 | 266.747 | 267.3232
2260.852 | 349.7123 | 266.743 | 267.3194
2261.852 | 349.758 | 266.732 | 267.3368
2262.852 | 349.8347 | 266.7365 | 267.3345
2263.852 | 349.8826 | 266.727 | 267.3362
2264.852 | 349.9467 | 266.7383 | 267.3371
2265.852 | 349.8905 | 266.7401 | 267.3431
2266.852 | 349.8697 | 266.7411 | 267.3494
2267.852 | 349.836 | 266.7522 | 267.3491
2268.852 | 349.823 | 266.7338 | 267.3549
2269.852 | 349.8671 | 266.7472 | 267.3553
2270.852 | 349.8625 | 266.7435 | 267.3586
2271.852 | 349.8576 | 266.7353 | 267.3657
2272.852 | 349.7758 | 266.7337 | 267.3634
2273.852 | 349.5979 | 266.7536 | 267.3654
2274.852 | 349.5806 | 266.7385 | 267.3687
2275.852 | 349.6246 | 266.7333 | 267.3764
2276.852 | 349.4939 | 266.741 | 267.3759
2277.852 | 349.4901 | 266.7454 | 267.3823
2278.852 | 349.5244 | 266.742 | 267.3905
2279.852 | 349.5268 | 266.7392 | 267.3913
2280.852 | 349.5738 | 266.7442 | 267.3947
2281.852 | 349.5574 | 266.7421 | 267.3949
2282.852 | 349.6171 | 266.7457 | 267.401
2283.852 | 349.5762 | 266.7311 | 267.4072
2284.852 | 349.6003 | 266.7415 | 267.4129
2285.852 | 349.6135 | 266.7394 | 267.4112
2286.852 | 349.6896 | 266.7392 | 267.4183
2287.852 | 349.6542 | 266.7373 | 267.4169
2288.852 | 349.6732 | 266.7319 | 267.4296
2289.852 | 349.68 | 266.7286 | 267.428
2290.852 | 349.8384 | 266.7376 | 267.4246
2291.852 | 349.8791 | 266.7396 | 267.4308
2292.852 | 349.8938 | 266.7272 | 267.4407
2293.852 | 349.8804 | 266.7416 | 267.4456
2294.852 | 349.8414 | 266.7317 | 267.4457
2295.852 | 349.944 | 266.7377 | 267.4511
2296.852 | 349.9701 | 266.7273 | 267.4647
2297.852 | 349.9714 | 266.7286 | 267.4576
2298.852 | 350.0358 | 266.7253 | 267.4624
2299.852 | 349.9482 | 266.7337 | 267.4652
2300.852 | 349.8273 | 266.7272 | 267.4685
2301.852 | 349.6954 | 266.7274 | 267.4659
2302.852 | 349.7631 | 266.7264 | 267.4767
2303.852 | 349.7954 | 266.7299 | 267.4779
2304.852 | 349.7704 | 266.7183 | 267.4826
2305.852 | 349.6872 | 266.7129 | 267.4789
2306.852 | 349.6205 | 266.7243 | 267.4867
2307.852 | 349.5465 | 266.7216 | 267.4945
2308.852 | 349.6164 | 266.7113 | 267.4949
2309.852 | 349.574 | 266.7209 | 267.4936
2310.852 | 349.636 | 266.7182 | 267.4962
2311.852 | 349.6284 | 266.7299 | 267.5084
2312.852 | 349.5916 | 266.725 | 267.5063
2313.852 | 349.5363 | 266.7255 | 267.5037
2314.852 | 349.6722 | 266.7103 | 267.5037
2315.852 | 349.8535 | 266.7177 | 267.5055
2316.852 | 349.982 | 266.7167 | 267.5087
2317.852 | 350.0187 | 266.725 | 267.5122
2318.852 | 349.9065 | 266.7232 | 267.5198
2319.852 | 349.8666 | 266.7218 | 267.5274
2320.852 | 349.7085 | 266.7168 | 267.5228
2321.852 | 349.6612 | 266.7193 | 267.5304
2322.852 | 349.7649 | 266.7122 | 267.5412
2323.852 | 350.0609 | 266.7348 | 267.5343
2324.852 | 350.0735 | 266.7272 | 267.5348
2325.852 | 350.1247 | 266.7216 | 267.5373
2326.852 | 349.9685 | 266.7238 | 267.5363
2327.852 | 349.8694 | 266.7205 | 267.5388
2328.852 | 349.8107 | 266.7266 | 267.551
2329.852 | 349.7726 | 266.7194 | 267.5478
2330.852 | 349.74 | 266.7094 | 267.5457
2331.852 | 349.781 | 266.73 | 267.5455
2332.852 | 349.7422 | 266.7215 | 267.5505
2333.852 | 349.7866 | 266.724 | 267.5598
2334.852 | 349.8074 | 266.7245 | 267.5589
2335.852 | 349.8582 | 266.7308 | 267.5591
2336.852 | 349.9105 | 266.7207 | 267.5577
2337.852 | 349.8197 | 266.7264 | 267.5671
2338.852 | 349.8682 | 266.7202 | 267.5704
2339.852 | 349.8862 | 266.732 | 267.5678
2340.852 | 349.9364 | 266.7455 | 267.572
2341.852 | 349.9349 | 266.7336 | 267.578
2342.852 | 349.8804 | 266.7421 | 267.5809
2343.852 | 349.8321 | 266.7359 | 267.5863
2344.852 | 349.8611 | 266.7166 | 267.5833
2345.852 | 349.6628 | 266.7542 | 267.5937
2346.852 | 349.5963 | 266.7422 | 267.6029
2347.852 | 349.4976 | 266.7551 | 267.5939
2348.852 | 349.4597 | 266.7498 | 267.6018
2349.852 | 349.3756 | 266.7671 | 267.5985
2350.852 | 349.3298 | 266.7607 | 267.6098
2351.852 | 349.3816 | 266.754 | 267.6188
2352.852 | 349.4671 | 266.7568 | 267.6161
2353.852 | 349.4917 | 266.7648 | 267.6215
2354.852 | 349.375 | 266.7613 | 267.6378
2355.852 | 349.3202 | 266.7635 | 267.6342
2356.852 | 349.3805 | 266.775 | 267.6348
2357.852 | 349.442 | 266.7721 | 267.6403
2358.852 | 349.3971 | 266.7781 | 267.6402
2359.852 | 349.4822 | 266.753 | 267.6421
2360.852 | 349.6418 | 266.7776 | 267.6453
2361.852 | 349.7413 | 266.7839 | 267.6565
2362.852 | 349.759 | 266.7802 | 267.6607
2363.852 | 349.9024 | 266.7626 | 267.6638
2364.852 | 350.0087 | 266.7785 | 267.6708
2365.852 | 350.0022 | 266.7852 | 267.6775
2366.852 | 349.9255 | 266.7622 | 267.6788
2367.852 | 349.9782 | 266.7849 | 267.6945
2368.852 | 350.0608 | 266.7835 | 267.695
2369.852 | 350.0711 | 266.7843 | 267.6946
2370.852 | 349.9477 | 266.7889 | 267.7043
2371.852 | 349.9111 | 266.783 | 267.707
2372.852 | 349.9416 | 266.7951 | 267.7161
2373.852 | 350.1295 | 266.7864 | 267.7183
2374.852 | 350.244 | 266.7911 | 267.7263
2375.852 | 350.2455 | 266.7855 | 267.7277
2376.852 | 350.2077 | 266.7976 | 267.7274
2377.852 | 350.1852 | 266.7988 | 267.7319
2378.852 | 350.2188 | 266.7907 | 267.7401
2379.852 | 350.2861 | 266.7826 | 267.7417
2380.852 | 350.3677 | 266.7807 | 267.7471
2381.852 | 350.2804 | 266.7891 | 267.749
2382.852 | 350.2396 | 266.7995 | 267.757
2383.852 | 350.1272 | 266.7964 | 267.7572
2384.852 | 350.0829 | 266.7882 | 267.7676
2385.852 | 350.0845 | 266.7872 | 267.7543
2386.852 | 350.0926 | 266.7764 | 267.7663
2387.852 | 350.1037 | 266.7937 | 267.7684
2388.852 | 350.1029 | 266.8075 | 267.7776
2389.852 | 350.2278 | 266.8086 | 267.7771
2390.852 | 350.2387 | 266.796 | 267.7801
2391.852 | 350.238 | 266.8062 | 267.7901
2392.852 | 350.2282 | 266.7887 | 267.7882
2393.852 | 350.2413 | 266.8002 | 267.7948
2394.852 | 350.2856 | 266.8036 | 267.7979
2395.852 | 350.364 | 266.7992 | 267.7956
2396.852 | 350.3155 | 266.8061 | 267.8044
2397.852 | 350.1918 | 266.8001 | 267.8049
2398.852 | 350.2079 | 266.8024 | 267.8091
2399.852 | 350.1959 | 266.8186 | 267.8106
2400.852 | 350.1992 | 266.802 | 267.8196
2401.852 | 350.2558 | 266.7959 | 267.817
2402.852 | 350.2336 | 266.7997 | 267.8163
2403.852 | 350.2443 | 266.7948 | 267.818
2404.852 | 350.1492 | 266.8071 | 267.8243
2405.852 | 350.0926 | 266.8118 | 267.8339
2406.852 | 350.1304 | 266.8125 | 267.8291
2407.852 | 350.0861 | 266.7969 | 267.8327
2408.852 | 350.0854 | 266.8031 | 267.8433
2409.852 | 350.0312 | 266.8032 | 267.8408
2410.852 | 349.9842 | 266.8015 | 267.8454
2411.852 | 349.8926 | 266.8117 | 267.843
2412.852 | 349.8549 | 266.8187 | 267.8474
2413.852 | 349.7642 | 266.8112 | 267.8425
2414.852 | 349.8414 | 266.8152 | 267.8508
2415.852 | 349.7963 | 266.8163 | 267.8581
2416.852 | 349.8187 | 266.8039 | 267.8599
2417.852 | 349.7673 | 266.8128 | 267.8553
2418.852 | 349.8381 | 266.8125 | 267.8578
2419.852 | 349.7821 | 266.8083 | 267.8618
2420.852 | 349.7513 | 266.8234 | 267.8754
2421.852 | 349.7342 | 266.8149 | 267.8658
2422.852 | 349.873 | 266.8124 | 267.8705
2423.852 | 349.8939 | 266.8006 | 267.876
2424.852 | 350.0066 | 266.8212 | 267.887
2425.852 | 350.0336 | 266.8184 | 267.8814
2426.852 | 349.9255 | 266.8175 | 267.8842
2427.852 | 349.8695 | 266.8115 | 267.888
2428.852 | 349.8926 | 266.8083 | 267.8904
2429.852 | 349.7741 | 266.8182 | 267.9001
2430.852 | 349.5413 | 266.8087 | 267.8935
2431.852 | 349.6124 | 266.806 | 267.8925
2432.852 | 349.6238 | 266.8153 | 267.8965
2433.852 | 349.5689 | 266.8206 | 267.8908
2434.852 | 349.4031 | 266.8121 | 267.8969
2435.852 | 349.3856 | 266.8094 | 267.9024
2436.852 | 349.4662 | 266.8099 | 267.9031
2437.852 | 349.5349 | 266.8228 | 267.8955
2438.852 | 349.6813 | 266.8154 | 267.904
2439.852 | 349.6185 | 266.8168 | 267.9063
2440.852 | 349.6396 | 266.8143 | 267.908
2441.852 | 349.6673 | 266.8092 | 267.9045
2442.852 | 349.6981 | 266.8081 | 267.9137
2443.852 | 349.7073 | 266.799 | 267.9088
2444.852 | 349.661 | 266.8079 | 267.9109
2445.852 | 349.572 | 266.8126 | 267.9145
2446.852 | 349.5888 | 266.8199 | 267.9185
2447.852 | 349.5499 | 266.8068 | 267.9119
2448.852 | 349.5486 | 266.808 | 267.9164
2449.852 | 349.5363 | 266.8006 | 267.9101
2450.852 | 349.6453 | 266.8054 | 267.9184
2451.852 | 349.7456 | 266.8036 | 267.9253
2452.852 | 349.7115 | 266.8051 | 267.9242
2453.852 | 349.7875 | 266.8071 | 267.9231
2454.852 | 349.776 | 266.8011 | 267.9281
2455.852 | 349.7517 | 266.8094 | 267.9166
2456.852 | 349.7068 | 266.7932 | 267.9242
2457.852 | 349.703 | 266.7985 | 267.9157
2458.852 | 349.729 | 266.7963 | 267.9201
2459.852 | 349.7678 | 266.802 | 267.9143
2460.852 | 349.7113 | 266.7872 | 267.9165
2461.852 | 349.6811 | 266.7934 | 267.9213
2462.852 | 349.6614 | 266.798 | 267.9175
2463.852 | 349.6046 | 266.8047 | 267.921
2464.852 | 349.5752 | 266.7966 | 267.9224
2465.852 | 349.5761 | 266.801 | 267.9103
2466.852 | 349.6044 | 266.7888 | 267.9137
2467.852 | 349.68 | 266.7918 | 267.9146
2468.852 | 349.6136 | 266.7807 | 267.9136
2469.852 | 349.6703 | 266.7883 | 267.9174
2470.852 | 349.7793 | 266.7988 | 267.9168
2471.852 | 349.7237 | 266.7838 | 267.9182
2472.852 | 349.8235 | 266.792 | 267.9099
2473.852 | 349.7738 | 266.7892 | 267.9078
2474.852 | 349.755 | 266.7636 | 267.9137
2475.852 | 349.7024 | 266.7727 | 267.9134
2476.852 | 349.7662 | 266.7828 | 267.911
2477.852 | 349.8075 | 266.7792 | 267.9106
2478.852 | 349.7922 | 266.7795 | 267.9118
2479.852 | 349.7502 | 266.7855 | 267.9186
2480.852 | 349.7447 | 266.777 | 267.9191
2481.852 | 349.7736 | 266.7677 | 267.9026
2482.852 | 349.8746 | 266.7634 | 267.9097
2483.852 | 349.8967 | 266.7807 | 267.907
2484.852 | 349.9425 | 266.7721 | 267.9038
2485.852 | 350.0546 | 266.767 | 267.9034
2486.852 | 350.044 | 266.7686 | 267.9117
2487.852 | 350.0518 | 266.7722 | 267.9058
2488.852 | 349.9825 | 266.7643 | 267.9055
2489.852 | 350.0061 | 266.7519 | 267.909
2490.852 | 350.0107 | 266.7591 | 267.9064
2491.852 | 349.9779 | 266.7634 | 267.8981
2492.852 | 349.9551 | 266.7625 | 267.8958
2493.852 | 350.0066 | 266.7564 | 267.8986
2494.852 | 350.0356 | 266.7585 | 267.897
2495.852 | 350.1538 | 266.7484 | 267.8899
2496.852 | 350.2696 | 266.7424 | 267.8938
2497.852 | 350.1773 | 266.7506 | 267.8951
2498.852 | 350.2235 | 266.7493 | 267.8901
2499.852 | 350.2327 | 266.7414 | 267.8889
2500.852 | 350.1763 | 266.7498 | 267.8936
2501.852 | 350.1918 | 266.7429 | 267.8846
2502.852 | 350.0891 | 266.7431 | 267.8934
2503.852 | 350.0381 | 266.7379 | 267.8785
2504.852 | 349.9898 | 266.7418 | 267.8849
2505.852 | 349.9235 | 266.7418 | 267.8856
2506.852 | 349.9259 | 266.7366 | 267.8899
2507.852 | 349.9482 | 266.7397 | 267.8823
2508.852 | 349.96 | 266.7375 | 267.8862
2509.852 | 350.0101 | 266.7366 | 267.8797
2510.852 | 349.9856 | 266.7337 | 267.8823
2511.852 | 349.9215 | 266.7311 | 267.8821
2512.852 | 349.9559 | 266.7349 | 267.8839
2513.852 | 349.9161 | 266.7274 | 267.8803
2514.852 | 349.7819 | 266.7213 | 267.8804
2515.852 | 349.6732 | 266.7241 | 267.8784
2516.852 | 349.6146 | 266.7206 | 267.8844
2517.852 | 349.6665 | 266.7345 | 267.8799
2518.852 | 349.5656 | 266.7338 | 267.8797
2519.852 | 349.5628 | 266.7327 | 267.8775
2520.852 | 349.625 | 266.7195 | 267.8768
2521.852 | 349.6879 | 266.7288 | 267.8796
2522.852 | 349.671 | 266.7172 | 267.8731
2523.852 | 349.6126 | 266.7318 | 267.8798
2524.852 | 349.5298 | 266.7279 | 267.8771
2525.852 | 349.558 | 266.7277 | 267.8803
2526.852 | 349.6135 | 266.7195 | 267.8806
2527.852 | 349.6836 | 266.7118 | 267.8896
2528.852 | 349.6729 | 266.7137 | 267.8896
2529.852 | 349.6884 | 266.719 | 267.8917
2530.852 | 349.6473 | 266.7157 | 267.8821
2531.852 | 349.5622 | 266.7106 | 267.8882
2532.852 | 349.6153 | 266.7114 | 267.8878
2533.852 | 349.683 | 266.7177 | 267.892
2534.852 | 349.6049 | 266.7057 | 267.8861
2535.852 | 349.5801 | 266.7126 | 267.8876
2536.852 | 349.6254 | 266.7149 | 267.8906
2537.852 | 349.6226 | 266.7085 | 267.904
2538.852 | 349.5977 | 266.7161 | 267.9005
2539.852 | 349.5703 | 266.7093 | 267.8945
2540.852 | 349.601 | 266.7183 | 267.894
2541.852 | 349.6701 | 266.7032 | 267.8985
2542.852 | 349.7248 | 266.7136 | 267.9023
2543.852 | 349.7584 | 266.7095 | 267.909
2544.852 | 349.7685 | 266.7099 | 267.9154
2545.852 | 349.7278 | 266.7079 | 267.912
2546.852 | 349.8221 | 266.707 | 267.9169
2547.852 | 349.8862 | 266.7038 | 267.9114
2548.852 | 349.8029 | 266.7145 | 267.91
2549.852 | 349.8575 | 266.7129 | 267.917
2550.852 | 349.9462 | 266.714 | 267.9183
2551.852 | 349.7846 | 266.7102 | 267.9147
2552.852 | 349.8028 | 266.7093 | 267.9185
2553.852 | 349.7672 | 266.7066 | 267.9162
2554.852 | 349.7684 | 266.716 | 267.9252
2555.852 | 349.7881 | 266.7004 | 267.9239
2556.852 | 349.7924 | 266.7113 | 267.9233
2557.852 | 349.8099 | 266.724 | 267.9248
2558.852 | 349.8435 | 266.7155 | 267.9261
2559.852 | 349.8794 | 266.7177 | 267.9306
2560.852 | 349.7337 | 266.7168 | 267.926
2561.852 | 349.7156 | 266.7154 | 267.9402
2562.852 | 349.741 | 266.7199 | 267.9394
2563.852 | 349.7111 | 266.7283 | 267.9371
2564.852 | 349.5678 | 266.7275 | 267.9379
2565.852 | 349.5456 | 266.7217 | 267.9416
2566.852 | 349.4755 | 266.7198 | 267.9469
2567.852 | 349.3181 | 266.7235 | 267.9416
2568.852 | 349.1914 | 266.7257 | 267.9424
2569.852 | 349.2374 | 266.7237 | 267.9425
2570.852 | 349.3987 | 266.7268 | 267.9488
2571.852 | 349.4202 | 266.7309 | 267.9485
2572.852 | 349.3859 | 266.7326 | 267.9453
2573.852 | 349.3708 | 266.7344 | 267.9537
2574.852 | 349.3336 | 266.7193 | 267.9467
2575.852 | 349.3947 | 266.7321 | 267.948
2576.852 | 349.47 | 266.7178 | 267.9515
2577.852 | 349.4011 | 266.7321 | 267.9516
2578.852 | 349.3972 | 266.7351 | 267.9483
2579.852 | 349.4737 | 266.7371 | 267.9554
2580.852 | 349.4713 | 266.741 | 267.9593
2581.852 | 349.4497 | 266.7351 | 267.9565
2582.852 | 349.4912 | 266.7395 | 267.9529
2583.852 | 349.5829 | 266.7421 | 267.9522
2584.852 | 349.6698 | 266.7436 | 267.9537
2585.852 | 349.7046 | 266.7365 | 267.9562
2586.852 | 349.7335 | 266.734 | 267.9521
2587.852 | 349.7546 | 266.7292 | 267.9578
2588.852 | 349.7791 | 266.7402 | 267.9547
2589.852 | 349.8872 | 266.7447 | 267.9547
2590.852 | 349.8967 | 266.7391 | 267.9497
2591.852 | 349.8056 | 266.735 | 267.9596
2592.852 | 349.7053 | 266.7505 | 267.9513
2593.852 | 349.5816 | 266.7413 | 267.9582
2594.852 | 349.61 | 266.7419 | 267.9532
2595.852 | 349.5957 | 266.7331 | 267.9462
2596.852 | 349.6254 | 266.7467 | 267.9515
2597.852 | 349.7145 | 266.7347 | 267.9528
2598.852 | 349.7184 | 266.7467 | 267.9508
2599.852 | 349.6805 | 266.7441 | 267.9473
2600.852 | 349.6209 | 266.7395 | 267.9417
2601.852 | 349.6946 | 266.7422 | 267.948
2602.852 | 349.8188 | 266.7346 | 267.9515
2603.852 | 349.8525 | 266.7303 | 267.9465
2604.852 | 349.8279 | 266.7356 | 267.9485
2605.852 | 349.893 | 266.7313 | 267.9452
2606.852 | 349.9523 | 266.7368 | 267.9439
2607.852 | 349.9047 | 266.737 | 267.9392
2608.852 | 349.9253 | 266.7339 | 267.9408
2609.852 | 349.8097 | 266.7456 | 267.9331
2610.852 | 349.7803 | 266.7408 | 267.9421
2611.852 | 349.6591 | 266.7487 | 267.94
2612.852 | 349.5475 | 266.7409 | 267.9412
2613.852 | 349.6141 | 266.7365 | 267.9329
2614.852 | 349.598 | 266.7385 | 267.9237
2615.852 | 349.6746 | 266.721 | 267.9365
2616.852 | 349.6924 | 266.7437 | 267.9385
2617.852 | 349.7134 | 266.7433 | 267.9334
2618.852 | 349.7257 | 266.7513 | 267.9403
2619.852 | 349.6833 | 266.748 | 267.9356
2620.852 | 349.6362 | 266.7381 | 267.9333
2621.852 | 349.759 | 266.7329 | 267.9319
2622.852 | 349.8755 | 266.7547 | 267.932
2623.852 | 349.8181 | 266.7428 | 267.9283
2624.852 | 349.6887 | 266.739 | 267.9337
2625.852 | 349.6992 | 266.7366 | 267.9291
2626.852 | 349.6851 | 266.7563 | 267.9257
2627.852 | 349.6533 | 266.754 | 267.9278
2628.852 | 349.6541 | 266.7629 | 267.9358
2629.852 | 349.6633 | 266.7424 | 267.9274
2630.852 | 349.7705 | 266.7533 | 267.9316
2631.852 | 349.8235 | 266.7502 | 267.9266
2632.852 | 349.9228 | 266.7631 | 267.929
2633.852 | 350.0106 | 266.7621 | 267.9276
2634.852 | 350.1009 | 266.7597 | 267.933
2635.852 | 350.0534 | 266.7667 | 267.9309
2636.852 | 350.0406 | 266.759 | 267.9246
2637.852 | 350.0341 | 266.7636 | 267.9241
2638.852 | 349.9693 | 266.7635 | 267.9193
2639.852 | 350.002 | 266.7627 | 267.926
2640.852 | 349.8916 | 266.7601 | 267.9225
2641.852 | 349.7404 | 266.7778 | 267.9272
2642.852 | 349.7717 | 266.7731 | 267.9311
2643.852 | 349.799 | 266.7692 | 267.9292
2644.852 | 349.7813 | 266.7757 | 267.9227
2645.852 | 349.7386 | 266.7671 | 267.9297
2646.852 | 349.7255 | 266.7706 | 267.931
2647.852 | 349.7112 | 266.7631 | 267.9285
2648.852 | 349.7495 | 266.7701 | 267.926
2649.852 | 349.6339 | 266.7654 | 267.9202
2650.852 | 349.6196 | 266.762 | 267.933
2651.852 | 349.6479 | 266.7666 | 267.9136
2652.852 | 349.7628 | 266.7716 | 267.9194
2653.852 | 349.6819 | 266.7657 | 267.9179
2654.852 | 349.677 | 266.7633 | 267.9153
2655.852 | 349.628 | 266.7713 | 267.9221
2656.852 | 349.5934 | 266.7763 | 267.9185
2657.852 | 349.6014 | 266.7605 | 267.9187
2658.852 | 349.541 | 266.768 | 267.923
2659.852 | 349.5299 | 266.7733 | 267.9229
2660.852 | 349.5205 | 266.7713 | 267.9149
2661.852 | 349.5568 | 266.7701 | 267.9143
2662.852 | 349.6044 | 266.7552 | 267.9218
2663.852 | 349.5551 | 266.7784 | 267.911
2664.852 | 349.5239 | 266.7601 | 267.9136
2665.852 | 349.4993 | 266.768 | 267.915
2666.852 | 349.3752 | 266.7634 | 267.9093
2667.852 | 349.3011 | 266.7609 | 267.9124
2668.852 | 349.2927 | 266.7461 | 267.915
2669.852 | 349.3489 | 266.7467 | 267.9095
2670.852 | 349.3533 | 266.754 | 267.9138
2671.852 | 349.3743 | 266.7403 | 267.9102
2672.852 | 349.4173 | 266.7384 | 267.9117
2673.852 | 349.4617 | 266.7356 | 267.9064
2674.852 | 349.4055 | 266.7205 | 267.9085
2675.852 | 349.2743 | 266.7209 | 267.9017
2676.852 | 349.3193 | 266.7223 | 267.8975
2677.852 | 349.3836 | 266.7163 | 267.9041
2678.852 | 349.5195 | 266.7001 | 267.8935
2679.852 | 349.5486 | 266.6912 | 267.9044
2680.852 | 349.4779 | 266.6931 | 267.9042
2681.852 | 349.5091 | 266.6903 | 267.9068
2682.852 | 349.4395 | 266.6766 | 267.9045
2683.852 | 349.5172 | 266.6762 | 267.9052
2684.852 | 349.4834 | 266.6591 | 267.9023
2685.852 | 349.4944 | 266.6607 | 267.899
2686.852 | 349.483 | 266.6514 | 267.8928
2687.852 | 349.4154 | 266.6479 | 267.8912
2688.852 | 349.3986 | 266.6397 | 267.8992
2689.852 | 349.3892 | 266.6398 | 267.8979
2690.852 | 349.4779 | 266.6241 | 267.8936
2691.852 | 349.4951 | 266.6139 | 267.8949
2692.852 | 349.5794 | 266.6053 | 267.889
2693.852 | 349.6957 | 266.6051 | 267.8852
2694.852 | 349.6629 | 266.5865 | 267.8808
2695.852 | 349.7527 | 266.5833 | 267.8806
2696.852 | 349.7027 | 266.5759 | 267.8811
2697.852 | 349.5986 | 266.5744 | 267.8754
2698.852 | 349.5221 | 266.5733 | 267.8743
2699.852 | 349.527 | 266.5559 | 267.869
2700.852 | 349.4641 | 266.5572 | 267.8778
2701.852 | 349.4549 | 266.5361 | 267.8678
2702.852 | 349.572 | 266.5402 | 267.8708
2703.852 | 349.7209 | 266.5297 | 267.8677
2704.852 | 349.6375 | 266.5087 | 267.8663
2705.852 | 349.5022 | 266.5172 | 267.8562
2706.852 | 349.694 | 266.5086 | 267.8594
2707.852 | 349.8854 | 266.5138 | 267.8567
2708.852 | 350.0505 | 266.4987 | 267.8552
2709.852 | 350.0462 | 266.4907 | 267.8513
2710.852 | 349.9452 | 266.4818 | 267.8427
2711.852 | 349.7948 | 266.4716 | 267.8476
2712.852 | 349.7234 | 266.4621 | 267.8472
2713.852 | 349.5823 | 266.4664 | 267.8391
2714.852 | 349.5745 | 266.4591 | 267.8479
2715.852 | 349.7619 | 266.4576 | 267.8434
2716.852 | 349.753 | 266.4397 | 267.8425
2717.852 | 349.816 | 266.4475 | 267.8417
2718.852 | 349.962 | 266.4285 | 267.8387
2719.852 | 349.9975 | 266.4264 | 267.8367
2720.852 | 350.0719 | 266.4147 | 267.8319
2721.852 | 350.0735 | 266.416 | 267.8251
2722.852 | 349.999 | 266.41 | 267.8249
2723.852 | 349.9773 | 266.4048 | 267.8277
2724.852 | 350.0401 | 266.3964 | 267.8241
2725.852 | 350.0396 | 266.3903 | 267.8187
2726.852 | 350.0479 | 266.3678 | 267.8242
2727.852 | 350.1177 | 266.3742 | 267.8229
2728.852 | 350.2257 | 266.3653 | 267.8207
2729.852 | 350.3498 | 266.3658 | 267.8313
2730.852 | 350.4861 | 266.3576 | 267.8223
2731.852 | 350.5293 | 266.3441 | 267.8196
2732.852 | 350.4447 | 266.3447 | 267.8197
2733.852 | 350.3984 | 266.3374 | 267.8163
2734.852 | 350.4226 | 266.3083 | 267.8091
2735.852 | 350.4023 | 266.3169 | 267.8178
2736.852 | 350.3044 | 266.306 | 267.8064
2737.852 | 350.33 | 266.3049 | 267.8069
2738.852 | 350.2684 | 266.2956 | 267.8084
2739.852 | 350.245 | 266.3017 | 267.7961
2740.852 | 350.1654 | 266.2915 | 267.8038
2741.852 | 350.1931 | 266.2846 | 267.8018
2742.852 | 350.2449 | 266.2731 | 267.8124
2743.852 | 350.1498 | 266.2705 | 267.8023
2744.852 | 350.1001 | 266.2593 | 267.806
2745.852 | 350.0237 | 266.2626 | 267.8022
2746.852 | 349.9526 | 266.2516 | 267.8012
2747.852 | 349.943 | 266.2303 | 267.8017
2748.852 | 349.9257 | 266.2385 | 267.7987
2749.852 | 349.9327 | 266.2276 | 267.8026
2750.852 | 349.9755 | 266.2247 | 267.7978
2751.852 | 349.9227 | 266.2348 | 267.8
2752.852 | 349.9115 | 266.2105 | 267.8046
2753.852 | 349.8287 | 266.2211 | 267.7979
2754.852 | 349.8619 | 266.2127 | 267.7909
2755.852 | 349.881 | 266.2092 | 267.8018
2756.852 | 349.8097 | 266.2084 | 267.7908
2757.852 | 349.6442 | 266.2018 | 267.8029
2758.852 | 349.4319 | 266.1963 | 267.7924
2759.852 | 349.3797 | 266.1901 | 267.7918
2760.852 | 349.369 | 266.1868 | 267.7946
2761.852 | 349.275 | 266.1873 | 267.7911
2762.852 | 349.3413 | 266.1767 | 267.7889
2763.852 | 349.3509 | 266.1761 | 267.7941
2764.852 | 349.4879 | 266.1596 | 267.7893
2765.852 | 349.6233 | 266.1631 | 267.7923
2766.852 | 349.6388 | 266.1588 | 267.7898
2767.852 | 349.7615 | 266.159 | 267.7883
2768.852 | 349.767 | 266.1533 | 267.7851
2769.852 | 349.738 | 266.1358 | 267.7835
2770.852 | 349.7495 | 266.1441 | 267.788
2771.852 | 349.8023 | 266.1258 | 267.7851
2772.852 | 349.9767 | 266.1153 | 267.7799
2773.852 | 350.0741 | 266.12 | 267.7771
2774.852 | 350.158 | 266.1153 | 267.775
2775.852 | 350.2234 | 266.0987 | 267.7817
2776.852 | 350.2004 | 266.107 | 267.7821
2777.852 | 350.2622 | 266.0958 | 267.7761
2778.852 | 350.249 | 266.0956 | 267.7658
2779.852 | 350.1688 | 266.0956 | 267.7779
2780.852 | 350.2265 | 266.0867 | 267.7746
2781.852 | 350.2094 | 266.0774 | 267.7689
2782.852 | 350.1655 | 266.0813 | 267.7687
2783.852 | 350.2097 | 266.0692 | 267.7669
2784.852 | 350.1825 | 266.0751 | 267.7714
2785.852 | 350.1459 | 266.0637 | 267.7735
2786.852 | 350.1097 | 266.0605 | 267.7666
2787.852 | 350.0983 | 266.0673 | 267.7599
2788.852 | 349.9991 | 266.0541 | 267.7618
2789.852 | 349.9672 | 266.0581 | 267.7591
2790.852 | 350.0185 | 266.0465 | 267.758
2791.852 | 350.0196 | 266.0359 | 267.7577
2792.852 | 350.0496 | 266.0353 | 267.7556
2793.852 | 349.9933 | 266.0206 | 267.7491
2794.852 | 349.9097 | 266.0277 | 267.7464
2795.852 | 349.8281 | 266.0213 | 267.7473
2796.852 | 349.8306 | 266.0165 | 267.7557
2797.852 | 349.8062 | 266.0127 | 267.753
2798.852 | 349.7668 | 266.0131 | 267.742
2799.852 | 349.8163 | 266.0121 | 267.738
2800.852 | 349.7062 | 265.9993 | 267.7495
2801.852 | 349.5513 | 265.9843 | 267.7471
2802.852 | 349.4249 | 265.9883 | 267.7413
2803.852 | 349.5062 | 265.9818 | 267.7361
2804.852 | 349.5891 | 265.988 | 267.7316
2805.852 | 349.5943 | 265.9734 | 267.7383
2806.852 | 349.6596 | 265.9753 | 267.7281
2807.852 | 349.6872 | 265.9717 | 267.7358
2808.852 | 349.6891 | 265.9601 | 267.7247
2809.852 | 349.7794 | 265.9614 | 267.7274
2810.852 | 349.829 | 265.9532 | 267.7237
2811.852 | 349.808 | 265.9464 | 267.7217
2812.852 | 349.7523 | 265.9444 | 267.7264
2813.852 | 349.661 | 265.934 | 267.7191
2814.852 | 349.726 | 265.9363 | 267.7184
2815.852 | 349.7535 | 265.9239 | 267.7218
2816.852 | 349.6855 | 265.924 | 267.711
2817.852 | 349.6736 | 265.8983 | 267.7142
2818.852 | 349.6323 | 265.9112 | 267.7086
2819.852 | 349.7035 | 265.9029 | 267.7064
2820.852 | 349.7244 | 265.9048 | 267.702
2821.852 | 349.7445 | 265.8981 | 267.7049
2822.852 | 349.8176 | 265.8946 | 267.692
2823.852 | 349.8705 | 265.8877 | 267.6959
2824.852 | 349.9528 | 265.8828 | 267.6928
2825.852 | 350.0081 | 265.8586 | 267.6835
2826.852 | 350.0971 | 265.8757 | 267.6804
2827.852 | 350.1338 | 265.854 | 267.6746
2828.852 | 350.201 | 265.8546 | 267.6652
2829.852 | 350.2675 | 265.8452 | 267.6709
2830.852 | 350.2646 | 265.844 | 267.6692
2831.852 | 350.2638 | 265.8279 | 267.6606
2832.852 | 350.2581 | 265.8266 | 267.6623
2833.852 | 350.1884 | 265.8217 | 267.6526
2834.852 | 350.1967 | 265.813 | 267.6557
2835.852 | 350.0968 | 265.8007 | 267.6446
2836.852 | 349.906 | 265.7967 | 267.6382
2864.4927 | 349.9176 | 265.6328 | 267.5466
2865.0552 | 349.9199 | 265.6315 | 267.5512
3006.7427 | 313.1783 | 263.4984 | 266.3939